## Notes
Because this is the temporary 2 channel set up, the spreadsheet has changed
range limits
cadmium column efficiency calculations
Nitrogen run was trouble free
Phosphate carrier was mis matched resulting in big dips before peaks
resulting in integration problems
so I tried 4 ways of dealing with the deep dips
2 factors
| baseline | integration out of the dip vs integration off a flat baseline ignoring dips
| signal | signal as area vs signal as height

## QAQC flat ht
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE NO3-N: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- dilution corrected | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- dilution corrected | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | original order
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | dye | S16 | 1 | 1 | TN | ug N/L | -0.00637 | -0.00462 | 0 | TP | ug P/L | 0.0299 | 0.592 | -2.03 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 2 | reworked calibration | -40.8001 | -40.8001 | -40.8001 | [6 empty] | 1
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | dye | S16 | 1 | 1 | TN | ug N/L | 0.0144 | 0.152 | 0 | TP | ug P/L | 0.0314 | 0.622 | -0.949 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 2 | reworked calibration | -38.155 | -38.155 | -38.155 | [6 empty] | 2
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.27 | 30 | 0 | TP | ug P/L | -0.00787 | -0.0671 | -30.3 | 98.3333 | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 2 | reworked calibration | -116.7122 | -116.7122 | -116.7122 | [6 empty] | 3
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.25 | 29.5 | 0 | TP | ug P/L | -0.00942 | -0.0814 | -31.5 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 2 | reworked calibration | -120.21 | -120.21 | -120.21 | [6 empty] | 4
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0162 | -0.271 | 0 | TP | ug P/L | -0.00923 | -0.0958 | -31.4 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 2 | reworked calibration | -119.7796 | -119.7796 | -119.7796 | [6 empty] | 5
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00707 | -0.00498 | 0 | TP | ug P/L | 0.0522 | 0.698 | 13.8 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 2 | reworked calibration | -4.3866 | -4.3866 | -4.3866 | [6 empty] | 6
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.237 | 4.65 | 150 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 2 | reworked calibration | 57.1988 | 57.1988 | 57.1988 | -61.8675 | FAIL |  |  |  |  | 7
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.237 | 4.59 | 150 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 2 | reworked calibration | 57.1988 | 57.1988 | 57.1988 | -61.8675 | FAIL |  |  |  |  | 8
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.66 | 21.8 | 1000 | TP | ug P/L | 0.164 | 3.25 | 100 |  | 1 | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 2 | reworked calibration | 84.0872 | 84.0872 | 84.0872 | -15.9128 | PASS |  |  |  |  | 9
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.67 | 22.6 | 1000 | TP | ug P/L | 0.16 | 3.15 | 100 |  | 1 | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 2 | reworked calibration | 83.6278 | 83.6278 | 83.6278 | -16.3722 | PASS |  |  |  |  | 10
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.867 | 11.4 | 500 | TP | ug P/L | 0.0979 | 1.72 | 50 |  | 1 | 1 |  | 500 | 500 | 500 | [6 empty] | 2 | reworked calibration | 50.735 | 50.735 | 50.735 | 1.47 | PASS |  |  |  |  | 11
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.86 | 11.4 | 500 | TP | ug P/L | 0.0977 | 1.73 | 50 |  | 1 | 1 |  | 500 | 500 | 500 | [6 empty] | 2 | reworked calibration | 50.5509 | 50.5509 | 50.5509 | 1.1017 | PASS |  |  |  |  | 12
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.478 | 6.3 | 250 | TP | ug P/L | 0.0729 | 0.978 | 25 |  | 1 | 1 |  | 250 | 250 | 250 | [6 empty] | 2 | reworked calibration | 23.8284 | 23.8284 | 23.8284 | -4.6862 | PASS |  |  |  |  | 13
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.474 | 6.29 | 250 | TP | ug P/L | 0.073 | 0.925 | 25 |  | 1 | 1 |  | 250 | 250 | 250 | [6 empty] | 2 | reworked calibration | 23.9517 | 23.9517 | 23.9517 | -4.1932 | PASS |  |  |  |  | 14
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.43 | 100 | TP | ug P/L | 0.0629 | 0.498 | 10 |  | 1 | 1 |  | 100 | 100 | 100 | [6 empty] | 2 | reworked calibration | 10.8694 | 10.8694 | 10.8694 | 8.6943 | PASS |  |  |  |  | 15
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.44 | 100 | TP | ug P/L | 0.0682 | 0.488 | 10 |  | 1 | 1 |  | 100 | 100 | 100 | [6 empty] | 2 | reworked calibration | 17.894 | 17.894 | 17.894 | 78.9403 | FAIL |  |  |  |  | 16
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.176 | 2.33 | 50 | TP | ug P/L | 0.0612 | 0.443 | 5 |  | 1 | 1 |  | 50 | 50 | 50 | [6 empty] | 2 | reworked calibration | 8.5416 | 8.5416 | 8.5416 | 70.8317 | FAIL |  |  |  |  | 17
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.166 | 2.28 | 50 | TP | ug P/L | 0.0591 | 0.353 | 5 |  | 1 | 1 |  | 50 | 50 | 50 | [6 empty] | 2 | reworked calibration | 5.6159 | 5.6159 | 5.6159 | 12.3186 | PASS |  |  |  |  | 18
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.131 | 1.8 | 25 | TP | ug P/L | 0.059 | 0.855 | 2.5 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 2 | reworked calibration | 5.4752 | 5.4752 | 5.4752 | 119.0093 | FAIL |  |  |  |  | 19
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.129 | 1.76 | 25 | TP | ug P/L | 0.0563 | 0.799 | 2.5 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 2 | reworked calibration | 1.629 | 1.629 | 1.629 | -34.8417 | FAIL |  |  |  |  | 20
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0914 | 1.27 | 0 | TP | ug P/L | 0.0562 | 0.806 | 0 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 2 | reworked calibration | 1.4847 | 1.4847 | 1.4847 | [6 empty] | 21
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0915 | 1.27 | 0 | TP | ug P/L | 0.0567 | 0.799 | 0 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 2 | reworked calibration | 2.2045 | 2.2045 | 2.2045 | [6 empty] | 22
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.2 | 29.8 | 1400 | TP | ug P/L | -0.0119 | -0.111 | -92.5 | 99.6644 | 1 | 1 |  | 1400 | 1400 | 1400 | [6 empty] | 2 | reworked calibration | -125.869 | -125.869 | -125.869 | [6 empty] | 23
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.22 | 29.7 | 1390 | TP | ug P/L | 0.0192 | 0.22 | -47.7 |  | 1 | 1 |  | 1390 | 1390 | 1390 | [6 empty] | 2 | reworked calibration | -60.4878 | -60.4878 | -60.4878 | [6 empty] | 24
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0172 | -0.298 | -75.6 | TP | ug P/L | -0.00934 | -0.0803 | -88.6 |  | 1 | 1 |  | -75.6 | -75.6 | -75.6 | [6 empty] | 2 | reworked calibration | -120.0287 | -120.0287 | -120.0287 | [6 empty] | 25
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00164 | 0.0147 | -60.5 | TP | ug P/L | 0.0552 | 0.782 | -0.858 |  | 1 | 1 |  | -60.5 | -60.5 | -60.5 | [6 empty] | 2 | reworked calibration | 0.035725 | 0.035725 | 0.035725 | [6 empty] | 26
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.24 | 30.5 | 1440 | TP | ug P/L | -0.0196 | -0.17 | -49 | 97.377 | 1 | 1 |  | 1440 | 1440 | 1440 | [6 empty] | 2 | reworked calibration | -143.9315 | -143.9315 | -143.9315 | [6 empty] | 32
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.21 | 29.7 | 1400 | TP | ug P/L | -0.015 | -0.122 | -48.7 |  | 1 | 1 |  | 1400 | 1400 | 1400 | [6 empty] | 2 | reworked calibration | -133.0515 | -133.0515 | -133.0515 | [6 empty] | 33
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0165 | -0.274 | -74.4 | TP | ug P/L | -0.022 | -0.227 | -49.2 |  | 1 | 1 |  | -74.4 | -74.4 | -74.4 | [6 empty] | 2 | reworked calibration | -149.7135 | -149.7135 | -149.7135 | [6 empty] | 34
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.26 | 242 | TP | ug P/L | 0.071 | 0.863 | -43.1 |  | 1 | 1 |  | 242 | 242 | 242 | -3.2 | PASS |  |  |  |  | 2 | reworked calibration | 21.4628 | 21.4628 | 21.4628 | -14.1488 | PASS |  |  |  |  | 35
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.0025 | -0.0316 | -62.7 | TP | ug P/L | 0.0484 | 0.653 | -44.6 |  | 1 | 1 |  | -62.7 | -62.7 | -62.7 | [6 empty] | 2 | reworked calibration | -10.1505 | -10.1505 | -10.1505 | [6 empty] | 36
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0167 | -0.267 | -74.1 | TP | ug P/L | 0.0114 | 0.176 | -47 |  | 1 | 1 |  | -74.1 | -74.1 | -74.1 | [6 empty] | 2 | reworked calibration | -75.7452 | -75.7452 | -75.7452 | [6 empty] | 37
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00439 | -0.035 | -62.9 | TP | ug P/L | 0.042 | 0.5 | -45 |  | 1 | 1 |  | -62.9 | -62.9 | -62.9 | [6 empty] | 2 | reworked calibration | -20.2676 | -20.2676 | -20.2676 | [6 empty] | 38
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 20 | 1 | 1 | TN | ug N/L | 0.084 | 1.2 | -3.48 | TP | ug P/L | 0.0482 | 0.56 | -44.6 |  | 1 | 1 |  | -3.48 | -3.48 | -3.48 | [6 empty] | 2 | reworked calibration | -10.4588 | -10.4588 | -10.4588 | [6 empty] | 39
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 21 | 1 | 1 | TN | ug N/L | 0.0936 | 1.22 | -2.22 | TP | ug P/L | 0.0476 | 0.617 | -44.6 |  | 1 | 1 |  | -2.22 | -2.22 | -2.22 | [6 empty] | 2 | reworked calibration | -11.387 | -11.387 | -11.387 | [6 empty] | 40
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 22 | 1 | 1 | TN | ug N/L | 0.0896 | 1.28 | 0.652 | TP | ug P/L | 0.0538 | 0.764 | -44.2 |  | 1 | 1 |  | 0.652 | 0.652 | 0.652 | [6 empty] | 2 | reworked calibration | -2.014 | -2.014 | -2.014 | [6 empty] | 41
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 23 | 1 | 1 | TN | ug N/L | 0.095 | 1.3 | 1.41 | TP | ug P/L | 0.0539 | 0.721 | -44.2 |  | 1 | 1 |  | 1.41 | 1.41 | 1.41 | [6 empty] | 2 | reworked calibration | -1.8668 | -1.8668 | -1.8668 | [6 empty] | 42
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 24 | 1 | 1 | TN | ug N/L | 0.0925 | 1.31 | 1.89 | TP | ug P/L | 0.0537 | 0.741 | -44.3 |  | 1 | 1 |  | 1.89 | 1.89 | 1.89 | [6 empty] | 2 | reworked calibration | -2.1613 | -2.1613 | -2.1613 | [6 empty] | 43
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 25 | 1 | 1 | TN | ug N/L | 0.71 | 9.48 | 398 | TP | ug P/L | 0.0817 | 1.27 | -42.4 |  | 1 | 1 |  | 398 | 398 | 398 | [6 empty] | 2 | reworked calibration | 34.1942 | 34.1942 | 34.1942 | [6 empty] | 44
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 26 | 1 | 1 | TN | ug N/L | 0.707 | 9.5 | 399 | TP | ug P/L | 0.0795 | 1.25 | -42.6 |  | 1 | 1 |  | 399 | 399 | 399 | [6 empty] | 2 | reworked calibration | 31.6939 | 31.6939 | 31.6939 | [6 empty] | 45
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 27 | 1 | 1 | TN | ug N/L | 0.709 | 9.47 | 398 | TP | ug P/L | 0.0827 | 1.27 | -42.4 |  | 1 | 1 |  | 398 | 398 | 398 | [6 empty] | 2 | reworked calibration | 35.3106 | 35.3106 | 35.3106 | [6 empty] | 46
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 28 | 1 | 1 | TN | ug N/L | 0.711 | 9.46 | 397 | TP | ug P/L | 0.0812 | 1.28 | -42.5 |  | 1 | 1 |  | 397 | 397 | 397 | [6 empty] | 2 | reworked calibration | 33.6312 | 33.6312 | 33.6312 | [6 empty] | 47
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 29 | 1 | 1 | TN | ug N/L | 0.709 | 9.42 | 395 | TP | ug P/L | 0.0823 | 1.19 | -42.4 |  | 1 | 1 |  | 395 | 395 | 395 | [6 empty] | 2 | reworked calibration | 34.8655 | 34.8655 | 34.8655 | [6 empty] | 48
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 30 | 1 | 1 | TN | ug N/L | 0.433 | 5.75 | 217 | TP | ug P/L | 0.0724 | 1.02 | -43 |  | 1 | 1 |  | 217 | 217 | 217 | [6 empty] | 2 | reworked calibration | 23.2103 | 23.2103 | 23.2103 | [6 empty] | 49
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | skp blank | 31 | 1 | 1 | TN | ug N/L | 0.436 | 5.89 | 224 | TP | ug P/L | 0.0746 | 1.06 | -42.9 |  | 1 | 1 |  | 224 | 224 | 224 | [6 empty] | 2 | reworked calibration | 25.9067 | 25.9067 | 25.9067 | [6 empty] | 50
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 32 | 1 | 1 | TN | ug N/L | 0.441 | 5.9 | 224 | TP | ug P/L | 0.0761 | 0.935 | -42.8 |  | 1 | 1 |  | 224 | 224 | 224 | [6 empty] | 2 | reworked calibration | 27.7103 | 27.7103 | 27.7103 | [6 empty] | 51
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 33 | 1 | 1 | TN | ug N/L | 0.437 | 5.85 | 222 | TP | ug P/L | 0.07 | 0.875 | -43.2 |  | 1 | 1 |  | 222 | 222 | 222 | [6 empty] | 2 | reworked calibration | 20.1995 | 20.1995 | 20.1995 | [6 empty] | 52
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 34 | 1 | 1 | TN | ug N/L | 0.446 | 5.92 | 225 | TP | ug P/L | 0.0694 | 0.862 | -43.2 |  | 1 | 1 |  | 225 | 225 | 225 | [6 empty] | 2 | reworked calibration | 19.4356 | 19.4356 | 19.4356 | [6 empty] | 53
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s1 | 35 | 1 | 1 | TN | ug N/L | 1.3 | 17.5 | 792 | TP | ug P/L | 0.0796 | 1.23 | -42.6 |  | 1 | 1 |  | 792 | 792 | 792 | [6 empty] | 2 | reworked calibration | 31.8088 | 31.8088 | 31.8088 | [6 empty] | 54
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s2 | 36 | 1 | 1 | TN | ug N/L | 0.538 | 7.16 | 285 | TP | ug P/L | 0.0657 | 0.687 | -43.5 |  | 1 | 1 |  | 285 | 285 | 285 | [6 empty] | 2 | reworked calibration | 14.6245 | 14.6245 | 14.6245 | [6 empty] | 55
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s3 | 37 | 1 | 1 | TN | ug N/L | 0.56 | 7.54 | 304 | TP | ug P/L | 0.0706 | 0.851 | -43.2 |  | 1 | 1 |  | 304 | 304 | 304 | [6 empty] | 2 | reworked calibration | 20.959 | 20.959 | 20.959 | [6 empty] | 56
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s4 | 38 | 1 | 1 | TN | ug N/L | 0.958 | 12.8 | 560 | TP | ug P/L | 0.0735 | 0.933 | -43 |  | 1 | 1 |  | 560 | 560 | 560 | [6 empty] | 2 | reworked calibration | 24.5661 | 24.5661 | 24.5661 | [6 empty] | 57
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s5 | 39 | 1 | 1 | TN | ug N/L | 3.36 | 45.1 | 2170 | TP | ug P/L | 0.0871 | 1.4 | -42.1 |  | 1 | 1 |  | 2170 | 2170 | 2170 | [6 empty] | 2 | reworked calibration | 40.0738 | 40.0738 | 40.0738 | [6 empty] | 58
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s6 | 40 | 1 | 1 | TN | ug N/L | 0.931 | 12.5 | 548 | TP | ug P/L | 0.0798 | 1.12 | -42.6 |  | 1 | 1 |  | 548 | 548 | 548 | [6 empty] | 2 | reworked calibration | 32.0384 | 32.0384 | 32.0384 | [6 empty] | 59
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s7 | 41 | 1 | 1 | TN | ug N/L | 0.886 | 11.9 | 515 | TP | ug P/L | 0.0833 | 1.24 | -42.3 |  | 1 | 1 |  | 515 | 515 | 515 | [6 empty] | 2 | reworked calibration | 35.9744 | 35.9744 | 35.9744 | [6 empty] | 60
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s8 | 42 | 1 | 1 | TN | ug N/L | 0.573 | 7.62 | 308 | TP | ug P/L | 0.0679 | 0.744 | -43.3 |  | 1 | 1 |  | 308 | 308 | 308 | [6 empty] | 2 | reworked calibration | 17.5058 | 17.5058 | 17.5058 | [6 empty] | 61
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s9 | 43 | 1 | 1 | TN | ug N/L | 1.6 | 21.5 | 988 | TP | ug P/L | 0.072 | 0.806 | -43.1 |  | 1 | 1 |  | 988 | 988 | 988 | [6 empty] | 2 | reworked calibration | 22.7135 | 22.7135 | 22.7135 | [6 empty] | 62
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | a10 | 44 | 1 | 1 | TN | ug N/L | 0.764 | 10.2 | 435 | TP | ug P/L | 0.0738 | 0.902 | -43 |  | 1 | 1 |  | 435 | 435 | 435 | [6 empty] | 2 | reworked calibration | 24.9332 | 24.9332 | 24.9332 | [6 empty] | 63
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup5 | 45 | 1 | 1 | TN | ug N/L | 3.39 | 45.6 | 2190 | TP | ug P/L | 0.0882 | 1.48 | -42 |  | 1 | 1 |  | 2190 | 2190 | 2190 |  |  | 0.91743 | PASS |  |  | 2 | reworked calibration | 41.2266 | 41.2266 | 41.2266 |  |  | 2.836 | PASS |  |  | 64
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk10 | 46 | 1 | 1 | TN | ug N/L | 1.07 | 14.4 | 639 | TP | ug P/L | 0.0958 | 1.64 | -41.5 |  | 1 | 1 |  | 639 | 639 | 639 |  |  |  |  | 87.99 | PASS | 2 | reworked calibration | 48.7766 | 48.7766 | 48.7766 |  |  |  |  | 100.2514 | PASS | 65
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.28 | 243 | TP | ug P/L | 0.071 | 0.861 | -43.1 |  | 1 | 1 |  | 243 | 243 | 243 | -2.8 | PASS |  |  |  |  | 2 | reworked calibration | 21.4628 | 21.4628 | 21.4628 | -14.1488 | PASS |  |  |  |  | 66
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00254 | 0.00654 | -60.9 | TP | ug P/L | 0.0496 | 0.646 | -44.5 |  | 1 | 1 |  | -60.9 | -60.9 | -60.9 | [6 empty] | 2 | reworked calibration | -8.3107 | -8.3107 | -8.3107 | [6 empty] | 67
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.016 | -0.252 | -73.3 | TP | ug P/L | -0.0104 | -0.111 | -48.4 |  | 1 | 1 |  | -73.3 | -73.3 | -73.3 | [6 empty] | 2 | reworked calibration | -122.437 | -122.437 | -122.437 | [6 empty] | 68
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s11 | 47 | 1 | 1 | TN | ug N/L | 1.62 | 21.9 | 1010 | TP | ug P/L | 0.0728 | 0.996 | -43 |  | 1 | 1 |  | 1010 | 1010 | 1010 | [6 empty] | 2 | reworked calibration | 23.7051 | 23.7051 | 23.7051 | [6 empty] | 69
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s12 | 48 | 1 | 1 | TN | ug N/L | 0.843 | 11.3 | 485 | TP | ug P/L | 0.0743 | 0.976 | -42.9 |  | 1 | 1 |  | 485 | 485 | 485 | [6 empty] | 2 | reworked calibration | 25.5425 | 25.5425 | 25.5425 | [6 empty] | 70
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s13 | 49 | 1 | 1 | TN | ug N/L | 1.6 | 21.5 | 990 | TP | ug P/L | 0.0699 | 0.814 | -43.2 |  | 1 | 1 |  | 990 | 990 | 990 | [6 empty] | 2 | reworked calibration | 20.0725 | 20.0725 | 20.0725 | [6 empty] | 71
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s14 | 50 | 1 | 1 | TN | ug N/L | 1.67 | 22.6 | 1040 | TP | ug P/L | 0.0869 | 1.34 | -42.1 |  | 1 | 1 |  | 1040 | 1040 | 1040 | [6 empty] | 2 | reworked calibration | 39.8625 | 39.8625 | 39.8625 | [6 empty] | 72
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s15 | 51 | 1 | 1 | TN | ug N/L | 0.54 | 7.23 | 289 | TP | ug P/L | 0.0693 | 0.803 | -43.2 |  | 1 | 1 |  | 289 | 289 | 289 | [6 empty] | 2 | reworked calibration | 19.3078 | 19.3078 | 19.3078 | [6 empty] | 73
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s16 | 52 | 1 | 1 | TN | ug N/L | 0.594 | 8.01 | 327 | TP | ug P/L | 0.0666 | 0.673 | -43.4 |  | 1 | 1 |  | 327 | 327 | 327 | [6 empty] | 2 | reworked calibration | 15.8105 | 15.8105 | 15.8105 | [6 empty] | 74
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s17 | 53 | 1 | 1 | TN | ug N/L | 0.798 | 10.7 | 457 | TP | ug P/L | 0.0704 | 0.825 | -43.2 |  | 1 | 1 |  | 457 | 457 | 457 | [6 empty] | 2 | reworked calibration | 20.7064 | 20.7064 | 20.7064 | [6 empty] | 75
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s18 | 54 | 1 | 1 | TN | ug N/L | 3.33 | 44.7 | 2150 | TP | ug P/L | 0.0853 | 1.25 | -42.2 |  | 1 | 1 |  | 2150 | 2150 | 2150 | [6 empty] | 2 | reworked calibration | 38.1546 | 38.1546 | 38.1546 | [6 empty] | 76
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s19 | 55 | 1 | 1 | TN | ug N/L | 2.53 | 34 | 1610 | TP | ug P/L | 0.0824 | 1.2 | -42.4 |  | 1 | 1 |  | 1610 | 1610 | 1610 | [6 empty] | 2 | reworked calibration | 34.977 | 34.977 | 34.977 | [6 empty] | 77
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s20 | 56 | 1 | 1 | TN | ug N/L | 2.49 | 33.3 | 1580 | TP | ug P/L | 0.0788 | 1.13 | -42.6 |  | 1 | 1 |  | 1580 | 1580 | 1580 | [6 empty] | 2 | reworked calibration | 30.8856 | 30.8856 | 30.8856 | [6 empty] | 78
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup15 | 57 | 1 | 1 | TN | ug N/L | 0.54 | 7.19 | 287 | TP | ug P/L | 0.0682 | 0.797 | -43.3 |  | 1 | 1 |  | 287 | 287 | 287 |  |  | 0.69444 | PASS |  |  | 2 | reworked calibration | 17.894 | 17.894 | 17.894 |  |  | 7.6005 | PASS |  |  | 79
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk20 | 58 | 1 | 1 | TN | ug N/L | 2.81 | 37.7 | 1800 | TP | ug P/L | 0.105 | 1.9 | -41 |  | 1 | 1 |  | 1800 | 1800 | 1800 |  |  |  |  | 106 | PASS | 2 | reworked calibration | 56.9462 | 56.9462 | 56.9462 |  |  |  |  | 109.9372 | PASS | 80
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.35 | 246 | TP | ug P/L | 0.0721 | 0.877 | -43.1 |  | 1 | 1 |  | 246 | 246 | 246 | -1.6 | PASS |  |  |  |  | 2 | reworked calibration | 22.8379 | 22.8379 | 22.8379 | -8.6484 | PASS |  |  |  |  | 81
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00264 | -0.0321 | -62.8 | TP | ug P/L | 0.0491 | 0.679 | -44.5 |  | 1 | 1 |  | -62.8 | -62.8 | -62.8 | [6 empty] | 2 | reworked calibration | -9.0751 | -9.0751 | -9.0751 | [6 empty] | 82
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00701 | 0.037 | -59.4 | TP | ug P/L | 0.00901 | 0.151 | -47.1 |  | 1 | 1 |  | -59.4 | -59.4 | -59.4 | [6 empty] | 2 | reworked calibration | -80.5731 | -80.5731 | -80.5731 | [6 empty] | 83
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s21 | 59 | 1 | 1 | TN | ug N/L | 0.554 | 7.45 | 299 | TP | ug P/L | 0.0688 | 0.687 | -43.3 |  | 1 | 1 |  | 299 | 299 | 299 | [6 empty] | 2 | reworked calibration | 18.6671 | 18.6671 | 18.6671 | [6 empty] | 84
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s22 | 60 | 1 | 1 | TN | ug N/L | 0.736 | 9.94 | 421 | TP | ug P/L | 0.0671 | 0.742 | -43.4 |  | 1 | 1 |  | 421 | 421 | 421 | [6 empty] | 2 | reworked calibration | 16.4651 | 16.4651 | 16.4651 | [6 empty] | 85
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s23 | 61 | 1 | 1 | TN | ug N/L | 0.615 | 8.28 | 340 | TP | ug P/L | 0.0712 | 0.88 | -43.1 |  | 1 | 1 |  | 340 | 340 | 340 | [6 empty] | 2 | reworked calibration | 21.714 | 21.714 | 21.714 | [6 empty] | 86
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s24 | 62 | 1 | 1 | TN | ug N/L | 1.04 | 14 | 622 | TP | ug P/L | 0.0657 | 0.596 | -43.5 |  | 1 | 1 |  | 622 | 622 | 622 | [6 empty] | 2 | reworked calibration | 14.6245 | 14.6245 | 14.6245 | [6 empty] | 87
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s25 | 63 | 1 | 1 | TN | ug N/L | 3.03 | 40.8 | 1950 | TP | ug P/L | 0.1 | 1.72 | -41.2 |  | 1 | 1 |  | 1950 | 1950 | 1950 | [6 empty] | 2 | reworked calibration | 52.638 | 52.638 | 52.638 | [6 empty] | 88
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s26 | 64 | 1 | 1 | TN | ug N/L | 1.17 | 15.8 | 708 | TP | ug P/L | 0.0748 | 0.951 | -42.9 |  | 1 | 1 |  | 708 | 708 | 708 | [6 empty] | 2 | reworked calibration | 26.1488 | 26.1488 | 26.1488 | [6 empty] | 89
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s27 | 65 | 1 | 1 | TN | ug N/L | 0.5 | 6.71 | 264 | TP | ug P/L | 0.0683 | 0.693 | -43.3 |  | 1 | 1 |  | 264 | 264 | 264 | [6 empty] | 2 | reworked calibration | 18.0232 | 18.0232 | 18.0232 | [6 empty] | 90
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s28 | 66 | 1 | 1 | TN | ug N/L | 2.56 | 34.5 | 1630 | TP | ug P/L | 0.144 | 2.72 | -38.4 |  | 1 | 1 |  | 1630 | 1630 | 1630 | [6 empty] | 2 | reworked calibration | 79.7817 | 79.7817 | 79.7817 | [6 empty] | 91
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s29 | 67 | 1 | 1 | TN | ug N/L | 2.18 | 29.3 | 1370 | TP | ug P/L | 0.0803 | 1.17 | -42.5 |  | 1 | 1 |  | 1370 | 1370 | 1370 | [6 empty] | 2 | reworked calibration | 32.6101 | 32.6101 | 32.6101 | [6 empty] | 92
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s30 | 68 | 1 | 1 | TN | ug N/L | 1.14 | 15.3 | 682 | TP | ug P/L | 0.0756 | 1.01 | -42.8 |  | 1 | 1 |  | 682 | 682 | 682 | [6 empty] | 2 | reworked calibration | 27.1122 | 27.1122 | 27.1122 | [6 empty] | 93
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup25 | 69 | 1 | 1 | TN | ug N/L | 3 | 40.5 | 1930 | TP | ug P/L | 0.0993 | 1.71 | -41.3 |  | 1 | 1 |  | 1930 | 1930 | 1930 |  |  | 1.0309 | PASS |  |  | 2 | reworked calibration | 52.0098 | 52.0098 | 52.0098 |  |  | 1.2006 | PASS |  |  | 94
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup30 | 70 | 1 | 1 | TN | ug N/L | 1.46 | 19.7 | 900 | TP | ug P/L | 0.098 | 1.72 | -41.4 |  | 1 | 1 |  | 900 | 900 | 900 |  |  |  |  | 96.2 | PASS | 2 | reworked calibration | 50.8269 | 50.8269 | 50.8269 |  |  |  |  | 99.9414 | PASS | 95
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.469 | 6.28 | 243 | TP | ug P/L | 0.0715 | 0.884 | -43.1 |  | 1 | 1 |  | 243 | 243 | 243 | -2.8 | PASS |  |  |  |  | 2 | reworked calibration | 22.0897 | 22.0897 | 22.0897 | -11.6411 | PASS |  |  |  |  | 96
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00521 | -0.0319 | -62.7 | TP | ug P/L | 0.0511 | 0.69 | -44.4 |  | 1 | 1 |  | -62.7 | -62.7 | -62.7 | [6 empty] | 2 | reworked calibration | -6.0365 | -6.0365 | -6.0365 | [6 empty] | 97
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0169 | -0.275 | -74.4 | TP | ug P/L | 0.0095 | 0.164 | -47.1 |  | 1 | 1 |  | -74.4 | -74.4 | -74.4 | [6 empty] | 2 | reworked calibration | -79.5774 | -79.5774 | -79.5774 | [6 empty] | 98
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s31 | 71 | 1 | 1 | TN | ug N/L | 0.491 | 6.8 | 268 | TP | ug P/L | 0.0653 | 0.68 | -43.5 |  | 1 | 1 |  | 268 | 268 | 268 | [6 empty] | 2 | reworked calibration | 14.0941 | 14.0941 | 14.0941 | [6 empty] | 99
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s32 | 72 | 1 | 1 | TN | ug N/L | 1.13 | 15.2 | 679 | TP | ug P/L | 0.0693 | 0.734 | -43.2 |  | 1 | 1 |  | 679 | 679 | 679 | [6 empty] | 2 | reworked calibration | 19.3078 | 19.3078 | 19.3078 | [6 empty] | 100
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s33 | 73 | 1 | 1 | TN | ug N/L | 0.853 | 11.5 | 495 | TP | ug P/L | 0.0791 | 1.08 | -42.6 |  | 1 | 1 |  | 495 | 495 | 495 | [6 empty] | 2 | reworked calibration | 31.2327 | 31.2327 | 31.2327 | [6 empty] | 101
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s34 | 74 | 1 | 1 | TN | ug N/L | 1.37 | 18.4 | 837 | TP | ug P/L | 0.0827 | 1.21 | -42.4 |  | 1 | 1 |  | 837 | 837 | 837 | [6 empty] | 2 | reworked calibration | 35.3106 | 35.3106 | 35.3106 | [6 empty] | 102
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s35 | 75 | 1 | 1 | TN | ug N/L | 0.745 | 10 | 425 | TP | ug P/L | 0.0705 | 0.828 | -43.2 |  | 1 | 1 |  | 425 | 425 | 425 | [6 empty] | 2 | reworked calibration | 20.8327 | 20.8327 | 20.8327 | [6 empty] | 103
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s36 | 76 | 1 | 1 | TN | ug N/L | 1.64 | 22.1 | 1020 | TP | ug P/L | 0.1 | 1.75 | -41.3 |  | 1 | 1 |  | 1020 | 1020 | 1020 | [6 empty] | 2 | reworked calibration | 52.638 | 52.638 | 52.638 | [6 empty] | 104
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s37 | 77 | 1 | 1 | TN | ug N/L | 2.57 | 34.9 | 1650 | TP | ug P/L | 0.0822 | 1.26 | -42.4 |  | 1 | 1 |  | 1650 | 1650 | 1650 | [6 empty] | 2 | reworked calibration | 34.7539 | 34.7539 | 34.7539 | [6 empty] | 105
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s38 | 78 | 1 | 1 | TN | ug N/L | 0.752 | 10.1 | 428 | TP | ug P/L | 0.0664 | 0.619 | -43.4 |  | 1 | 1 |  | 428 | 428 | 428 | [6 empty] | 2 | reworked calibration | 15.5479 | 15.5479 | 15.5479 | [6 empty] | 106
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s39 | 79 | 1 | 1 | TN | ug N/L | 1.17 | 15.8 | 707 | TP | ug P/L | 0.0662 | 0.682 | -43.4 |  | 1 | 1 |  | 707 | 707 | 707 | [6 empty] | 2 | reworked calibration | 15.2847 | 15.2847 | 15.2847 | [6 empty] | 107
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s40 | 80 | 1 | 1 | TN | ug N/L | 1.76 | 23.7 | 1100 | TP | ug P/L | 0.0725 | 0.916 | -43 |  | 1 | 1 |  | 1100 | 1100 | 1100 | [6 empty] | 2 | reworked calibration | 23.3342 | 23.3342 | 23.3342 | [6 empty] | 108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup35 | 81 | 1 | 1 | TN | ug N/L | 0.746 | 10.1 | 430 | TP | ug P/L | 0.0691 | 0.712 | -43.3 |  | 1 | 1 |  | 430 | 430 | 430 |  |  | 1.1696 | PASS |  |  | 2 | reworked calibration | 19.0519 | 19.0519 | 19.0519 |  |  | 8.9301 | PASS |  |  | 109
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk40 | 82 | 1 | 1 | TN | ug N/L | 2.09 | 28.1 | 1310 | TP | ug P/L | 0.1 | 1.74 | -41.3 |  | 1 | 1 |  | 1310 | 1310 | 1310 |  |  |  |  | 97.1 | PASS | 2 | reworked calibration | 52.638 | 52.638 | 52.638 |  |  |  |  | 122.4791 | FAIL | 110
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.3 | 244 | TP | ug P/L | 0.0725 | 0.92 | -43 |  | 1 | 1 |  | 244 | 244 | 244 | -2.4 | PASS |  |  |  |  | 2 | reworked calibration | 23.3342 | 23.3342 | 23.3342 | -6.6633 | PASS |  |  |  |  | 111
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00528 | -0.0552 | -63.9 | TP | ug P/L | 0.0486 | 0.652 | -44.6 |  | 1 | 1 |  | -63.9 | -63.9 | -63.9 | [6 empty] | 2 | reworked calibration | -9.8426 | -9.8426 | -9.8426 | [6 empty] | 112
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0144 | -0.226 | -72.1 | TP | ug P/L | 0.00937 | 0.15 | -47.1 |  | 1 | 1 |  | -72.1 | -72.1 | -72.1 | [6 empty] | 2 | reworked calibration | -79.8413 | -79.8413 | -79.8413 | [6 empty] | 113
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s41 | 83 | 1 | 1 | TN | ug N/L | 1.16 | 15.7 | 703 | TP | ug P/L | 0.0666 | 0.719 | -43.4 |  | 1 | 1 |  | 703 | 703 | 703 | [6 empty] | 2 | reworked calibration | 15.8105 | 15.8105 | 15.8105 | [6 empty] | 114
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s42 | 84 | 1 | 1 | TN | ug N/L | 0.5 | 6.77 | 266 | TP | ug P/L | 0.067 | 0.674 | -43.4 |  | 1 | 1 |  | 266 | 266 | 266 | [6 empty] | 2 | reworked calibration | 16.3344 | 16.3344 | 16.3344 | [6 empty] | 115
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s43 | 85 | 1 | 1 | TN | ug N/L | 1.02 | 13.8 | 610 | TP | ug P/L | 0.0791 | 0.986 | -42.6 |  | 1 | 1 |  | 610 | 610 | 610 | [6 empty] | 2 | reworked calibration | 31.2327 | 31.2327 | 31.2327 | [6 empty] | 116
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s44 | 86 | 1 | 1 | TN | ug N/L | 2.55 | 34.5 | 1640 | TP | ug P/L | 0.0796 | 1.08 | -42.6 |  | 1 | 1 |  | 1640 | 1640 | 1640 | [6 empty] | 2 | reworked calibration | 31.8088 | 31.8088 | 31.8088 | [6 empty] | 117
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s45 | 87 | 1 | 1 | TN | ug N/L | 1.46 | 19.7 | 901 | TP | ug P/L | 0.0698 | 0.876 | -43.2 |  | 1 | 1 |  | 901 | 901 | 901 | [6 empty] | 2 | reworked calibration | 19.9454 | 19.9454 | 19.9454 | [6 empty] | 118
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s46 | 88 | 1 | 1 | TN | ug N/L | 2.42 | 32.6 | 1540 | TP | ug P/L | 0.139 | 2.74 | -38.8 |  | 1 | 1 |  | 1540 | 1540 | 1540 | [6 empty] | 2 | reworked calibration | 77.9209 | 77.9209 | 77.9209 | [6 empty] | 119
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s47 | 89 | 1 | 1 | TN | ug N/L | 0.713 | 9.56 | 402 | TP | ug P/L | 0.0683 | 0.766 | -43.3 |  | 1 | 1 |  | 402 | 402 | 402 | [6 empty] | 2 | reworked calibration | 18.0232 | 18.0232 | 18.0232 | [6 empty] | 120
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s48 | 90 | 1 | 1 | TN | ug N/L | 0.975 | 13.2 | 581 | TP | ug P/L | 0.0786 | 1.12 | -42.6 |  | 1 | 1 |  | 581 | 581 | 581 | [6 empty] | 2 | reworked calibration | 30.6535 | 30.6535 | 30.6535 | [6 empty] | 121
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s49 | 91 | 1 | 1 | TN | ug N/L | 1.1 | 15 | 666 | TP | ug P/L | 0.0713 | 0.807 | -43.1 |  | 1 | 1 |  | 666 | 666 | 666 | [6 empty] | 2 | reworked calibration | 21.8393 | 21.8393 | 21.8393 | [6 empty] | 122
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s50 | 92 | 1 | 1 | TN | ug N/L | 0.825 | 11.1 | 478 | TP | ug P/L | 0.0713 | 0.984 | -43.1 |  | 1 | 1 |  | 478 | 478 | 478 | [6 empty] | 2 | reworked calibration | 21.8393 | 21.8393 | 21.8393 | [6 empty] | 123
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup45 | 93 | 1 | 1 | TN | ug N/L | 1.45 | 19.7 | 899 | TP | ug P/L | 0.0702 | 0.778 | -43.2 |  | 1 | 1 |  | 899 | 899 | 899 |  |  | 0.22222 | PASS |  |  | 2 | reworked calibration | 20.4532 | 20.4532 | 20.4532 |  |  | 2.514 | PASS |  |  | 124
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk50 | 94 | 1 | 1 | TN | ug N/L | 1.2 | 16.2 | 726 | TP | ug P/L | 0.101 | 1.85 | -41.2 |  | 1 | 1 |  | 726 | 726 | 726 |  |  |  |  | 106.46 | PASS | 2 | reworked calibration | 53.5247 | 53.5247 | 53.5247 |  |  |  |  | 132.0941 | FAIL | 125
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.471 | 6.35 | 246 | TP | ug P/L | 0.0719 | 0.911 | -43.1 |  | 1 | 1 |  | 246 | 246 | 246 | -1.6 | PASS |  |  |  |  | 2 | reworked calibration | 22.589 | 22.589 | 22.589 | -9.6439 | PASS |  |  |  |  | 126
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00351 | -0.0324 | -62.8 | TP | ug P/L | 0.0546 | 0.756 | -44.2 |  | 1 | 1 |  | -62.8 | -62.8 | -62.8 | [6 empty] | 2 | reworked calibration | -0.83971 | -0.83971 | -0.83971 | [6 empty] | 127
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.016 | -0.27 | -74.2 | TP | ug P/L | 0.00692 | 0.093 | -47.3 |  | 1 | 1 |  | -74.2 | -74.2 | -74.2 | [6 empty] | 2 | reworked calibration | -84.8537 | -84.8537 | -84.8537 | [6 empty] | 128
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal1500 | 1 | 1 | 1 | TN | ug N/L | 2.37 | 31.9 | 1500 | TP | ug P/L | 0.258 | 4.53 | -31.3 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 2 | reworked calibration | 37.0761 | 37.0761 | 37.0761 | -75.2826 | FAIL |  |  |  |  | 129
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal1000 | 3 | 1 | 1 | TN | ug N/L | 1.62 | 21.9 | 1010 | TP | ug P/L | 0.157 | 3.07 | -37.6 |  | 1 | 1 |  | 1010 | 1010 | 1010 | [6 empty] | 2 | reworked calibration | 83.1514 | 83.1514 | 83.1514 | -16.8486 | PASS |  |  |  |  | 130
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call500 | 5 | 1 | 1 | TN | ug N/L | 0.848 | 11.5 | 495 | TP | ug P/L | 0.0965 | 1.81 | -41.5 |  | 1 | 1 |  | 495 | 495 | 495 | [6 empty] | 2 | reworked calibration | 49.4356 | 49.4356 | 49.4356 | -1.1289 | PASS |  |  |  |  | 131
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call250 | 7 | 1 | 1 | TN | ug N/L | 0.471 | 6.35 | 246 | TP | ug P/L | 0.0697 | 0.881 | -43.2 |  | 1 | 1 |  | 246 | 246 | 246 | [6 empty] | 2 | reworked calibration | 19.8181 | 19.8181 | 19.8181 | -20.7276 | FAIL |  |  |  |  | 132
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal100 | 9 | 1 | 1 | TN | ug N/L | 0.254 | 3.45 | 105 | TP | ug P/L | 0.0595 | 0.468 | -43.9 |  | 1 | 1 |  | 105 | 105 | 105 | [6 empty] | 2 | reworked calibration | 6.1775 | 6.1775 | 6.1775 | -38.2253 | FAIL |  |  |  |  | 133
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.172 | 2.35 | 52.2 | TP | ug P/L | 0.0563 | 0.351 | -44.1 |  | 1 | 1 |  | 52.2 | 52.2 | 52.2 | [6 empty] | 2 | reworked calibration | 1.629 | 1.629 | 1.629 | -67.4209 | FAIL |  |  |  |  | 134
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.135 | 1.81 | 26.3 | TP | ug P/L | 0.0576 | 0.817 | -44 |  | 1 | 1 |  | 26.3 | 26.3 | 26.3 | [6 empty] | 2 | reworked calibration | 3.4923 | 3.4923 | 3.4923 | 39.6916 | FAIL |  |  |  |  | 135
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0958 | 1.29 | 1.07 | TP | ug P/L | 0.056 | 0.797 | -44.1 |  | 1 | 1 |  | 1.07 | 1.07 | 1.07 | [6 empty] | 2 | reworked calibration | 1.1959 | 1.1959 | 1.1959 | [6 empty] | 136
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrite | S1 | 1 | 1 | TN | ug N/L | 2.19 | 30.1 | 1420 | TP | ug P/L | -0.0089 | -0.0746 | -48.3 | 99.6678 | 1 | 1 |  | 1420 | 1420 | 1420 | [6 empty] | 2 | reworked calibration | -119.0332 | -119.0332 | -119.0332 | [6 empty] | 137
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrate | S2 | 1 | 1 | TN | ug N/L | 2.21 | 30 | 1410 | TP | ug P/L | -0.0099 | -0.0977 | -48.4 |  | 1 | 1 |  | 1410 | 1410 | 1410 | [6 empty] | 2 | reworked calibration | -121.2992 | -121.2992 | -121.2992 | [6 empty] | 138
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s51 | 95 | 1 | 1 | TN | ug N/L | 0.609 | 8.22 | 337 | TP | ug P/L | 0.0687 | 0.784 | -43.3 |  | 1 | 1 |  | 337 | 337 | 337 | [6 empty] | 2 | reworked calibration | 18.5385 | 18.5385 | 18.5385 | [6 empty] | 139
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s52 | 96 | 1 | 1 | TN | ug N/L | 0.837 | 11.3 | 486 | TP | ug P/L | 0.0662 | 0.673 | -43.4 |  | 1 | 1 |  | 486 | 486 | 486 | [6 empty] | 2 | reworked calibration | 15.2847 | 15.2847 | 15.2847 | [6 empty] | 140
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s53 | 97 | 1 | 1 | TN | ug N/L | 4.93 | 67.1 | 3290 | TP | ug P/L | 0.0989 | 1.77 | -41.3 |  | 1 | 1 |  | 3290 | 3290 | 3290 | [6 empty] | 2 | reworked calibration | 51.6481 | 51.6481 | 51.6481 | [6 empty] | 141
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s54 | 98 | 1 | 1 | TN | ug N/L | 0.626 | 8.5 | 351 | TP | ug P/L | 0.0718 | 0.86 | -43.1 |  | 1 | 1 |  | 351 | 351 | 351 | [6 empty] | 2 | reworked calibration | 22.4644 | 22.4644 | 22.4644 | [6 empty] | 142
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s55 | 99 | 1 | 1 | TN | ug N/L | 0.728 | 9.89 | 418 | TP | ug P/L | 0.0711 | 0.805 | -43.1 |  | 1 | 1 |  | 418 | 418 | 418 | [6 empty] | 2 | reworked calibration | 21.5884 | 21.5884 | 21.5884 | [6 empty] | 143
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s56 | 100 | 1 | 1 | TN | ug N/L | 1.13 | 15.4 | 686 | TP | ug P/L | 0.0679 | 0.731 | -43.3 |  | 1 | 1 |  | 686 | 686 | 686 | [6 empty] | 2 | reworked calibration | 17.5058 | 17.5058 | 17.5058 | [6 empty] | 144
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s57 | 101 | 1 | 1 | TN | ug N/L | 2.45 | 33.1 | 1560 | TP | ug P/L | 0.0812 | 1.23 | -42.5 |  | 1 | 1 |  | 1560 | 1560 | 1560 | [6 empty] | 2 | reworked calibration | 33.6312 | 33.6312 | 33.6312 | [6 empty] | 145
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s58 | 102 | 1 | 1 | TN | ug N/L | 2.77 | 37.4 | 1780 | TP | ug P/L | 0.0885 | 1.41 | -42 |  | 1 | 1 |  | 1780 | 1780 | 1780 | [6 empty] | 2 | reworked calibration | 41.5383 | 41.5383 | 41.5383 | [6 empty] | 146
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s59 | 103 | 1 | 1 | TN | ug N/L | 1.08 | 14.7 | 654 | TP | ug P/L | 0.0682 | 0.79 | -43.3 |  | 1 | 1 |  | 654 | 654 | 654 | [6 empty] | 2 | reworked calibration | 17.894 | 17.894 | 17.894 | [6 empty] | 147
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s61 | 104 | 1 | 1 | TN | ug N/L | 3.12 | 42.4 | 2030 | TP | ug P/L | 0.0997 | 1.75 | -41.3 |  | 1 | 1 |  | 2030 | 2030 | 2030 | [6 empty] | 2 | reworked calibration | 52.3695 | 52.3695 | 52.3695 | [6 empty] | 148
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup55 | 105 | 1 | 1 | TN | ug N/L | 0.729 | 9.84 | 416 | TP | ug P/L | 0.0701 | 0.797 | -43.2 |  | 1 | 1 |  | 416 | 416 | 416 |  |  | 0.47962 | PASS |  |  | 2 | reworked calibration | 20.3264 | 20.3264 | 20.3264 |  |  | 6.0218 | PASS |  |  | 149
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk60 | 106 | 1 | 1 | TN | ug N/L | 3.47 | 47.2 | 2280 | TP | ug P/L | 0.136 | 2.71 | -39 |  | 1 | 1 |  | 2280 | 2280 | 2280 |  |  |  |  | 122.8 | FAIL | 2 | reworked calibration | 76.6538 | 76.6538 | 76.6538 |  |  |  |  | 104.8025 | PASS | 150
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.462 | 6.23 | 240 | TP | ug P/L | 0.0707 | 0.884 | -43.2 |  | 1 | 1 |  | 240 | 240 | 240 | -4 | PASS |  |  |  |  | 2 | reworked calibration | 21.0851 | 21.0851 | 21.0851 | -15.6594 | PASS |  |  |  |  | 151
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00622 | -0.0368 | -63 | TP | ug P/L | 0.0521 | 0.724 | -44.4 |  | 1 | 1 |  | -63 | -63 | -63 | [6 empty] | 2 | reworked calibration | -4.536 | -4.536 | -4.536 | [6 empty] | 152
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0176 | -0.254 | -73.4 | TP | ug P/L | 0.00727 | 0.147 | -47.3 |  | 1 | 1 |  | -73.4 | -73.4 | -73.4 | [6 empty] | 2 | reworked calibration | -84.133 | -84.133 | -84.133 | [6 empty] | 153
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s61 | 107 | 1 | 1 | TN | ug N/L | 1.51 | 20.4 | 936 | TP | ug P/L | 0.0784 | 1.07 | -42.7 |  | 1 | 1 |  | 936 | 936 | 936 | [6 empty] | 2 | reworked calibration | 30.4209 | 30.4209 | 30.4209 | [6 empty] | 154
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s62 | 108 | 1 | 1 | TN | ug N/L | 1.14 | 15.5 | 692 | TP | ug P/L | 0.0682 | 0.734 | -43.3 |  | 1 | 1 |  | 692 | 692 | 692 | [6 empty] | 2 | reworked calibration | 17.894 | 17.894 | 17.894 | [6 empty] | 155
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s63 | 109 | 1 | 1 | TN | ug N/L | 0.577 | 7.82 | 317 | TP | ug P/L | 0.0681 | 0.777 | -43.3 |  | 1 | 1 |  | 317 | 317 | 317 | [6 empty] | 2 | reworked calibration | 17.7648 | 17.7648 | 17.7648 | [6 empty] | 156
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s64 | 110 | 1 | 1 | TN | ug N/L | 1.56 | 21.2 | 971 | TP | ug P/L | 0.0731 | 0.866 | -43 |  | 1 | 1 |  | 971 | 971 | 971 | [6 empty] | 2 | reworked calibration | 24.0748 | 24.0748 | 24.0748 | [6 empty] | 157
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s65 | 111 | 1 | 1 | TN | ug N/L | 4.36 | 59.2 | 2890 | TP | ug P/L | 0.0874 | 1.36 | -42.1 |  | 1 | 1 |  | 2890 | 2890 | 2890 | [6 empty] | 2 | reworked calibration | 40.3897 | 40.3897 | 40.3897 | [6 empty] | 158
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s66 | 112 | 1 | 1 | TN | ug N/L | 2.17 | 29.5 | 1380 | TP | ug P/L | 0.0812 | 1.19 | -42.5 |  | 1 | 1 |  | 1380 | 1380 | 1380 | [6 empty] | 2 | reworked calibration | 33.6312 | 33.6312 | 33.6312 | [6 empty] | 159
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s67 | 113 | 1 | 1 | TN | ug N/L | 1.03 | 13.9 | 614 | TP | ug P/L | 0.0733 | 0.896 | -43 |  | 1 | 1 |  | 614 | 614 | 614 | [6 empty] | 2 | reworked calibration | 24.3207 | 24.3207 | 24.3207 | [6 empty] | 160
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s68 | 114 | 1 | 1 | TN | ug N/L | 3.14 | 42.6 | 2040 | TP | ug P/L | 0.101 | 1.77 | -41.2 |  | 1 | 1 |  | 2040 | 2040 | 2040 | [6 empty] | 2 | reworked calibration | 53.5247 | 53.5247 | 53.5247 | [6 empty] | 161
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s69 | 115 | 1 | 1 | TN | ug N/L | 0.668 | 9.05 | 378 | TP | ug P/L | 0.0719 | 0.815 | -43.1 |  | 1 | 1 |  | 378 | 378 | 378 | [6 empty] | 2 | reworked calibration | 22.589 | 22.589 | 22.589 | [6 empty] | 162
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s70 | 116 | 1 | 1 | TN | ug N/L | 1.88 | 25.7 | 1190 | TP | ug P/L | 0.0843 | 1.27 | -42.3 |  | 1 | 1 |  | 1190 | 1190 | 1190 | [6 empty] | 2 | reworked calibration | 37.0708 | 37.0708 | 37.0708 | [6 empty] | 163
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup65 | 117 | 1 | 1 | TN | ug N/L | 4.38 | 59.6 | 2910 | TP | ug P/L | 0.0893 | 1.44 | -42 |  | 1 | 1 |  | 2910 | 2910 | 2910 |  |  | 0.68966 | PASS |  |  | 2 | reworked calibration | 42.3642 | 42.3642 | 42.3642 |  |  | 4.7721 | PASS |  |  | 164
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk70 | 118 | 1 | 1 | TN | ug N/L | 2.14 | 29 | 1360 | TP | ug P/L | 0.107 | 1.95 | -40.8 |  | 1 | 1 |  | 1360 | 1360 | 1360 |  |  |  |  | 81.6 | PASS | 2 | reworked calibration | 58.5816 | 58.5816 | 58.5816 |  |  |  |  | 91.9017 | PASS | 165
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.459 | 6.21 | 239 | TP | ug P/L | 0.0699 | 0.891 | -43.2 |  | 1 | 1 |  | 239 | 239 | 239 | -4.4 | PASS |  |  |  |  | 2 | reworked calibration | 20.0725 | 20.0725 | 20.0725 | -19.7099 | PASS |  |  |  |  | 166
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00409 | -0.0102 | -61.7 | TP | ug P/L | 0.0543 | 0.74 | -44.2 |  | 1 | 1 |  | -61.7 | -61.7 | -61.7 | [6 empty] | 2 | reworked calibration | -1.2791 | -1.2791 | -1.2791 | [6 empty] | 167
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0154 | -0.262 | -73.9 | TP | ug P/L | 0.00746 | 0.124 | -47.2 |  | 1 | 1 |  | -73.9 | -73.9 | -73.9 | [6 empty] | 2 | reworked calibration | -83.7424 | -83.7424 | -83.7424 | [6 empty] | 168
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s71 | 119 | 1 | 1 | TN | ug N/L | 1.83 | 24.9 | 1150 | TP | ug P/L | 0.0893 | 1.45 | -42 |  | 1 | 1 |  | 1150 | 1150 | 1150 | [6 empty] | 2 | reworked calibration | 42.3642 | 42.3642 | 42.3642 | [6 empty] | 169
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s72 | 120 | 1 | 1 | TN | ug N/L | 0.94 | 12.8 | 558 | TP | ug P/L | 0.0768 | 1.07 | -42.8 |  | 1 | 1 |  | 558 | 558 | 558 | [6 empty] | 2 | reworked calibration | 28.5423 | 28.5423 | 28.5423 | [6 empty] | 170
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s73 | 121 | 1 | 1 | TN | ug N/L | 1.17 | 16 | 718 | TP | ug P/L | 0.0724 | 0.985 | -43 |  | 1 | 1 |  | 718 | 718 | 718 | [6 empty] | 2 | reworked calibration | 23.2103 | 23.2103 | 23.2103 | [6 empty] | 171
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s74 | 122 | 1 | 1 | TN | ug N/L | 1.25 | 16.9 | 762 | TP | ug P/L | 0.0769 | 1.04 | -42.8 |  | 1 | 1 |  | 762 | 762 | 762 | [6 empty] | 2 | reworked calibration | 28.6606 | 28.6606 | 28.6606 | [6 empty] | 172
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s75 | 123 | 1 | 1 | TN | ug N/L | 1.66 | 22.6 | 1040 | TP | ug P/L | 0.095 | 1.58 | -41.6 |  | 1 | 1 |  | 1040 | 1040 | 1040 | [6 empty] | 2 | reworked calibration | 48.016 | 48.016 | 48.016 | [6 empty] | 173
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s76 | 124 | 1 | 1 | TN | ug N/L | 1.14 | 15.5 | 691 | TP | ug P/L | 0.0704 | 0.834 | -43.2 |  | 1 | 1 |  | 691 | 691 | 691 | [6 empty] | 2 | reworked calibration | 20.7064 | 20.7064 | 20.7064 | [6 empty] | 174
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s77 | 125 | 1 | 1 | TN | ug N/L | 1.87 | 25.5 | 1190 | TP | ug P/L | 0.122 | 2.28 | -39.9 |  | 1 | 1 |  | 1190 | 1190 | 1190 | [6 empty] | 2 | reworked calibration | 69.2471 | 69.2471 | 69.2471 | [6 empty] | 175
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s78 | 126 | 1 | 1 | TN | ug N/L | 1.34 | 18.1 | 819 | TP | ug P/L | 0.0775 | 1.04 | -42.7 |  | 1 | 1 |  | 819 | 819 | 819 | [6 empty] | 2 | reworked calibration | 29.3681 | 29.3681 | 29.3681 | [6 empty] | 176
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s79 | 127 | 1 | 1 | TN | ug N/L | 1.01 | 13.8 | 607 | TP | ug P/L | 0.0676 | 0.658 | -43.4 |  | 1 | 1 |  | 607 | 607 | 607 | [6 empty] | 2 | reworked calibration | 17.1165 | 17.1165 | 17.1165 | [6 empty] | 177
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s80 | 128 | 1 | 1 | TN | ug N/L | 3 | 40.8 | 1950 | TP | ug P/L | 0.0868 | 1.31 | -42.1 |  | 1 | 1 |  | 1950 | 1950 | 1950 | [6 empty] | 2 | reworked calibration | 39.7567 | 39.7567 | 39.7567 | [6 empty] | 178
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup75 | 129 | 1 | 1 | TN | ug N/L | 1.93 | 26.2 | 1220 | TP | ug P/L | 0.121 | 2.28 | -39.9 |  | 1 | 1 |  | 1220 | 1220 | 1220 |  |  | 15.9292 | PASS |  |  | 2 | reworked calibration | 68.624 | 68.624 | 68.624 |  |  | 35.336 | FAIL |  |  | 179
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk80 | 130 | 1 | 1 | TN | ug N/L | 3.25 | 44.1 | 2120 | TP | ug P/L | 0.107 | 1.99 | -40.9 |  | 1 | 1 |  | 2120 | 2120 | 2120 |  |  |  |  | 89.2 | PASS | 2 | reworked calibration | 58.5816 | 58.5816 | 58.5816 |  |  |  |  | 81.1579 | PASS | 180
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.472 | 6.36 | 247 | TP | ug P/L | 0.0729 | 0.883 | -43 |  | 1 | 1 |  | 247 | 247 | 247 | -1.2 | PASS |  |  |  |  | 2 | reworked calibration | 23.8284 | 23.8284 | 23.8284 | -4.6862 | PASS |  |  |  |  | 181
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00739 | -0.0299 | -62.6 | TP | ug P/L | 0.0494 | 0.671 | -44.5 |  | 1 | 1 |  | -62.6 | -62.6 | -62.6 | [6 empty] | 2 | reworked calibration | -8.6161 | -8.6161 | -8.6161 | [6 empty] | 182
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0154 | -0.231 | -72.4 | TP | ug P/L | 0.00635 | 0.0879 | -47.3 |  | 1 | 1 |  | -72.4 | -72.4 | -72.4 | [6 empty] | 2 | reworked calibration | -86.0306 | -86.0306 | -86.0306 | [6 empty] | 183
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s81 | 131 | 1 | 1 | TN | ug N/L | 1.14 | 15.6 | 695 | TP | ug P/L | 0.0713 | 0.81 | -43.1 |  | 1 | 1 |  | 695 | 695 | 695 | [6 empty] | 2 | reworked calibration | 21.8393 | 21.8393 | 21.8393 | [6 empty] | 184
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s82 | 132 | 1 | 1 | TN | ug N/L | 0.583 | 7.92 | 322 | TP | ug P/L | 0.0787 | 1.09 | -42.6 |  | 1 | 1 |  | 322 | 322 | 322 | [6 empty] | 2 | reworked calibration | 30.7696 | 30.7696 | 30.7696 | [6 empty] | 185
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s83 | 133 | 1 | 1 | TN | ug N/L | 1.14 | 15.4 | 690 | TP | ug P/L | 0.0685 | 0.697 | -43.3 |  | 1 | 1 |  | 690 | 690 | 690 | [6 empty] | 2 | reworked calibration | 18.2811 | 18.2811 | 18.2811 | [6 empty] | 186
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s84 | 134 | 1 | 1 | TN | ug N/L | 1.42 | 19.3 | 880 | TP | ug P/L | 0.0886 | 1.42 | -42 |  | 1 | 1 |  | 880 | 880 | 880 | [6 empty] | 2 | reworked calibration | 41.642 | 41.642 | 41.642 | [6 empty] | 187
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s85 | 135 | 1 | 1 | TN | ug N/L | 0.496 | 6.71 | 263 | TP | ug P/L | 0.0659 | 0.614 | -43.5 |  | 1 | 1 |  | 263 | 263 | 263 | [6 empty] | 2 | reworked calibration | 14.8889 | 14.8889 | 14.8889 | [6 empty] | 188
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s86 | 136 | 1 | 1 | TN | ug N/L | 0.56 | 7.61 | 307 | TP | ug P/L | 0.0753 | 1.02 | -42.9 |  | 1 | 1 |  | 307 | 307 | 307 | [6 empty] | 2 | reworked calibration | 26.7518 | 26.7518 | 26.7518 | [6 empty] | 189
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s87 | 137 | 1 | 1 | TN | ug N/L | 0.629 | 8.57 | 354 | TP | ug P/L | 0.0731 | 0.885 | -43 |  | 1 | 1 |  | 354 | 354 | 354 | [6 empty] | 2 | reworked calibration | 24.0748 | 24.0748 | 24.0748 | [6 empty] | 190
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s88 | 138 | 1 | 1 | TN | ug N/L | 0.546 | 7.39 | 297 | TP | ug P/L | 0.0759 | 0.962 | -42.8 |  | 1 | 1 |  | 297 | 297 | 297 | [6 empty] | 2 | reworked calibration | 27.4714 | 27.4714 | 27.4714 | [6 empty] | 191
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s89 | 139 | 1 | 1 | TN | ug N/L | 1.16 | 15.8 | 708 | TP | ug P/L | 0.0786 | 1.07 | -42.6 |  | 1 | 1 |  | 708 | 708 | 708 | [6 empty] | 2 | reworked calibration | 30.6535 | 30.6535 | 30.6535 | [6 empty] | 192
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s90 | 140 | 1 | 1 | TN | ug N/L | 0.613 | 8.31 | 341 | TP | ug P/L | 0.0661 | 0.638 | -43.5 |  | 1 | 1 |  | 341 | 341 | 341 | [6 empty] | 2 | reworked calibration | 15.1529 | 15.1529 | 15.1529 | [6 empty] | 193
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup85 | 141 | 1 | 1 | TN | ug N/L | 0.503 | 6.8 | 268 | TP | ug P/L | 0.0709 | 0.675 | -43.1 |  | 1 | 1 |  | 268 | 268 | 268 |  |  | 1.8832 | PASS |  |  | 2 | reworked calibration | 21.337 | 21.337 | 21.337 |  |  | 35.5995 | FAIL |  |  | 194
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk90 | 142 | 1 | 1 | TN | ug N/L | 0.961 | 13 | 571 | TP | ug P/L | 0.0892 | 1.38 | -42 |  | 1 | 1 |  | 571 | 571 | 571 |  |  |  |  | 97.71 | PASS | 2 | reworked calibration | 42.2614 | 42.2614 | 42.2614 |  |  |  |  | 112.6603 | PASS | 195
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.466 | 6.36 | 247 | TP | ug P/L | 0.072 | 0.878 | -43.1 |  | 1 | 1 |  | 247 | 247 | 247 | -1.2 | PASS |  |  |  |  | 2 | reworked calibration | 22.7135 | 22.7135 | 22.7135 | -9.1459 | PASS |  |  |  |  | 196
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.0081 | 0.0293 | -59.8 | TP | ug P/L | 0.0503 | 0.665 | -44.5 |  | 1 | 1 |  | -59.8 | -59.8 | -59.8 | [6 empty] | 2 | reworked calibration | -7.2459 | -7.2459 | -7.2459 | [6 empty] | 197
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00635 | 0.0636 | -58.1 | TP | ug P/L | 0.00846 | 0.133 | -47.2 |  | 1 | 1 |  | -58.1 | -58.1 | -58.1 | [6 empty] | 2 | reworked calibration | -81.6942 | -81.6942 | -81.6942 | [6 empty] | 198
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s91 | 143 | 1 | 1 | TN | ug N/L | 1.83 | 24.9 | 1150 | TP | ug P/L | 0.0965 | 1.66 | -41.5 |  | 1 | 1 |  | 1150 | 1150 | 1150 | [6 empty] | 2 | reworked calibration | 49.4356 | 49.4356 | 49.4356 | [6 empty] | 199
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s92 | 144 | 1 | 1 | TN | ug N/L | 3.77 | 51.3 | 2480 | TP | ug P/L | 0.0855 | 1.29 | -42.2 |  | 1 | 1 |  | 2480 | 2480 | 2480 | [6 empty] | 2 | reworked calibration | 38.3698 | 38.3698 | 38.3698 | [6 empty] | 200
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s93 | 145 | 1 | 1 | TN | ug N/L | 0.522 | 7.09 | 282 | TP | ug P/L | 0.0714 | 0.795 | -43.1 |  | 1 | 1 |  | 282 | 282 | 282 | [6 empty] | 2 | reworked calibration | 21.9646 | 21.9646 | 21.9646 | [6 empty] | 201
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s94 | 146 | 1 | 1 | TN | ug N/L | 3.47 | 47 | 2270 | TP | ug P/L | 0.0877 | 1.32 | -42.1 |  | 1 | 1 |  | 2270 | 2270 | 2270 | [6 empty] | 2 | reworked calibration | 40.7044 | 40.7044 | 40.7044 | [6 empty] | 202
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s95 | 147 | 1 | 1 | TN | ug N/L | 1.43 | 19.3 | 881 | TP | ug P/L | 0.0808 | 1.1 | -42.5 |  | 1 | 1 |  | 881 | 881 | 881 | [6 empty] | 2 | reworked calibration | 33.1787 | 33.1787 | 33.1787 | [6 empty] | 203
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s96 | 148 | 1 | 1 | TN | ug N/L | 0.643 | 8.69 | 360 | TP | ug P/L | 0.0679 | 0.774 | -43.3 |  | 1 | 1 |  | 360 | 360 | 360 | [6 empty] | 2 | reworked calibration | 17.5058 | 17.5058 | 17.5058 | [6 empty] | 204
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s97 | 149 | 1 | 1 | TN | ug N/L | 0.752 | 10.2 | 432 | TP | ug P/L | 0.0703 | 0.823 | -43.2 |  | 1 | 1 |  | 432 | 432 | 432 | [6 empty] | 2 | reworked calibration | 20.5798 | 20.5798 | 20.5798 | [6 empty] | 205
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s98 | 150 | 1 | 1 | TN | ug N/L | 0.673 | 9.07 | 379 | TP | ug P/L | 0.075 | 0.964 | -42.9 |  | 1 | 1 |  | 379 | 379 | 379 | [6 empty] | 2 | reworked calibration | 26.3904 | 26.3904 | 26.3904 | [6 empty] | 206
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s99 | 151 | 1 | 1 | TN | ug N/L | 0.665 | 9.02 | 376 | TP | ug P/L | 0.0677 | 0.618 | -43.3 |  | 1 | 1 |  | 376 | 376 | 376 | [6 empty] | 2 | reworked calibration | 17.2464 | 17.2464 | 17.2464 | [6 empty] | 207
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s100 | 152 | 1 | 1 | TN | ug N/L | 1.32 | 17.8 | 808 | TP | ug P/L | 0.0732 | 0.872 | -43 |  | 1 | 1 |  | 808 | 808 | 808 | [6 empty] | 2 | reworked calibration | 24.1978 | 24.1978 | 24.1978 | [6 empty] | 208
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup95 | 153 | 1 | 1 | TN | ug N/L | 1.47 | 19.8 | 905 | TP | ug P/L | 0.0787 | 1.1 | -42.6 |  | 1 | 1 |  | 905 | 905 | 905 |  |  | 2.6876 | PASS |  |  | 2 | reworked calibration | 30.7696 | 30.7696 | 30.7696 |  |  | 7.5344 | PASS |  |  | 209
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk100 | 154 | 1 | 1 | TN | ug N/L | 1.66 | 22.4 | 1030 | TP | ug P/L | 0.102 | 1.67 | -41.1 |  | 1 | 1 |  | 1030 | 1030 | 1030 |  |  |  |  | 99.1 | PASS | 2 | reworked calibration | 54.3989 | 54.3989 | 54.3989 |  |  |  |  | 126.2444 | FAIL | 210
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.473 | 6.38 | 248 | TP | ug P/L | 0.0701 | 0.872 | -43.2 |  | 1 | 1 |  | 248 | 248 | 248 | -0.8 | PASS |  |  |  |  | 2 | reworked calibration | 20.3264 | 20.3264 | 20.3264 | -18.6943 | PASS |  |  |  |  | 211
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00967 | 0.041 | -59.2 | TP | ug P/L | 0.0522 | 0.707 | -44.3 |  | 1 | 1 |  | -59.2 | -59.2 | -59.2 | [6 empty] | 2 | reworked calibration | -4.3866 | -4.3866 | -4.3866 | [6 empty] | 212
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00841 | 0.057 | -58.5 | TP | ug P/L | 0.00661 | 0.072 | -47.3 |  | 1 | 1 |  | -58.5 | -58.5 | -58.5 | [6 empty] | 2 | reworked calibration | -85.4932 | -85.4932 | -85.4932 | [6 empty] | 213
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s101 | 155 | 1 | 1 | TN | ug N/L | 0.534 | 7.24 | 289 | TP | ug P/L | 0.065 | 0.694 | -43.5 |  | 1 | 1 |  | 289 | 289 | 289 | [6 empty] | 2 | reworked calibration | 13.6949 | 13.6949 | 13.6949 | [6 empty] | 214
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s102 | 156 | 1 | 1 | TN | ug N/L | 0.787 | 10.7 | 460 | TP | ug P/L | 0.0679 | 0.719 | -43.3 |  | 1 | 1 |  | 460 | 460 | 460 | [6 empty] | 2 | reworked calibration | 17.5058 | 17.5058 | 17.5058 | [6 empty] | 215
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s103 | 157 | 1 | 1 | TN | ug N/L | 1.22 | 16.6 | 747 | TP | ug P/L | 0.0813 | 1.25 | -42.5 |  | 1 | 1 |  | 747 | 747 | 747 | [6 empty] | 2 | reworked calibration | 33.7441 | 33.7441 | 33.7441 | [6 empty] | 216
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s104 | 158 | 1 | 1 | TN | ug N/L | 0.545 | 9.26 | 387 | TP | ug P/L | 0.0839 | 0.908 | 24.3 |  | 1 | 1 |  | 387 | 387 | 387 | [6 empty] | 2 | reworked calibration | 36.6337 | 36.6337 | 36.6337 | [6 empty] | 217
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s105 | 159 | 1 | 1 | TN | ug N/L | 0.016 | 0.0587 | -58.4 | TP | ug P/L | 0.124 | 0.672 | -39.8 |  | 1 | 3 | air spikes | -58.4 | -58.4 | -58.4 | [6 empty] | 3 | air spikes | 70.4558 | 70.4558 | 70.4558 | [6 empty] | 218
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s106 | 160 | 1 | 1 | TN | ug N/L | 3.28 | 36.6 | 1740 | TP | ug P/L | 0.223 | 2.26 | -33.5 |  | 1 | 3 | air spikes | 1740 | 1740 | 1740 | [6 empty] | 3 | air spikes | 67.5389 | 67.5389 | 67.5389 | [6 empty] | 219
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s107 | 161 | 1 | 1 | TN | ug N/L | -2.01 | -50.9 | -2450 | TP | ug P/L | 0.265 | 1.99 | -30.9 |  | 1 | 3 | air spikes | -2450 | -2450 | -2450 | [6 empty] | 3 | air spikes | 29.1385 | 29.1385 | 29.1385 | [6 empty] | 220
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dups105 | 162 | 1 | 1 | TN | ug N/L | -2.1 | -59.8 | -2850 | TP | ug P/L | -0.171 | -1.89 | -59 |  | 1 | 3 | air spikes | -2850 | -2850 | -2850 | [6 empty] | 3 | air spikes | -650.2421 | -650.2421 | -650.2421 | [6 empty] | 221
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spks107 | 163 | 1 | 1 | TN | ug N/L | 3.44 | 7.38 | 296 | TP | ug P/L | 0.29 | 2.64 | -29.3 |  | 1 | 3 | air spikes | 296 | 296 | 296 | [6 empty] | 3 | air spikes | -4.2301 | -4.2301 | -4.2301 | [6 empty] | 222
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s108 | 164 | 1 | 1 | TN | ug N/L | -2.83 | -83.1 | -3890 | TP | ug P/L | 0.161 | 0.0967 | -37.4 |  | 1 | 3 | air spikes | -3890 | -3890 | -3890 | [6 empty] | 3 | air spikes | 83.7615 | 83.7615 | 83.7615 | [6 empty] | 223
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 1.78 | -6.12 | -356 | TP | ug P/L | 0.119 | 1.84 | -40 |  | 1 | 3 | air spikes | -356 | -356 | -356 | [6 empty] | 3 | air spikes | 67.34 | 67.34 | 67.34 | [6 empty] | 224
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.788 | -30.1 | -1490 | TP | ug P/L | -0.114 | -1.53 | -55.2 |  | 1 | 3 | air spikes | -1490 | -1490 | -1490 | [6 empty] | 3 | air spikes | -425.8565 | -425.8565 | -425.8565 | [6 empty] | 225
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.77 | -12.6 | -665 | TP | ug P/L | 0.0697 | -0.315 | -43.2 |  | 1 | 3 | air spikes | -665 | -665 | -665 | [6 empty] | 3 | air spikes | 19.8181 | 19.8181 | 19.8181 | [6 empty] | 226
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal1500 | 1 | 1 | 1 | TN | ug N/L | 2.91 | 4.73 | 167 | TP | ug P/L | 0.348 | 5.62 | -25.7 |  | 1 | 3 | air spikes | 167 | 167 | 167 | [6 empty] | 3 | air spikes | -111.8552 | -111.8552 | -111.8552 | [6 empty] | 227
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal1000 | 3 | 1 | 1 | TN | ug N/L | -2.58 | -57 | -2720 | TP | ug P/L | 0.232 | 4.25 | -32.9 |  | 1 | 3 | air spikes | -2720 | -2720 | -2720 | [6 empty] | 3 | air spikes | 61.1741 | 61.1741 | 61.1741 | [6 empty] | 228
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call500 | 5 | 1 | 1 | TN | ug N/L | 1.98 | 3.67 | 116 | TP | ug P/L | 0.253 | 2.78 | -31.6 |  | 1 | 3 | air spikes | 116 | 116 | 116 | [6 empty] | 3 | air spikes | 42.3692 | 42.3692 | 42.3692 | [6 empty] | 229
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call250 | 7 | 1 | 1 | TN | ug N/L | 1.61 | -10.7 | -572 | TP | ug P/L | 0.115 | 1.85 | -40.3 |  | 1 | 3 | air spikes | -572 | -572 | -572 | [6 empty] | 3 | air spikes | 64.6213 | 64.6213 | 64.6213 | [6 empty] | 230
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal100 | 9 | 1 | 1 | TN | ug N/L | -1.98 | -22.1 | -1110 | TP | ug P/L | -0.0785 | -1.33 | -52.9 |  | 1 | 3 | air spikes | -1110 | -1110 | -1110 | [6 empty] | 3 | air spikes | -306.7144 | -306.7144 | -306.7144 | [6 empty] | 231
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | 2 | 3.26 | 96.1 | TP | ug P/L | -0.0457 | -0.32 | -50.7 |  | 1 | 3 | air spikes | 96.1 | 96.1 | 96.1 | [6 empty] | 3 | air spikes | -210.6923 | -210.6923 | -210.6923 | [6 empty] | 232
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | 4.66 | 29.1 | 1370 | TP | ug P/L | 0.119 | -0.0565 | -40.1 |  | 1 | 3 | air spikes | 1370 | 1370 | 1370 | [6 empty] | 3 | air spikes | 67.34 | 67.34 | 67.34 | [6 empty] | 233
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | 2.46 | 14.2 | 630 | TP | ug P/L | -0.00144 | 0.0124 | -47.8 |  | 1 | 3 | air spikes | 630 | 630 | 630 | [6 empty] | 3 | air spikes | -102.5243 | -102.5243 | -102.5243 | [6 empty] | 234
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | 2.95 | 12 | 520 | TP | ug P/L | -0.052 | -0.591 | -51.1 |  | 1 | 3 | air spikes | 520 | 520 | 520 | [6 empty] | 3 | air spikes | -228.0879 | -228.0879 | -228.0879 | [6 empty] | 235
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | 9.64 | 69.9 | 3440 | TP | ug P/L | -0.0399 | -0.284 | -50.3 |  | 1 | 3 | air spikes | 3440 | 3440 | 3440 | [6 empty] | 3 | air spikes | -195.1177 | -195.1177 | -195.1177 | [6 empty] | 236
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.158 | 2.03 | 36.9 | TP | ug P/L | 0.0595 | 0.838 | -43.9 |  | 1 | 1 |  | 36.9 | 36.9 | 36.9 | -26.2 | FAIL |  |  |  |  | 2 | reworked calibration | 6.1775 | 6.1775 | 6.1775 | 23.5493 | FAIL |  |  |  |  | 237
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.142 | 1.94 | 32.4 | TP | ug P/L | 0.0584 | 0.817 | -43.9 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 | 29.6 | FAIL |  |  |  |  | 2 | reworked calibration | 4.6284 | 4.6284 | 4.6284 | 85.1365 | FAIL |  |  |  |  | 238
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0919 | 1.28 | 0.497 | TP | ug P/L | 0.052 | 0.686 | -44.4 |  | 1 | 1 |  | 0.497 | 0.497 | 0.497 | [6 empty] | 2 | reworked calibration | -4.6855 | -4.6855 | -4.6855 | [6 empty] | 239
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrite | S1 | 1 | 1 | TN | ug N/L | 2.27 | 30.8 | 1450 | TP | ug P/L | 0.00874 | 0.143 | -47.2 | 97.7273 | 1 | 1 |  | 1450 | 1450 | 1450 | [6 empty] | 2 | reworked calibration | -81.123 | -81.123 | -81.123 | [6 empty] | 240
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrate | S2 | 1 | 1 | TN | ug N/L | 2.22 | 30.1 | 1410 | TP | ug P/L | 0.00884 | 0.101 | -47.1 |  | 1 | 1 |  | 1410 | 1410 | 1410 | [6 empty] | 2 | reworked calibration | -80.9192 | -80.9192 | -80.9192 | [6 empty] | 241
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0179 | -0.37 | -79 | TP | ug P/L | 0.00942 | 0.145 | -47.1 |  | 1 | 1 |  | -79 | -79 | -79 | [6 empty] | 2 | reworked calibration | -79.7398 | -79.7398 | -79.7398 | [6 empty] | 242
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s105 rerun | 159 | 1 | 1 | TN | ug N/L | 1.11 | 15.1 | 674 | TP | ug P/L | 0.0696 | 0.782 | -43.2 |  | 1 | 1 |  | 674 | 674 | 674 | [6 empty] | 2 | reworked calibration | 19.6907 | 19.6907 | 19.6907 | [6 empty] | 243
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s106 rerun | 160 | 1 | 1 | TN | ug N/L | 0.842 | 11.6 | 501 | TP | ug P/L | 0.0675 | 0.736 | -43.4 |  | 1 | 1 |  | 501 | 501 | 501 | [6 empty] | 2 | reworked calibration | 16.9865 | 16.9865 | 16.9865 | [6 empty] | 244
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s107 rerun | 161 | 1 | 1 | TN | ug N/L | 1.1 | 15.2 | 678 | TP | ug P/L | 0.0705 | 0.822 | -43.2 |  | 1 | 1 |  | 678 | 678 | 678 | [6 empty] | 2 | reworked calibration | 20.8327 | 20.8327 | 20.8327 | [6 empty] | 245
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dups105 rerun | 162 | 1 | 1 | TN | ug N/L | 1.1 | 15.3 | 682 | TP | ug P/L | 0.0698 | 0.792 | -43.2 |  | 1 | 1 |  | 682 | 682 | 682 | [6 empty] | 2 | reworked calibration | 19.9454 | 19.9454 | 19.9454 | [6 empty] | 246
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk107 rerun | 163 | 1 | 1 | TN | ug N/L | 1.42 | 19.7 | 897 | TP | ug P/L | 0.0958 | 1.58 | -41.5 |  | 1 | 1 |  | 897 | 897 | 897 | [6 empty] | 2 | reworked calibration | 48.7766 | 48.7766 | 48.7766 | [6 empty] | 247
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s108 rerun | 164 | 1 | 1 | TN | ug N/L | 0.68 | 9.51 | 400 | TP | ug P/L | 0.0719 | 0.863 | -43.1 |  | 1 | 1 |  | 400 | 400 | 400 | [6 empty] | 2 | reworked calibration | 22.589 | 22.589 | 22.589 | [6 empty] | 248
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal 1500 | 1 | 1 | 1 | TN | ug N/L | 2.32 | 32 | 1510 | TP | ug P/L | 0.24 | 4.61 | -32.4 |  | 1 | 1 |  | 1510 | 1510 | 1510 | [6 empty] | 2 | reworked calibration | 616.0527 | 616.0527 | 616.0527 | 310.7018 | FAIL |  |  |  |  | 249
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal 50/5 | 11 | 1 | 1 | TN | ug N/L | 0.159 | 2.28 | 49 | TP | ug P/L | 0.0592 | 0.389 | -43.9 |  | 1 | 1 |  | 49 | 49 | 49 | [6 empty] | 2 | reworked calibration | 3.7014 | 3.7014 | 3.7014 | -25.9712 | FAIL |  |  |  |  | 250
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | chk 25/250 | 7 | 1 | 1 | TN | ug N/L | 0.451 | 6.34 | 246 | TP | ug P/L | 0.0695 | 0.878 | -43.2 |  | 1 | 1 |  | 246 | 246 | 246 | [6 empty] | 2 | reworked calibration | 15.3339 | 15.3339 | 15.3339 | -38.6646 | FAIL |  |  |  |  | 251
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal 25/2.5 | 13 | 1 | 1 | TN | ug N/L | 0.123 | 1.78 | 24.7 | TP | ug P/L | 0.0562 | 0.772 | -44.1 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | [6 empty] | 2 | reworked calibration | 2.4457 | 2.4457 | 2.4457 | -2.1719 | PASS |  |  |  |  | 252
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate | S4 | 1 | 1 | TN | ug N/L | -0.00258 | -0.0295 | -62.6 | TP | ug P/L | 0.0457 | 0.615 | -44.8 |  | 1 | 1 |  | -62.6 | -62.6 | -62.6 | [6 empty] | 2 | reworked calibration | -12.6309 | -12.6309 | -12.6309 | [6 empty] | 253

## QAQC flat area
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE NO3-N: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- dilution corrected | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- dilution corrected | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | dye | S16 | 1 | 1 | TN | ug N/L | -0.00637 | -0.00462 | 0 | TP | ug P/L | 0.0299 | 0.592 | -2.03 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 3.15 | 3.15 | 3.15
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | dye | S16 | 1 | 1 | TN | ug N/L | 0.0144 | 0.152 | 0 | TP | ug P/L | 0.0314 | 0.622 | -0.949 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 3.4574 | 3.4574 | 3.4574
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.27 | 30 | 0 | TP | ug P/L | -0.00787 | -0.0671 | -30.3 | 98.3333 | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 0.035505 | 0.035505 | 0.035505
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.25 | 29.5 | 0 | TP | ug P/L | -0.00942 | -0.0814 | -31.5 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 0.045065 | 0.045065 | 0.045065
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0162 | -0.271 | 0 | TP | ug P/L | -0.00923 | -0.0958 | -31.4 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 0.058001 | 0.058001 | 0.058001
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00707 | -0.00498 | 0 | TP | ug P/L | 0.0522 | 0.698 | 13.8 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 4.3006 | 4.3006 | 4.3006
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.237 | 4.65 | 150 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 1 |  | 175.6533 | 175.6533 | 175.6533
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.237 | 4.59 | 150 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 1 |  | 171.1813 | 171.1813 | 171.1813
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.66 | 21.8 | 1000 | TP | ug P/L | 0.164 | 3.25 | 100 |  | 1 | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 1 |  | 86.3333 | 86.3333 | 86.3333
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.67 | 22.6 | 1000 | TP | ug P/L | 0.16 | 3.15 | 100 |  | 1 | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 1 |  | 81.1548 | 81.1548 | 81.1548
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.867 | 11.4 | 500 | TP | ug P/L | 0.0979 | 1.72 | 50 |  | 1 | 1 |  | 500 | 500 | 500 | [6 empty] | 1 |  | 24.6273 | 24.6273 | 24.6273
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.86 | 11.4 | 500 | TP | ug P/L | 0.0977 | 1.73 | 50 |  | 1 | 1 |  | 500 | 500 | 500 | [6 empty] | 1 |  | 24.9089 | 24.9089 | 24.9089
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.478 | 6.3 | 250 | TP | ug P/L | 0.0729 | 0.978 | 25 |  | 1 | 1 |  | 250 | 250 | 250 | [6 empty] | 1 |  | 8.2054 | 8.2054 | 8.2054 | -67.1785 | FAIL
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.474 | 6.29 | 250 | TP | ug P/L | 0.073 | 0.925 | 25 |  | 1 | 1 |  | 250 | 250 | 250 | [6 empty] | 1 |  | 7.3699 | 7.3699 | 7.3699 | -70.5205 | FAIL
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.43 | 100 | TP | ug P/L | 0.0629 | 0.498 | 10 |  | 1 | 1 |  | 100 | 100 | 100 | [6 empty] | 1 |  | 2.2804 | 2.2804 | 2.2804
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.44 | 100 | TP | ug P/L | 0.0682 | 0.488 | 10 |  | 1 | 1 |  | 100 | 100 | 100 | [6 empty] | 1 |  | 2.1962 | 2.1962 | 2.1962
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.176 | 2.33 | 50 | TP | ug P/L | 0.0612 | 0.443 | 5 |  | 1 | 1 |  | 50 | 50 | 50 | [6 empty] | 1 |  | 1.8372 | 1.8372 | 1.8372
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.166 | 2.28 | 50 | TP | ug P/L | 0.0591 | 0.353 | 5 |  | 1 | 1 |  | 50 | 50 | 50 | [6 empty] | 1 |  | 1.2164 | 1.2164 | 1.2164
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.131 | 1.8 | 25 | TP | ug P/L | 0.059 | 0.855 | 2.5 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 6.3354 | 6.3354 | 6.3354
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.129 | 1.76 | 25 | TP | ug P/L | 0.0563 | 0.799 | 2.5 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 5.5643 | 5.5643 | 5.5643
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0914 | 1.27 | 0 | TP | ug P/L | 0.0562 | 0.806 | 0 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 5.6579 | 5.6579 | 5.6579
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0915 | 1.27 | 0 | TP | ug P/L | 0.0567 | 0.799 | 0 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 5.5643 | 5.5643 | 5.5643
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.2 | 29.8 | 1400 | TP | ug P/L | -0.0119 | -0.111 | -92.5 | 99.6644 | 1 | 1 |  | 1400 | 1400 | 1400 | [6 empty] | 1 |  | 0.07526 | 0.07526 | 0.07526
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.22 | 29.7 | 1390 | TP | ug P/L | 0.0192 | 0.22 | -47.7 |  | 1 | 1 |  | 1390 | 1390 | 1390 | [6 empty] | 1 |  | 0.5367 | 0.5367 | 0.5367
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0172 | -0.298 | -75.6 | TP | ug P/L | -0.00934 | -0.0803 | -88.6 |  | 1 | 1 |  | -75.6 | -75.6 | -75.6 | [6 empty] | 1 |  | 0.044213 | 0.044213 | 0.044213
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00164 | 0.0147 | -60.5 | TP | ug P/L | 0.0552 | 0.782 | -0.858 |  | 1 | 1 |  | -60.5 | -60.5 | -60.5 | [6 empty] | 1 |  | 5.3401 | 5.3401 | 5.3401
0
0
0
0
0
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.24 | 30.5 | 1440 | TP | ug P/L | -0.0196 | -0.17 | -49 | 97.377 | 1 | 1 |  | 1440 | 1440 | 1440 | [6 empty] | 1 |  | 0.17732 | 0.17732 | 0.17732
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.21 | 29.7 | 1400 | TP | ug P/L | -0.015 | -0.122 | -48.7 |  | 1 | 1 |  | 1400 | 1400 | 1400 | [6 empty] | 1 |  | 0.090059 | 0.090059 | 0.090059
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0165 | -0.274 | -74.4 | TP | ug P/L | -0.022 | -0.227 | -49.2 |  | 1 | 1 |  | -74.4 | -74.4 | -74.4 | [6 empty] | 1 |  | 0.32888 | 0.32888 | 0.32888
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.26 | 242 | TP | ug P/L | 0.071 | 0.863 | -43.1 |  | 1 | 1 |  | 242 | 242 | 242 | -3.2 | PASS |  |  |  |  | 1 |  | 6.4496 | 6.4496 | 6.4496 | -74.2015 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.0025 | -0.0316 | -62.7 | TP | ug P/L | 0.0484 | 0.653 | -44.6 |  | 1 | 1 |  | -62.7 | -62.7 | -62.7 | [6 empty] | 1 |  | 3.7901 | 3.7901 | 3.7901
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0167 | -0.267 | -74.1 | TP | ug P/L | 0.0114 | 0.176 | -47 |  | 1 | 1 |  | -74.1 | -74.1 | -74.1 | [6 empty] | 1 |  | 0.37421 | 0.37421 | 0.37421
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00439 | -0.035 | -62.9 | TP | ug P/L | 0.042 | 0.5 | -45 |  | 1 | 1 |  | -62.9 | -62.9 | -62.9 | [6 empty] | 1 |  | 2.2974 | 2.2974 | 2.2974
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 20 | 1 | 1 | TN | ug N/L | 0.084 | 1.2 | -3.48 | TP | ug P/L | 0.0482 | 0.56 | -44.6 |  | 1 | 1 |  | -3.48 | -3.48 | -3.48 | [6 empty] | 1 |  | 2.8381 | 2.8381 | 2.8381
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 21 | 1 | 1 | TN | ug N/L | 0.0936 | 1.22 | -2.22 | TP | ug P/L | 0.0476 | 0.617 | -44.6 |  | 1 | 1 |  | -2.22 | -2.22 | -2.22 | [6 empty] | 1 |  | 3.4052 | 3.4052 | 3.4052
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 22 | 1 | 1 | TN | ug N/L | 0.0896 | 1.28 | 0.652 | TP | ug P/L | 0.0538 | 0.764 | -44.2 |  | 1 | 1 |  | 0.652 | 0.652 | 0.652 | [6 empty] | 1 |  | 5.1078 | 5.1078 | 5.1078
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 23 | 1 | 1 | TN | ug N/L | 0.095 | 1.3 | 1.41 | TP | ug P/L | 0.0539 | 0.721 | -44.2 |  | 1 | 1 |  | 1.41 | 1.41 | 1.41 | [6 empty] | 1 |  | 4.574 | 4.574 | 4.574
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 24 | 1 | 1 | TN | ug N/L | 0.0925 | 1.31 | 1.89 | TP | ug P/L | 0.0537 | 0.741 | -44.3 |  | 1 | 1 |  | 1.89 | 1.89 | 1.89 | [6 empty] | 1 |  | 4.8186 | 4.8186 | 4.8186
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 25 | 1 | 1 | TN | ug N/L | 0.71 | 9.48 | 398 | TP | ug P/L | 0.0817 | 1.27 | -42.4 |  | 1 | 1 |  | 398 | 398 | 398 | [6 empty] | 1 |  | 13.6154 | 13.6154 | 13.6154
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 26 | 1 | 1 | TN | ug N/L | 0.707 | 9.5 | 399 | TP | ug P/L | 0.0795 | 1.25 | -42.6 |  | 1 | 1 |  | 399 | 399 | 399 | [6 empty] | 1 |  | 13.2013 | 13.2013 | 13.2013
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 27 | 1 | 1 | TN | ug N/L | 0.709 | 9.47 | 398 | TP | ug P/L | 0.0827 | 1.27 | -42.4 |  | 1 | 1 |  | 398 | 398 | 398 | [6 empty] | 1 |  | 13.6154 | 13.6154 | 13.6154
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 28 | 1 | 1 | TN | ug N/L | 0.711 | 9.46 | 397 | TP | ug P/L | 0.0812 | 1.28 | -42.5 |  | 1 | 1 |  | 397 | 397 | 397 | [6 empty] | 1 |  | 13.8249 | 13.8249 | 13.8249
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 29 | 1 | 1 | TN | ug N/L | 0.709 | 9.42 | 395 | TP | ug P/L | 0.0823 | 1.19 | -42.4 |  | 1 | 1 |  | 395 | 395 | 395 | [6 empty] | 1 |  | 11.9974 | 11.9974 | 11.9974
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 30 | 1 | 1 | TN | ug N/L | 0.433 | 5.75 | 217 | TP | ug P/L | 0.0724 | 1.02 | -43 |  | 1 | 1 |  | 217 | 217 | 217 | [6 empty] | 1 |  | 8.8994 | 8.8994 | 8.8994
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | skp blank | 31 | 1 | 1 | TN | ug N/L | 0.436 | 5.89 | 224 | TP | ug P/L | 0.0746 | 1.06 | -42.9 |  | 1 | 1 |  | 224 | 224 | 224 | [6 empty] | 1 |  | 9.5867 | 9.5867 | 9.5867
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 32 | 1 | 1 | TN | ug N/L | 0.441 | 5.9 | 224 | TP | ug P/L | 0.0761 | 0.935 | -42.8 |  | 1 | 1 |  | 224 | 224 | 224 | [6 empty] | 1 |  | 7.5241 | 7.5241 | 7.5241
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 33 | 1 | 1 | TN | ug N/L | 0.437 | 5.85 | 222 | TP | ug P/L | 0.07 | 0.875 | -43.2 |  | 1 | 1 |  | 222 | 222 | 222 | [6 empty] | 1 |  | 6.6229 | 6.6229 | 6.6229
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 34 | 1 | 1 | TN | ug N/L | 0.446 | 5.92 | 225 | TP | ug P/L | 0.0694 | 0.862 | -43.2 |  | 1 | 1 |  | 225 | 225 | 225 | [6 empty] | 1 |  | 6.4353 | 6.4353 | 6.4353
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s1 | 35 | 1 | 1 | TN | ug N/L | 1.3 | 17.5 | 792 | TP | ug P/L | 0.0796 | 1.23 | -42.6 |  | 1 | 1 |  | 792 | 792 | 792 | [6 empty] | 1 |  | 12.7936 | 12.7936 | 12.7936
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s2 | 36 | 1 | 1 | TN | ug N/L | 0.538 | 7.16 | 285 | TP | ug P/L | 0.0657 | 0.687 | -43.5 |  | 1 | 1 |  | 285 | 285 | 285 | [6 empty] | 1 |  | 4.1728 | 4.1728 | 4.1728
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s3 | 37 | 1 | 1 | TN | ug N/L | 0.56 | 7.54 | 304 | TP | ug P/L | 0.0706 | 0.851 | -43.2 |  | 1 | 1 |  | 304 | 304 | 304 | [6 empty] | 1 |  | 6.2786 | 6.2786 | 6.2786
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s4 | 38 | 1 | 1 | TN | ug N/L | 0.958 | 12.8 | 560 | TP | ug P/L | 0.0735 | 0.933 | -43 |  | 1 | 1 |  | 560 | 560 | 560 | [6 empty] | 1 |  | 7.4931 | 7.4931 | 7.4931
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s5 | 39 | 1 | 1 | TN | ug N/L | 3.36 | 45.1 | 2170 | TP | ug P/L | 0.0871 | 1.4 | -42.1 |  | 1 | 1 |  | 2170 | 2170 | 2170 | [6 empty] | 1 |  | 16.4634 | 16.4634 | 16.4634
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s6 | 40 | 1 | 1 | TN | ug N/L | 0.931 | 12.5 | 548 | TP | ug P/L | 0.0798 | 1.12 | -42.6 |  | 1 | 1 |  | 548 | 548 | 548 | [6 empty] | 1 |  | 10.6657 | 10.6657 | 10.6657
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s7 | 41 | 1 | 1 | TN | ug N/L | 0.886 | 11.9 | 515 | TP | ug P/L | 0.0833 | 1.24 | -42.3 |  | 1 | 1 |  | 515 | 515 | 515 | [6 empty] | 1 |  | 12.9966 | 12.9966 | 12.9966
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s8 | 42 | 1 | 1 | TN | ug N/L | 0.573 | 7.62 | 308 | TP | ug P/L | 0.0679 | 0.744 | -43.3 |  | 1 | 1 |  | 308 | 308 | 308 | [6 empty] | 1 |  | 4.8558 | 4.8558 | 4.8558
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s9 | 43 | 1 | 1 | TN | ug N/L | 1.6 | 21.5 | 988 | TP | ug P/L | 0.072 | 0.806 | -43.1 |  | 1 | 1 |  | 988 | 988 | 988 | [6 empty] | 1 |  | 5.6579 | 5.6579 | 5.6579
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | a10 | 44 | 1 | 1 | TN | ug N/L | 0.764 | 10.2 | 435 | TP | ug P/L | 0.0738 | 0.902 | -43 |  | 1 | 1 |  | 435 | 435 | 435 | [6 empty] | 1 |  | 7.0213 | 7.0213 | 7.0213
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup5 | 45 | 1 | 1 | TN | ug N/L | 3.39 | 45.6 | 2190 | TP | ug P/L | 0.0882 | 1.48 | -42 |  | 1 | 1 |  | 2190 | 2190 | 2190 |  |  | 0.91743 | PASS |  |  | 1 |  | 18.3506 | 18.3506 | 18.3506 |  |  | 10.8415 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk10 | 46 | 1 | 1 | TN | ug N/L | 1.07 | 14.4 | 639 | TP | ug P/L | 0.0958 | 1.64 | -41.5 |  | 1 | 1 |  | 639 | 639 | 639 |  |  |  |  | 87.99 | PASS | 1 |  | 22.4325 | 22.4325 | 22.4325 |  |  |  |  | 63.8882 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.28 | 243 | TP | ug P/L | 0.071 | 0.861 | -43.1 |  | 1 | 1 |  | 243 | 243 | 243 | -2.8 | PASS |  |  |  |  | 1 |  | 6.421 | 6.421 | 6.421 | -74.3162 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00254 | 0.00654 | -60.9 | TP | ug P/L | 0.0496 | 0.646 | -44.5 |  | 1 | 1 |  | -60.9 | -60.9 | -60.9 | [6 empty] | 1 |  | 3.7137 | 3.7137 | 3.7137
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.016 | -0.252 | -73.3 | TP | ug P/L | -0.0104 | -0.111 | -48.4 |  | 1 | 1 |  | -73.3 | -73.3 | -73.3 | [6 empty] | 1 |  | 0.07526 | 0.07526 | 0.07526
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s11 | 47 | 1 | 1 | TN | ug N/L | 1.62 | 21.9 | 1010 | TP | ug P/L | 0.0728 | 0.996 | -43 |  | 1 | 1 |  | 1010 | 1010 | 1010 | [6 empty] | 1 |  | 8.4994 | 8.4994 | 8.4994
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s12 | 48 | 1 | 1 | TN | ug N/L | 0.843 | 11.3 | 485 | TP | ug P/L | 0.0743 | 0.976 | -42.9 |  | 1 | 1 |  | 485 | 485 | 485 | [6 empty] | 1 |  | 8.173 | 8.173 | 8.173
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s13 | 49 | 1 | 1 | TN | ug N/L | 1.6 | 21.5 | 990 | TP | ug P/L | 0.0699 | 0.814 | -43.2 |  | 1 | 1 |  | 990 | 990 | 990 | [6 empty] | 1 |  | 5.7659 | 5.7659 | 5.7659
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s14 | 50 | 1 | 1 | TN | ug N/L | 1.67 | 22.6 | 1040 | TP | ug P/L | 0.0869 | 1.34 | -42.1 |  | 1 | 1 |  | 1040 | 1040 | 1040 | [6 empty] | 1 |  | 15.1153 | 15.1153 | 15.1153
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s15 | 51 | 1 | 1 | TN | ug N/L | 0.54 | 7.23 | 289 | TP | ug P/L | 0.0693 | 0.803 | -43.2 |  | 1 | 1 |  | 289 | 289 | 289 | [6 empty] | 1 |  | 5.6177 | 5.6177 | 5.6177
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s16 | 52 | 1 | 1 | TN | ug N/L | 0.594 | 8.01 | 327 | TP | ug P/L | 0.0666 | 0.673 | -43.4 |  | 1 | 1 |  | 327 | 327 | 327 | [6 empty] | 1 |  | 4.013 | 4.013 | 4.013
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s17 | 53 | 1 | 1 | TN | ug N/L | 0.798 | 10.7 | 457 | TP | ug P/L | 0.0704 | 0.825 | -43.2 |  | 1 | 1 |  | 457 | 457 | 457 | [6 empty] | 1 |  | 5.916 | 5.916 | 5.916
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s18 | 54 | 1 | 1 | TN | ug N/L | 3.33 | 44.7 | 2150 | TP | ug P/L | 0.0853 | 1.25 | -42.2 |  | 1 | 1 |  | 2150 | 2150 | 2150 | [6 empty] | 1 |  | 13.2013 | 13.2013 | 13.2013
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s19 | 55 | 1 | 1 | TN | ug N/L | 2.53 | 34 | 1610 | TP | ug P/L | 0.0824 | 1.2 | -42.4 |  | 1 | 1 |  | 1610 | 1610 | 1610 | [6 empty] | 1 |  | 12.194 | 12.194 | 12.194
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s20 | 56 | 1 | 1 | TN | ug N/L | 2.49 | 33.3 | 1580 | TP | ug P/L | 0.0788 | 1.13 | -42.6 |  | 1 | 1 |  | 1580 | 1580 | 1580 | [6 empty] | 1 |  | 10.8511 | 10.8511 | 10.8511
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup15 | 57 | 1 | 1 | TN | ug N/L | 0.54 | 7.19 | 287 | TP | ug P/L | 0.0682 | 0.797 | -43.3 |  | 1 | 1 |  | 287 | 287 | 287 |  |  | 0.69444 | PASS |  |  | 1 |  | 5.5377 | 5.5377 | 5.5377 |  |  | 1.4347 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk20 | 58 | 1 | 1 | TN | ug N/L | 2.81 | 37.7 | 1800 | TP | ug P/L | 0.105 | 1.9 | -41 |  | 1 | 1 |  | 1800 | 1800 | 1800 |  |  |  |  | 106 | PASS | 1 |  | 29.9404 | 29.9404 | 29.9404 |  |  |  |  | 79.3512 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.35 | 246 | TP | ug P/L | 0.0721 | 0.877 | -43.1 |  | 1 | 1 |  | 246 | 246 | 246 | -1.6 | PASS |  |  |  |  | 1 |  | 6.652 | 6.652 | 6.652 | -73.3918 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00264 | -0.0321 | -62.8 | TP | ug P/L | 0.0491 | 0.679 | -44.5 |  | 1 | 1 |  | -62.8 | -62.8 | -62.8 | [6 empty] | 1 |  | 4.0811 | 4.0811 | 4.0811
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00701 | 0.037 | -59.4 | TP | ug P/L | 0.00901 | 0.151 | -47.1 |  | 1 | 1 |  | -59.4 | -59.4 | -59.4 | [6 empty] | 1 |  | 0.29571 | 0.29571 | 0.29571
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s21 | 59 | 1 | 1 | TN | ug N/L | 0.554 | 7.45 | 299 | TP | ug P/L | 0.0688 | 0.687 | -43.3 |  | 1 | 1 |  | 299 | 299 | 299 | [6 empty] | 1 |  | 4.1728 | 4.1728 | 4.1728
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s22 | 60 | 1 | 1 | TN | ug N/L | 0.736 | 9.94 | 421 | TP | ug P/L | 0.0671 | 0.742 | -43.4 |  | 1 | 1 |  | 421 | 421 | 421 | [6 empty] | 1 |  | 4.831 | 4.831 | 4.831
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s23 | 61 | 1 | 1 | TN | ug N/L | 0.615 | 8.28 | 340 | TP | ug P/L | 0.0712 | 0.88 | -43.1 |  | 1 | 1 |  | 340 | 340 | 340 | [6 empty] | 1 |  | 6.6958 | 6.6958 | 6.6958
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s24 | 62 | 1 | 1 | TN | ug N/L | 1.04 | 14 | 622 | TP | ug P/L | 0.0657 | 0.596 | -43.5 |  | 1 | 1 |  | 622 | 622 | 622 | [6 empty] | 1 |  | 3.1902 | 3.1902 | 3.1902
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s25 | 63 | 1 | 1 | TN | ug N/L | 3.03 | 40.8 | 1950 | TP | ug P/L | 0.1 | 1.72 | -41.2 |  | 1 | 1 |  | 1950 | 1950 | 1950 | [6 empty] | 1 |  | 24.6273 | 24.6273 | 24.6273
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s26 | 64 | 1 | 1 | TN | ug N/L | 1.17 | 15.8 | 708 | TP | ug P/L | 0.0748 | 0.951 | -42.9 |  | 1 | 1 |  | 708 | 708 | 708 | [6 empty] | 1 |  | 7.7741 | 7.7741 | 7.7741
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s27 | 65 | 1 | 1 | TN | ug N/L | 0.5 | 6.71 | 264 | TP | ug P/L | 0.0683 | 0.693 | -43.3 |  | 1 | 1 |  | 264 | 264 | 264 | [6 empty] | 1 |  | 4.2422 | 4.2422 | 4.2422
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s28 | 66 | 1 | 1 | TN | ug N/L | 2.56 | 34.5 | 1630 | TP | ug P/L | 0.144 | 2.72 | -38.4 |  | 1 | 1 |  | 1630 | 1630 | 1630 | [6 empty] | 1 |  | 60.7127 | 60.7127 | 60.7127
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s29 | 67 | 1 | 1 | TN | ug N/L | 2.18 | 29.3 | 1370 | TP | ug P/L | 0.0803 | 1.17 | -42.5 |  | 1 | 1 |  | 1370 | 1370 | 1370 | [6 empty] | 1 |  | 11.6089 | 11.6089 | 11.6089
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s30 | 68 | 1 | 1 | TN | ug N/L | 1.14 | 15.3 | 682 | TP | ug P/L | 0.0756 | 1.01 | -42.8 |  | 1 | 1 |  | 682 | 682 | 682 | [6 empty] | 1 |  | 8.7316 | 8.7316 | 8.7316
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup25 | 69 | 1 | 1 | TN | ug N/L | 3 | 40.5 | 1930 | TP | ug P/L | 0.0993 | 1.71 | -41.3 |  | 1 | 1 |  | 1930 | 1930 | 1930 |  |  | 1.0309 | PASS |  |  | 1 |  | 24.3474 | 24.3474 | 24.3474 |  |  | 1.1433 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup30 | 70 | 1 | 1 | TN | ug N/L | 1.46 | 19.7 | 900 | TP | ug P/L | 0.098 | 1.72 | -41.4 |  | 1 | 1 |  | 900 | 900 | 900 |  |  |  |  | 96.2 | PASS | 1 |  | 24.6273 | 24.6273 | 24.6273 |  |  |  |  | 66.0455 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.469 | 6.28 | 243 | TP | ug P/L | 0.0715 | 0.884 | -43.1 |  | 1 | 1 |  | 243 | 243 | 243 | -2.8 | PASS |  |  |  |  | 1 |  | 6.7544 | 6.7544 | 6.7544 | -72.9823 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00521 | -0.0319 | -62.7 | TP | ug P/L | 0.0511 | 0.69 | -44.4 |  | 1 | 1 |  | -62.7 | -62.7 | -62.7 | [6 empty] | 1 |  | 4.2075 | 4.2075 | 4.2075
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0169 | -0.275 | -74.4 | TP | ug P/L | 0.0095 | 0.164 | -47.1 |  | 1 | 1 |  | -74.4 | -74.4 | -74.4 | [6 empty] | 1 |  | 0.33528 | 0.33528 | 0.33528
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s31 | 71 | 1 | 1 | TN | ug N/L | 0.491 | 6.8 | 268 | TP | ug P/L | 0.0653 | 0.68 | -43.5 |  | 1 | 1 |  | 268 | 268 | 268 | [6 empty] | 1 |  | 4.0925 | 4.0925 | 4.0925
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s32 | 72 | 1 | 1 | TN | ug N/L | 1.13 | 15.2 | 679 | TP | ug P/L | 0.0693 | 0.734 | -43.2 |  | 1 | 1 |  | 679 | 679 | 679 | [6 empty] | 1 |  | 4.7322 | 4.7322 | 4.7322
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s33 | 73 | 1 | 1 | TN | ug N/L | 0.853 | 11.5 | 495 | TP | ug P/L | 0.0791 | 1.08 | -42.6 |  | 1 | 1 |  | 495 | 495 | 495 | [6 empty] | 1 |  | 9.9399 | 9.9399 | 9.9399
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s34 | 74 | 1 | 1 | TN | ug N/L | 1.37 | 18.4 | 837 | TP | ug P/L | 0.0827 | 1.21 | -42.4 |  | 1 | 1 |  | 837 | 837 | 837 | [6 empty] | 1 |  | 12.3923 | 12.3923 | 12.3923
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s35 | 75 | 1 | 1 | TN | ug N/L | 0.745 | 10 | 425 | TP | ug P/L | 0.0705 | 0.828 | -43.2 |  | 1 | 1 |  | 425 | 425 | 425 | [6 empty] | 1 |  | 5.9573 | 5.9573 | 5.9573
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s36 | 76 | 1 | 1 | TN | ug N/L | 1.64 | 22.1 | 1020 | TP | ug P/L | 0.1 | 1.75 | -41.3 |  | 1 | 1 |  | 1020 | 1020 | 1020 | [6 empty] | 1 |  | 25.4768 | 25.4768 | 25.4768
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s37 | 77 | 1 | 1 | TN | ug N/L | 2.57 | 34.9 | 1650 | TP | ug P/L | 0.0822 | 1.26 | -42.4 |  | 1 | 1 |  | 1650 | 1650 | 1650 | [6 empty] | 1 |  | 13.4075 | 13.4075 | 13.4075
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s38 | 78 | 1 | 1 | TN | ug N/L | 0.752 | 10.1 | 428 | TP | ug P/L | 0.0664 | 0.619 | -43.4 |  | 1 | 1 |  | 428 | 428 | 428 | [6 empty] | 1 |  | 3.426 | 3.426 | 3.426
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s39 | 79 | 1 | 1 | TN | ug N/L | 1.17 | 15.8 | 707 | TP | ug P/L | 0.0662 | 0.682 | -43.4 |  | 1 | 1 |  | 707 | 707 | 707 | [6 empty] | 1 |  | 4.1154 | 4.1154 | 4.1154
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s40 | 80 | 1 | 1 | TN | ug N/L | 1.76 | 23.7 | 1100 | TP | ug P/L | 0.0725 | 0.916 | -43 |  | 1 | 1 |  | 1100 | 1100 | 1100 | [6 empty] | 1 |  | 7.2325 | 7.2325 | 7.2325
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup35 | 81 | 1 | 1 | TN | ug N/L | 0.746 | 10.1 | 430 | TP | ug P/L | 0.0691 | 0.712 | -43.3 |  | 1 | 1 |  | 430 | 430 | 430 |  |  | 1.1696 | PASS |  |  | 1 |  | 4.466 | 4.466 | 4.466 |  |  | 28.6156 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk40 | 82 | 1 | 1 | TN | ug N/L | 2.09 | 28.1 | 1310 | TP | ug P/L | 0.1 | 1.74 | -41.3 |  | 1 | 1 |  | 1310 | 1310 | 1310 |  |  |  |  | 97.1 | PASS | 1 |  | 25.192 | 25.192 | 25.192 |  |  |  |  | 74.3574 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.3 | 244 | TP | ug P/L | 0.0725 | 0.92 | -43 |  | 1 | 1 |  | 244 | 244 | 244 | -2.4 | PASS |  |  |  |  | 1 |  | 7.2934 | 7.2934 | 7.2934 | -70.8265 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00528 | -0.0552 | -63.9 | TP | ug P/L | 0.0486 | 0.652 | -44.6 |  | 1 | 1 |  | -63.9 | -63.9 | -63.9 | [6 empty] | 1 |  | 3.7792 | 3.7792 | 3.7792
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0144 | -0.226 | -72.1 | TP | ug P/L | 0.00937 | 0.15 | -47.1 |  | 1 | 1 |  | -72.1 | -72.1 | -72.1 | [6 empty] | 1 |  | 0.29278 | 0.29278 | 0.29278
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s41 | 83 | 1 | 1 | TN | ug N/L | 1.16 | 15.7 | 703 | TP | ug P/L | 0.0666 | 0.719 | -43.4 |  | 1 | 1 |  | 703 | 703 | 703 | [6 empty] | 1 |  | 4.5499 | 4.5499 | 4.5499
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s42 | 84 | 1 | 1 | TN | ug N/L | 0.5 | 6.77 | 266 | TP | ug P/L | 0.067 | 0.674 | -43.4 |  | 1 | 1 |  | 266 | 266 | 266 | [6 empty] | 1 |  | 4.0243 | 4.0243 | 4.0243
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s43 | 85 | 1 | 1 | TN | ug N/L | 1.02 | 13.8 | 610 | TP | ug P/L | 0.0791 | 0.986 | -42.6 |  | 1 | 1 |  | 610 | 610 | 610 | [6 empty] | 1 |  | 8.3354 | 8.3354 | 8.3354
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s44 | 86 | 1 | 1 | TN | ug N/L | 2.55 | 34.5 | 1640 | TP | ug P/L | 0.0796 | 1.08 | -42.6 |  | 1 | 1 |  | 1640 | 1640 | 1640 | [6 empty] | 1 |  | 9.9399 | 9.9399 | 9.9399
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s45 | 87 | 1 | 1 | TN | ug N/L | 1.46 | 19.7 | 901 | TP | ug P/L | 0.0698 | 0.876 | -43.2 |  | 1 | 1 |  | 901 | 901 | 901 | [6 empty] | 1 |  | 6.6375 | 6.6375 | 6.6375
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s46 | 88 | 1 | 1 | TN | ug N/L | 2.42 | 32.6 | 1540 | TP | ug P/L | 0.139 | 2.74 | -38.8 |  | 1 | 1 |  | 1540 | 1540 | 1540 | [6 empty] | 1 |  | 61.5978 | 61.5978 | 61.5978
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s47 | 89 | 1 | 1 | TN | ug N/L | 0.713 | 9.56 | 402 | TP | ug P/L | 0.0683 | 0.766 | -43.3 |  | 1 | 1 |  | 402 | 402 | 402 | [6 empty] | 1 |  | 5.1334 | 5.1334 | 5.1334
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s48 | 90 | 1 | 1 | TN | ug N/L | 0.975 | 13.2 | 581 | TP | ug P/L | 0.0786 | 1.12 | -42.6 |  | 1 | 1 |  | 581 | 581 | 581 | [6 empty] | 1 |  | 10.6657 | 10.6657 | 10.6657
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s49 | 91 | 1 | 1 | TN | ug N/L | 1.1 | 15 | 666 | TP | ug P/L | 0.0713 | 0.807 | -43.1 |  | 1 | 1 |  | 666 | 666 | 666 | [6 empty] | 1 |  | 5.6714 | 5.6714 | 5.6714
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s50 | 92 | 1 | 1 | TN | ug N/L | 0.825 | 11.1 | 478 | TP | ug P/L | 0.0713 | 0.984 | -43.1 |  | 1 | 1 |  | 478 | 478 | 478 | [6 empty] | 1 |  | 8.3028 | 8.3028 | 8.3028
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup45 | 93 | 1 | 1 | TN | ug N/L | 1.45 | 19.7 | 899 | TP | ug P/L | 0.0702 | 0.778 | -43.2 |  | 1 | 1 |  | 899 | 899 | 899 |  |  | 0.22222 | PASS |  |  | 1 |  | 5.2881 | 5.2881 | 5.2881 |  |  | 22.6307 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk50 | 94 | 1 | 1 | TN | ug N/L | 1.2 | 16.2 | 726 | TP | ug P/L | 0.101 | 1.85 | -41.2 |  | 1 | 1 |  | 726 | 726 | 726 |  |  |  |  | 106.46 | PASS | 1 |  | 28.4125 | 28.4125 | 28.4125 |  |  |  |  | 83.28 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.471 | 6.35 | 246 | TP | ug P/L | 0.0719 | 0.911 | -43.1 |  | 1 | 1 |  | 246 | 246 | 246 | -1.6 | PASS |  |  |  |  | 1 |  | 7.1567 | 7.1567 | 7.1567 | -71.3732 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00351 | -0.0324 | -62.8 | TP | ug P/L | 0.0546 | 0.756 | -44.2 |  | 1 | 1 |  | -62.8 | -62.8 | -62.8 | [6 empty] | 1 |  | 5.0063 | 5.0063 | 5.0063
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.016 | -0.27 | -74.2 | TP | ug P/L | 0.00692 | 0.093 | -47.3 |  | 1 | 1 |  | -74.2 | -74.2 | -74.2 | [6 empty] | 1 |  | 0.15213 | 0.15213 | 0.15213
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal1500 | 1 | 1 | 1 | TN | ug N/L | 2.37 | 31.9 | 1500 | TP | ug P/L | 0.258 | 4.53 | -31.3 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 1 |  | 166.7669 | 166.7669 | 166.7669
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal1000 | 3 | 1 | 1 | TN | ug N/L | 1.62 | 21.9 | 1010 | TP | ug P/L | 0.157 | 3.07 | -37.6 |  | 1 | 1 |  | 1010 | 1010 | 1010 | [6 empty] | 1 |  | 77.1273 | 77.1273 | 77.1273
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call500 | 5 | 1 | 1 | TN | ug N/L | 0.848 | 11.5 | 495 | TP | ug P/L | 0.0965 | 1.81 | -41.5 |  | 1 | 1 |  | 495 | 495 | 495 | [6 empty] | 1 |  | 27.219 | 27.219 | 27.219
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call250 | 7 | 1 | 1 | TN | ug N/L | 0.471 | 6.35 | 246 | TP | ug P/L | 0.0697 | 0.881 | -43.2 |  | 1 | 1 |  | 246 | 246 | 246 | [6 empty] | 1 |  | 6.7105 | 6.7105 | 6.7105 | -73.1582 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal100 | 9 | 1 | 1 | TN | ug N/L | 0.254 | 3.45 | 105 | TP | ug P/L | 0.0595 | 0.468 | -43.9 |  | 1 | 1 |  | 105 | 105 | 105 | [6 empty] | 1 |  | 2.0326 | 2.0326 | 2.0326
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.172 | 2.35 | 52.2 | TP | ug P/L | 0.0563 | 0.351 | -44.1 |  | 1 | 1 |  | 52.2 | 52.2 | 52.2 | [6 empty] | 1 |  | 1.2041 | 1.2041 | 1.2041
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.135 | 1.81 | 26.3 | TP | ug P/L | 0.0576 | 0.817 | -44 |  | 1 | 1 |  | 26.3 | 26.3 | 26.3 | [6 empty] | 1 |  | 5.8066 | 5.8066 | 5.8066
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0958 | 1.29 | 1.07 | TP | ug P/L | 0.056 | 0.797 | -44.1 |  | 1 | 1 |  | 1.07 | 1.07 | 1.07 | [6 empty] | 1 |  | 5.5377 | 5.5377 | 5.5377
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrite | S1 | 1 | 1 | TN | ug N/L | 2.19 | 30.1 | 1420 | TP | ug P/L | -0.0089 | -0.0746 | -48.3 | 99.6678 | 1 | 1 |  | 1420 | 1420 | 1420 | [6 empty] | 1 |  | 0.04011 | 0.04011 | 0.04011
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrate | S2 | 1 | 1 | TN | ug N/L | 2.21 | 30 | 1410 | TP | ug P/L | -0.0099 | -0.0977 | -48.4 |  | 1 | 1 |  | 1410 | 1410 | 1410 | [6 empty] | 1 |  | 0.059956 | 0.059956 | 0.059956
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s51 | 95 | 1 | 1 | TN | ug N/L | 0.609 | 8.22 | 337 | TP | ug P/L | 0.0687 | 0.784 | -43.3 |  | 1 | 1 |  | 337 | 337 | 337 | [6 empty] | 1 |  | 5.3663 | 5.3663 | 5.3663
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s52 | 96 | 1 | 1 | TN | ug N/L | 0.837 | 11.3 | 486 | TP | ug P/L | 0.0662 | 0.673 | -43.4 |  | 1 | 1 |  | 486 | 486 | 486 | [6 empty] | 1 |  | 4.013 | 4.013 | 4.013
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s53 | 97 | 1 | 1 | TN | ug N/L | 4.93 | 67.1 | 3290 | TP | ug P/L | 0.0989 | 1.77 | -41.3 |  | 1 | 1 |  | 3290 | 3290 | 3290 | [6 empty] | 1 |  | 26.0511 | 26.0511 | 26.0511
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s54 | 98 | 1 | 1 | TN | ug N/L | 0.626 | 8.5 | 351 | TP | ug P/L | 0.0718 | 0.86 | -43.1 |  | 1 | 1 |  | 351 | 351 | 351 | [6 empty] | 1 |  | 6.4067 | 6.4067 | 6.4067
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s55 | 99 | 1 | 1 | TN | ug N/L | 0.728 | 9.89 | 418 | TP | ug P/L | 0.0711 | 0.805 | -43.1 |  | 1 | 1 |  | 418 | 418 | 418 | [6 empty] | 1 |  | 5.6445 | 5.6445 | 5.6445
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s56 | 100 | 1 | 1 | TN | ug N/L | 1.13 | 15.4 | 686 | TP | ug P/L | 0.0679 | 0.731 | -43.3 |  | 1 | 1 |  | 686 | 686 | 686 | [6 empty] | 1 |  | 4.6955 | 4.6955 | 4.6955
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s57 | 101 | 1 | 1 | TN | ug N/L | 2.45 | 33.1 | 1560 | TP | ug P/L | 0.0812 | 1.23 | -42.5 |  | 1 | 1 |  | 1560 | 1560 | 1560 | [6 empty] | 1 |  | 12.7936 | 12.7936 | 12.7936
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s58 | 102 | 1 | 1 | TN | ug N/L | 2.77 | 37.4 | 1780 | TP | ug P/L | 0.0885 | 1.41 | -42 |  | 1 | 1 |  | 1780 | 1780 | 1780 | [6 empty] | 1 |  | 16.6937 | 16.6937 | 16.6937
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s59 | 103 | 1 | 1 | TN | ug N/L | 1.08 | 14.7 | 654 | TP | ug P/L | 0.0682 | 0.79 | -43.3 |  | 1 | 1 |  | 654 | 654 | 654 | [6 empty] | 1 |  | 5.445 | 5.445 | 5.445
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s61 | 104 | 1 | 1 | TN | ug N/L | 3.12 | 42.4 | 2030 | TP | ug P/L | 0.0997 | 1.75 | -41.3 |  | 1 | 1 |  | 2030 | 2030 | 2030 | [6 empty] | 1 |  | 25.4768 | 25.4768 | 25.4768
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup55 | 105 | 1 | 1 | TN | ug N/L | 0.729 | 9.84 | 416 | TP | ug P/L | 0.0701 | 0.797 | -43.2 |  | 1 | 1 |  | 416 | 416 | 416 |  |  | 0.47962 | PASS |  |  | 1 |  | 5.5377 | 5.5377 | 5.5377 |  |  | 1.9106 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk60 | 106 | 1 | 1 | TN | ug N/L | 3.47 | 47.2 | 2280 | TP | ug P/L | 0.136 | 2.71 | -39 |  | 1 | 1 |  | 2280 | 2280 | 2280 |  |  |  |  | 122.8 | FAIL | 1 |  | 60.2726 | 60.2726 | 60.2726 |  |  |  |  | 145.2104 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.462 | 6.23 | 240 | TP | ug P/L | 0.0707 | 0.884 | -43.2 |  | 1 | 1 |  | 240 | 240 | 240 | -4 | PASS |  |  |  |  | 1 |  | 6.7544 | 6.7544 | 6.7544 | -72.9823 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00622 | -0.0368 | -63 | TP | ug P/L | 0.0521 | 0.724 | -44.4 |  | 1 | 1 |  | -63 | -63 | -63 | [6 empty] | 1 |  | 4.6103 | 4.6103 | 4.6103
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0176 | -0.254 | -73.4 | TP | ug P/L | 0.00727 | 0.147 | -47.3 |  | 1 | 1 |  | -73.4 | -73.4 | -73.4 | [6 empty] | 1 |  | 0.28408 | 0.28408 | 0.28408
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s61 | 107 | 1 | 1 | TN | ug N/L | 1.51 | 20.4 | 936 | TP | ug P/L | 0.0784 | 1.07 | -42.7 |  | 1 | 1 |  | 936 | 936 | 936 | [6 empty] | 1 |  | 9.7625 | 9.7625 | 9.7625
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s62 | 108 | 1 | 1 | TN | ug N/L | 1.14 | 15.5 | 692 | TP | ug P/L | 0.0682 | 0.734 | -43.3 |  | 1 | 1 |  | 692 | 692 | 692 | [6 empty] | 1 |  | 4.7322 | 4.7322 | 4.7322
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s63 | 109 | 1 | 1 | TN | ug N/L | 0.577 | 7.82 | 317 | TP | ug P/L | 0.0681 | 0.777 | -43.3 |  | 1 | 1 |  | 317 | 317 | 317 | [6 empty] | 1 |  | 5.2751 | 5.2751 | 5.2751
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s64 | 110 | 1 | 1 | TN | ug N/L | 1.56 | 21.2 | 971 | TP | ug P/L | 0.0731 | 0.866 | -43 |  | 1 | 1 |  | 971 | 971 | 971 | [6 empty] | 1 |  | 6.4927 | 6.4927 | 6.4927
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s65 | 111 | 1 | 1 | TN | ug N/L | 4.36 | 59.2 | 2890 | TP | ug P/L | 0.0874 | 1.36 | -42.1 |  | 1 | 1 |  | 2890 | 2890 | 2890 | [6 empty] | 1 |  | 15.5583 | 15.5583 | 15.5583
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s66 | 112 | 1 | 1 | TN | ug N/L | 2.17 | 29.5 | 1380 | TP | ug P/L | 0.0812 | 1.19 | -42.5 |  | 1 | 1 |  | 1380 | 1380 | 1380 | [6 empty] | 1 |  | 11.9974 | 11.9974 | 11.9974
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s67 | 113 | 1 | 1 | TN | ug N/L | 1.03 | 13.9 | 614 | TP | ug P/L | 0.0733 | 0.896 | -43 |  | 1 | 1 |  | 614 | 614 | 614 | [6 empty] | 1 |  | 6.9318 | 6.9318 | 6.9318
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s68 | 114 | 1 | 1 | TN | ug N/L | 3.14 | 42.6 | 2040 | TP | ug P/L | 0.101 | 1.77 | -41.2 |  | 1 | 1 |  | 2040 | 2040 | 2040 | [6 empty] | 1 |  | 26.0511 | 26.0511 | 26.0511
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s69 | 115 | 1 | 1 | TN | ug N/L | 0.668 | 9.05 | 378 | TP | ug P/L | 0.0719 | 0.815 | -43.1 |  | 1 | 1 |  | 378 | 378 | 378 | [6 empty] | 1 |  | 5.7795 | 5.7795 | 5.7795
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s70 | 116 | 1 | 1 | TN | ug N/L | 1.88 | 25.7 | 1190 | TP | ug P/L | 0.0843 | 1.27 | -42.3 |  | 1 | 1 |  | 1190 | 1190 | 1190 | [6 empty] | 1 |  | 13.6154 | 13.6154 | 13.6154
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup65 | 117 | 1 | 1 | TN | ug N/L | 4.38 | 59.6 | 2910 | TP | ug P/L | 0.0893 | 1.44 | -42 |  | 1 | 1 |  | 2910 | 2910 | 2910 |  |  | 0.68966 | PASS |  |  | 1 |  | 17.3942 | 17.3942 | 17.3942 |  |  | 11.1428 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk70 | 118 | 1 | 1 | TN | ug N/L | 2.14 | 29 | 1360 | TP | ug P/L | 0.107 | 1.95 | -40.8 |  | 1 | 1 |  | 1360 | 1360 | 1360 |  |  |  |  | 81.6 | PASS | 1 |  | 31.5084 | 31.5084 | 31.5084 |  |  |  |  | 74.7227 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.459 | 6.21 | 239 | TP | ug P/L | 0.0699 | 0.891 | -43.2 |  | 1 | 1 |  | 239 | 239 | 239 | -4.4 | PASS |  |  |  |  | 1 |  | 6.8576 | 6.8576 | 6.8576 | -72.5696 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00409 | -0.0102 | -61.7 | TP | ug P/L | 0.0543 | 0.74 | -44.2 |  | 1 | 1 |  | -61.7 | -61.7 | -61.7 | [6 empty] | 1 |  | 4.8062 | 4.8062 | 4.8062
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0154 | -0.262 | -73.9 | TP | ug P/L | 0.00746 | 0.124 | -47.2 |  | 1 | 1 |  | -73.9 | -73.9 | -73.9 | [6 empty] | 1 |  | 0.22217 | 0.22217 | 0.22217
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s71 | 119 | 1 | 1 | TN | ug N/L | 1.83 | 24.9 | 1150 | TP | ug P/L | 0.0893 | 1.45 | -42 |  | 1 | 1 |  | 1150 | 1150 | 1150 | [6 empty] | 1 |  | 17.6309 | 17.6309 | 17.6309
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s72 | 120 | 1 | 1 | TN | ug N/L | 0.94 | 12.8 | 558 | TP | ug P/L | 0.0768 | 1.07 | -42.8 |  | 1 | 1 |  | 558 | 558 | 558 | [6 empty] | 1 |  | 9.7625 | 9.7625 | 9.7625
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s73 | 121 | 1 | 1 | TN | ug N/L | 1.17 | 16 | 718 | TP | ug P/L | 0.0724 | 0.985 | -43 |  | 1 | 1 |  | 718 | 718 | 718 | [6 empty] | 1 |  | 8.3191 | 8.3191 | 8.3191
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s74 | 122 | 1 | 1 | TN | ug N/L | 1.25 | 16.9 | 762 | TP | ug P/L | 0.0769 | 1.04 | -42.8 |  | 1 | 1 |  | 762 | 762 | 762 | [6 empty] | 1 |  | 9.2399 | 9.2399 | 9.2399
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s75 | 123 | 1 | 1 | TN | ug N/L | 1.66 | 22.6 | 1040 | TP | ug P/L | 0.095 | 1.58 | -41.6 |  | 1 | 1 |  | 1040 | 1040 | 1040 | [6 empty] | 1 |  | 20.8537 | 20.8537 | 20.8537
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s76 | 124 | 1 | 1 | TN | ug N/L | 1.14 | 15.5 | 691 | TP | ug P/L | 0.0704 | 0.834 | -43.2 |  | 1 | 1 |  | 691 | 691 | 691 | [6 empty] | 1 |  | 6.0403 | 6.0403 | 6.0403
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s77 | 125 | 1 | 1 | TN | ug N/L | 1.87 | 25.5 | 1190 | TP | ug P/L | 0.122 | 2.28 | -39.9 |  | 1 | 1 |  | 1190 | 1190 | 1190 | [6 empty] | 1 |  | 42.8615 | 42.8615 | 42.8615
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s78 | 126 | 1 | 1 | TN | ug N/L | 1.34 | 18.1 | 819 | TP | ug P/L | 0.0775 | 1.04 | -42.7 |  | 1 | 1 |  | 819 | 819 | 819 | [6 empty] | 1 |  | 9.2399 | 9.2399 | 9.2399
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s79 | 127 | 1 | 1 | TN | ug N/L | 1.01 | 13.8 | 607 | TP | ug P/L | 0.0676 | 0.658 | -43.4 |  | 1 | 1 |  | 607 | 607 | 607 | [6 empty] | 1 |  | 3.8453 | 3.8453 | 3.8453
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s80 | 128 | 1 | 1 | TN | ug N/L | 3 | 40.8 | 1950 | TP | ug P/L | 0.0868 | 1.31 | -42.1 |  | 1 | 1 |  | 1950 | 1950 | 1950 | [6 empty] | 1 |  | 14.4629 | 14.4629 | 14.4629
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup75 | 129 | 1 | 1 | TN | ug N/L | 1.93 | 26.2 | 1220 | TP | ug P/L | 0.121 | 2.28 | -39.9 |  | 1 | 1 |  | 1220 | 1220 | 1220 |  |  | 15.9292 | PASS |  |  | 1 |  | 42.8615 | 42.8615 | 42.8615 |  |  | 69.0815 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk80 | 130 | 1 | 1 | TN | ug N/L | 3.25 | 44.1 | 2120 | TP | ug P/L | 0.107 | 1.99 | -40.9 |  | 1 | 1 |  | 2120 | 2120 | 2120 |  |  |  |  | 89.2 | PASS | 1 |  | 32.7916 | 32.7916 | 32.7916 |  |  |  |  | 76.594 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.472 | 6.36 | 247 | TP | ug P/L | 0.0729 | 0.883 | -43 |  | 1 | 1 |  | 247 | 247 | 247 | -1.2 | PASS |  |  |  |  | 1 |  | 6.7398 | 6.7398 | 6.7398 | -73.041 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00739 | -0.0299 | -62.6 | TP | ug P/L | 0.0494 | 0.671 | -44.5 |  | 1 | 1 |  | -62.6 | -62.6 | -62.6 | [6 empty] | 1 |  | 3.9904 | 3.9904 | 3.9904
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0154 | -0.231 | -72.4 | TP | ug P/L | 0.00635 | 0.0879 | -47.3 |  | 1 | 1 |  | -72.4 | -72.4 | -72.4 | [6 empty] | 1 |  | 0.14208 | 0.14208 | 0.14208
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s81 | 131 | 1 | 1 | TN | ug N/L | 1.14 | 15.6 | 695 | TP | ug P/L | 0.0713 | 0.81 | -43.1 |  | 1 | 1 |  | 695 | 695 | 695 | [6 empty] | 1 |  | 5.7118 | 5.7118 | 5.7118
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s82 | 132 | 1 | 1 | TN | ug N/L | 0.583 | 7.92 | 322 | TP | ug P/L | 0.0787 | 1.09 | -42.6 |  | 1 | 1 |  | 322 | 322 | 322 | [6 empty] | 1 |  | 10.119 | 10.119 | 10.119
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s83 | 133 | 1 | 1 | TN | ug N/L | 1.14 | 15.4 | 690 | TP | ug P/L | 0.0685 | 0.697 | -43.3 |  | 1 | 1 |  | 690 | 690 | 690 | [6 empty] | 1 |  | 4.2889 | 4.2889 | 4.2889
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s84 | 134 | 1 | 1 | TN | ug N/L | 1.42 | 19.3 | 880 | TP | ug P/L | 0.0886 | 1.42 | -42 |  | 1 | 1 |  | 880 | 880 | 880 | [6 empty] | 1 |  | 16.9256 | 16.9256 | 16.9256
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s85 | 135 | 1 | 1 | TN | ug N/L | 0.496 | 6.71 | 263 | TP | ug P/L | 0.0659 | 0.614 | -43.5 |  | 1 | 1 |  | 263 | 263 | 263 | [6 empty] | 1 |  | 3.374 | 3.374 | 3.374
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s86 | 136 | 1 | 1 | TN | ug N/L | 0.56 | 7.61 | 307 | TP | ug P/L | 0.0753 | 1.02 | -42.9 |  | 1 | 1 |  | 307 | 307 | 307 | [6 empty] | 1 |  | 8.8994 | 8.8994 | 8.8994
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s87 | 137 | 1 | 1 | TN | ug N/L | 0.629 | 8.57 | 354 | TP | ug P/L | 0.0731 | 0.885 | -43 |  | 1 | 1 |  | 354 | 354 | 354 | [6 empty] | 1 |  | 6.7691 | 6.7691 | 6.7691
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s88 | 138 | 1 | 1 | TN | ug N/L | 0.546 | 7.39 | 297 | TP | ug P/L | 0.0759 | 0.962 | -42.8 |  | 1 | 1 |  | 297 | 297 | 297 | [6 empty] | 1 |  | 7.9484 | 7.9484 | 7.9484
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s89 | 139 | 1 | 1 | TN | ug N/L | 1.16 | 15.8 | 708 | TP | ug P/L | 0.0786 | 1.07 | -42.6 |  | 1 | 1 |  | 708 | 708 | 708 | [6 empty] | 1 |  | 9.7625 | 9.7625 | 9.7625
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s90 | 140 | 1 | 1 | TN | ug N/L | 0.613 | 8.31 | 341 | TP | ug P/L | 0.0661 | 0.638 | -43.5 |  | 1 | 1 |  | 341 | 341 | 341 | [6 empty] | 1 |  | 3.6272 | 3.6272 | 3.6272
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup85 | 141 | 1 | 1 | TN | ug N/L | 0.503 | 6.8 | 268 | TP | ug P/L | 0.0709 | 0.675 | -43.1 |  | 1 | 1 |  | 268 | 268 | 268 |  |  | 1.8832 | PASS |  |  | 1 |  | 4.0356 | 4.0356 | 4.0356 |  |  | 17.8577 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk90 | 142 | 1 | 1 | TN | ug N/L | 0.961 | 13 | 571 | TP | ug P/L | 0.0892 | 1.38 | -42 |  | 1 | 1 |  | 571 | 571 | 571 |  |  |  |  | 97.71 | PASS | 1 |  | 16.0076 | 16.0076 | 16.0076 |  |  |  |  | 51.1224 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.466 | 6.36 | 247 | TP | ug P/L | 0.072 | 0.878 | -43.1 |  | 1 | 1 |  | 247 | 247 | 247 | -1.2 | PASS |  |  |  |  | 1 |  | 6.6666 | 6.6666 | 6.6666 | -73.3335 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.0081 | 0.0293 | -59.8 | TP | ug P/L | 0.0503 | 0.665 | -44.5 |  | 1 | 1 |  | -59.8 | -59.8 | -59.8 | [6 empty] | 1 |  | 3.9231 | 3.9231 | 3.9231
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00635 | 0.0636 | -58.1 | TP | ug P/L | 0.00846 | 0.133 | -47.2 |  | 1 | 1 |  | -58.1 | -58.1 | -58.1 | [6 empty] | 1 |  | 0.24538 | 0.24538 | 0.24538
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s91 | 143 | 1 | 1 | TN | ug N/L | 1.83 | 24.9 | 1150 | TP | ug P/L | 0.0965 | 1.66 | -41.5 |  | 1 | 1 |  | 1150 | 1150 | 1150 | [6 empty] | 1 |  | 22.9716 | 22.9716 | 22.9716
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s92 | 144 | 1 | 1 | TN | ug N/L | 3.77 | 51.3 | 2480 | TP | ug P/L | 0.0855 | 1.29 | -42.2 |  | 1 | 1 |  | 2480 | 2480 | 2480 | [6 empty] | 1 |  | 14.0359 | 14.0359 | 14.0359
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s93 | 145 | 1 | 1 | TN | ug N/L | 0.522 | 7.09 | 282 | TP | ug P/L | 0.0714 | 0.795 | -43.1 |  | 1 | 1 |  | 282 | 282 | 282 | [6 empty] | 1 |  | 5.5111 | 5.5111 | 5.5111
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s94 | 146 | 1 | 1 | TN | ug N/L | 3.47 | 47 | 2270 | TP | ug P/L | 0.0877 | 1.32 | -42.1 |  | 1 | 1 |  | 2270 | 2270 | 2270 | [6 empty] | 1 |  | 14.6787 | 14.6787 | 14.6787
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s95 | 147 | 1 | 1 | TN | ug N/L | 1.43 | 19.3 | 881 | TP | ug P/L | 0.0808 | 1.1 | -42.5 |  | 1 | 1 |  | 881 | 881 | 881 | [6 empty] | 1 |  | 10.2996 | 10.2996 | 10.2996
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s96 | 148 | 1 | 1 | TN | ug N/L | 0.643 | 8.69 | 360 | TP | ug P/L | 0.0679 | 0.774 | -43.3 |  | 1 | 1 |  | 360 | 360 | 360 | [6 empty] | 1 |  | 5.2363 | 5.2363 | 5.2363
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s97 | 149 | 1 | 1 | TN | ug N/L | 0.752 | 10.2 | 432 | TP | ug P/L | 0.0703 | 0.823 | -43.2 |  | 1 | 1 |  | 432 | 432 | 432 | [6 empty] | 1 |  | 5.8886 | 5.8886 | 5.8886
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s98 | 150 | 1 | 1 | TN | ug N/L | 0.673 | 9.07 | 379 | TP | ug P/L | 0.075 | 0.964 | -42.9 |  | 1 | 1 |  | 379 | 379 | 379 | [6 empty] | 1 |  | 7.9803 | 7.9803 | 7.9803
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s99 | 151 | 1 | 1 | TN | ug N/L | 0.665 | 9.02 | 376 | TP | ug P/L | 0.0677 | 0.618 | -43.3 |  | 1 | 1 |  | 376 | 376 | 376 | [6 empty] | 1 |  | 3.4156 | 3.4156 | 3.4156
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s100 | 152 | 1 | 1 | TN | ug N/L | 1.32 | 17.8 | 808 | TP | ug P/L | 0.0732 | 0.872 | -43 |  | 1 | 1 |  | 808 | 808 | 808 | [6 empty] | 1 |  | 6.5794 | 6.5794 | 6.5794
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup95 | 153 | 1 | 1 | TN | ug N/L | 1.47 | 19.8 | 905 | TP | ug P/L | 0.0787 | 1.1 | -42.6 |  | 1 | 1 |  | 905 | 905 | 905 |  |  | 2.6876 | PASS |  |  | 1 |  | 10.2996 | 10.2996 | 10.2996 |  |  | 0 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk100 | 154 | 1 | 1 | TN | ug N/L | 1.66 | 22.4 | 1030 | TP | ug P/L | 0.102 | 1.67 | -41.1 |  | 1 | 1 |  | 1030 | 1030 | 1030 |  |  |  |  | 99.1 | PASS | 1 |  | 23.2436 | 23.2436 | 23.2436 |  |  |  |  | 68.981 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.473 | 6.38 | 248 | TP | ug P/L | 0.0701 | 0.872 | -43.2 |  | 1 | 1 |  | 248 | 248 | 248 | -0.8 | PASS |  |  |  |  | 1 |  | 6.5794 | 6.5794 | 6.5794 | -73.6824 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00967 | 0.041 | -59.2 | TP | ug P/L | 0.0522 | 0.707 | -44.3 |  | 1 | 1 |  | -59.2 | -59.2 | -59.2 | [6 empty] | 1 |  | 4.4065 | 4.4065 | 4.4065
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00841 | 0.057 | -58.5 | TP | ug P/L | 0.00661 | 0.072 | -47.3 |  | 1 | 1 |  | -58.5 | -58.5 | -58.5 | [6 empty] | 1 |  | 0.11344 | 0.11344 | 0.11344
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s101 | 155 | 1 | 1 | TN | ug N/L | 0.534 | 7.24 | 289 | TP | ug P/L | 0.065 | 0.694 | -43.5 |  | 1 | 1 |  | 289 | 289 | 289 | [6 empty] | 1 |  | 4.2539 | 4.2539 | 4.2539
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s102 | 156 | 1 | 1 | TN | ug N/L | 0.787 | 10.7 | 460 | TP | ug P/L | 0.0679 | 0.719 | -43.3 |  | 1 | 1 |  | 460 | 460 | 460 | [6 empty] | 1 |  | 4.5499 | 4.5499 | 4.5499
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s103 | 157 | 1 | 1 | TN | ug N/L | 1.22 | 16.6 | 747 | TP | ug P/L | 0.0813 | 1.25 | -42.5 |  | 1 | 1 |  | 747 | 747 | 747 | [6 empty] | 1 |  | 13.2013 | 13.2013 | 13.2013
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s104 | 158 | 1 | 1 | TN | ug N/L | 0.545 | 9.26 | 387 | TP | ug P/L | 0.908 | 24.3 | -42.3 |  | 1 | 1 |  | 387 | 387 | 387 | [6 empty] | 1 |  | 4742.5196 | 4742.5196 | 4742.5196
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s105 | 159 | 1 | 1 | TN | ug N/L | 0.016 | 0.0587 | -58.4 | TP | ug P/L | 0.124 | 0.672 | -39.8 |  | 1 | 3 | air spikes | -58.4 | -58.4 | -58.4 | [6 empty] | 3 | air spikes | 4.0017 | 4.0017 | 4.0017
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s106 | 160 | 1 | 1 | TN | ug N/L | 3.28 | 36.6 | 1740 | TP | ug P/L | 0.223 | 2.26 | -33.5 |  | 1 | 3 | air spikes | 1740 | 1740 | 1740 | [6 empty] | 3 | air spikes | 42.1237 | 42.1237 | 42.1237
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s107 | 161 | 1 | 1 | TN | ug N/L | -2.01 | -50.9 | -2450 | TP | ug P/L | 0.265 | 1.99 | -30.9 |  | 1 | 3 | air spikes | -2450 | -2450 | -2450 | [6 empty] | 3 | air spikes | 32.7916 | 32.7916 | 32.7916
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dups105 | 162 | 1 | 1 | TN | ug N/L | -2.1 | -59.8 | -2850 | TP | ug P/L | -0.171 | -1.89 | -59 |  | 1 | 3 | air spikes | -2850 | -2850 | -2850 |  |  | 276.9874 | FAIL |  |  | 3 | air spikes | 27.6622 | 27.6622 | 27.6622 |  |  | 143.5013 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spks107 | 163 | 1 | 1 | TN | ug N/L | 3.44 | 7.38 | 296 | TP | ug P/L | 0.29 | 2.64 | -29.3 |  | 1 | 3 | air spikes | 296 | 296 | 296 |  |  |  |  | 1101.36 | FAIL | 3 | air spikes | 57.2363 | 57.2363 | 57.2363 |  |  |  |  | 103.5026 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s108 | 164 | 1 | 1 | TN | ug N/L | -2.83 | -83.1 | -3890 | TP | ug P/L | 0.161 | 0.0967 | -37.4 |  | 1 | 3 | air spikes | -3890 | -3890 | -3890 | -1656 | FAIL |  |  |  |  | 3 | air spikes | 0.15968 | 0.15968 | 0.15968 | -99.3613 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 1.78 | -6.12 | -356 | TP | ug P/L | 0.119 | 1.84 | -40 |  | 1 | 3 | air spikes | -356 | -356 | -356 | [6 empty] | 3 | air spikes | 28.1117 | 28.1117 | 28.1117
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.788 | -30.1 | -1490 | TP | ug P/L | -0.114 | -1.53 | -55.2 |  | 1 | 3 | air spikes | -1490 | -1490 | -1490 | [6 empty] | 3 | air spikes | 17.9879 | 17.9879 | 17.9879
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.77 | -12.6 | -665 | TP | ug P/L | 0.0697 | -0.315 | -43.2 |  | 1 | 3 | air spikes | -665 | -665 | -665 | [6 empty] | 3 | air spikes | 0.66508 | 0.66508 | 0.66508
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal1500 | 1 | 1 | 1 | TN | ug N/L | 2.91 | 4.73 | 167 | TP | ug P/L | 0.348 | 5.62 | -25.7 |  | 1 | 3 | air spikes | 167 | 167 | 167 | [6 empty] | 3 | air spikes | 255.9535 | 255.9535 | 255.9535
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal1000 | 3 | 1 | 1 | TN | ug N/L | -2.58 | -57 | -2720 | TP | ug P/L | 0.232 | 4.25 | -32.9 |  | 1 | 3 | air spikes | -2720 | -2720 | -2720 | [6 empty] | 3 | air spikes | 146.9293 | 146.9293 | 146.9293
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call500 | 5 | 1 | 1 | TN | ug N/L | 1.98 | 3.67 | 116 | TP | ug P/L | 0.253 | 2.78 | -31.6 |  | 1 | 3 | air spikes | 116 | 116 | 116 | [6 empty] | 3 | air spikes | 63.3873 | 63.3873 | 63.3873
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call250 | 7 | 1 | 1 | TN | ug N/L | 1.61 | -10.7 | -572 | TP | ug P/L | 0.115 | 1.85 | -40.3 |  | 1 | 3 | air spikes | -572 | -572 | -572 | [6 empty] | 3 | air spikes | 28.4125 | 28.4125 | 28.4125
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal100 | 9 | 1 | 1 | TN | ug N/L | -1.98 | -22.1 | -1110 | TP | ug P/L | -0.0785 | -1.33 | -52.9 |  | 1 | 3 | air spikes | -1110 | -1110 | -1110 | [6 empty] | 3 | air spikes | 13.5104 | 13.5104 | 13.5104
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | 2 | 3.26 | 96.1 | TP | ug P/L | -0.0457 | -0.32 | -50.7 |  | 1 | 3 | air spikes | 96.1 | 96.1 | 96.1 | [6 empty] | 3 | air spikes | 0.6879 | 0.6879 | 0.6879
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | 4.66 | 29.1 | 1370 | TP | ug P/L | 0.119 | -0.0565 | -40.1 |  | 1 | 3 | air spikes | 1370 | 1370 | 1370 | [6 empty] | 3 | air spikes | 0.030533 | 0.030533 | 0.030533
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | 2.46 | 14.2 | 630 | TP | ug P/L | -0.00144 | 0.0124 | -47.8 |  | 1 | 3 | air spikes | 630 | 630 | 630 | [6 empty] | 3 | air spikes | 0.042092 | 0.042092 | 0.042092
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | 2.95 | 12 | 520 | TP | ug P/L | -0.052 | -0.591 | -51.1 |  | 1 | 3 | air spikes | 520 | 520 | 520 | [6 empty] | 3 | air spikes | 2.5242 | 2.5242 | 2.5242
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | 9.64 | 69.9 | 3440 | TP | ug P/L | -0.0399 | -0.284 | -50.3 |  | 1 | 3 | air spikes | 3440 | 3440 | 3440 | [6 empty] | 3 | air spikes | 0.53249 | 0.53249 | 0.53249
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.158 | 2.03 | 36.9 | TP | ug P/L | 0.0595 | 0.838 | -43.9 |  | 1 | 1 |  | 36.9 | 36.9 | 36.9 | [6 empty] | 1 |  | 6.096 | 6.096 | 6.096
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.142 | 1.94 | 32.4 | TP | ug P/L | 0.0584 | 0.817 | -43.9 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 | [6 empty] | 1 |  | 5.8066 | 5.8066 | 5.8066 | -76.7734 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0919 | 1.28 | 0.497 | TP | ug P/L | 0.052 | 0.686 | -44.4 |  | 1 | 1 |  | 0.497 | 0.497 | 0.497 | [6 empty] | 1 |  | 4.1613 | 4.1613 | 4.1613
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrite | S1 | 1 | 1 | TN | ug N/L | 2.27 | 30.8 | 1450 | TP | ug P/L | 0.00874 | 0.143 | -47.2 | 97.7273 | 1 | 1 |  | 1450 | 1450 | 1450 | [6 empty] | 1 |  | 0.2727 | 0.2727 | 0.2727
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | nitrate | S2 | 1 | 1 | TN | ug N/L | 2.22 | 30.1 | 1410 | TP | ug P/L | 0.00884 | 0.101 | -47.1 |  | 1 | 1 |  | 1410 | 1410 | 1410 | [6 empty] | 1 |  | 0.16873 | 0.16873 | 0.16873
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0179 | -0.37 | -79 | TP | ug P/L | 0.00942 | 0.145 | -47.1 |  | 1 | 1 |  | -79 | -79 | -79 | [6 empty] | 1 |  | 0.27836 | 0.27836 | 0.27836
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s105 rerun | 159 | 1 | 1 | TN | ug N/L | 1.11 | 15.1 | 674 | TP | ug P/L | 0.0696 | 0.782 | -43.2 |  | 1 | 1 |  | 674 | 674 | 674 | [6 empty] | 1 |  | 5.3401 | 5.3401 | 5.3401
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s106 rerun | 160 | 1 | 1 | TN | ug N/L | 0.842 | 11.6 | 501 | TP | ug P/L | 0.0675 | 0.736 | -43.4 |  | 1 | 1 |  | 501 | 501 | 501 | [6 empty] | 1 |  | 4.7568 | 4.7568 | 4.7568
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s107 rerun | 161 | 1 | 1 | TN | ug N/L | 1.1 | 15.2 | 678 | TP | ug P/L | 0.0705 | 0.822 | -43.2 |  | 1 | 1 |  | 678 | 678 | 678 | [6 empty] | 1 |  | 5.8749 | 5.8749 | 5.8749
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dups105 rerun | 162 | 1 | 1 | TN | ug N/L | 1.1 | 15.3 | 682 | TP | ug P/L | 0.0698 | 0.792 | -43.2 |  | 1 | 1 |  | 682 | 682 | 682 | [6 empty] | 1 |  | 5.4714 | 5.4714 | 5.4714
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk107 rerun | 163 | 1 | 1 | TN | ug N/L | 1.42 | 19.7 | 897 | TP | ug P/L | 0.0958 | 1.58 | -41.5 |  | 1 | 1 |  | 897 | 897 | 897 | [6 empty] | 1 |  | 20.8537 | 20.8537 | 20.8537
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | s108 rerun | 164 | 1 | 1 | TN | ug N/L | 0.68 | 9.51 | 400 | TP | ug P/L | 0.0719 | 0.863 | -43.1 |  | 1 | 1 |  | 400 | 400 | 400 | [6 empty] | 1 |  | 6.4496 | 6.4496 | 6.4496
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal 1500 | 1 | 1 | 1 | TN | ug N/L | 2.32 | 32 | 1510 | TP | ug P/L | 0.24 | 4.61 | -32.4 |  | 1 | 1 |  | 1510 | 1510 | 1510 | [6 empty] | 1 |  | 172.6655 | 172.6655 | 172.6655
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal 50/5 | 11 | 1 | 1 | TN | ug N/L | 0.159 | 2.28 | 49 | TP | ug P/L | 0.0592 | 0.389 | -43.9 |  | 1 | 1 |  | 49 | 49 | 49 | [6 empty] | 1 |  | 1.4492 | 1.4492 | 1.4492
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | chk 25/250 | 7 | 1 | 1 | TN | ug N/L | 0.451 | 6.34 | 246 | TP | ug P/L | 0.0695 | 0.878 | -43.2 |  | 1 | 1 |  | 246 | 246 | 246 | [6 empty] | 1 |  | 6.6666 | 6.6666 | 6.6666 | -73.3335 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal 25/2.5 | 13 | 1 | 1 | TN | ug N/L | 0.123 | 1.78 | 24.7 | TP | ug P/L | 0.0562 | 0.772 | -44.1 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | [6 empty] | 1 |  | 5.2104 | 5.2104 | 5.2104
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | bisulfate | S4 | 1 | 1 | TN | ug N/L | -0.00258 | -0.0295 | -62.6 | TP | ug P/L | 0.0457 | 0.615 | -44.8 |  | 1 | 1 |  | -62.6 | -62.6 | -62.6 | [6 empty] | 1 |  | 3.3844 | 3.3844 | 3.3844

## QAQC dip ht
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE NO3-N: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- dilution corrected | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- dilution corrected | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | dye | S16 | 1 | 1 | TN | ug N/L | -0.00637 | -0.00462 | 0 | TP | ug P/L | 0.0308 | 0.664 | 0.166 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -104.1857 | -104.1857 | -104.1857
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | dye | S16 | 1 | 1 | TN | ug N/L | 0.0144 | 0.152 | 0 | TP | ug P/L | 0.0324 | 0.702 | 1.33 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -101.0527 | -101.0527 | -101.0527
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.27 | 30 | 0 | TP | ug P/L | -0.00742 | -0.058 | -21.6 | 98.3333 | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -184.3688 | -184.3688 | -184.3688
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.25 | 29.5 | 0 | TP | ug P/L | -0.00878 | -0.0697 | -22 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -187.411 | -187.411 | -187.411
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0162 | -0.271 | 0 | TP | ug P/L | 0.00998 | 0.228 | -13 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -146.5932 | -146.5932 | -146.5932
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00707 | -0.00498 | 0 | TP | ug P/L | 0.0865 | 1.42 | 22.9 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -5.6927 | -5.6927 | -5.6927
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.238 | 4.76 | 150 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 1 |  | 152.0196 | 152.0196 | 152.0196
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.236 | 4.54 | 150 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 1 |  | 150.987 | 150.987 | 150.987
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.66 | 21.8 | 1000 | TP | ug P/L | 0.165 | 3.31 | 100 |  | 1 | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 1 |  | 96.1402 | 96.1402 | 96.1402
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.67 | 22.6 | 1000 | TP | ug P/L | 0.16 | 3.2 | 100 |  | 1 | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 1 |  | 90.944 | 90.944 | 90.944
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.867 | 11.4 | 500 | TP | ug P/L | 0.098 | 1.73 | 50 |  | 1 | 1 |  | 500 | 500 | 500 | [6 empty] | 1 |  | 11.93 | 11.93 | 11.93
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.86 | 11.4 | 500 | TP | ug P/L | 0.0979 | 1.75 | 50 |  | 1 | 1 |  | 500 | 500 | 500 | [6 empty] | 1 |  | 11.7807 | 11.7807 | 11.7807
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.478 | 6.3 | 250 | TP | ug P/L | 0.0926 | 1.65 | 25 |  | 1 | 1 |  | 250 | 250 | 250 | [6 empty] | 1 |  | 3.7706 | 3.7706 | 3.7706 | -84.9176 | FAIL
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.474 | 6.29 | 250 | TP | ug P/L | 0.089 | 1.48 | 25 |  | 1 | 1 |  | 250 | 250 | 250 | [6 empty] | 1 |  | -1.7827 | -1.7827 | -1.7827 | -107.1308 | FAIL
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.43 | 100 | TP | ug P/L | 0.0942 | 1.65 | 10 |  | 1 | 1 |  | 100 | 100 | 100 | [6 empty] | 1 |  | 6.2095 | 6.2095 | 6.2095
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.44 | 100 | TP | ug P/L | 0.095 | 1.5 | 10 |  | 1 | 1 |  | 100 | 100 | 100 | [6 empty] | 1 |  | 7.4222 | 7.4222 | 7.4222
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.176 | 2.33 | 50 | TP | ug P/L | 0.0894 | 1.49 | 5 |  | 1 | 1 |  | 50 | 50 | 50 | [6 empty] | 1 |  | -1.1612 | -1.1612 | -1.1612
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.166 | 2.28 | 50 | TP | ug P/L | 0.092 | 1.54 | 5 |  | 1 | 1 |  | 50 | 50 | 50 | [6 empty] | 1 |  | 2.8514 | 2.8514 | 2.8514
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.131 | 1.8 | 25 | TP | ug P/L | 0.0934 | 1.57 | 2.5 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 4.9923 | 4.9923 | 4.9923
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.129 | 1.76 | 25 | TP | ug P/L | 0.0915 | 1.59 | 2.5 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 2.0834 | 2.0834 | 2.0834
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0914 | 1.27 | 0 | TP | ug P/L | 0.0916 | 1.54 | 0 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 2.2371 | 2.2371 | 2.2371
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0915 | 1.27 | 0 | TP | ug P/L | 0.0945 | 1.64 | 0 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 6.6648 | 6.6648 | 6.6648
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.2 | 29.8 | 1400 | TP | ug P/L | -0.0111 | -0.0948 | -89.5 | 99.6644 | 1 | 1 |  | 1400 | 1400 | 1400 | [6 empty] | 1 |  | -192.6305 | -192.6305 | -192.6305
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.22 | 29.7 | 1390 | TP | ug P/L | 0.0192 | 0.22 | -71.1 |  | 1 | 1 |  | 1390 | 1390 | 1390 | [6 empty] | 1 |  | -127.4379 | -127.4379 | -127.4379
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0172 | -0.298 | -75.6 | TP | ug P/L | -0.00889 | -0.0728 | -88.2 |  | 1 | 1 |  | -75.6 | -75.6 | -75.6 | [6 empty] | 1 |  | -187.6576 | -187.6576 | -187.6576
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00164 | 0.0147 | -60.5 | TP | ug P/L | 0.0932 | 1.68 | 10.2 |  | 1 | 1 |  | -60.5 | -60.5 | -60.5 | [6 empty] | 1 |  | 4.6873 | 4.6873 | 4.6873
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.24 | 30.5 | 1440 | TP | ug P/L | -0.0195 | -0.165 | -58.6 | 97.377 | 1 | 1 |  | 1440 | 1440 | 1440 | [6 empty] | 1 |  | -211.8447 | -211.8447 | -211.8447
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.21 | 29.7 | 1400 | TP | ug P/L | -0.015 | -0.122 | -55.7 |  | 1 | 1 |  | 1400 | 1400 | 1400 | [6 empty] | 1 |  | -201.4897 | -201.4897 | -201.4897
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0165 | -0.274 | -74.4 | TP | ug P/L | -0.0204 | -0.175 | -59.3 |  | 1 | 1 |  | -74.4 | -74.4 | -74.4 | [6 empty] | 1 |  | -213.9327 | -213.9327 | -213.9327
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.26 | 242 | TP | ug P/L | 0.0918 | 1.74 | 50.3 |  | 1 | 1 |  | 242 | 242 | 242 | -3.2 | PASS |  |  |  |  | 1 |  | 2.5444 | 2.5444 | 2.5444 | -89.8224 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.0025 | -0.0316 | -62.7 | TP | ug P/L | 0.0922 | 1.65 | 45.6 |  | 1 | 1 |  | -62.7 | -62.7 | -62.7 | [6 empty] | 1 |  | 3.1581 | 3.1581 | 3.1581
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0167 | -0.267 | -74.1 | TP | ug P/L | 0.0141 | 0.187 | -35.8 |  | 1 | 1 |  | -74.1 | -74.1 | -74.1 | [6 empty] | 1 |  | -137.9598 | -137.9598 | -137.9598
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00439 | -0.035 | -62.9 | TP | ug P/L | 0.0866 | 1.4 | 33.2 |  | 1 | 1 |  | -62.9 | -62.9 | -62.9 | [6 empty] | 1 |  | -5.5355 | -5.5355 | -5.5355
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | 20 | 1 | 1 | TN | ug N/L | 0.084 | 1.2 | -3.48 | TP | ug P/L | 0.0937 | 1.62 | 44.4 |  | 1 | 1 |  | -3.48 | -3.48 | -3.48 | [6 empty] | 1 |  | 5.4493 | 5.4493 | 5.4493
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | 21 | 1 | 1 | TN | ug N/L | 0.0936 | 1.22 | -2.22 | TP | ug P/L | 0.0875 | 1.47 | 37.2 |  | 1 | 1 |  | -2.22 | -2.22 | -2.22 | [6 empty] | 1 |  | -4.1234 | -4.1234 | -4.1234
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | 22 | 1 | 1 | TN | ug N/L | 0.0896 | 1.28 | 0.652 | TP | ug P/L | 0.0928 | 1.68 | 47.4 |  | 1 | 1 |  | 0.652 | 0.652 | 0.652 | [6 empty] | 1 |  | 4.0764 | 4.0764 | 4.0764
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | 23 | 1 | 1 | TN | ug N/L | 0.095 | 1.3 | 1.41 | TP | ug P/L | 0.0875 | 1.45 | 35.9 |  | 1 | 1 |  | 1.41 | 1.41 | 1.41 | [6 empty] | 1 |  | -4.1234 | -4.1234 | -4.1234
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | 24 | 1 | 1 | TN | ug N/L | 0.0925 | 1.31 | 1.89 | TP | ug P/L | 0.0877 | 1.55 | 40.9 |  | 1 | 1 |  | 1.89 | 1.89 | 1.89 | [6 empty] | 1 |  | -3.8104 | -3.8104 | -3.8104
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | digest | 25 | 1 | 1 | TN | ug N/L | 0.71 | 9.48 | 398 | TP | ug P/L | 0.0906 | 1.6 | 43.5 |  | 1 | 1 |  | 398 | 398 | 398 | [6 empty] | 1 |  | 0.69666 | 0.69666 | 0.69666
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | digest | 26 | 1 | 1 | TN | ug N/L | 0.707 | 9.5 | 399 | TP | ug P/L | 0.0889 | 1.6 | 43.6 |  | 1 | 1 |  | 399 | 399 | 399 | [6 empty] | 1 |  | -1.9383 | -1.9383 | -1.9383
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | digest | 27 | 1 | 1 | TN | ug N/L | 0.709 | 9.47 | 398 | TP | ug P/L | 0.0927 | 1.66 | 46.4 |  | 1 | 1 |  | 398 | 398 | 398 | [6 empty] | 1 |  | 3.9236 | 3.9236 | 3.9236
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | digest | 28 | 1 | 1 | TN | ug N/L | 0.711 | 9.46 | 397 | TP | ug P/L | 0.0949 | 1.89 | 57.1 |  | 1 | 1 |  | 397 | 397 | 397 | [6 empty] | 1 |  | 7.2709 | 7.2709 | 7.2709
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | digest | 29 | 1 | 1 | TN | ug N/L | 0.709 | 9.42 | 395 | TP | ug P/L | 0.0944 | 1.68 | 47.1 |  | 1 | 1 |  | 395 | 395 | 395 | [6 empty] | 1 |  | 6.5131 | 6.5131 | 6.5131
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk blank | 30 | 1 | 1 | TN | ug N/L | 0.433 | 5.75 | 217 | TP | ug P/L | 0.0889 | 1.65 | 46 |  | 1 | 1 |  | 217 | 217 | 217 | [6 empty] | 1 |  | -1.9383 | -1.9383 | -1.9383
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | skp blank | 31 | 1 | 1 | TN | ug N/L | 0.436 | 5.89 | 224 | TP | ug P/L | 0.0925 | 1.75 | 50.4 |  | 1 | 1 |  | 224 | 224 | 224 | [6 empty] | 1 |  | 3.6176 | 3.6176 | 3.6176
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk blank | 32 | 1 | 1 | TN | ug N/L | 0.441 | 5.9 | 224 | TP | ug P/L | 0.0954 | 1.73 | 49.4 |  | 1 | 1 |  | 224 | 224 | 224 | [6 empty] | 1 |  | 8.0269 | 8.0269 | 8.0269
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk blank | 33 | 1 | 1 | TN | ug N/L | 0.437 | 5.85 | 222 | TP | ug P/L | 0.0907 | 1.7 | 48.4 |  | 1 | 1 |  | 222 | 222 | 222 | [6 empty] | 1 |  | 0.85102 | 0.85102 | 0.85102
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk blank | 34 | 1 | 1 | TN | ug N/L | 0.446 | 5.92 | 225 | TP | ug P/L | 0.0879 | 1.64 | 45.4 |  | 1 | 1 |  | 225 | 225 | 225 | [6 empty] | 1 |  | -3.4977 | -3.4977 | -3.4977
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s1 | 35 | 1 | 1 | TN | ug N/L | 1.3 | 17.5 | 792 | TP | ug P/L | 0.0911 | 1.66 | 46.2 |  | 1 | 1 |  | 792 | 792 | 792 | [6 empty] | 1 |  | 1.4678 | 1.4678 | 1.4678
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s2 | 36 | 1 | 1 | TN | ug N/L | 0.538 | 7.16 | 285 | TP | ug P/L | 0.0918 | 1.71 | 48.7 |  | 1 | 1 |  | 285 | 285 | 285 | [6 empty] | 1 |  | 2.5444 | 2.5444 | 2.5444
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s3 | 37 | 1 | 1 | TN | ug N/L | 0.56 | 7.54 | 304 | TP | ug P/L | 0.0962 | 1.82 | 53.7 |  | 1 | 1 |  | 304 | 304 | 304 | [6 empty] | 1 |  | 9.2329 | 9.2329 | 9.2329
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s4 | 38 | 1 | 1 | TN | ug N/L | 0.958 | 12.8 | 560 | TP | ug P/L | 0.0943 | 1.76 | 51.2 |  | 1 | 1 |  | 560 | 560 | 560 | [6 empty] | 1 |  | 6.3614 | 6.3614 | 6.3614
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s5 | 39 | 1 | 1 | TN | ug N/L | 3.36 | 45.1 | 2170 | TP | ug P/L | 0.0934 | 1.67 | 46.7 |  | 1 | 1 |  | 2170 | 2170 | 2170 | [6 empty] | 1 |  | 4.9923 | 4.9923 | 4.9923
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s6 | 40 | 1 | 1 | TN | ug N/L | 0.931 | 12.5 | 548 | TP | ug P/L | 0.0931 | 1.7 | 48.1 |  | 1 | 1 |  | 548 | 548 | 548 | [6 empty] | 1 |  | 4.5347 | 4.5347 | 4.5347
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s7 | 41 | 1 | 1 | TN | ug N/L | 0.886 | 11.9 | 515 | TP | ug P/L | 0.0948 | 1.73 | 49.5 |  | 1 | 1 |  | 515 | 515 | 515 | [6 empty] | 1 |  | 7.1195 | 7.1195 | 7.1195
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s8 | 42 | 1 | 1 | TN | ug N/L | 0.573 | 7.62 | 308 | TP | ug P/L | 0.0926 | 1.74 | 50 |  | 1 | 1 |  | 308 | 308 | 308 | [6 empty] | 1 |  | 3.7706 | 3.7706 | 3.7706
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s9 | 43 | 1 | 1 | TN | ug N/L | 1.6 | 21.5 | 988 | TP | ug P/L | 0.097 | 1.84 | 54.6 |  | 1 | 1 |  | 988 | 988 | 988 | [6 empty] | 1 |  | 10.4344 | 10.4344 | 10.4344
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | a10 | 44 | 1 | 1 | TN | ug N/L | 0.764 | 10.2 | 435 | TP | ug P/L | 0.0946 | 1.7 | 48.3 |  | 1 | 1 |  | 435 | 435 | 435 | [6 empty] | 1 |  | 6.8164 | 6.8164 | 6.8164
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup5 | 45 | 1 | 1 | TN | ug N/L | 3.39 | 45.6 | 2190 | TP | ug P/L | 0.0931 | 1.69 | 48 |  | 1 | 1 |  | 2190 | 2190 | 2190 |  |  | 0.91743 | PASS |  |  | 1 |  | 4.5347 | 4.5347 | 4.5347 |  |  | 9.6067 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk10 | 46 | 1 | 1 | TN | ug N/L | 1.07 | 14.4 | 639 | TP | ug P/L | 0.099 | 1.81 | 53.5 |  | 1 | 1 |  | 639 | 639 | 639 |  |  |  |  | 87.99 | PASS | 1 |  | 13.4185 | 13.4185 | 13.4185 |  |  |  |  | 27.75 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.28 | 243 | TP | ug P/L | 0.0914 | 1.7 | 48.2 |  | 1 | 1 |  | 243 | 243 | 243 | -2.8 | PASS |  |  |  |  | 1 |  | 1.9296 | 1.9296 | 1.9296 | -92.2816 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00254 | 0.00654 | -60.9 | TP | ug P/L | 0.0916 | 1.59 | 42.7 |  | 1 | 1 |  | -60.9 | -60.9 | -60.9 | [6 empty] | 1 |  | 2.2371 | 2.2371 | 2.2371
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.016 | -0.252 | -73.3 | TP | ug P/L | 0.00864 | 0.13 | -39.4 |  | 1 | 1 |  | -73.3 | -73.3 | -73.3 | [6 empty] | 1 |  | -149.4268 | -149.4268 | -149.4268
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s11 | 47 | 1 | 1 | TN | ug N/L | 1.62 | 21.9 | 1010 | TP | ug P/L | 0.0924 | 1.8 | 52.9 |  | 1 | 1 |  | 1010 | 1010 | 1010 | [6 empty] | 1 |  | 3.4645 | 3.4645 | 3.4645
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s12 | 48 | 1 | 1 | TN | ug N/L | 0.843 | 11.3 | 485 | TP | ug P/L | 0.093 | 1.71 | 48.7 |  | 1 | 1 |  | 485 | 485 | 485 | [6 empty] | 1 |  | 4.382 | 4.382 | 4.382
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s13 | 49 | 1 | 1 | TN | ug N/L | 1.6 | 21.5 | 990 | TP | ug P/L | 0.0933 | 1.74 | 50.3 |  | 1 | 1 |  | 990 | 990 | 990 | [6 empty] | 1 |  | 4.8398 | 4.8398 | 4.8398
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s14 | 50 | 1 | 1 | TN | ug N/L | 1.67 | 22.6 | 1040 | TP | ug P/L | 0.0942 | 1.67 | 46.7 |  | 1 | 1 |  | 1040 | 1040 | 1040 | [6 empty] | 1 |  | 6.2095 | 6.2095 | 6.2095
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s15 | 51 | 1 | 1 | TN | ug N/L | 0.54 | 7.23 | 289 | TP | ug P/L | 0.0898 | 1.68 | 47.3 |  | 1 | 1 |  | 289 | 289 | 289 | [6 empty] | 1 |  | -0.54078 | -0.54078 | -0.54078
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s16 | 52 | 1 | 1 | TN | ug N/L | 0.594 | 8.01 | 327 | TP | ug P/L | 0.0935 | 1.8 | 52.7 |  | 1 | 1 |  | 327 | 327 | 327 | [6 empty] | 1 |  | 5.1447 | 5.1447 | 5.1447
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s17 | 53 | 1 | 1 | TN | ug N/L | 0.798 | 10.7 | 457 | TP | ug P/L | 0.0945 | 1.78 | 51.9 |  | 1 | 1 |  | 457 | 457 | 457 | [6 empty] | 1 |  | 6.6648 | 6.6648 | 6.6648
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s18 | 54 | 1 | 1 | TN | ug N/L | 3.33 | 44.7 | 2150 | TP | ug P/L | 0.0962 | 1.71 | 48.9 |  | 1 | 1 |  | 2150 | 2150 | 2150 | [6 empty] | 1 |  | 9.2329 | 9.2329 | 9.2329
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s19 | 55 | 1 | 1 | TN | ug N/L | 2.53 | 34 | 1610 | TP | ug P/L | 0.0935 | 1.65 | 45.8 |  | 1 | 1 |  | 1610 | 1610 | 1610 | [6 empty] | 1 |  | 5.1447 | 5.1447 | 5.1447
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s20 | 56 | 1 | 1 | TN | ug N/L | 2.49 | 33.3 | 1580 | TP | ug P/L | 0.0934 | 1.71 | 48.7 |  | 1 | 1 |  | 1580 | 1580 | 1580 | [6 empty] | 1 |  | 4.9923 | 4.9923 | 4.9923
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup15 | 57 | 1 | 1 | TN | ug N/L | 0.54 | 7.19 | 287 | TP | ug P/L | 0.0899 | 1.64 | 45.3 |  | 1 | 1 |  | 287 | 287 | 287 |  |  | 0.69444 | PASS |  |  | 1 |  | -0.38586 | -0.38586 | -0.38586 |  |  | 33.4382 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk20 | 58 | 1 | 1 | TN | ug N/L | 2.81 | 37.7 | 1800 | TP | ug P/L | 0.106 | 1.97 | 60.7 |  | 1 | 1 |  | 1800 | 1800 | 1800 |  |  |  |  | 106 | PASS | 1 |  | 23.6416 | 23.6416 | 23.6416 |  |  |  |  | 76.9615 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.35 | 246 | TP | ug P/L | 0.0924 | 1.7 | 48.1 |  | 1 | 1 |  | 246 | 246 | 246 | -1.6 | PASS |  |  |  |  | 1 |  | 3.4645 | 3.4645 | 3.4645 | -86.1422 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00264 | -0.0321 | -62.8 | TP | ug P/L | 0.0865 | 1.47 | 37 |  | 1 | 1 |  | -62.8 | -62.8 | -62.8 | [6 empty] | 1 |  | -5.6927 | -5.6927 | -5.6927
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00701 | 0.037 | -59.4 | TP | ug P/L | 0.0123 | 0.214 | -34.1 |  | 1 | 1 |  | -59.4 | -59.4 | -59.4 | [6 empty] | 1 |  | -141.717 | -141.717 | -141.717
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s21 | 59 | 1 | 1 | TN | ug N/L | 0.554 | 7.45 | 299 | TP | ug P/L | 0.0964 | 1.85 | 55.3 |  | 1 | 1 |  | 299 | 299 | 299 | [6 empty] | 1 |  | 9.5337 | 9.5337 | 9.5337
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s22 | 60 | 1 | 1 | TN | ug N/L | 0.736 | 9.94 | 421 | TP | ug P/L | 0.0929 | 1.79 | 52.6 |  | 1 | 1 |  | 421 | 421 | 421 | [6 empty] | 1 |  | 4.2293 | 4.2293 | 4.2293
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s23 | 61 | 1 | 1 | TN | ug N/L | 0.615 | 8.28 | 340 | TP | ug P/L | 0.0914 | 1.7 | 48 |  | 1 | 1 |  | 340 | 340 | 340 | [6 empty] | 1 |  | 1.9296 | 1.9296 | 1.9296
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s24 | 62 | 1 | 1 | TN | ug N/L | 1.04 | 14 | 622 | TP | ug P/L | 0.0976 | 1.86 | 55.7 |  | 1 | 1 |  | 622 | 622 | 622 | [6 empty] | 1 |  | 11.3326 | 11.3326 | 11.3326
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s25 | 63 | 1 | 1 | TN | ug N/L | 3.03 | 40.8 | 1950 | TP | ug P/L | 0.101 | 1.75 | 50.8 |  | 1 | 1 |  | 1950 | 1950 | 1950 | [6 empty] | 1 |  | 16.3745 | 16.3745 | 16.3745
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s26 | 64 | 1 | 1 | TN | ug N/L | 1.17 | 15.8 | 708 | TP | ug P/L | 0.0963 | 1.81 | 53.7 |  | 1 | 1 |  | 708 | 708 | 708 | [6 empty] | 1 |  | 9.3833 | 9.3833 | 9.3833
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s27 | 65 | 1 | 1 | TN | ug N/L | 0.5 | 6.71 | 264 | TP | ug P/L | 0.0945 | 1.75 | 50.7 |  | 1 | 1 |  | 264 | 264 | 264 | [6 empty] | 1 |  | 6.6648 | 6.6648 | 6.6648
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s28 | 66 | 1 | 1 | TN | ug N/L | 2.56 | 34.5 | 1630 | TP | ug P/L | 0.149 | 3.02 | 103 |  | 1 | 1 |  | 1630 | 1630 | 1630 | [6 empty] | 1 |  | 78.8945 | 78.8945 | 78.8945
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s29 | 67 | 1 | 1 | TN | ug N/L | 2.18 | 29.3 | 1370 | TP | ug P/L | 0.0938 | 1.71 | 48.9 |  | 1 | 1 |  | 1370 | 1370 | 1370 | [6 empty] | 1 |  | 5.6015 | 5.6015 | 5.6015
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s30 | 68 | 1 | 1 | TN | ug N/L | 1.14 | 15.3 | 682 | TP | ug P/L | 0.0936 | 1.75 | 50.5 |  | 1 | 1 |  | 682 | 682 | 682 | [6 empty] | 1 |  | 5.297 | 5.297 | 5.297
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup25 | 69 | 1 | 1 | TN | ug N/L | 3 | 40.5 | 1930 | TP | ug P/L | 0.1 | 1.75 | 50.7 |  | 1 | 1 |  | 1930 | 1930 | 1930 |  |  | 1.0309 | PASS |  |  | 1 |  | 14.9 | 14.9 | 14.9 |  |  | 9.4293 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup30 | 70 | 1 | 1 | TN | ug N/L | 1.46 | 19.7 | 900 | TP | ug P/L | 0.0988 | 1.78 | 51.8 |  | 1 | 1 |  | 900 | 900 | 900 |  |  |  |  | 96.2 | PASS | 1 |  | 13.1213 | 13.1213 | 13.1213 |  |  |  |  | 32.6094 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.469 | 6.28 | 243 | TP | ug P/L | 0.0919 | 1.69 | 47.6 |  | 1 | 1 |  | 243 | 243 | 243 | -2.8 | PASS |  |  |  |  | 1 |  | 2.6979 | 2.6979 | 2.6979 | -89.2084 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00521 | -0.0319 | -62.7 | TP | ug P/L | 0.0914 | 1.67 | 46.6 |  | 1 | 1 |  | -62.7 | -62.7 | -62.7 | [6 empty] | 1 |  | 1.9296 | 1.9296 | 1.9296
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0169 | -0.275 | -74.4 | TP | ug P/L | 0.0107 | 0.18 | -36.2 |  | 1 | 1 |  | -74.4 | -74.4 | -74.4 | [6 empty] | 1 |  | -145.0759 | -145.0759 | -145.0759
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s31 | 71 | 1 | 1 | TN | ug N/L | 0.491 | 6.8 | 268 | TP | ug P/L | 0.0921 | 1.73 | 49.6 |  | 1 | 1 |  | 268 | 268 | 268 | [6 empty] | 1 |  | 3.0047 | 3.0047 | 3.0047
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s32 | 72 | 1 | 1 | TN | ug N/L | 1.13 | 15.2 | 679 | TP | ug P/L | 0.094 | 1.78 | 52.1 |  | 1 | 1 |  | 679 | 679 | 679 | [6 empty] | 1 |  | 5.9056 | 5.9056 | 5.9056
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s33 | 73 | 1 | 1 | TN | ug N/L | 0.853 | 11.5 | 495 | TP | ug P/L | 0.0967 | 1.79 | 52.5 |  | 1 | 1 |  | 495 | 495 | 495 | [6 empty] | 1 |  | 9.9844 | 9.9844 | 9.9844
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s34 | 74 | 1 | 1 | TN | ug N/L | 1.37 | 18.4 | 837 | TP | ug P/L | 0.0959 | 1.72 | 49.2 |  | 1 | 1 |  | 837 | 837 | 837 | [6 empty] | 1 |  | 8.7812 | 8.7812 | 8.7812
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s35 | 75 | 1 | 1 | TN | ug N/L | 0.745 | 10 | 425 | TP | ug P/L | 0.094 | 1.78 | 52 |  | 1 | 1 |  | 425 | 425 | 425 | [6 empty] | 1 |  | 5.9056 | 5.9056 | 5.9056
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s36 | 76 | 1 | 1 | TN | ug N/L | 1.64 | 22.1 | 1020 | TP | ug P/L | 0.101 | 1.82 | 53.7 |  | 1 | 1 |  | 1020 | 1020 | 1020 | [6 empty] | 1 |  | 16.3745 | 16.3745 | 16.3745
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s37 | 77 | 1 | 1 | TN | ug N/L | 2.57 | 34.9 | 1650 | TP | ug P/L | 0.0937 | 1.72 | 49 |  | 1 | 1 |  | 1650 | 1650 | 1650 | [6 empty] | 1 |  | 5.4493 | 5.4493 | 5.4493
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s38 | 78 | 1 | 1 | TN | ug N/L | 0.752 | 10.1 | 428 | TP | ug P/L | 0.0925 | 1.72 | 49.3 |  | 1 | 1 |  | 428 | 428 | 428 | [6 empty] | 1 |  | 3.6176 | 3.6176 | 3.6176
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s39 | 79 | 1 | 1 | TN | ug N/L | 1.17 | 15.8 | 707 | TP | ug P/L | 0.0958 | 1.86 | 55.6 |  | 1 | 1 |  | 707 | 707 | 707 | [6 empty] | 1 |  | 8.6305 | 8.6305 | 8.6305
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s40 | 80 | 1 | 1 | TN | ug N/L | 1.76 | 23.7 | 1100 | TP | ug P/L | 0.093 | 1.73 | 49.4 |  | 1 | 1 |  | 1100 | 1100 | 1100 | [6 empty] | 1 |  | 4.382 | 4.382 | 4.382
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup35 | 81 | 1 | 1 | TN | ug N/L | 0.746 | 10.1 | 430 | TP | ug P/L | 0.0971 | 1.86 | 55.6 |  | 1 | 1 |  | 430 | 430 | 430 |  |  | 1.1696 | PASS |  |  | 1 |  | 10.5843 | 10.5843 | 10.5843 |  |  | 56.7455 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk40 | 82 | 1 | 1 | TN | ug N/L | 2.09 | 28.1 | 1310 | TP | ug P/L | 0.101 | 1.81 | 53.3 |  | 1 | 1 |  | 1310 | 1310 | 1310 |  |  |  |  | 97.1 | PASS | 1 |  | 16.3745 | 16.3745 | 16.3745 |  |  |  |  | 49.6074 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.3 | 244 | TP | ug P/L | 0.0913 | 1.67 | 47 |  | 1 | 1 |  | 244 | 244 | 244 | -2.4 | PASS |  |  |  |  | 1 |  | 1.7757 | 1.7757 | 1.7757 | -92.8971 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00528 | -0.0552 | -63.9 | TP | ug P/L | 0.0899 | 1.6 | 43.3 |  | 1 | 1 |  | -63.9 | -63.9 | -63.9 | [6 empty] | 1 |  | -0.38586 | -0.38586 | -0.38586
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0144 | -0.226 | -72.1 | TP | ug P/L | 0.0111 | 0.153 | -37.9 |  | 1 | 1 |  | -72.1 | -72.1 | -72.1 | [6 empty] | 1 |  | -144.2345 | -144.2345 | -144.2345
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s41 | 83 | 1 | 1 | TN | ug N/L | 1.16 | 15.7 | 703 | TP | ug P/L | 0.0937 | 1.83 | 54.5 |  | 1 | 1 |  | 703 | 703 | 703 | [6 empty] | 1 |  | 5.4493 | 5.4493 | 5.4493
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s42 | 84 | 1 | 1 | TN | ug N/L | 0.5 | 6.77 | 266 | TP | ug P/L | 0.095 | 1.81 | 53.6 |  | 1 | 1 |  | 266 | 266 | 266 | [6 empty] | 1 |  | 7.4222 | 7.4222 | 7.4222
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s43 | 85 | 1 | 1 | TN | ug N/L | 1.02 | 13.8 | 610 | TP | ug P/L | 0.0994 | 1.8 | 53.1 |  | 1 | 1 |  | 610 | 610 | 610 | [6 empty] | 1 |  | 14.0119 | 14.0119 | 14.0119
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s44 | 86 | 1 | 1 | TN | ug N/L | 2.55 | 34.5 | 1640 | TP | ug P/L | 0.0971 | 1.75 | 50.7 |  | 1 | 1 |  | 1640 | 1640 | 1640 | [6 empty] | 1 |  | 10.5843 | 10.5843 | 10.5843
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s45 | 87 | 1 | 1 | TN | ug N/L | 1.46 | 19.7 | 901 | TP | ug P/L | 0.0916 | 1.69 | 47.9 |  | 1 | 1 |  | 901 | 901 | 901 | [6 empty] | 1 |  | 2.2371 | 2.2371 | 2.2371
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s46 | 88 | 1 | 1 | TN | ug N/L | 2.42 | 32.6 | 1540 | TP | ug P/L | 0.139 | 2.78 | 93.9 |  | 1 | 1 |  | 1540 | 1540 | 1540 | [6 empty] | 1 |  | 67.2033 | 67.2033 | 67.2033
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s47 | 89 | 1 | 1 | TN | ug N/L | 0.713 | 9.56 | 402 | TP | ug P/L | 0.0944 | 1.82 | 53.7 |  | 1 | 1 |  | 402 | 402 | 402 | [6 empty] | 1 |  | 6.5131 | 6.5131 | 6.5131
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s48 | 90 | 1 | 1 | TN | ug N/L | 0.975 | 13.2 | 581 | TP | ug P/L | 0.107 | 2.18 | 69.9 |  | 1 | 1 |  | 581 | 581 | 581 | [6 empty] | 1 |  | 25.074 | 25.074 | 25.074
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s49 | 91 | 1 | 1 | TN | ug N/L | 1.1 | 15 | 666 | TP | ug P/L | 0.0962 | 1.8 | 52.8 |  | 1 | 1 |  | 666 | 666 | 666 | [6 empty] | 1 |  | 9.2329 | 9.2329 | 9.2329
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s50 | 92 | 1 | 1 | TN | ug N/L | 0.825 | 11.1 | 478 | TP | ug P/L | 0.0917 | 1.76 | 51.3 |  | 1 | 1 |  | 478 | 478 | 478 | [6 empty] | 1 |  | 2.3908 | 2.3908 | 2.3908
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup45 | 93 | 1 | 1 | TN | ug N/L | 1.45 | 19.7 | 899 | TP | ug P/L | 0.0951 | 1.82 | 53.7 |  | 1 | 1 |  | 899 | 899 | 899 |  |  | 0.22222 | PASS |  |  | 1 |  | 7.5735 | 7.5735 | 7.5735 |  |  | 108.7876 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk50 | 94 | 1 | 1 | TN | ug N/L | 1.2 | 16.2 | 726 | TP | ug P/L | 0.101 | 1.86 | 56 |  | 1 | 1 |  | 726 | 726 | 726 |  |  |  |  | 106.46 | PASS | 1 |  | 16.3745 | 16.3745 | 16.3745 |  |  |  |  | 57.5722 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.471 | 6.35 | 246 | TP | ug P/L | 0.0921 | 1.75 | 50.8 |  | 1 | 1 |  | 246 | 246 | 246 | -1.6 | PASS |  |  |  |  | 1 |  | 3.0047 | 3.0047 | 3.0047 | -87.981 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00351 | -0.0324 | -62.8 | TP | ug P/L | 0.0885 | 1.53 | 39.7 |  | 1 | 1 |  | -62.8 | -62.8 | -62.8 | [6 empty] | 1 |  | -2.5612 | -2.5612 | -2.5612
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.016 | -0.27 | -74.2 | TP | ug P/L | 0.0105 | 0.169 | -36.9 |  | 1 | 1 |  | -74.2 | -74.2 | -74.2 | [6 empty] | 1 |  | -145.497 | -145.497 | -145.497
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal1500 | 1 | 1 | 1 | TN | ug N/L | 2.37 | 31.9 | 1500 | TP | ug P/L | 0.259 | 4.6 | 145 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 1 |  | 161.1657 | 161.1657 | 161.1657
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal1000 | 3 | 1 | 1 | TN | ug N/L | 1.62 | 21.9 | 1010 | TP | ug P/L | 0.159 | 3.17 | 108 |  | 1 | 1 |  | 1010 | 1010 | 1010 | [6 empty] | 1 |  | 89.8837 | 89.8837 | 89.8837
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | call500 | 5 | 1 | 1 | TN | ug N/L | 0.848 | 11.5 | 495 | TP | ug P/L | 0.0962 | 1.8 | 52.8 |  | 1 | 1 |  | 495 | 495 | 495 | [6 empty] | 1 |  | 9.2329 | 9.2329 | 9.2329
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | call250 | 7 | 1 | 1 | TN | ug N/L | 0.471 | 6.35 | 246 | TP | ug P/L | 0.0908 | 1.68 | 47.3 |  | 1 | 1 |  | 246 | 246 | 246 | [6 empty] | 1 |  | 1.0053 | 1.0053 | 1.0053 | -95.9787 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal100 | 9 | 1 | 1 | TN | ug N/L | 0.254 | 3.45 | 105 | TP | ug P/L | 0.0907 | 1.76 | 51.3 |  | 1 | 1 |  | 105 | 105 | 105 | [6 empty] | 1 |  | 0.85102 | 0.85102 | 0.85102
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.172 | 2.35 | 52.2 | TP | ug P/L | 0.0944 | 1.88 | 56.8 |  | 1 | 1 |  | 52.2 | 52.2 | 52.2 | [6 empty] | 1 |  | 6.5131 | 6.5131 | 6.5131
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.135 | 1.81 | 26.3 | TP | ug P/L | 0.0929 | 1.76 | 51.1 |  | 1 | 1 |  | 26.3 | 26.3 | 26.3 | [6 empty] | 1 |  | 4.2293 | 4.2293 | 4.2293
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0958 | 1.29 | 1.07 | TP | ug P/L | 0.0887 | 1.56 | 41.2 |  | 1 | 1 |  | 1.07 | 1.07 | 1.07 | [6 empty] | 1 |  | -2.2496 | -2.2496 | -2.2496
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrite | S1 | 1 | 1 | TN | ug N/L | 2.19 | 30.1 | 1420 | TP | ug P/L | 0.00586 | 0.0804 | -42.5 | 99.6678 | 1 | 1 |  | 1420 | 1420 | 1420 | [6 empty] | 1 |  | -155.3457 | -155.3457 | -155.3457
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrate | S2 | 1 | 1 | TN | ug N/L | 2.21 | 30 | 1410 | TP | ug P/L | 0.00889 | 0.119 | -40.1 |  | 1 | 1 |  | 1410 | 1410 | 1410 | [6 empty] | 1 |  | -148.8972 | -148.8972 | -148.8972
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s51 | 95 | 1 | 1 | TN | ug N/L | 0.609 | 8.22 | 337 | TP | ug P/L | 0.0929 | 1.8 | 53 |  | 1 | 1 |  | 337 | 337 | 337 | [6 empty] | 1 |  | 4.2293 | 4.2293 | 4.2293
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s52 | 96 | 1 | 1 | TN | ug N/L | 0.837 | 11.3 | 486 | TP | ug P/L | 0.0938 | 1.83 | 54.4 |  | 1 | 1 |  | 486 | 486 | 486 | [6 empty] | 1 |  | 5.6015 | 5.6015 | 5.6015
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s53 | 97 | 1 | 1 | TN | ug N/L | 4.93 | 67.1 | 3290 | TP | ug P/L | 0.1 | 1.82 | 54.1 |  | 1 | 1 |  | 3290 | 3290 | 3290 | [6 empty] | 1 |  | 14.9 | 14.9 | 14.9
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s54 | 98 | 1 | 1 | TN | ug N/L | 0.626 | 8.5 | 351 | TP | ug P/L | 0.0937 | 1.79 | 52.4 |  | 1 | 1 |  | 351 | 351 | 351 | [6 empty] | 1 |  | 5.4493 | 5.4493 | 5.4493
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s55 | 99 | 1 | 1 | TN | ug N/L | 0.728 | 9.89 | 418 | TP | ug P/L | 0.0916 | 1.67 | 46.7 |  | 1 | 1 |  | 418 | 418 | 418 | [6 empty] | 1 |  | 2.2371 | 2.2371 | 2.2371
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s56 | 100 | 1 | 1 | TN | ug N/L | 1.13 | 15.4 | 686 | TP | ug P/L | 0.0931 | 1.67 | 46.9 |  | 1 | 1 |  | 686 | 686 | 686 | [6 empty] | 1 |  | 4.5347 | 4.5347 | 4.5347
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s57 | 101 | 1 | 1 | TN | ug N/L | 2.45 | 33.1 | 1560 | TP | ug P/L | 0.0929 | 1.7 | 48.4 |  | 1 | 1 |  | 1560 | 1560 | 1560 | [6 empty] | 1 |  | 4.2293 | 4.2293 | 4.2293
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s58 | 102 | 1 | 1 | TN | ug N/L | 2.77 | 37.4 | 1780 | TP | ug P/L | 0.0983 | 1.77 | 51.6 |  | 1 | 1 |  | 1780 | 1780 | 1780 | [6 empty] | 1 |  | 12.3773 | 12.3773 | 12.3773
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s59 | 103 | 1 | 1 | TN | ug N/L | 1.08 | 14.7 | 654 | TP | ug P/L | 0.091 | 1.71 | 48.8 |  | 1 | 1 |  | 654 | 654 | 654 | [6 empty] | 1 |  | 1.3137 | 1.3137 | 1.3137
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s61 | 104 | 1 | 1 | TN | ug N/L | 3.12 | 42.4 | 2030 | TP | ug P/L | 0.1 | 1.79 | 52.3 |  | 1 | 1 |  | 2030 | 2030 | 2030 | [6 empty] | 1 |  | 14.9 | 14.9 | 14.9
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup55 | 105 | 1 | 1 | TN | ug N/L | 0.729 | 9.84 | 416 | TP | ug P/L | 0.094 | 1.72 | 49.1 |  | 1 | 1 |  | 416 | 416 | 416 |  |  | 0.47962 | PASS |  |  | 1 |  | 5.9056 | 5.9056 | 5.9056 |  |  | 90.1046 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk60 | 106 | 1 | 1 | TN | ug N/L | 3.47 | 47.2 | 2280 | TP | ug P/L | 0.136 | 2.68 | 90.3 |  | 1 | 1 |  | 2280 | 2280 | 2280 |  |  |  |  | 122.8 | FAIL | 1 |  | 63.559 | 63.559 | 63.559 |  |  |  |  | 200.9921 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.462 | 6.23 | 240 | TP | ug P/L | 0.0925 | 1.75 | 50.8 |  | 1 | 1 |  | 240 | 240 | 240 | -4 | PASS |  |  |  |  | 1 |  | 3.6176 | 3.6176 | 3.6176 | -85.5297 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00622 | -0.0368 | -63 | TP | ug P/L | 0.0853 | 1.42 | 34.6 |  | 1 | 1 |  | -63 | -63 | -63 | [6 empty] | 1 |  | -7.5851 | -7.5851 | -7.5851
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0176 | -0.254 | -73.4 | TP | ug P/L | 0.00903 | 0.146 | -38.3 |  | 1 | 1 |  | -73.4 | -73.4 | -73.4 | [6 empty] | 1 |  | -148.6008 | -148.6008 | -148.6008
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s61 | 107 | 1 | 1 | TN | ug N/L | 1.51 | 20.4 | 936 | TP | ug P/L | 0.0951 | 1.79 | 52.4 |  | 1 | 1 |  | 936 | 936 | 936 | [6 empty] | 1 |  | 7.5735 | 7.5735 | 7.5735
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s62 | 108 | 1 | 1 | TN | ug N/L | 1.14 | 15.5 | 692 | TP | ug P/L | 0.0941 | 1.77 | 51.7 |  | 1 | 1 |  | 692 | 692 | 692 | [6 empty] | 1 |  | 6.0576 | 6.0576 | 6.0576
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s63 | 109 | 1 | 1 | TN | ug N/L | 0.577 | 7.82 | 317 | TP | ug P/L | 0.0933 | 1.78 | 52 |  | 1 | 1 |  | 317 | 317 | 317 | [6 empty] | 1 |  | 4.8398 | 4.8398 | 4.8398
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s64 | 110 | 1 | 1 | TN | ug N/L | 1.56 | 21.2 | 971 | TP | ug P/L | 0.098 | 1.86 | 55.6 |  | 1 | 1 |  | 971 | 971 | 971 | [6 empty] | 1 |  | 11.93 | 11.93 | 11.93
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s65 | 111 | 1 | 1 | TN | ug N/L | 4.36 | 59.2 | 2890 | TP | ug P/L | 0.0967 | 1.71 | 48.8 |  | 1 | 1 |  | 2890 | 2890 | 2890 | [6 empty] | 1 |  | 9.9844 | 9.9844 | 9.9844
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s66 | 112 | 1 | 1 | TN | ug N/L | 2.17 | 29.5 | 1380 | TP | ug P/L | 0.0951 | 1.75 | 50.8 |  | 1 | 1 |  | 1380 | 1380 | 1380 | [6 empty] | 1 |  | 7.5735 | 7.5735 | 7.5735
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s67 | 113 | 1 | 1 | TN | ug N/L | 1.03 | 13.9 | 614 | TP | ug P/L | 0.0968 | 1.79 | 52.6 |  | 1 | 1 |  | 614 | 614 | 614 | [6 empty] | 1 |  | 10.1345 | 10.1345 | 10.1345
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s68 | 114 | 1 | 1 | TN | ug N/L | 3.14 | 42.6 | 2040 | TP | ug P/L | 0.103 | 1.87 | 56.2 |  | 1 | 1 |  | 2040 | 2040 | 2040 | [6 empty] | 1 |  | 19.3024 | 19.3024 | 19.3024
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s69 | 115 | 1 | 1 | TN | ug N/L | 0.668 | 9.05 | 378 | TP | ug P/L | 0.0963 | 1.81 | 53.6 |  | 1 | 1 |  | 378 | 378 | 378 | [6 empty] | 1 |  | 9.3833 | 9.3833 | 9.3833
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s70 | 116 | 1 | 1 | TN | ug N/L | 1.88 | 25.7 | 1190 | TP | ug P/L | 0.0956 | 1.7 | 48.4 |  | 1 | 1 |  | 1190 | 1190 | 1190 | [6 empty] | 1 |  | 8.3288 | 8.3288 | 8.3288
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup65 | 117 | 1 | 1 | TN | ug N/L | 4.38 | 59.6 | 2910 | TP | ug P/L | 0.0948 | 1.68 | 47.1 |  | 1 | 1 |  | 2910 | 2910 | 2910 |  |  | 0.68966 | PASS |  |  | 1 |  | 7.1195 | 7.1195 | 7.1195 |  |  | 33.4999 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk70 | 118 | 1 | 1 | TN | ug N/L | 2.14 | 29 | 1360 | TP | ug P/L | 0.108 | 2.04 | 64 |  | 1 | 1 |  | 1360 | 1360 | 1360 |  |  |  |  | 81.6 | PASS | 1 |  | 26.4994 | 26.4994 | 26.4994 |  |  |  |  | 75.332 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.459 | 6.21 | 239 | TP | ug P/L | 0.0896 | 1.67 | 47 |  | 1 | 1 |  | 239 | 239 | 239 | -4.4 | PASS |  |  |  |  | 1 |  | -0.85084 | -0.85084 | -0.85084 | -103.4034 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00409 | -0.0102 | -61.7 | TP | ug P/L | 0.0895 | 1.61 | 43.7 |  | 1 | 1 |  | -61.7 | -61.7 | -61.7 | [6 empty] | 1 |  | -1.006 | -1.006 | -1.006
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0154 | -0.262 | -73.9 | TP | ug P/L | 0.0111 | 0.178 | -36.3 |  | 1 | 1 |  | -73.9 | -73.9 | -73.9 | [6 empty] | 1 |  | -144.2345 | -144.2345 | -144.2345
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s71 | 119 | 1 | 1 | TN | ug N/L | 1.83 | 24.9 | 1150 | TP | ug P/L | 0.0957 | 1.73 | 49.8 |  | 1 | 1 |  | 1150 | 1150 | 1150 | [6 empty] | 1 |  | 8.4797 | 8.4797 | 8.4797
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s72 | 120 | 1 | 1 | TN | ug N/L | 0.94 | 12.8 | 558 | TP | ug P/L | 0.0933 | 1.72 | 49.1 |  | 1 | 1 |  | 558 | 558 | 558 | [6 empty] | 1 |  | 4.8398 | 4.8398 | 4.8398
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s73 | 121 | 1 | 1 | TN | ug N/L | 1.17 | 16 | 718 | TP | ug P/L | 0.0923 | 1.76 | 51 |  | 1 | 1 |  | 718 | 718 | 718 | [6 empty] | 1 |  | 3.3113 | 3.3113 | 3.3113
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s74 | 122 | 1 | 1 | TN | ug N/L | 1.25 | 16.9 | 762 | TP | ug P/L | 0.0948 | 1.78 | 51.9 |  | 1 | 1 |  | 762 | 762 | 762 | [6 empty] | 1 |  | 7.1195 | 7.1195 | 7.1195
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s75 | 123 | 1 | 1 | TN | ug N/L | 1.66 | 22.6 | 1040 | TP | ug P/L | 0.1 | 1.8 | 52.9 |  | 1 | 1 |  | 1040 | 1040 | 1040 | [6 empty] | 1 |  | 14.9 | 14.9 | 14.9
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s76 | 124 | 1 | 1 | TN | ug N/L | 1.14 | 15.5 | 691 | TP | ug P/L | 0.0931 | 1.73 | 49.6 |  | 1 | 1 |  | 691 | 691 | 691 | [6 empty] | 1 |  | 4.5347 | 4.5347 | 4.5347
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s77 | 125 | 1 | 1 | TN | ug N/L | 1.87 | 25.5 | 1190 | TP | ug P/L | 0.124 | 2.39 | 78.9 |  | 1 | 1 |  | 1190 | 1190 | 1190 | [6 empty] | 1 |  | 48.3502 | 48.3502 | 48.3502
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s78 | 126 | 1 | 1 | TN | ug N/L | 1.34 | 18.1 | 819 | TP | ug P/L | 0.0939 | 1.68 | 47.5 |  | 1 | 1 |  | 819 | 819 | 819 | [6 empty] | 1 |  | 5.7536 | 5.7536 | 5.7536
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s79 | 127 | 1 | 1 | TN | ug N/L | 1.01 | 13.8 | 607 | TP | ug P/L | 0.0943 | 1.78 | 51.9 |  | 1 | 1 |  | 607 | 607 | 607 | [6 empty] | 1 |  | 6.3614 | 6.3614 | 6.3614
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s80 | 128 | 1 | 1 | TN | ug N/L | 3 | 40.8 | 1950 | TP | ug P/L | 0.0965 | 1.73 | 49.5 |  | 1 | 1 |  | 1950 | 1950 | 1950 | [6 empty] | 1 |  | 9.684 | 9.684 | 9.684
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup75 | 129 | 1 | 1 | TN | ug N/L | 1.93 | 26.2 | 1220 | TP | ug P/L | 0.122 | 2.35 | 77.4 |  | 1 | 1 |  | 1220 | 1220 | 1220 |  |  | 15.9292 | PASS |  |  | 1 |  | 45.7172 | 45.7172 | 45.7172 |  |  | 101.678 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk80 | 130 | 1 | 1 | TN | ug N/L | 3.25 | 44.1 | 2120 | TP | ug P/L | 0.107 | 2.01 | 62.5 |  | 1 | 1 |  | 2120 | 2120 | 2120 |  |  |  |  | 89.2 | PASS | 1 |  | 25.074 | 25.074 | 25.074 |  |  |  |  | 64.0674 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.472 | 6.36 | 247 | TP | ug P/L | 0.0942 | 1.76 | 51.2 |  | 1 | 1 |  | 247 | 247 | 247 | -1.2 | PASS |  |  |  |  | 1 |  | 6.2095 | 6.2095 | 6.2095 | -75.1619 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00739 | -0.0299 | -62.6 | TP | ug P/L | 0.0905 | 1.66 | 46.1 |  | 1 | 1 |  | -62.6 | -62.6 | -62.6 | [6 empty] | 1 |  | 0.54222 | 0.54222 | 0.54222
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0154 | -0.231 | -72.4 | TP | ug P/L | 0.00907 | 0.142 | -38.6 |  | 1 | 1 |  | -72.4 | -72.4 | -72.4 | [6 empty] | 1 |  | -148.5161 | -148.5161 | -148.5161
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s81 | 131 | 1 | 1 | TN | ug N/L | 1.14 | 15.6 | 695 | TP | ug P/L | 0.0965 | 1.76 | 51.2 |  | 1 | 1 |  | 695 | 695 | 695 | [6 empty] | 1 |  | 9.684 | 9.684 | 9.684
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s82 | 132 | 1 | 1 | TN | ug N/L | 0.583 | 7.92 | 322 | TP | ug P/L | 0.0961 | 1.77 | 51.7 |  | 1 | 1 |  | 322 | 322 | 322 | [6 empty] | 1 |  | 9.0824 | 9.0824 | 9.0824
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s83 | 133 | 1 | 1 | TN | ug N/L | 1.14 | 15.4 | 690 | TP | ug P/L | 0.0935 | 1.71 | 48.9 |  | 1 | 1 |  | 690 | 690 | 690 | [6 empty] | 1 |  | 5.1447 | 5.1447 | 5.1447
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s84 | 134 | 1 | 1 | TN | ug N/L | 1.42 | 19.3 | 880 | TP | ug P/L | 0.0967 | 1.77 | 51.5 |  | 1 | 1 |  | 880 | 880 | 880 | [6 empty] | 1 |  | 9.9844 | 9.9844 | 9.9844
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s85 | 135 | 1 | 1 | TN | ug N/L | 0.496 | 6.71 | 263 | TP | ug P/L | 0.0943 | 1.81 | 53.6 |  | 1 | 1 |  | 263 | 263 | 263 | [6 empty] | 1 |  | 6.3614 | 6.3614 | 6.3614
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s86 | 136 | 1 | 1 | TN | ug N/L | 0.56 | 7.61 | 307 | TP | ug P/L | 0.0944 | 1.8 | 53.2 |  | 1 | 1 |  | 307 | 307 | 307 | [6 empty] | 1 |  | 6.5131 | 6.5131 | 6.5131
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s87 | 137 | 1 | 1 | TN | ug N/L | 0.629 | 8.57 | 354 | TP | ug P/L | 0.0922 | 1.65 | 45.8 |  | 1 | 1 |  | 354 | 354 | 354 | [6 empty] | 1 |  | 3.1581 | 3.1581 | 3.1581
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s88 | 138 | 1 | 1 | TN | ug N/L | 0.546 | 7.39 | 297 | TP | ug P/L | 0.0952 | 1.74 | 50.1 |  | 1 | 1 |  | 297 | 297 | 297 | [6 empty] | 1 |  | 7.7247 | 7.7247 | 7.7247
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s89 | 139 | 1 | 1 | TN | ug N/L | 1.16 | 15.8 | 708 | TP | ug P/L | 0.0964 | 1.78 | 52.1 |  | 1 | 1 |  | 708 | 708 | 708 | [6 empty] | 1 |  | 9.5337 | 9.5337 | 9.5337
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s90 | 140 | 1 | 1 | TN | ug N/L | 0.613 | 8.31 | 341 | TP | ug P/L | 0.0936 | 1.72 | 49 |  | 1 | 1 |  | 341 | 341 | 341 | [6 empty] | 1 |  | 5.297 | 5.297 | 5.297
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup85 | 141 | 1 | 1 | TN | ug N/L | 0.503 | 6.8 | 268 | TP | ug P/L | 0.1 | 1.88 | 56.7 |  | 1 | 1 |  | 268 | 268 | 268 |  |  | 1.8832 | PASS |  |  | 1 |  | 14.9 | 14.9 | 14.9 |  |  | 80.3207 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk90 | 142 | 1 | 1 | TN | ug N/L | 0.961 | 13 | 571 | TP | ug P/L | 0.0982 | 1.74 | 49.9 |  | 1 | 1 |  | 571 | 571 | 571 |  |  |  |  | 97.71 | PASS | 1 |  | 12.2282 | 12.2282 | 12.2282 |  |  |  |  | 28.9476 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.466 | 6.36 | 247 | TP | ug P/L | 0.0916 | 1.67 | 47 |  | 1 | 1 |  | 247 | 247 | 247 | -1.2 | PASS |  |  |  |  | 1 |  | 2.2371 | 2.2371 | 2.2371 | -91.0515 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.0081 | 0.0293 | -59.8 | TP | ug P/L | 0.0867 | 1.44 | 35.6 |  | 1 | 1 |  | -59.8 | -59.8 | -59.8 | [6 empty] | 1 |  | -5.3783 | -5.3783 | -5.3783
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00635 | 0.0636 | -58.1 | TP | ug P/L | 0.0117 | 0.216 | -33.9 |  | 1 | 1 |  | -58.1 | -58.1 | -58.1 | [6 empty] | 1 |  | -142.9745 | -142.9745 | -142.9745
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s91 | 143 | 1 | 1 | TN | ug N/L | 1.83 | 24.9 | 1150 | TP | ug P/L | 0.0998 | 1.8 | 52.9 |  | 1 | 1 |  | 1150 | 1150 | 1150 | [6 empty] | 1 |  | 14.6043 | 14.6043 | 14.6043
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s92 | 144 | 1 | 1 | TN | ug N/L | 3.77 | 51.3 | 2480 | TP | ug P/L | 0.0967 | 1.72 | 49.2 |  | 1 | 1 |  | 2480 | 2480 | 2480 | [6 empty] | 1 |  | 9.9844 | 9.9844 | 9.9844
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s93 | 145 | 1 | 1 | TN | ug N/L | 0.522 | 7.09 | 282 | TP | ug P/L | 0.0972 | 1.82 | 53.8 |  | 1 | 1 |  | 282 | 282 | 282 | [6 empty] | 1 |  | 10.7341 | 10.7341 | 10.7341
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s94 | 146 | 1 | 1 | TN | ug N/L | 3.47 | 47 | 2270 | TP | ug P/L | 0.0983 | 1.76 | 51.1 |  | 1 | 1 |  | 2270 | 2270 | 2270 | [6 empty] | 1 |  | 12.3773 | 12.3773 | 12.3773
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s95 | 147 | 1 | 1 | TN | ug N/L | 1.43 | 19.3 | 881 | TP | ug P/L | 0.0975 | 1.74 | 50.1 |  | 1 | 1 |  | 881 | 881 | 881 | [6 empty] | 1 |  | 11.1831 | 11.1831 | 11.1831
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s96 | 148 | 1 | 1 | TN | ug N/L | 0.643 | 8.69 | 360 | TP | ug P/L | 0.0935 | 1.77 | 51.4 |  | 1 | 1 |  | 360 | 360 | 360 | [6 empty] | 1 |  | 5.1447 | 5.1447 | 5.1447
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s97 | 149 | 1 | 1 | TN | ug N/L | 0.752 | 10.2 | 432 | TP | ug P/L | 0.0925 | 1.75 | 50.6 |  | 1 | 1 |  | 432 | 432 | 432 | [6 empty] | 1 |  | 3.6176 | 3.6176 | 3.6176
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s98 | 150 | 1 | 1 | TN | ug N/L | 0.673 | 9.07 | 379 | TP | ug P/L | 0.0945 | 1.78 | 52.1 |  | 1 | 1 |  | 379 | 379 | 379 | [6 empty] | 1 |  | 6.6648 | 6.6648 | 6.6648
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s99 | 151 | 1 | 1 | TN | ug N/L | 0.665 | 9.02 | 376 | TP | ug P/L | 0.101 | 2 | 62.3 |  | 1 | 1 |  | 376 | 376 | 376 | [6 empty] | 1 |  | 16.3745 | 16.3745 | 16.3745
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s100 | 152 | 1 | 1 | TN | ug N/L | 1.32 | 17.8 | 808 | TP | ug P/L | 0.0962 | 1.79 | 52.5 |  | 1 | 1 |  | 808 | 808 | 808 | [6 empty] | 1 |  | 9.2329 | 9.2329 | 9.2329
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup95 | 153 | 1 | 1 | TN | ug N/L | 1.47 | 19.8 | 905 | TP | ug P/L | 0.0943 | 1.7 | 48.2 |  | 1 | 1 |  | 905 | 905 | 905 |  |  | 2.6876 | PASS |  |  | 1 |  | 6.3614 | 6.3614 | 6.3614 |  |  | 54.9656 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk100 | 154 | 1 | 1 | TN | ug N/L | 1.66 | 22.4 | 1030 | TP | ug P/L | 0.108 | 1.93 | 59 |  | 1 | 1 |  | 1030 | 1030 | 1030 |  |  |  |  | 99.1 | PASS | 1 |  | 26.4994 | 26.4994 | 26.4994 |  |  |  |  | 71.7157 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.473 | 6.38 | 248 | TP | ug P/L | 0.0919 | 1.75 | 50.8 |  | 1 | 1 |  | 248 | 248 | 248 | -0.8 | PASS |  |  |  |  | 1 |  | 2.6979 | 2.6979 | 2.6979 | -89.2084 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00967 | 0.041 | -59.2 | TP | ug P/L | 0.0894 | 1.56 | 41.7 |  | 1 | 1 |  | -59.2 | -59.2 | -59.2 | [6 empty] | 1 |  | -1.1612 | -1.1612 | -1.1612
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00841 | 0.057 | -58.5 | TP | ug P/L | 0.00999 | 0.135 | -39.1 |  | 1 | 1 |  | -58.5 | -58.5 | -58.5 | [6 empty] | 1 |  | -146.5721 | -146.5721 | -146.5721
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s101 | 155 | 1 | 1 | TN | ug N/L | 0.534 | 7.24 | 289 | TP | ug P/L | 0.0917 | 1.79 | 52.7 |  | 1 | 1 |  | 289 | 289 | 289 | [6 empty] | 1 |  | 2.3908 | 2.3908 | 2.3908
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s102 | 156 | 1 | 1 | TN | ug N/L | 0.787 | 10.7 | 460 | TP | ug P/L | 0.0936 | 1.76 | 51.2 |  | 1 | 1 |  | 460 | 460 | 460 | [6 empty] | 1 |  | 5.297 | 5.297 | 5.297
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s103 | 157 | 1 | 1 | TN | ug N/L | 1.22 | 16.6 | 747 | TP | ug P/L | 0.0929 | 1.74 | 49.9 |  | 1 | 1 |  | 747 | 747 | 747 | [6 empty] | 1 |  | 4.2293 | 4.2293 | 4.2293
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s104 | 158 | 1 | 1 | TN | ug N/L | 0.545 | 9.26 | 387 | TP | ug P/L | 0.955 | 26.9 | 2 |  | 1 | 1 |  | 387 | 387 | 387 | [6 empty] | 1 |  | -1287.3077 | -1287.3077 | -1287.3077
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s105 | 159 | 1 | 1 | TN | ug N/L | 0.016 | 0.0587 | -58.4 | TP | ug P/L | 0.144 | 1.5 | 38.6 |  | 1 | 3 | air spikes | -58.4 | -58.4 | -58.4 | [6 empty] | 3 | air spikes | 73.1366 | 73.1366 | 73.1366
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s106 | 160 | 1 | 1 | TN | ug N/L | 3.28 | 36.6 | 1740 | TP | ug P/L | 0.224 | 2.42 | 80 |  | 1 | 3 | air spikes | 1740 | 1740 | 1740 | [6 empty] | 3 | air spikes | 144.2022 | 144.2022 | 144.2022
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s107 | 161 | 1 | 1 | TN | ug N/L | -2.01 | -50.9 | -2450 | TP | ug P/L | 0.265 | 2.18 | 70 |  | 1 | 3 | air spikes | -2450 | -2450 | -2450 | [6 empty] | 3 | air spikes | 163.2102 | 163.2102 | 163.2102
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dups105 | 162 | 1 | 1 | TN | ug N/L | -2.1 | -59.8 | -2850 | TP | ug P/L | 0.156 | -1.05 | -122 |  | 1 | 3 | air spikes | -2850 | -2850 | -2850 |  |  | 276.9874 | FAIL |  |  | 3 | air spikes | 86.6606 | 86.6606 | 86.6606 |  |  | 176.9588 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spks107 | 163 | 1 | 1 | TN | ug N/L | 3.44 | 7.38 | 296 | TP | ug P/L | 0.292 | 2.83 | 95.9 |  | 1 | 3 | air spikes | 296 | 296 | 296 |  |  |  |  | 1101.36 | FAIL | 3 | air spikes | 169.2834 | 169.2834 | 169.2834 |  |  |  |  | 41.2208 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s108 | 164 | 1 | 1 | TN | ug N/L | -2.83 | -83.1 | -3890 | TP | ug P/L | 0.246 | 0.259 | -31.2 |  | 1 | 3 | air spikes | -3890 | -3890 | -3890 | -1656 | FAIL |  |  |  |  | 3 | air spikes | 155.8688 | 155.8688 | 155.8688 | 523.4754 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 1.78 | -6.12 | -356 | TP | ug P/L | 0.123 | 2.07 | 65.1 |  | 1 | 3 | air spikes | -356 | -356 | -356 | [6 empty] | 3 | air spikes | 47.0372 | 47.0372 | 47.0372
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.788 | -30.1 | -1490 | TP | ug P/L | -0.114 | -1.53 | -159 |  | 1 | 3 | air spikes | -1490 | -1490 | -1490 | [6 empty] | 3 | air spikes | -462.136 | -462.136 | -462.136
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.77 | -12.6 | -665 | TP | ug P/L | 0.0857 | 0.146 | -38.3 |  | 1 | 3 | air spikes | -665 | -665 | -665 | [6 empty] | 3 | air spikes | -6.9532 | -6.9532 | -6.9532
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal1500 | 1 | 1 | 1 | TN | ug N/L | 2.91 | 4.73 | 167 | TP | ug P/L | 0.35 | 5.8 | 160 |  | 1 | 3 | air spikes | 167 | 167 | 167 | [6 empty] | 3 | air spikes | 165.025 | 165.025 | 165.025
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal1000 | 3 | 1 | 1 | TN | ug N/L | -2.58 | -57 | -2720 | TP | ug P/L | 0.235 | 4.44 | 142 |  | 1 | 3 | air spikes | -2720 | -2720 | -2720 | [6 empty] | 3 | air spikes | 150.4602 | 150.4602 | 150.4602
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | call500 | 5 | 1 | 1 | TN | ug N/L | 1.98 | 3.67 | 116 | TP | ug P/L | 0.214 | 1.09 | 16.9 |  | 1 | 3 | air spikes | 116 | 116 | 116 | [6 empty] | 3 | air spikes | 137.776 | 137.776 | 137.776
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | call250 | 7 | 1 | 1 | TN | ug N/L | 1.61 | -10.7 | -572 | TP | ug P/L | 0.116 | 1.96 | 60.3 |  | 1 | 3 | air spikes | -572 | -572 | -572 | [6 empty] | 3 | air spikes | 37.6494 | 37.6494 | 37.6494
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal100 | 9 | 1 | 1 | TN | ug N/L | -1.98 | -22.1 | -1110 | TP | ug P/L | 0.15 | 2.16 | 69.2 |  | 1 | 3 | air spikes | -1110 | -1110 | -1110 | [6 empty] | 3 | air spikes | 80.025 | 80.025 | 80.025
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | 2 | 3.26 | 96.1 | TP | ug P/L | -0.0399 | -0.234 | -63.2 |  | 1 | 3 | air spikes | 96.1 | 96.1 | 96.1 | [6 empty] | 3 | air spikes | -260.57 | -260.57 | -260.57
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | 4.66 | 29.1 | 1370 | TP | ug P/L | 0.118 | 0.104 | -41 |  | 1 | 3 | air spikes | 1370 | 1370 | 1370 | [6 empty] | 3 | air spikes | 40.3668 | 40.3668 | 40.3668
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | 2.46 | 14.2 | 630 | TP | ug P/L | -0.00146 | -0.00226 | -47.9 |  | 1 | 3 | air spikes | 630 | 630 | 630 | [6 empty] | 3 | air spikes | -171.1903 | -171.1903 | -171.1903
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | 2.95 | 12 | 520 | TP | ug P/L | -0.047 | -0.25 | -64.3 |  | 1 | 3 | air spikes | 520 | 520 | 520 | [6 empty] | 3 | air spikes | -278.2136 | -278.2136 | -278.2136
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | 9.64 | 69.9 | 3440 | TP | ug P/L | -0.0307 | -0.173 | -59.1 |  | 1 | 3 | air spikes | 3440 | 3440 | 3440 | [6 empty] | 3 | air spikes | -238.2341 | -238.2341 | -238.2341
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.158 | 2.03 | 36.9 | TP | ug P/L | 0.0949 | 1.9 | 57.7 |  | 1 | 1 |  | 36.9 | 36.9 | 36.9 | [6 empty] | 1 |  | 7.2709 | 7.2709 | 7.2709
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.142 | 1.94 | 32.4 | TP | ug P/L | 0.0942 | 1.77 | 51.7 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 | [6 empty] | 1 |  | 6.2095 | 6.2095 | 6.2095 | -75.1619 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0919 | 1.28 | 0.497 | TP | ug P/L | 0.0876 | 1.54 | 40.2 |  | 1 | 1 |  | 0.497 | 0.497 | 0.497 | [6 empty] | 1 |  | -3.9669 | -3.9669 | -3.9669
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrite | S1 | 1 | 1 | TN | ug N/L | 2.27 | 30.8 | 1450 | TP | ug P/L | 0.0151 | 0.215 | -34 | 97.7273 | 1 | 1 |  | 1450 | 1450 | 1450 | [6 empty] | 1 |  | -135.8823 | -135.8823 | -135.8823
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrate | S2 | 1 | 1 | TN | ug N/L | 2.22 | 30.1 | 1410 | TP | ug P/L | 0.0165 | 0.248 | -31.9 |  | 1 | 1 |  | 1410 | 1410 | 1410 | [6 empty] | 1 |  | -132.9856 | -132.9856 | -132.9856
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0179 | -0.37 | -79 | TP | ug P/L | 0.0117 | 0.183 | -36 |  | 1 | 1 |  | -79 | -79 | -79 | [6 empty] | 1 |  | -142.9745 | -142.9745 | -142.9745
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s105 rerun | 159 | 1 | 1 | TN | ug N/L | 1.11 | 15.1 | 674 | TP | ug P/L | 0.0963 | 1.87 | 56 |  | 1 | 1 |  | 674 | 674 | 674 | [6 empty] | 1 |  | 9.3833 | 9.3833 | 9.3833
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s106 rerun | 160 | 1 | 1 | TN | ug N/L | 0.842 | 11.6 | 501 | TP | ug P/L | 0.0921 | 1.75 | 50.6 |  | 1 | 1 |  | 501 | 501 | 501 | [6 empty] | 1 |  | 3.0047 | 3.0047 | 3.0047
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s107 rerun | 161 | 1 | 1 | TN | ug N/L | 1.1 | 15.2 | 678 | TP | ug P/L | 0.0943 | 1.83 | 54.2 |  | 1 | 1 |  | 678 | 678 | 678 | [6 empty] | 1 |  | 6.3614 | 6.3614 | 6.3614
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dups105 rerun | 162 | 1 | 1 | TN | ug N/L | 1.1 | 15.3 | 682 | TP | ug P/L | 0.0916 | 1.65 | 46 |  | 1 | 1 |  | 682 | 682 | 682 | [6 empty] | 1 |  | 2.2371 | 2.2371 | 2.2371
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk107 rerun | 163 | 1 | 1 | TN | ug N/L | 1.42 | 19.7 | 897 | TP | ug P/L | 0.0989 | 1.72 | 49.4 |  | 1 | 1 |  | 897 | 897 | 897 | [6 empty] | 1 |  | 13.27 | 13.27 | 13.27
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s108 rerun | 164 | 1 | 1 | TN | ug N/L | 0.68 | 9.51 | 400 | TP | ug P/L | 0.0941 | 1.8 | 52.9 |  | 1 | 1 |  | 400 | 400 | 400 | [6 empty] | 1 |  | 6.0576 | 6.0576 | 6.0576
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal 1500 | 1 | 1 | 1 | TN | ug N/L | 2.32 | 32 | 1510 | TP | ug P/L | 0.24 | 4.63 | 145 |  | 1 | 1 |  | 1510 | 1510 | 1510 | [6 empty] | 1 |  | 153.024 | 153.024 | 153.024
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal 50/5 | 11 | 1 | 1 | TN | ug N/L | 0.159 | 2.28 | 49 | TP | ug P/L | 0.0933 | 1.74 | 50.3 |  | 1 | 1 |  | 49 | 49 | 49 | [6 empty] | 1 |  | 4.8398 | 4.8398 | 4.8398
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | chk 25/250 | 7 | 1 | 1 | TN | ug N/L | 0.451 | 6.34 | 246 | TP | ug P/L | 0.0887 | 1.6 | 43.4 |  | 1 | 1 |  | 246 | 246 | 246 | [6 empty] | 1 |  | -2.2496 | -2.2496 | -2.2496 | -108.9984 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal 25/2.5 | 13 | 1 | 1 | TN | ug N/L | 0.123 | 1.78 | 24.7 | TP | ug P/L | 0.0968 | 1.84 | 54.8 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | [6 empty] | 1 |  | 10.1345 | 10.1345 | 10.1345
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate | S4 | 1 | 1 | TN | ug N/L | -0.00258 | -0.0295 | -62.6 | TP | ug P/L | 0.0878 | 1.6 | 43.4 |  | 1 | 1 |  | -62.6 | -62.6 | -62.6 | [6 empty] | 1 |  | -3.654 | -3.654 | -3.654

## QAQC dip area
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE NO3-N: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- dilution corrected | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- dilution corrected | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | dye | S16 | 1 | 1 | TN | ug N/L | -0.00637 | -0.00462 | 0 | TP | ug P/L | 0.0308 | 0.664 | 0.166 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -45.6906 | -45.6906 | -45.6906
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | dye | S16 | 1 | 1 | TN | ug N/L | 0.0144 | 0.152 | 0 | TP | ug P/L | 0.0324 | 0.702 | 1.33 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -43.5464 | -43.5464 | -43.5464
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.27 | 30 | 0 | TP | ug P/L | -0.00742 | -0.058 | -21.6 | 98.3333 | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -87.3041 | -87.3041 | -87.3041
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.25 | 29.5 | 0 | TP | ug P/L | -0.00878 | -0.0697 | -22 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -87.9921 | -87.9921 | -87.9921
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0162 | -0.271 | 0 | TP | ug P/L | 0.00998 | 0.228 | -13 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -70.6219 | -70.6219 | -70.6219
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00707 | -0.00498 | 0 | TP | ug P/L | 0.0865 | 1.42 | 22.9 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | -3.8941 | -3.8941 | -3.8941
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.238 | 4.76 | 150 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 1 |  | 159.0104 | 159.0104 | 159.0104
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.236 | 4.54 | 150 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 1 |  | 149.3716 | 149.3716 | 149.3716
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.66 | 21.8 | 1000 | TP | ug P/L | 0.165 | 3.31 | 100 |  | 1 | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 1 |  | 92.6458 | 92.6458 | 92.6458
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.67 | 22.6 | 1000 | TP | ug P/L | 0.16 | 3.2 | 100 |  | 1 | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 1 |  | 87.3384 | 87.3384 | 87.3384
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.867 | 11.4 | 500 | TP | ug P/L | 0.098 | 1.73 | 50 |  | 1 | 1 |  | 500 | 500 | 500 | [6 empty] | 1 |  | 12.7193 | 12.7193 | 12.7193
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.86 | 11.4 | 500 | TP | ug P/L | 0.0979 | 1.75 | 50 |  | 1 | 1 |  | 500 | 500 | 500 | [6 empty] | 1 |  | 13.7806 | 13.7806 | 13.7806
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.478 | 6.3 | 250 | TP | ug P/L | 0.0926 | 1.65 | 25 |  | 1 | 1 |  | 250 | 250 | 250 | [6 empty] | 1 |  | 8.4612 | 8.4612 | 8.4612 | -66.1551 | FAIL
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.474 | 6.29 | 250 | TP | ug P/L | 0.089 | 1.48 | 25 |  | 1 | 1 |  | 250 | 250 | 250 | [6 empty] | 1 |  | -0.65474 | -0.65474 | -0.65474 | -102.6189 | FAIL
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.43 | 100 | TP | ug P/L | 0.0942 | 1.65 | 10 |  | 1 | 1 |  | 100 | 100 | 100 | [6 empty] | 1 |  | 8.4612 | 8.4612 | 8.4612
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.44 | 100 | TP | ug P/L | 0.095 | 1.5 | 10 |  | 1 | 1 |  | 100 | 100 | 100 | [6 empty] | 1 |  | 0.4225 | 0.4225 | 0.4225
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.176 | 2.33 | 50 | TP | ug P/L | 0.0894 | 1.49 | 5 |  | 1 | 1 |  | 50 | 50 | 50 | [6 empty] | 1 |  | -0.11596 | -0.11596 | -0.11596
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.166 | 2.28 | 50 | TP | ug P/L | 0.092 | 1.54 | 5 |  | 1 | 1 |  | 50 | 50 | 50 | [6 empty] | 1 |  | 2.5732 | 2.5732 | 2.5732
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.131 | 1.8 | 25 | TP | ug P/L | 0.0934 | 1.57 | 2.5 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 4.1828 | 4.1828 | 4.1828
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.129 | 1.76 | 25 | TP | ug P/L | 0.0915 | 1.59 | 2.5 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 5.2543 | 5.2543 | 5.2543
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0914 | 1.27 | 0 | TP | ug P/L | 0.0916 | 1.54 | 0 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 2.5732 | 2.5732 | 2.5732
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0915 | 1.27 | 0 | TP | ug P/L | 0.0945 | 1.64 | 0 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 7.9275 | 7.9275 | 7.9275
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.2 | 29.8 | 1400 | TP | ug P/L | -0.0111 | -0.0948 | -89.5 | 99.6644 | 1 | 1 |  | 1400 | 1400 | 1400 | [6 empty] | 1 |  | -89.4696 | -89.4696 | -89.4696
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.22 | 29.7 | 1390 | TP | ug P/L | 0.0192 | 0.22 | -71.1 |  | 1 | 1 |  | 1390 | 1390 | 1390 | [6 empty] | 1 |  | -71.085 | -71.085 | -71.085
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0172 | -0.298 | -75.6 | TP | ug P/L | -0.00889 | -0.0728 | -88.2 |  | 1 | 1 |  | -75.6 | -75.6 | -75.6 | [6 empty] | 1 |  | -88.1745 | -88.1745 | -88.1745
2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00164 | 0.0147 | -60.5 | TP | ug P/L | 0.0932 | 1.68 | 10.2 |  | 1 | 1 |  | -60.5 | -60.5 | -60.5 | [6 empty] | 1 |  | 10.0604 | 10.0604 | 10.0604
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrite in water | S1 | 1 | 1 | TN | ug N/L | 2.24 | 30.5 | 1440 | TP | ug P/L | -0.0195 | -0.165 | -58.6 | 97.377 | 1 | 1 |  | 1440 | 1440 | 1440 | [6 empty] | 1 |  | -93.6123 | -93.6123 | -93.6123
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrate in water | S2 | 1 | 1 | TN | ug N/L | 2.21 | 29.7 | 1400 | TP | ug P/L | -0.015 | -0.122 | -55.7 |  | 1 | 1 |  | 1400 | 1400 | 1400 | [6 empty] | 1 |  | -91.0729 | -91.0729 | -91.0729
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0165 | -0.274 | -74.4 | TP | ug P/L | -0.0204 | -0.175 | -59.3 |  | 1 | 1 |  | -74.4 | -74.4 | -74.4 | [6 empty] | 1 |  | -94.2037 | -94.2037 | -94.2037
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.26 | 242 | TP | ug P/L | 0.0918 | 1.74 | 50.3 |  | 1 | 1 |  | 242 | 242 | 242 | -3.2 | PASS |  |  |  |  | 1 |  | 13.2501 | 13.2501 | 13.2501 | -46.9995 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.0025 | -0.0316 | -62.7 | TP | ug P/L | 0.0922 | 1.65 | 45.6 |  | 1 | 1 |  | -62.7 | -62.7 | -62.7 | [6 empty] | 1 |  | 8.4612 | 8.4612 | 8.4612
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0167 | -0.267 | -74.1 | TP | ug P/L | 0.0141 | 0.187 | -35.8 |  | 1 | 1 |  | -74.1 | -74.1 | -74.1 | [6 empty] | 1 |  | -72.9974 | -72.9974 | -72.9974
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00439 | -0.035 | -62.9 | TP | ug P/L | 0.0866 | 1.4 | 33.2 |  | 1 | 1 |  | -62.9 | -62.9 | -62.9 | [6 empty] | 1 |  | -4.9764 | -4.9764 | -4.9764
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | 20 | 1 | 1 | TN | ug N/L | 0.084 | 1.2 | -3.48 | TP | ug P/L | 0.0937 | 1.62 | 44.4 |  | 1 | 1 |  | -3.48 | -3.48 | -3.48 | [6 empty] | 1 |  | 6.8592 | 6.8592 | 6.8592
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | 21 | 1 | 1 | TN | ug N/L | 0.0936 | 1.22 | -2.22 | TP | ug P/L | 0.0875 | 1.47 | 37.2 |  | 1 | 1 |  | -2.22 | -2.22 | -2.22 | [6 empty] | 1 |  | -1.1938 | -1.1938 | -1.1938
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | 22 | 1 | 1 | TN | ug N/L | 0.0896 | 1.28 | 0.652 | TP | ug P/L | 0.0928 | 1.68 | 47.4 |  | 1 | 1 |  | 0.652 | 0.652 | 0.652 | [6 empty] | 1 |  | 10.0604 | 10.0604 | 10.0604
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | 23 | 1 | 1 | TN | ug N/L | 0.095 | 1.3 | 1.41 | TP | ug P/L | 0.0875 | 1.45 | 35.9 |  | 1 | 1 |  | 1.41 | 1.41 | 1.41 | [6 empty] | 1 |  | -2.273 | -2.273 | -2.273
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | 24 | 1 | 1 | TN | ug N/L | 0.0925 | 1.31 | 1.89 | TP | ug P/L | 0.0877 | 1.55 | 40.9 |  | 1 | 1 |  | 1.89 | 1.89 | 1.89 | [6 empty] | 1 |  | 3.11 | 3.11 | 3.11
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | digest | 25 | 1 | 1 | TN | ug N/L | 0.71 | 9.48 | 398 | TP | ug P/L | 0.0906 | 1.6 | 43.5 |  | 1 | 1 |  | 398 | 398 | 398 | [6 empty] | 1 |  | 5.7896 | 5.7896 | 5.7896
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | digest | 26 | 1 | 1 | TN | ug N/L | 0.707 | 9.5 | 399 | TP | ug P/L | 0.0889 | 1.6 | 43.6 |  | 1 | 1 |  | 399 | 399 | 399 | [6 empty] | 1 |  | 5.7896 | 5.7896 | 5.7896
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | digest | 27 | 1 | 1 | TN | ug N/L | 0.709 | 9.47 | 398 | TP | ug P/L | 0.0927 | 1.66 | 46.4 |  | 1 | 1 |  | 398 | 398 | 398 | [6 empty] | 1 |  | 8.9946 | 8.9946 | 8.9946
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | digest | 28 | 1 | 1 | TN | ug N/L | 0.711 | 9.46 | 397 | TP | ug P/L | 0.0949 | 1.89 | 57.1 |  | 1 | 1 |  | 397 | 397 | 397 | [6 empty] | 1 |  | 21.1744 | 21.1744 | 21.1744
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | digest | 29 | 1 | 1 | TN | ug N/L | 0.709 | 9.42 | 395 | TP | ug P/L | 0.0944 | 1.68 | 47.1 |  | 1 | 1 |  | 395 | 395 | 395 | [6 empty] | 1 |  | 10.0604 | 10.0604 | 10.0604
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk blank | 30 | 1 | 1 | TN | ug N/L | 0.433 | 5.75 | 217 | TP | ug P/L | 0.0889 | 1.65 | 46 |  | 1 | 1 |  | 217 | 217 | 217 | [6 empty] | 1 |  | 8.4612 | 8.4612 | 8.4612
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | skp blank | 31 | 1 | 1 | TN | ug N/L | 0.436 | 5.89 | 224 | TP | ug P/L | 0.0925 | 1.75 | 50.4 |  | 1 | 1 |  | 224 | 224 | 224 | [6 empty] | 1 |  | 13.7806 | 13.7806 | 13.7806
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk blank | 32 | 1 | 1 | TN | ug N/L | 0.441 | 5.9 | 224 | TP | ug P/L | 0.0954 | 1.73 | 49.4 |  | 1 | 1 |  | 224 | 224 | 224 | [6 empty] | 1 |  | 12.7193 | 12.7193 | 12.7193
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk blank | 33 | 1 | 1 | TN | ug N/L | 0.437 | 5.85 | 222 | TP | ug P/L | 0.0907 | 1.7 | 48.4 |  | 1 | 1 |  | 222 | 222 | 222 | [6 empty] | 1 |  | 11.1249 | 11.1249 | 11.1249
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk blank | 34 | 1 | 1 | TN | ug N/L | 0.446 | 5.92 | 225 | TP | ug P/L | 0.0879 | 1.64 | 45.4 |  | 1 | 1 |  | 225 | 225 | 225 | [6 empty] | 1 |  | 7.9275 | 7.9275 | 7.9275
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s1 | 35 | 1 | 1 | TN | ug N/L | 1.3 | 17.5 | 792 | TP | ug P/L | 0.0911 | 1.66 | 46.2 |  | 1 | 1 |  | 792 | 792 | 792 | [6 empty] | 1 |  | 8.9946 | 8.9946 | 8.9946
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s2 | 36 | 1 | 1 | TN | ug N/L | 0.538 | 7.16 | 285 | TP | ug P/L | 0.0918 | 1.71 | 48.7 |  | 1 | 1 |  | 285 | 285 | 285 | [6 empty] | 1 |  | 11.6567 | 11.6567 | 11.6567
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s3 | 37 | 1 | 1 | TN | ug N/L | 0.56 | 7.54 | 304 | TP | ug P/L | 0.0962 | 1.82 | 53.7 |  | 1 | 1 |  | 304 | 304 | 304 | [6 empty] | 1 |  | 17.4853 | 17.4853 | 17.4853
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s4 | 38 | 1 | 1 | TN | ug N/L | 0.958 | 12.8 | 560 | TP | ug P/L | 0.0943 | 1.76 | 51.2 |  | 1 | 1 |  | 560 | 560 | 560 | [6 empty] | 1 |  | 14.3108 | 14.3108 | 14.3108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s5 | 39 | 1 | 1 | TN | ug N/L | 3.36 | 45.1 | 2170 | TP | ug P/L | 0.0934 | 1.67 | 46.7 |  | 1 | 1 |  | 2170 | 2170 | 2170 | [6 empty] | 1 |  | 9.5276 | 9.5276 | 9.5276
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s6 | 40 | 1 | 1 | TN | ug N/L | 0.931 | 12.5 | 548 | TP | ug P/L | 0.0931 | 1.7 | 48.1 |  | 1 | 1 |  | 548 | 548 | 548 | [6 empty] | 1 |  | 11.1249 | 11.1249 | 11.1249
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s7 | 41 | 1 | 1 | TN | ug N/L | 0.886 | 11.9 | 515 | TP | ug P/L | 0.0948 | 1.73 | 49.5 |  | 1 | 1 |  | 515 | 515 | 515 | [6 empty] | 1 |  | 12.7193 | 12.7193 | 12.7193
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s8 | 42 | 1 | 1 | TN | ug N/L | 0.573 | 7.62 | 308 | TP | ug P/L | 0.0926 | 1.74 | 50 |  | 1 | 1 |  | 308 | 308 | 308 | [6 empty] | 1 |  | 13.2501 | 13.2501 | 13.2501
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s9 | 43 | 1 | 1 | TN | ug N/L | 1.6 | 21.5 | 988 | TP | ug P/L | 0.097 | 1.84 | 54.6 |  | 1 | 1 |  | 988 | 988 | 988 | [6 empty] | 1 |  | 18.5409 | 18.5409 | 18.5409
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | a10 | 44 | 1 | 1 | TN | ug N/L | 0.764 | 10.2 | 435 | TP | ug P/L | 0.0946 | 1.7 | 48.3 |  | 1 | 1 |  | 435 | 435 | 435 | [6 empty] | 1 |  | 11.1249 | 11.1249 | 11.1249
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup5 | 45 | 1 | 1 | TN | ug N/L | 3.39 | 45.6 | 2190 | TP | ug P/L | 0.0931 | 1.69 | 48 |  | 1 | 1 |  | 2190 | 2190 | 2190 |  |  | 0.91743 | PASS |  |  | 1 |  | 10.5928 | 10.5928 | 10.5928 |  |  | 10.5878 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk10 | 46 | 1 | 1 | TN | ug N/L | 1.07 | 14.4 | 639 | TP | ug P/L | 0.099 | 1.81 | 53.5 |  | 1 | 1 |  | 639 | 639 | 639 |  |  |  |  | 87.99 | PASS | 1 |  | 16.957 | 16.957 | 16.957 |  |  |  |  | 25.0241 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.28 | 243 | TP | ug P/L | 0.0914 | 1.7 | 48.2 |  | 1 | 1 |  | 243 | 243 | 243 | -2.8 | PASS |  |  |  |  | 1 |  | 11.1249 | 11.1249 | 11.1249 | -55.5004 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00254 | 0.00654 | -60.9 | TP | ug P/L | 0.0916 | 1.59 | 42.7 |  | 1 | 1 |  | -60.9 | -60.9 | -60.9 | [6 empty] | 1 |  | 5.2543 | 5.2543 | 5.2543
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.016 | -0.252 | -73.3 | TP | ug P/L | 0.00864 | 0.13 | -39.4 |  | 1 | 1 |  | -73.3 | -73.3 | -73.3 | [6 empty] | 1 |  | -76.3089 | -76.3089 | -76.3089
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s11 | 47 | 1 | 1 | TN | ug N/L | 1.62 | 21.9 | 1010 | TP | ug P/L | 0.0924 | 1.8 | 52.9 |  | 1 | 1 |  | 1010 | 1010 | 1010 | [6 empty] | 1 |  | 16.4284 | 16.4284 | 16.4284
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s12 | 48 | 1 | 1 | TN | ug N/L | 0.843 | 11.3 | 485 | TP | ug P/L | 0.093 | 1.71 | 48.7 |  | 1 | 1 |  | 485 | 485 | 485 | [6 empty] | 1 |  | 11.6567 | 11.6567 | 11.6567
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s13 | 49 | 1 | 1 | TN | ug N/L | 1.6 | 21.5 | 990 | TP | ug P/L | 0.0933 | 1.74 | 50.3 |  | 1 | 1 |  | 990 | 990 | 990 | [6 empty] | 1 |  | 13.2501 | 13.2501 | 13.2501
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s14 | 50 | 1 | 1 | TN | ug N/L | 1.67 | 22.6 | 1040 | TP | ug P/L | 0.0942 | 1.67 | 46.7 |  | 1 | 1 |  | 1040 | 1040 | 1040 | [6 empty] | 1 |  | 9.5276 | 9.5276 | 9.5276
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s15 | 51 | 1 | 1 | TN | ug N/L | 0.54 | 7.23 | 289 | TP | ug P/L | 0.0898 | 1.68 | 47.3 |  | 1 | 1 |  | 289 | 289 | 289 | [6 empty] | 1 |  | 10.0604 | 10.0604 | 10.0604
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s16 | 52 | 1 | 1 | TN | ug N/L | 0.594 | 8.01 | 327 | TP | ug P/L | 0.0935 | 1.8 | 52.7 |  | 1 | 1 |  | 327 | 327 | 327 | [6 empty] | 1 |  | 16.4284 | 16.4284 | 16.4284
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s17 | 53 | 1 | 1 | TN | ug N/L | 0.798 | 10.7 | 457 | TP | ug P/L | 0.0945 | 1.78 | 51.9 |  | 1 | 1 |  | 457 | 457 | 457 | [6 empty] | 1 |  | 15.3702 | 15.3702 | 15.3702
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s18 | 54 | 1 | 1 | TN | ug N/L | 3.33 | 44.7 | 2150 | TP | ug P/L | 0.0962 | 1.71 | 48.9 |  | 1 | 1 |  | 2150 | 2150 | 2150 | [6 empty] | 1 |  | 11.6567 | 11.6567 | 11.6567
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s19 | 55 | 1 | 1 | TN | ug N/L | 2.53 | 34 | 1610 | TP | ug P/L | 0.0935 | 1.65 | 45.8 |  | 1 | 1 |  | 1610 | 1610 | 1610 | [6 empty] | 1 |  | 8.4612 | 8.4612 | 8.4612
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s20 | 56 | 1 | 1 | TN | ug N/L | 2.49 | 33.3 | 1580 | TP | ug P/L | 0.0934 | 1.71 | 48.7 |  | 1 | 1 |  | 1580 | 1580 | 1580 | [6 empty] | 1 |  | 11.6567 | 11.6567 | 11.6567
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup15 | 57 | 1 | 1 | TN | ug N/L | 0.54 | 7.19 | 287 | TP | ug P/L | 0.0899 | 1.64 | 45.3 |  | 1 | 1 |  | 287 | 287 | 287 |  |  | 0.69444 | PASS |  |  | 1 |  | 7.9275 | 7.9275 | 7.9275 |  |  | 23.7142 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk20 | 58 | 1 | 1 | TN | ug N/L | 2.81 | 37.7 | 1800 | TP | ug P/L | 0.106 | 1.97 | 60.7 |  | 1 | 1 |  | 1800 | 1800 | 1800 |  |  |  |  | 106 | PASS | 1 |  | 25.3714 | 25.3714 | 25.3714 |  |  |  |  | 57.3959 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.35 | 246 | TP | ug P/L | 0.0924 | 1.7 | 48.1 |  | 1 | 1 |  | 246 | 246 | 246 | -1.6 | PASS |  |  |  |  | 1 |  | 11.1249 | 11.1249 | 11.1249 | -55.5004 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00264 | -0.0321 | -62.8 | TP | ug P/L | 0.0865 | 1.47 | 37 |  | 1 | 1 |  | -62.8 | -62.8 | -62.8 | [6 empty] | 1 |  | -1.1938 | -1.1938 | -1.1938
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00701 | 0.037 | -59.4 | TP | ug P/L | 0.0123 | 0.214 | -34.1 |  | 1 | 1 |  | -59.4 | -59.4 | -59.4 | [6 empty] | 1 |  | -71.4324 | -71.4324 | -71.4324
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s21 | 59 | 1 | 1 | TN | ug N/L | 0.554 | 7.45 | 299 | TP | ug P/L | 0.0964 | 1.85 | 55.3 |  | 1 | 1 |  | 299 | 299 | 299 | [6 empty] | 1 |  | 19.0682 | 19.0682 | 19.0682
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s22 | 60 | 1 | 1 | TN | ug N/L | 0.736 | 9.94 | 421 | TP | ug P/L | 0.0929 | 1.79 | 52.6 |  | 1 | 1 |  | 421 | 421 | 421 | [6 empty] | 1 |  | 15.8995 | 15.8995 | 15.8995
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s23 | 61 | 1 | 1 | TN | ug N/L | 0.615 | 8.28 | 340 | TP | ug P/L | 0.0914 | 1.7 | 48 |  | 1 | 1 |  | 340 | 340 | 340 | [6 empty] | 1 |  | 11.1249 | 11.1249 | 11.1249
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s24 | 62 | 1 | 1 | TN | ug N/L | 1.04 | 14 | 622 | TP | ug P/L | 0.0976 | 1.86 | 55.7 |  | 1 | 1 |  | 622 | 622 | 622 | [6 empty] | 1 |  | 19.5952 | 19.5952 | 19.5952
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s25 | 63 | 1 | 1 | TN | ug N/L | 3.03 | 40.8 | 1950 | TP | ug P/L | 0.101 | 1.75 | 50.8 |  | 1 | 1 |  | 1950 | 1950 | 1950 | [6 empty] | 1 |  | 13.7806 | 13.7806 | 13.7806
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s26 | 64 | 1 | 1 | TN | ug N/L | 1.17 | 15.8 | 708 | TP | ug P/L | 0.0963 | 1.81 | 53.7 |  | 1 | 1 |  | 708 | 708 | 708 | [6 empty] | 1 |  | 16.957 | 16.957 | 16.957
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s27 | 65 | 1 | 1 | TN | ug N/L | 0.5 | 6.71 | 264 | TP | ug P/L | 0.0945 | 1.75 | 50.7 |  | 1 | 1 |  | 264 | 264 | 264 | [6 empty] | 1 |  | 13.7806 | 13.7806 | 13.7806
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s28 | 66 | 1 | 1 | TN | ug N/L | 2.56 | 34.5 | 1630 | TP | ug P/L | 0.149 | 3.02 | 103 |  | 1 | 1 |  | 1630 | 1630 | 1630 | [6 empty] | 1 |  | 78.5706 | 78.5706 | 78.5706
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s29 | 67 | 1 | 1 | TN | ug N/L | 2.18 | 29.3 | 1370 | TP | ug P/L | 0.0938 | 1.71 | 48.9 |  | 1 | 1 |  | 1370 | 1370 | 1370 | [6 empty] | 1 |  | 11.6567 | 11.6567 | 11.6567
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s30 | 68 | 1 | 1 | TN | ug N/L | 1.14 | 15.3 | 682 | TP | ug P/L | 0.0936 | 1.75 | 50.5 |  | 1 | 1 |  | 682 | 682 | 682 | [6 empty] | 1 |  | 13.7806 | 13.7806 | 13.7806
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup25 | 69 | 1 | 1 | TN | ug N/L | 3 | 40.5 | 1930 | TP | ug P/L | 0.1 | 1.75 | 50.7 |  | 1 | 1 |  | 1930 | 1930 | 1930 |  |  | 1.0309 | PASS |  |  | 1 |  | 13.7806 | 13.7806 | 13.7806 |  |  | 0 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup30 | 70 | 1 | 1 | TN | ug N/L | 1.46 | 19.7 | 900 | TP | ug P/L | 0.0988 | 1.78 | 51.8 |  | 1 | 1 |  | 900 | 900 | 900 |  |  |  |  | 96.2 | PASS | 1 |  | 15.3702 | 15.3702 | 15.3702 |  |  |  |  | 7.8955 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.469 | 6.28 | 243 | TP | ug P/L | 0.0919 | 1.69 | 47.6 |  | 1 | 1 |  | 243 | 243 | 243 | -2.8 | PASS |  |  |  |  | 1 |  | 10.5928 | 10.5928 | 10.5928 | -57.6288 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00521 | -0.0319 | -62.7 | TP | ug P/L | 0.0914 | 1.67 | 46.6 |  | 1 | 1 |  | -62.7 | -62.7 | -62.7 | [6 empty] | 1 |  | 9.5276 | 9.5276 | 9.5276
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0169 | -0.275 | -74.4 | TP | ug P/L | 0.0107 | 0.18 | -36.2 |  | 1 | 1 |  | -74.4 | -74.4 | -74.4 | [6 empty] | 1 |  | -73.4035 | -73.4035 | -73.4035
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s31 | 71 | 1 | 1 | TN | ug N/L | 0.491 | 6.8 | 268 | TP | ug P/L | 0.0921 | 1.73 | 49.6 |  | 1 | 1 |  | 268 | 268 | 268 | [6 empty] | 1 |  | 12.7193 | 12.7193 | 12.7193
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s32 | 72 | 1 | 1 | TN | ug N/L | 1.13 | 15.2 | 679 | TP | ug P/L | 0.094 | 1.78 | 52.1 |  | 1 | 1 |  | 679 | 679 | 679 | [6 empty] | 1 |  | 15.3702 | 15.3702 | 15.3702
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s33 | 73 | 1 | 1 | TN | ug N/L | 0.853 | 11.5 | 495 | TP | ug P/L | 0.0967 | 1.79 | 52.5 |  | 1 | 1 |  | 495 | 495 | 495 | [6 empty] | 1 |  | 15.8995 | 15.8995 | 15.8995
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s34 | 74 | 1 | 1 | TN | ug N/L | 1.37 | 18.4 | 837 | TP | ug P/L | 0.0959 | 1.72 | 49.2 |  | 1 | 1 |  | 837 | 837 | 837 | [6 empty] | 1 |  | 12.1881 | 12.1881 | 12.1881
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s35 | 75 | 1 | 1 | TN | ug N/L | 0.745 | 10 | 425 | TP | ug P/L | 0.094 | 1.78 | 52 |  | 1 | 1 |  | 425 | 425 | 425 | [6 empty] | 1 |  | 15.3702 | 15.3702 | 15.3702
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s36 | 76 | 1 | 1 | TN | ug N/L | 1.64 | 22.1 | 1020 | TP | ug P/L | 0.101 | 1.82 | 53.7 |  | 1 | 1 |  | 1020 | 1020 | 1020 | [6 empty] | 1 |  | 17.4853 | 17.4853 | 17.4853
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s37 | 77 | 1 | 1 | TN | ug N/L | 2.57 | 34.9 | 1650 | TP | ug P/L | 0.0937 | 1.72 | 49 |  | 1 | 1 |  | 1650 | 1650 | 1650 | [6 empty] | 1 |  | 12.1881 | 12.1881 | 12.1881
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s38 | 78 | 1 | 1 | TN | ug N/L | 0.752 | 10.1 | 428 | TP | ug P/L | 0.0925 | 1.72 | 49.3 |  | 1 | 1 |  | 428 | 428 | 428 | [6 empty] | 1 |  | 12.1881 | 12.1881 | 12.1881
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s39 | 79 | 1 | 1 | TN | ug N/L | 1.17 | 15.8 | 707 | TP | ug P/L | 0.0958 | 1.86 | 55.6 |  | 1 | 1 |  | 707 | 707 | 707 | [6 empty] | 1 |  | 19.5952 | 19.5952 | 19.5952
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s40 | 80 | 1 | 1 | TN | ug N/L | 1.76 | 23.7 | 1100 | TP | ug P/L | 0.093 | 1.73 | 49.4 |  | 1 | 1 |  | 1100 | 1100 | 1100 | [6 empty] | 1 |  | 12.7193 | 12.7193 | 12.7193
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup35 | 81 | 1 | 1 | TN | ug N/L | 0.746 | 10.1 | 430 | TP | ug P/L | 0.0971 | 1.86 | 55.6 |  | 1 | 1 |  | 430 | 430 | 430 |  |  | 1.1696 | PASS |  |  | 1 |  | 19.5952 | 19.5952 | 19.5952 |  |  | 24.1666 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk40 | 82 | 1 | 1 | TN | ug N/L | 2.09 | 28.1 | 1310 | TP | ug P/L | 0.101 | 1.81 | 53.3 |  | 1 | 1 |  | 1310 | 1310 | 1310 |  |  |  |  | 97.1 | PASS | 1 |  | 16.957 | 16.957 | 16.957 |  |  |  |  | 18.6465 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.47 | 6.3 | 244 | TP | ug P/L | 0.0913 | 1.67 | 47 |  | 1 | 1 |  | 244 | 244 | 244 | -2.4 | PASS |  |  |  |  | 1 |  | 9.5276 | 9.5276 | 9.5276 | -61.8894 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00528 | -0.0552 | -63.9 | TP | ug P/L | 0.0899 | 1.6 | 43.3 |  | 1 | 1 |  | -63.9 | -63.9 | -63.9 | [6 empty] | 1 |  | 5.7896 | 5.7896 | 5.7896
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0144 | -0.226 | -72.1 | TP | ug P/L | 0.0111 | 0.153 | -37.9 |  | 1 | 1 |  | -72.1 | -72.1 | -72.1 | [6 empty] | 1 |  | -74.9714 | -74.9714 | -74.9714
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s41 | 83 | 1 | 1 | TN | ug N/L | 1.16 | 15.7 | 703 | TP | ug P/L | 0.0937 | 1.83 | 54.5 |  | 1 | 1 |  | 703 | 703 | 703 | [6 empty] | 1 |  | 18.0132 | 18.0132 | 18.0132
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s42 | 84 | 1 | 1 | TN | ug N/L | 0.5 | 6.77 | 266 | TP | ug P/L | 0.095 | 1.81 | 53.6 |  | 1 | 1 |  | 266 | 266 | 266 | [6 empty] | 1 |  | 16.957 | 16.957 | 16.957
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s43 | 85 | 1 | 1 | TN | ug N/L | 1.02 | 13.8 | 610 | TP | ug P/L | 0.0994 | 1.8 | 53.1 |  | 1 | 1 |  | 610 | 610 | 610 | [6 empty] | 1 |  | 16.4284 | 16.4284 | 16.4284
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s44 | 86 | 1 | 1 | TN | ug N/L | 2.55 | 34.5 | 1640 | TP | ug P/L | 0.0971 | 1.75 | 50.7 |  | 1 | 1 |  | 1640 | 1640 | 1640 | [6 empty] | 1 |  | 13.7806 | 13.7806 | 13.7806
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s45 | 87 | 1 | 1 | TN | ug N/L | 1.46 | 19.7 | 901 | TP | ug P/L | 0.0916 | 1.69 | 47.9 |  | 1 | 1 |  | 901 | 901 | 901 | [6 empty] | 1 |  | 10.5928 | 10.5928 | 10.5928
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s46 | 88 | 1 | 1 | TN | ug N/L | 2.42 | 32.6 | 1540 | TP | ug P/L | 0.139 | 2.78 | 93.9 |  | 1 | 1 |  | 1540 | 1540 | 1540 | [6 empty] | 1 |  | 66.7198 | 66.7198 | 66.7198
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s47 | 89 | 1 | 1 | TN | ug N/L | 0.713 | 9.56 | 402 | TP | ug P/L | 0.0944 | 1.82 | 53.7 |  | 1 | 1 |  | 402 | 402 | 402 | [6 empty] | 1 |  | 17.4853 | 17.4853 | 17.4853
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s48 | 90 | 1 | 1 | TN | ug N/L | 0.975 | 13.2 | 581 | TP | ug P/L | 0.107 | 2.18 | 69.9 |  | 1 | 1 |  | 581 | 581 | 581 | [6 empty] | 1 |  | 36.2917 | 36.2917 | 36.2917
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s49 | 91 | 1 | 1 | TN | ug N/L | 1.1 | 15 | 666 | TP | ug P/L | 0.0962 | 1.8 | 52.8 |  | 1 | 1 |  | 666 | 666 | 666 | [6 empty] | 1 |  | 16.4284 | 16.4284 | 16.4284
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s50 | 92 | 1 | 1 | TN | ug N/L | 0.825 | 11.1 | 478 | TP | ug P/L | 0.0917 | 1.76 | 51.3 |  | 1 | 1 |  | 478 | 478 | 478 | [6 empty] | 1 |  | 14.3108 | 14.3108 | 14.3108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup45 | 93 | 1 | 1 | TN | ug N/L | 1.45 | 19.7 | 899 | TP | ug P/L | 0.0951 | 1.82 | 53.7 |  | 1 | 1 |  | 899 | 899 | 899 |  |  | 0.22222 | PASS |  |  | 1 |  | 17.4853 | 17.4853 | 17.4853 |  |  | 49.0951 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk50 | 94 | 1 | 1 | TN | ug N/L | 1.2 | 16.2 | 726 | TP | ug P/L | 0.101 | 1.86 | 56 |  | 1 | 1 |  | 726 | 726 | 726 |  |  |  |  | 106.46 | PASS | 1 |  | 19.5952 | 19.5952 | 19.5952 |  |  |  |  | 23.0972 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.471 | 6.35 | 246 | TP | ug P/L | 0.0921 | 1.75 | 50.8 |  | 1 | 1 |  | 246 | 246 | 246 | -1.6 | PASS |  |  |  |  | 1 |  | 13.7806 | 13.7806 | 13.7806 | -44.8775 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00351 | -0.0324 | -62.8 | TP | ug P/L | 0.0885 | 1.53 | 39.7 |  | 1 | 1 |  | -62.8 | -62.8 | -62.8 | [6 empty] | 1 |  | 2.036 | 2.036 | 2.036
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.016 | -0.27 | -74.2 | TP | ug P/L | 0.0105 | 0.169 | -36.9 |  | 1 | 1 |  | -74.2 | -74.2 | -74.2 | [6 empty] | 1 |  | -74.042 | -74.042 | -74.042
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal1500 | 1 | 1 | 1 | TN | ug N/L | 2.37 | 31.9 | 1500 | TP | ug P/L | 0.259 | 4.6 | 145 |  | 1 | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 1 |  | 152.0156 | 152.0156 | 152.0156
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal1000 | 3 | 1 | 1 | TN | ug N/L | 1.62 | 21.9 | 1010 | TP | ug P/L | 0.159 | 3.17 | 108 |  | 1 | 1 |  | 1010 | 1010 | 1010 | [6 empty] | 1 |  | 85.8842 | 85.8842 | 85.8842
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | call500 | 5 | 1 | 1 | TN | ug N/L | 0.848 | 11.5 | 495 | TP | ug P/L | 0.0962 | 1.8 | 52.8 |  | 1 | 1 |  | 495 | 495 | 495 | [6 empty] | 1 |  | 16.4284 | 16.4284 | 16.4284
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | call250 | 7 | 1 | 1 | TN | ug N/L | 0.471 | 6.35 | 246 | TP | ug P/L | 0.0908 | 1.68 | 47.3 |  | 1 | 1 |  | 246 | 246 | 246 | [6 empty] | 1 |  | 10.0604 | 10.0604 | 10.0604 | -59.7585 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal100 | 9 | 1 | 1 | TN | ug N/L | 0.254 | 3.45 | 105 | TP | ug P/L | 0.0907 | 1.76 | 51.3 |  | 1 | 1 |  | 105 | 105 | 105 | [6 empty] | 1 |  | 14.3108 | 14.3108 | 14.3108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.172 | 2.35 | 52.2 | TP | ug P/L | 0.0944 | 1.88 | 56.8 |  | 1 | 1 |  | 52.2 | 52.2 | 52.2 | [6 empty] | 1 |  | 20.6483 | 20.6483 | 20.6483
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.135 | 1.81 | 26.3 | TP | ug P/L | 0.0929 | 1.76 | 51.1 |  | 1 | 1 |  | 26.3 | 26.3 | 26.3 | [6 empty] | 1 |  | 14.3108 | 14.3108 | 14.3108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0958 | 1.29 | 1.07 | TP | ug P/L | 0.0887 | 1.56 | 41.2 |  | 1 | 1 |  | 1.07 | 1.07 | 1.07 | [6 empty] | 1 |  | 3.6466 | 3.6466 | 3.6466
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrite | S1 | 1 | 1 | TN | ug N/L | 2.19 | 30.1 | 1420 | TP | ug P/L | 0.00586 | 0.0804 | -42.5 | 99.6678 | 1 | 1 |  | 1420 | 1420 | 1420 | [6 empty] | 1 |  | -79.1988 | -79.1988 | -79.1988
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrate | S2 | 1 | 1 | TN | ug N/L | 2.21 | 30 | 1410 | TP | ug P/L | 0.00889 | 0.119 | -40.1 |  | 1 | 1 |  | 1410 | 1410 | 1410 | [6 empty] | 1 |  | -76.9491 | -76.9491 | -76.9491
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s51 | 95 | 1 | 1 | TN | ug N/L | 0.609 | 8.22 | 337 | TP | ug P/L | 0.0929 | 1.8 | 53 |  | 1 | 1 |  | 337 | 337 | 337 | [6 empty] | 1 |  | 16.4284 | 16.4284 | 16.4284
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s52 | 96 | 1 | 1 | TN | ug N/L | 0.837 | 11.3 | 486 | TP | ug P/L | 0.0938 | 1.83 | 54.4 |  | 1 | 1 |  | 486 | 486 | 486 | [6 empty] | 1 |  | 18.0132 | 18.0132 | 18.0132
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s53 | 97 | 1 | 1 | TN | ug N/L | 4.93 | 67.1 | 3290 | TP | ug P/L | 0.1 | 1.82 | 54.1 |  | 1 | 1 |  | 3290 | 3290 | 3290 | [6 empty] | 1 |  | 17.4853 | 17.4853 | 17.4853
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s54 | 98 | 1 | 1 | TN | ug N/L | 0.626 | 8.5 | 351 | TP | ug P/L | 0.0937 | 1.79 | 52.4 |  | 1 | 1 |  | 351 | 351 | 351 | [6 empty] | 1 |  | 15.8995 | 15.8995 | 15.8995
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s55 | 99 | 1 | 1 | TN | ug N/L | 0.728 | 9.89 | 418 | TP | ug P/L | 0.0916 | 1.67 | 46.7 |  | 1 | 1 |  | 418 | 418 | 418 | [6 empty] | 1 |  | 9.5276 | 9.5276 | 9.5276
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s56 | 100 | 1 | 1 | TN | ug N/L | 1.13 | 15.4 | 686 | TP | ug P/L | 0.0931 | 1.67 | 46.9 |  | 1 | 1 |  | 686 | 686 | 686 | [6 empty] | 1 |  | 9.5276 | 9.5276 | 9.5276
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s57 | 101 | 1 | 1 | TN | ug N/L | 2.45 | 33.1 | 1560 | TP | ug P/L | 0.0929 | 1.7 | 48.4 |  | 1 | 1 |  | 1560 | 1560 | 1560 | [6 empty] | 1 |  | 11.1249 | 11.1249 | 11.1249
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s58 | 102 | 1 | 1 | TN | ug N/L | 2.77 | 37.4 | 1780 | TP | ug P/L | 0.0983 | 1.77 | 51.6 |  | 1 | 1 |  | 1780 | 1780 | 1780 | [6 empty] | 1 |  | 14.8407 | 14.8407 | 14.8407
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s59 | 103 | 1 | 1 | TN | ug N/L | 1.08 | 14.7 | 654 | TP | ug P/L | 0.091 | 1.71 | 48.8 |  | 1 | 1 |  | 654 | 654 | 654 | [6 empty] | 1 |  | 11.6567 | 11.6567 | 11.6567
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s61 | 104 | 1 | 1 | TN | ug N/L | 3.12 | 42.4 | 2030 | TP | ug P/L | 0.1 | 1.79 | 52.3 |  | 1 | 1 |  | 2030 | 2030 | 2030 | [6 empty] | 1 |  | 15.8995 | 15.8995 | 15.8995
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup55 | 105 | 1 | 1 | TN | ug N/L | 0.729 | 9.84 | 416 | TP | ug P/L | 0.094 | 1.72 | 49.1 |  | 1 | 1 |  | 416 | 416 | 416 |  |  | 0.47962 | PASS |  |  | 1 |  | 12.1881 | 12.1881 | 12.1881 |  |  | 24.5029 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk60 | 106 | 1 | 1 | TN | ug N/L | 3.47 | 47.2 | 2280 | TP | ug P/L | 0.136 | 2.68 | 90.3 |  | 1 | 1 |  | 2280 | 2280 | 2280 |  |  |  |  | 122.8 | FAIL | 1 |  | 61.728 | 61.728 | 61.728 |  |  |  |  | 189.4868 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.462 | 6.23 | 240 | TP | ug P/L | 0.0925 | 1.75 | 50.8 |  | 1 | 1 |  | 240 | 240 | 240 | -4 | PASS |  |  |  |  | 1 |  | 13.7806 | 13.7806 | 13.7806 | -44.8775 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00622 | -0.0368 | -63 | TP | ug P/L | 0.0853 | 1.42 | 34.6 |  | 1 | 1 |  | -63 | -63 | -63 | [6 empty] | 1 |  | -3.8941 | -3.8941 | -3.8941
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0176 | -0.254 | -73.4 | TP | ug P/L | 0.00903 | 0.146 | -38.3 |  | 1 | 1 |  | -73.4 | -73.4 | -73.4 | [6 empty] | 1 |  | -75.3783 | -75.3783 | -75.3783
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s61 | 107 | 1 | 1 | TN | ug N/L | 1.51 | 20.4 | 936 | TP | ug P/L | 0.0951 | 1.79 | 52.4 |  | 1 | 1 |  | 936 | 936 | 936 | [6 empty] | 1 |  | 15.8995 | 15.8995 | 15.8995
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s62 | 108 | 1 | 1 | TN | ug N/L | 1.14 | 15.5 | 692 | TP | ug P/L | 0.0941 | 1.77 | 51.7 |  | 1 | 1 |  | 692 | 692 | 692 | [6 empty] | 1 |  | 14.8407 | 14.8407 | 14.8407
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s63 | 109 | 1 | 1 | TN | ug N/L | 0.577 | 7.82 | 317 | TP | ug P/L | 0.0933 | 1.78 | 52 |  | 1 | 1 |  | 317 | 317 | 317 | [6 empty] | 1 |  | 15.3702 | 15.3702 | 15.3702
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s64 | 110 | 1 | 1 | TN | ug N/L | 1.56 | 21.2 | 971 | TP | ug P/L | 0.098 | 1.86 | 55.6 |  | 1 | 1 |  | 971 | 971 | 971 | [6 empty] | 1 |  | 19.5952 | 19.5952 | 19.5952
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s65 | 111 | 1 | 1 | TN | ug N/L | 4.36 | 59.2 | 2890 | TP | ug P/L | 0.0967 | 1.71 | 48.8 |  | 1 | 1 |  | 2890 | 2890 | 2890 | [6 empty] | 1 |  | 11.6567 | 11.6567 | 11.6567
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s66 | 112 | 1 | 1 | TN | ug N/L | 2.17 | 29.5 | 1380 | TP | ug P/L | 0.0951 | 1.75 | 50.8 |  | 1 | 1 |  | 1380 | 1380 | 1380 | [6 empty] | 1 |  | 13.7806 | 13.7806 | 13.7806
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s67 | 113 | 1 | 1 | TN | ug N/L | 1.03 | 13.9 | 614 | TP | ug P/L | 0.0968 | 1.79 | 52.6 |  | 1 | 1 |  | 614 | 614 | 614 | [6 empty] | 1 |  | 15.8995 | 15.8995 | 15.8995
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s68 | 114 | 1 | 1 | TN | ug N/L | 3.14 | 42.6 | 2040 | TP | ug P/L | 0.103 | 1.87 | 56.2 |  | 1 | 1 |  | 2040 | 2040 | 2040 | [6 empty] | 1 |  | 20.1219 | 20.1219 | 20.1219
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s69 | 115 | 1 | 1 | TN | ug N/L | 0.668 | 9.05 | 378 | TP | ug P/L | 0.0963 | 1.81 | 53.6 |  | 1 | 1 |  | 378 | 378 | 378 | [6 empty] | 1 |  | 16.957 | 16.957 | 16.957
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s70 | 116 | 1 | 1 | TN | ug N/L | 1.88 | 25.7 | 1190 | TP | ug P/L | 0.0956 | 1.7 | 48.4 |  | 1 | 1 |  | 1190 | 1190 | 1190 | [6 empty] | 1 |  | 11.1249 | 11.1249 | 11.1249
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup65 | 117 | 1 | 1 | TN | ug N/L | 4.38 | 59.6 | 2910 | TP | ug P/L | 0.0948 | 1.68 | 47.1 |  | 1 | 1 |  | 2910 | 2910 | 2910 |  |  | 0.68966 | PASS |  |  | 1 |  | 10.0604 | 10.0604 | 10.0604 |  |  | 14.7009 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk70 | 118 | 1 | 1 | TN | ug N/L | 2.14 | 29 | 1360 | TP | ug P/L | 0.108 | 2.04 | 64 |  | 1 | 1 |  | 1360 | 1360 | 1360 |  |  |  |  | 81.6 | PASS | 1 |  | 29.0271 | 29.0271 | 29.0271 |  |  |  |  | 74.5113 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.459 | 6.21 | 239 | TP | ug P/L | 0.0896 | 1.67 | 47 |  | 1 | 1 |  | 239 | 239 | 239 | -4.4 | PASS |  |  |  |  | 1 |  | 9.5276 | 9.5276 | 9.5276 | -61.8894 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00409 | -0.0102 | -61.7 | TP | ug P/L | 0.0895 | 1.61 | 43.7 |  | 1 | 1 |  | -61.7 | -61.7 | -61.7 | [6 empty] | 1 |  | 6.3246 | 6.3246 | 6.3246
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0154 | -0.262 | -73.9 | TP | ug P/L | 0.0111 | 0.178 | -36.3 |  | 1 | 1 |  | -73.9 | -73.9 | -73.9 | [6 empty] | 1 |  | -73.5196 | -73.5196 | -73.5196
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s71 | 119 | 1 | 1 | TN | ug N/L | 1.83 | 24.9 | 1150 | TP | ug P/L | 0.0957 | 1.73 | 49.8 |  | 1 | 1 |  | 1150 | 1150 | 1150 | [6 empty] | 1 |  | 12.7193 | 12.7193 | 12.7193
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s72 | 120 | 1 | 1 | TN | ug N/L | 0.94 | 12.8 | 558 | TP | ug P/L | 0.0933 | 1.72 | 49.1 |  | 1 | 1 |  | 558 | 558 | 558 | [6 empty] | 1 |  | 12.1881 | 12.1881 | 12.1881
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s73 | 121 | 1 | 1 | TN | ug N/L | 1.17 | 16 | 718 | TP | ug P/L | 0.0923 | 1.76 | 51 |  | 1 | 1 |  | 718 | 718 | 718 | [6 empty] | 1 |  | 14.3108 | 14.3108 | 14.3108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s74 | 122 | 1 | 1 | TN | ug N/L | 1.25 | 16.9 | 762 | TP | ug P/L | 0.0948 | 1.78 | 51.9 |  | 1 | 1 |  | 762 | 762 | 762 | [6 empty] | 1 |  | 15.3702 | 15.3702 | 15.3702
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s75 | 123 | 1 | 1 | TN | ug N/L | 1.66 | 22.6 | 1040 | TP | ug P/L | 0.1 | 1.8 | 52.9 |  | 1 | 1 |  | 1040 | 1040 | 1040 | [6 empty] | 1 |  | 16.4284 | 16.4284 | 16.4284
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s76 | 124 | 1 | 1 | TN | ug N/L | 1.14 | 15.5 | 691 | TP | ug P/L | 0.0931 | 1.73 | 49.6 |  | 1 | 1 |  | 691 | 691 | 691 | [6 empty] | 1 |  | 12.7193 | 12.7193 | 12.7193
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s77 | 125 | 1 | 1 | TN | ug N/L | 1.87 | 25.5 | 1190 | TP | ug P/L | 0.124 | 2.39 | 78.9 |  | 1 | 1 |  | 1190 | 1190 | 1190 | [6 empty] | 1 |  | 47.0718 | 47.0718 | 47.0718
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s78 | 126 | 1 | 1 | TN | ug N/L | 1.34 | 18.1 | 819 | TP | ug P/L | 0.0939 | 1.68 | 47.5 |  | 1 | 1 |  | 819 | 819 | 819 | [6 empty] | 1 |  | 10.0604 | 10.0604 | 10.0604
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s79 | 127 | 1 | 1 | TN | ug N/L | 1.01 | 13.8 | 607 | TP | ug P/L | 0.0943 | 1.78 | 51.9 |  | 1 | 1 |  | 607 | 607 | 607 | [6 empty] | 1 |  | 15.3702 | 15.3702 | 15.3702
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s80 | 128 | 1 | 1 | TN | ug N/L | 3 | 40.8 | 1950 | TP | ug P/L | 0.0965 | 1.73 | 49.5 |  | 1 | 1 |  | 1950 | 1950 | 1950 | [6 empty] | 1 |  | 12.7193 | 12.7193 | 12.7193
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup75 | 129 | 1 | 1 | TN | ug N/L | 1.93 | 26.2 | 1220 | TP | ug P/L | 0.122 | 2.35 | 77.4 |  | 1 | 1 |  | 1220 | 1220 | 1220 |  |  | 15.9292 | PASS |  |  | 1 |  | 45.0292 | 45.0292 | 45.0292 |  |  | 93.0749 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk80 | 130 | 1 | 1 | TN | ug N/L | 3.25 | 44.1 | 2120 | TP | ug P/L | 0.107 | 2.01 | 62.5 |  | 1 | 1 |  | 2120 | 2120 | 2120 |  |  |  |  | 89.2 | PASS | 1 |  | 27.4622 | 27.4622 | 27.4622 |  |  |  |  | 61.718 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.472 | 6.36 | 247 | TP | ug P/L | 0.0942 | 1.76 | 51.2 |  | 1 | 1 |  | 247 | 247 | 247 | -1.2 | PASS |  |  |  |  | 1 |  | 14.3108 | 14.3108 | 14.3108 | -42.7567 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.00739 | -0.0299 | -62.6 | TP | ug P/L | 0.0905 | 1.66 | 46.1 |  | 1 | 1 |  | -62.6 | -62.6 | -62.6 | [6 empty] | 1 |  | 8.9946 | 8.9946 | 8.9946
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0154 | -0.231 | -72.4 | TP | ug P/L | 0.00907 | 0.142 | -38.6 |  | 1 | 1 |  | -72.4 | -72.4 | -72.4 | [6 empty] | 1 |  | -75.6109 | -75.6109 | -75.6109
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s81 | 131 | 1 | 1 | TN | ug N/L | 1.14 | 15.6 | 695 | TP | ug P/L | 0.0965 | 1.76 | 51.2 |  | 1 | 1 |  | 695 | 695 | 695 | [6 empty] | 1 |  | 14.3108 | 14.3108 | 14.3108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s82 | 132 | 1 | 1 | TN | ug N/L | 0.583 | 7.92 | 322 | TP | ug P/L | 0.0961 | 1.77 | 51.7 |  | 1 | 1 |  | 322 | 322 | 322 | [6 empty] | 1 |  | 14.8407 | 14.8407 | 14.8407
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s83 | 133 | 1 | 1 | TN | ug N/L | 1.14 | 15.4 | 690 | TP | ug P/L | 0.0935 | 1.71 | 48.9 |  | 1 | 1 |  | 690 | 690 | 690 | [6 empty] | 1 |  | 11.6567 | 11.6567 | 11.6567
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s84 | 134 | 1 | 1 | TN | ug N/L | 1.42 | 19.3 | 880 | TP | ug P/L | 0.0967 | 1.77 | 51.5 |  | 1 | 1 |  | 880 | 880 | 880 | [6 empty] | 1 |  | 14.8407 | 14.8407 | 14.8407
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s85 | 135 | 1 | 1 | TN | ug N/L | 0.496 | 6.71 | 263 | TP | ug P/L | 0.0943 | 1.81 | 53.6 |  | 1 | 1 |  | 263 | 263 | 263 | [6 empty] | 1 |  | 16.957 | 16.957 | 16.957
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s86 | 136 | 1 | 1 | TN | ug N/L | 0.56 | 7.61 | 307 | TP | ug P/L | 0.0944 | 1.8 | 53.2 |  | 1 | 1 |  | 307 | 307 | 307 | [6 empty] | 1 |  | 16.4284 | 16.4284 | 16.4284
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s87 | 137 | 1 | 1 | TN | ug N/L | 0.629 | 8.57 | 354 | TP | ug P/L | 0.0922 | 1.65 | 45.8 |  | 1 | 1 |  | 354 | 354 | 354 | [6 empty] | 1 |  | 8.4612 | 8.4612 | 8.4612
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s88 | 138 | 1 | 1 | TN | ug N/L | 0.546 | 7.39 | 297 | TP | ug P/L | 0.0952 | 1.74 | 50.1 |  | 1 | 1 |  | 297 | 297 | 297 | [6 empty] | 1 |  | 13.2501 | 13.2501 | 13.2501
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s89 | 139 | 1 | 1 | TN | ug N/L | 1.16 | 15.8 | 708 | TP | ug P/L | 0.0964 | 1.78 | 52.1 |  | 1 | 1 |  | 708 | 708 | 708 | [6 empty] | 1 |  | 15.3702 | 15.3702 | 15.3702
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s90 | 140 | 1 | 1 | TN | ug N/L | 0.613 | 8.31 | 341 | TP | ug P/L | 0.0936 | 1.72 | 49 |  | 1 | 1 |  | 341 | 341 | 341 | [6 empty] | 1 |  | 12.1881 | 12.1881 | 12.1881
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup85 | 141 | 1 | 1 | TN | ug N/L | 0.503 | 6.8 | 268 | TP | ug P/L | 0.1 | 1.88 | 56.7 |  | 1 | 1 |  | 268 | 268 | 268 |  |  | 1.8832 | PASS |  |  | 1 |  | 20.6483 | 20.6483 | 20.6483 |  |  | 19.6318 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk90 | 142 | 1 | 1 | TN | ug N/L | 0.961 | 13 | 571 | TP | ug P/L | 0.0982 | 1.74 | 49.9 |  | 1 | 1 |  | 571 | 571 | 571 |  |  |  |  | 97.71 | PASS | 1 |  | 13.2501 | 13.2501 | 13.2501 |  |  |  |  | 5.5729 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.466 | 6.36 | 247 | TP | ug P/L | 0.0916 | 1.67 | 47 |  | 1 | 1 |  | 247 | 247 | 247 | -1.2 | PASS |  |  |  |  | 1 |  | 9.5276 | 9.5276 | 9.5276 | -61.8894 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.0081 | 0.0293 | -59.8 | TP | ug P/L | 0.0867 | 1.44 | 35.6 |  | 1 | 1 |  | -59.8 | -59.8 | -59.8 | [6 empty] | 1 |  | -2.813 | -2.813 | -2.813
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00635 | 0.0636 | -58.1 | TP | ug P/L | 0.0117 | 0.216 | -33.9 |  | 1 | 1 |  | -58.1 | -58.1 | -58.1 | [6 empty] | 1 |  | -71.3166 | -71.3166 | -71.3166
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s91 | 143 | 1 | 1 | TN | ug N/L | 1.83 | 24.9 | 1150 | TP | ug P/L | 0.0998 | 1.8 | 52.9 |  | 1 | 1 |  | 1150 | 1150 | 1150 | [6 empty] | 1 |  | 16.4284 | 16.4284 | 16.4284
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s92 | 144 | 1 | 1 | TN | ug N/L | 3.77 | 51.3 | 2480 | TP | ug P/L | 0.0967 | 1.72 | 49.2 |  | 1 | 1 |  | 2480 | 2480 | 2480 | [6 empty] | 1 |  | 12.1881 | 12.1881 | 12.1881
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s93 | 145 | 1 | 1 | TN | ug N/L | 0.522 | 7.09 | 282 | TP | ug P/L | 0.0972 | 1.82 | 53.8 |  | 1 | 1 |  | 282 | 282 | 282 | [6 empty] | 1 |  | 17.4853 | 17.4853 | 17.4853
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s94 | 146 | 1 | 1 | TN | ug N/L | 3.47 | 47 | 2270 | TP | ug P/L | 0.0983 | 1.76 | 51.1 |  | 1 | 1 |  | 2270 | 2270 | 2270 | [6 empty] | 1 |  | 14.3108 | 14.3108 | 14.3108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s95 | 147 | 1 | 1 | TN | ug N/L | 1.43 | 19.3 | 881 | TP | ug P/L | 0.0975 | 1.74 | 50.1 |  | 1 | 1 |  | 881 | 881 | 881 | [6 empty] | 1 |  | 13.2501 | 13.2501 | 13.2501
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s96 | 148 | 1 | 1 | TN | ug N/L | 0.643 | 8.69 | 360 | TP | ug P/L | 0.0935 | 1.77 | 51.4 |  | 1 | 1 |  | 360 | 360 | 360 | [6 empty] | 1 |  | 14.8407 | 14.8407 | 14.8407
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s97 | 149 | 1 | 1 | TN | ug N/L | 0.752 | 10.2 | 432 | TP | ug P/L | 0.0925 | 1.75 | 50.6 |  | 1 | 1 |  | 432 | 432 | 432 | [6 empty] | 1 |  | 13.7806 | 13.7806 | 13.7806
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s98 | 150 | 1 | 1 | TN | ug N/L | 0.673 | 9.07 | 379 | TP | ug P/L | 0.0945 | 1.78 | 52.1 |  | 1 | 1 |  | 379 | 379 | 379 | [6 empty] | 1 |  | 15.3702 | 15.3702 | 15.3702
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s99 | 151 | 1 | 1 | TN | ug N/L | 0.665 | 9.02 | 376 | TP | ug P/L | 0.101 | 2 | 62.3 |  | 1 | 1 |  | 376 | 376 | 376 | [6 empty] | 1 |  | 26.94 | 26.94 | 26.94
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s100 | 152 | 1 | 1 | TN | ug N/L | 1.32 | 17.8 | 808 | TP | ug P/L | 0.0962 | 1.79 | 52.5 |  | 1 | 1 |  | 808 | 808 | 808 | [6 empty] | 1 |  | 15.8995 | 15.8995 | 15.8995
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dup95 | 153 | 1 | 1 | TN | ug N/L | 1.47 | 19.8 | 905 | TP | ug P/L | 0.0943 | 1.7 | 48.2 |  | 1 | 1 |  | 905 | 905 | 905 |  |  | 2.6876 | PASS |  |  | 1 |  | 11.1249 | 11.1249 | 11.1249 |  |  | 17.4377 | PASS
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk100 | 154 | 1 | 1 | TN | ug N/L | 1.66 | 22.4 | 1030 | TP | ug P/L | 0.108 | 1.93 | 59 |  | 1 | 1 |  | 1030 | 1030 | 1030 |  |  |  |  | 99.1 | PASS | 1 |  | 23.2754 | 23.2754 | 23.2754 |  |  |  |  | 31.8313 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 0.473 | 6.38 | 248 | TP | ug P/L | 0.0919 | 1.75 | 50.8 |  | 1 | 1 |  | 248 | 248 | 248 | -0.8 | PASS |  |  |  |  | 1 |  | 13.7806 | 13.7806 | 13.7806 | -44.8775 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | 0.00967 | 0.041 | -59.2 | TP | ug P/L | 0.0894 | 1.56 | 41.7 |  | 1 | 1 |  | -59.2 | -59.2 | -59.2 | [6 empty] | 1 |  | 3.6466 | 3.6466 | 3.6466
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.00841 | 0.057 | -58.5 | TP | ug P/L | 0.00999 | 0.135 | -39.1 |  | 1 | 1 |  | -58.5 | -58.5 | -58.5 | [6 empty] | 1 |  | -76.018 | -76.018 | -76.018
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s101 | 155 | 1 | 1 | TN | ug N/L | 0.534 | 7.24 | 289 | TP | ug P/L | 0.0917 | 1.79 | 52.7 |  | 1 | 1 |  | 289 | 289 | 289 | [6 empty] | 1 |  | 15.8995 | 15.8995 | 15.8995
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s102 | 156 | 1 | 1 | TN | ug N/L | 0.787 | 10.7 | 460 | TP | ug P/L | 0.0936 | 1.76 | 51.2 |  | 1 | 1 |  | 460 | 460 | 460 | [6 empty] | 1 |  | 14.3108 | 14.3108 | 14.3108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s103 | 157 | 1 | 1 | TN | ug N/L | 1.22 | 16.6 | 747 | TP | ug P/L | 0.0929 | 1.74 | 49.9 |  | 1 | 1 |  | 747 | 747 | 747 | [6 empty] | 1 |  | 13.2501 | 13.2501 | 13.2501
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s104 | 158 | 1 | 1 | TN | ug N/L | 0.545 | 9.26 | 387 | TP | ug P/L | 0.955 | 26.9 | 48.2 |  | 1 | 1 |  | 387 | 387 | 387 | [6 empty] | 1 |  | 341.9001 | 341.9001 | 341.9001
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s105 | 159 | 1 | 1 | TN | ug N/L | 0.016 | 0.0587 | -58.4 | TP | ug P/L | 0.144 | 1.5 | 38.6 |  | 1 | 3 | air spikes | -58.4 | -58.4 | -58.4 | [6 empty] | 3 | air spikes | 0.4225 | 0.4225 | 0.4225
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s106 | 160 | 1 | 1 | TN | ug N/L | 3.28 | 36.6 | 1740 | TP | ug P/L | 0.224 | 2.42 | 80 |  | 1 | 3 | air spikes | 1740 | 1740 | 1740 | [6 empty] | 3 | air spikes | 48.6003 | 48.6003 | 48.6003
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s107 | 161 | 1 | 1 | TN | ug N/L | -2.01 | -50.9 | -2450 | TP | ug P/L | 0.265 | 2.18 | 70 |  | 1 | 3 | air spikes | -2450 | -2450 | -2450 | [6 empty] | 3 | air spikes | 36.2917 | 36.2917 | 36.2917
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dups105 | 162 | 1 | 1 | TN | ug N/L | -2.1 | -59.8 | -2850 | TP | ug P/L | 0.156 | -1.05 | -122 |  | 1 | 3 | air spikes | -2850 | -2850 | -2850 |  |  | 276.9874 | FAIL |  |  | 3 | air spikes | -147.183 | -147.183 | -147.183 |  |  | 243.0815 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spks107 | 163 | 1 | 1 | TN | ug N/L | 3.44 | 7.38 | 296 | TP | ug P/L | 0.292 | 2.83 | 95.9 |  | 1 | 3 | air spikes | 296 | 296 | 296 |  |  |  |  | 1101.36 | FAIL | 3 | air spikes | 69.2038 | 69.2038 | 69.2038 |  |  |  |  | 138.569 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s108 | 164 | 1 | 1 | TN | ug N/L | -2.83 | -83.1 | -3890 | TP | ug P/L | 0.246 | 0.259 | -31.2 |  | 1 | 3 | air spikes | -3890 | -3890 | -3890 | -1656 | FAIL |  |  |  |  | 3 | air spikes | -68.8293 | -68.8293 | -68.8293 | -375.317 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | ch25/250 | 7 | 1 | 1 | TN | ug N/L | 1.78 | -6.12 | -356 | TP | ug P/L | 0.123 | 2.07 | 65.1 |  | 1 | 3 | air spikes | -356 | -356 | -356 | [6 empty] | 3 | air spikes | 30.589 | 30.589 | 30.589
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate carrier | S4 | 1 | 1 | TN | ug N/L | -0.788 | -30.1 | -1490 | TP | ug P/L | -0.114 | -1.53 | -159 |  | 1 | 3 | air spikes | -1490 | -1490 | -1490 | [6 empty] | 3 | air spikes | -177.28 | -177.28 | -177.28
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.77 | -12.6 | -665 | TP | ug P/L | 0.0857 | 0.146 | -38.3 |  | 1 | 3 | air spikes | -665 | -665 | -665 | [6 empty] | 3 | air spikes | -75.3783 | -75.3783 | -75.3783
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal1500 | 1 | 1 | 1 | TN | ug N/L | 2.91 | 4.73 | 167 | TP | ug P/L | 0.35 | 5.8 | 160 |  | 1 | 3 | air spikes | 167 | 167 | 167 | [6 empty] | 3 | air spikes | 202.4924 | 202.4924 | 202.4924
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal1000 | 3 | 1 | 1 | TN | ug N/L | -2.58 | -57 | -2720 | TP | ug P/L | 0.235 | 4.44 | 142 |  | 1 | 3 | air spikes | -2720 | -2720 | -2720 | [6 empty] | 3 | air spikes | 144.9394 | 144.9394 | 144.9394
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | call500 | 5 | 1 | 1 | TN | ug N/L | 1.98 | 3.67 | 116 | TP | ug P/L | 0.214 | 1.09 | 16.9 |  | 1 | 3 | air spikes | 116 | 116 | 116 | [6 empty] | 3 | air spikes | -21.9151 | -21.9151 | -21.9151
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | call250 | 7 | 1 | 1 | TN | ug N/L | 1.61 | -10.7 | -572 | TP | ug P/L | 0.116 | 1.96 | 60.3 |  | 1 | 3 | air spikes | -572 | -572 | -572 | [6 empty] | 3 | air spikes | 24.8479 | 24.8479 | 24.8479
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal100 | 9 | 1 | 1 | TN | ug N/L | -1.98 | -22.1 | -1110 | TP | ug P/L | 0.15 | 2.16 | 69.2 |  | 1 | 3 | air spikes | -1110 | -1110 | -1110 | [6 empty] | 3 | air spikes | 35.2577 | 35.2577 | 35.2577
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | 2 | 3.26 | 96.1 | TP | ug P/L | -0.0399 | -0.234 | -63.2 |  | 1 | 3 | air spikes | 96.1 | 96.1 | 96.1 | [6 empty] | 3 | air spikes | -97.6995 | -97.6995 | -97.6995
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | 4.66 | 29.1 | 1370 | TP | ug P/L | 0.118 | 0.104 | -41 |  | 1 | 3 | air spikes | 1370 | 1370 | 1370 | [6 empty] | 3 | air spikes | -77.8228 | -77.8228 | -77.8228
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | 2.46 | 14.2 | 630 | TP | ug P/L | -0.00146 | -0.00226 | -47.9 |  | 1 | 3 | air spikes | 630 | 630 | 630 | [6 empty] | 3 | air spikes | -84.0324 | -84.0324 | -84.0324
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | 2.95 | 12 | 520 | TP | ug P/L | -0.047 | -0.25 | -64.3 |  | 1 | 3 | air spikes | 520 | 520 | 520 | [6 empty] | 3 | air spikes | -98.6494 | -98.6494 | -98.6494
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | 9.64 | 69.9 | 3440 | TP | ug P/L | -0.0307 | -0.173 | -59.1 |  | 1 | 3 | air spikes | 3440 | 3440 | 3440 | [6 empty] | 3 | air spikes | -94.0854 | -94.0854 | -94.0854
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.158 | 2.03 | 36.9 | TP | ug P/L | 0.0949 | 1.9 | 57.7 |  | 1 | 1 |  | 36.9 | 36.9 | 36.9 | [6 empty] | 1 |  | 21.7001 | 21.7001 | 21.7001
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.142 | 1.94 | 32.4 | TP | ug P/L | 0.0942 | 1.77 | 51.7 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 | [6 empty] | 1 |  | 14.8407 | 14.8407 | 14.8407 | -40.6372 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0919 | 1.28 | 0.497 | TP | ug P/L | 0.0876 | 1.54 | 40.2 |  | 1 | 1 |  | 0.497 | 0.497 | 0.497 | [6 empty] | 1 |  | 2.5732 | 2.5732 | 2.5732
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrite | S1 | 1 | 1 | TN | ug N/L | 2.27 | 30.8 | 1450 | TP | ug P/L | 0.0151 | 0.215 | -34 | 97.7273 | 1 | 1 |  | 1450 | 1450 | 1450 | [6 empty] | 1 |  | -71.3745 | -71.3745 | -71.3745
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | nitrate | S2 | 1 | 1 | TN | ug N/L | 2.22 | 30.1 | 1410 | TP | ug P/L | 0.0165 | 0.248 | -31.9 |  | 1 | 1 |  | 1410 | 1410 | 1410 | [6 empty] | 1 |  | -69.465 | -69.465 | -69.465
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | water | S3 | 1 | 1 | TN | ug N/L | -0.0179 | -0.37 | -79 | TP | ug P/L | 0.0117 | 0.183 | -36 |  | 1 | 1 |  | -79 | -79 | -79 | [6 empty] | 1 |  | -73.2294 | -73.2294 | -73.2294
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s105 rerun | 159 | 1 | 1 | TN | ug N/L | 1.11 | 15.1 | 674 | TP | ug P/L | 0.0963 | 1.87 | 56 |  | 1 | 1 |  | 674 | 674 | 674 | [6 empty] | 1 |  | 20.1219 | 20.1219 | 20.1219
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s106 rerun | 160 | 1 | 1 | TN | ug N/L | 0.842 | 11.6 | 501 | TP | ug P/L | 0.0921 | 1.75 | 50.6 |  | 1 | 1 |  | 501 | 501 | 501 | [6 empty] | 1 |  | 13.7806 | 13.7806 | 13.7806
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s107 rerun | 161 | 1 | 1 | TN | ug N/L | 1.1 | 15.2 | 678 | TP | ug P/L | 0.0943 | 1.83 | 54.2 |  | 1 | 1 |  | 678 | 678 | 678 | [6 empty] | 1 |  | 18.0132 | 18.0132 | 18.0132
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | dups105 rerun | 162 | 1 | 1 | TN | ug N/L | 1.1 | 15.3 | 682 | TP | ug P/L | 0.0916 | 1.65 | 46 |  | 1 | 1 |  | 682 | 682 | 682 | [6 empty] | 1 |  | 8.4612 | 8.4612 | 8.4612
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | spk107 rerun | 163 | 1 | 1 | TN | ug N/L | 1.42 | 19.7 | 897 | TP | ug P/L | 0.0989 | 1.72 | 49.4 |  | 1 | 1 |  | 897 | 897 | 897 | [6 empty] | 1 |  | 12.1881 | 12.1881 | 12.1881
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | s108 rerun | 164 | 1 | 1 | TN | ug N/L | 0.68 | 9.51 | 400 | TP | ug P/L | 0.0941 | 1.8 | 52.9 |  | 1 | 1 |  | 400 | 400 | 400 | [6 empty] | 1 |  | 16.4284 | 16.4284 | 16.4284
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal 1500 | 1 | 1 | 1 | TN | ug N/L | 2.32 | 32 | 1510 | TP | ug P/L | 0.24 | 4.63 | 145 |  | 1 | 1 |  | 1510 | 1510 | 1510 | [6 empty] | 1 |  | 153.3333 | 153.3333 | 153.3333
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal 50/5 | 11 | 1 | 1 | TN | ug N/L | 0.159 | 2.28 | 49 | TP | ug P/L | 0.0933 | 1.74 | 50.3 |  | 1 | 1 |  | 49 | 49 | 49 | [6 empty] | 1 |  | 13.2501 | 13.2501 | 13.2501
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | chk 25/250 | 7 | 1 | 1 | TN | ug N/L | 0.451 | 6.34 | 246 | TP | ug P/L | 0.0887 | 1.6 | 43.4 |  | 1 | 1 |  | 246 | 246 | 246 | [6 empty] | 1 |  | 5.7896 | 5.7896 | 5.7896 | -76.8416 | FAIL
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal 25/2.5 | 13 | 1 | 1 | TN | ug N/L | 0.123 | 1.78 | 24.7 | TP | ug P/L | 0.0968 | 1.84 | 54.8 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | [6 empty] | 1 |  | 18.5409 | 18.5409 | 18.5409
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | bisulfate | S4 | 1 | 1 | TN | ug N/L | -0.00258 | -0.0295 | -62.6 | TP | ug P/L | 0.0878 | 1.6 | 43.4 |  | 1 | 1 |  | -62.6 | -62.6 | -62.6 | [6 empty] | 1 |  | 5.7896 | 5.7896 | 5.7896

## compare cals
(-2209.9*X^2)+(1446.7*X)+(-71.879)
(-6275.4*X^2)+(2148.1*X)+(-99.418)
|  |  | FLAT | [15 empty] | DIP
original order | Known P conc |  | Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration
21 | 0 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0914 | 1.27 | 0 | TP | ug P/L | 0.0562 | 0.806 | 0 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0914 | 1.27 | 0 | TP | ug P/L | 0.0916 | 1.54 | 0
22 | 0 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0915 | 1.27 | 0 | TP | ug P/L | 0.0567 | 0.799 | 0 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 0 | 15 | 1 | 1 | TN | ug N/L | 0.0915 | 1.27 | 0 | TP | ug P/L | 0.0945 | 1.64 | 0
136 | 0 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0958 | 1.29 | 1.07 | TP | ug P/L | 0.056 | 0.797 | -44.1 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0958 | 1.29 | 1.07 | TP | ug P/L | 0.0887 | 1.56 | 41.2
239 | 0 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0919 | 1.28 | 0.497 | TP | ug P/L | 0.052 | 0.686 | -44.4 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal0 | 15 | 1 | 1 | TN | ug N/L | 0.0919 | 1.28 | 0.497 | TP | ug P/L | 0.0876 | 1.54 | 40.2
19 | 2.5 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.131 | 1.8 | 25 | TP | ug P/L | 0.059 | 0.855 | 2.5 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.131 | 1.8 | 25 | TP | ug P/L | 0.0934 | 1.57 | 2.5
20 | 2.5 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.129 | 1.76 | 25 | TP | ug P/L | 0.0563 | 0.799 | 2.5 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 2.5 | 13 | 1 | 1 | TN | ug N/L | 0.129 | 1.76 | 25 | TP | ug P/L | 0.0915 | 1.59 | 2.5
135 | 2.5 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.135 | 1.81 | 26.3 | TP | ug P/L | 0.0576 | 0.817 | -44 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.135 | 1.81 | 26.3 | TP | ug P/L | 0.0929 | 1.76 | 51.1
238 | 2.5 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.142 | 1.94 | 32.4 | TP | ug P/L | 0.0584 | 0.817 | -43.9 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal25 | 13 | 1 | 1 | TN | ug N/L | 0.142 | 1.94 | 32.4 | TP | ug P/L | 0.0942 | 1.77 | 51.7
252 | 2.5 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal 25/2.5 | 13 | 1 | 1 | TN | ug N/L | 0.123 | 1.78 | 24.7 | TP | ug P/L | 0.0562 | 0.772 | -44.1 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal 25/2.5 | 13 | 1 | 1 | TN | ug N/L | 0.123 | 1.78 | 24.7 | TP | ug P/L | 0.0968 | 1.84 | 54.8
17 | 5 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.176 | 2.33 | 50 | TP | ug P/L | 0.0612 | 0.443 | 5 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.176 | 2.33 | 50 | TP | ug P/L | 0.0894 | 1.49 | 5
18 | 5 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.166 | 2.28 | 50 | TP | ug P/L | 0.0591 | 0.353 | 5 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 5 | 11 | 1 | 1 | TN | ug N/L | 0.166 | 2.28 | 50 | TP | ug P/L | 0.092 | 1.54 | 5
134 | 5 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.172 | 2.35 | 52.2 | TP | ug P/L | 0.0563 | 0.351 | -44.1 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.172 | 2.35 | 52.2 | TP | ug P/L | 0.0944 | 1.88 | 56.8
237 | 5 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.158 | 2.03 | 36.9 | TP | ug P/L | 0.0595 | 0.838 | -43.9 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal50 | 11 | 1 | 1 | TN | ug N/L | 0.158 | 2.03 | 36.9 | TP | ug P/L | 0.0949 | 1.9 | 57.7
250 | 5 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal 50/5 | 11 | 1 | 1 | TN | ug N/L | 0.159 | 2.28 | 49 | TP | ug P/L | 0.0592 | 0.389 | -43.9 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal 50/5 | 11 | 1 | 1 | TN | ug N/L | 0.159 | 2.28 | 49 | TP | ug P/L | 0.0933 | 1.74 | 50.3
15 | 10 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.43 | 100 | TP | ug P/L | 0.0629 | 0.498 | 10 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.43 | 100 | TP | ug P/L | 0.0942 | 1.65 | 10
16 | 10 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.44 | 100 | TP | ug P/L | 0.0682 | 0.488 | 10 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 10 | 9 | 1 | 1 | TN | ug N/L | 0.257 | 3.44 | 100 | TP | ug P/L | 0.095 | 1.5 | 10
133 | 10 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal100 | 9 | 1 | 1 | TN | ug N/L | 0.254 | 3.45 | 105 | TP | ug P/L | 0.0595 | 0.468 | -43.9 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal100 | 9 | 1 | 1 | TN | ug N/L | 0.254 | 3.45 | 105 | TP | ug P/L | 0.0907 | 1.76 | 51.3
13 | 25 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.478 | 6.3 | 250 | TP | ug P/L | 0.0729 | 0.978 | 25 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.478 | 6.3 | 250 | TP | ug P/L | 0.0926 | 1.65 | 25
14 | 25 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.474 | 6.29 | 250 | TP | ug P/L | 0.073 | 0.925 | 25 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 25 | 7 | 1 | 1 | TN | ug N/L | 0.474 | 6.29 | 250 | TP | ug P/L | 0.089 | 1.48 | 25
132 | 25 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call250 | 7 | 1 | 1 | TN | ug N/L | 0.471 | 6.35 | 246 | TP | ug P/L | 0.0697 | 0.881 | -43.2 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | call250 | 7 | 1 | 1 | TN | ug N/L | 0.471 | 6.35 | 246 | TP | ug P/L | 0.0908 | 1.68 | 47.3
251 | 25 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | chk 25/250 | 7 | 1 | 1 | TN | ug N/L | 0.451 | 6.34 | 246 | TP | ug P/L | 0.0695 | 0.878 | -43.2 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | chk 25/250 | 7 | 1 | 1 | TN | ug N/L | 0.451 | 6.34 | 246 | TP | ug P/L | 0.0887 | 1.6 | 43.4
11 | 50 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.867 | 11.4 | 500 | TP | ug P/L | 0.0979 | 1.72 | 50 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.867 | 11.4 | 500 | TP | ug P/L | 0.098 | 1.73 | 50
12 | 50 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.86 | 11.4 | 500 | TP | ug P/L | 0.0977 | 1.73 | 50 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 50 | 5 | 1 | 1 | TN | ug N/L | 0.86 | 11.4 | 500 | TP | ug P/L | 0.0979 | 1.75 | 50
131 | 50 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call500 | 5 | 1 | 1 | TN | ug N/L | 0.848 | 11.5 | 495 | TP | ug P/L | 0.0965 | 1.81 | -41.5 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | call500 | 5 | 1 | 1 | TN | ug N/L | 0.848 | 11.5 | 495 | TP | ug P/L | 0.0962 | 1.8 | 52.8
9 | 100 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.66 | 21.8 | 1000 | TP | ug P/L | 0.164 | 3.25 | 100 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.66 | 21.8 | 1000 | TP | ug P/L | 0.165 | 3.31 | 100
10 | 100 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.67 | 22.6 | 1000 | TP | ug P/L | 0.16 | 3.15 | 100 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 100 | 3 | 1 | 1 | TN | ug N/L | 1.67 | 22.6 | 1000 | TP | ug P/L | 0.16 | 3.2 | 100
130 | 100 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal1000 | 3 | 1 | 1 | TN | ug N/L | 1.62 | 21.9 | 1010 | TP | ug P/L | 0.157 | 3.07 | -37.6 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal1000 | 3 | 1 | 1 | TN | ug N/L | 1.62 | 21.9 | 1010 | TP | ug P/L | 0.159 | 3.17 | 108
7 | 150 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.237 | 4.65 | 150 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.238 | 4.76 | 150
8 | 150 |  | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit flat.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.237 | 4.59 | 150 | 2023-02-08 | OM_2-8-2023_11-07-17AM BRN edit dip.omn | cal std 150 | 1 | 1 | 1 | TN | ug N/L | 2.41 | 31.6 | 1500 | TP | ug P/L | 0.236 | 4.54 | 150
129 | 150 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal1500 | 1 | 1 | 1 | TN | ug N/L | 2.37 | 31.9 | 1500 | TP | ug P/L | 0.258 | 4.53 | -31.3 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal1500 | 1 | 1 | 1 | TN | ug N/L | 2.37 | 31.9 | 1500 | TP | ug P/L | 0.259 | 4.6 | 145
249 | 150 |  | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | cal 1500 | 1 | 1 | 1 | TN | ug N/L | 2.32 | 32 | 1510 | TP | ug P/L | 0.24 | 4.61 | -32.4 | 2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit dip.omn | cal 1500 | 1 | 1 | 1 | TN | ug N/L | 2.32 | 32 | 1510 | TP | ug P/L | 0.24 | 4.63 | 145
[5 empty] | All data weighted | [6 empty] | Restricted to 50 and below weighted
| index | conc | signal |  | Call: | [6 empty] | Call:
| 15 | 0 | 0.0562 |  | lm(formula = x ~ poly(y, 2, raw = TRUE), data = data1, weights = 1/x) | [6 empty] | lm(formula = x ~ poly(y, 2, raw = TRUE), data = data1, weights = 1/x)
| 15 | 0 | 0.0567
| 15 | 0 | 0.056 |  | Weighted Residuals: | [6 empty] | Weighted Residuals:
| 15 | 0 | 0.052 |  | Min 1Q Median 3Q Max | [6 empty] | Min 1Q Median 3Q Max
| 13 | 2.5 | 0.059 |  | -2.8551 -0.3862 0.3018 1.1500 4.9304 | [6 empty] | -2.2613 -0.4022 0.6802 1.6198 2.5297
| 13 | 2.5 | 0.0563
| 13 | 2.5 | 0.0576 |  | Coefficients: | [6 empty] | Coefficients:
| 13 | 2.5 | 0.0584 |  | Estimate Std. Error t value Pr(>¦t¦) | [6 empty] | Estimate Std. Error t value Pr(>¦t¦)
| 13 | 2.5 | 0.0562 |  | (Intercept) -60.911 7.628 -7.985 1.07e-08 *** | [6 empty] | (Intercept) 7.718 24.238 0.318 0.7533
| 11 | 5 | 0.0612 |  | poly(y, 2, raw = TRUE)1 1184.601 170.214 6.959 1.44e-07 *** | [6 empty] | poly(y, 2, raw = TRUE)1 -920.018 729.722 -1.261 0.2212
| 11 | 5 | 0.0591 |  | poly(y, 2, raw = TRUE)2 -1259.750 642.962 -1.959 0.0601 . | [6 empty] | poly(y, 2, raw = TRUE)2 14394.775 5378.010 2.677 0.0141 *
| 11 | 5 | 0.0563 |  | --- | [6 empty] | ---
| 11 | 5 | 0.0595 |  | Signif. codes: 0 ‘***’ 0.001 ‘**’ 0.01 ‘*’ 0.05 ‘.’ 0.1 ‘ ’ 1 | [6 empty] | Signif. codes: 0 ‘***’ 0.001 ‘**’ 0.01 ‘*’ 0.05 ‘.’ 0.1 ‘ ’ 1
| 11 | 5 | 0.0592
| 9 | 10 | 0.0629 |  | Residual standard error: 1.528 on 28 degrees of freedom | [6 empty] | Residual standard error: 1.474 on 21 degrees of freedom
| 9 | 10 | 0.0682 |  | Multiple R-squared: 0.9435, Adjusted R-squared: 0.9395 | [6 empty] | Multiple R-squared: 0.8389, Adjusted R-squared: 0.8236
| 9 | 10 | 0.0595 |  | F-statistic: 233.8 on 2 and 28 DF, p-value: < 2.2e-16 | [6 empty] | F-statistic: 54.67 on 2 and 21 DF, p-value: 4.728e-09
| 7 | 25 | 0.0729
| 7 | 25 | 0.073
| 7 | 25 | 0.0697 |  |  |  | (-1259.750*X^2)+(1184.601*X)+(-60.911) | [5 empty] | (14394.775*X^2)+(-920.018*X)+(7.718)
| 7 | 25 | 0.0695
| 5 | 50 | 0.0979
| 5 | 50 | 0.0977
| 5 | 50 | 0.0965
| 3 | 100 | 0.164
| 3 | 100 | 0.16
| 3 | 100 | 0.157
| 1 | 150 | 0.237
| 1 | 150 | 0.237
| 1 | 150 | 0.258
| 1 | 150 | 0.24

## method checks
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE NO3-N: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- dilution corrected | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- dilution corrected | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ
BLANKS-digested type I reagent grade water
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 20 | 1 | 1 | TN | ug N/L | 0.084 | 1.2 | -3.48 | TP | ug P/L | 0.0482 | 0.56 | -44.6 |  | 1 | 1 |  | -3.48 | -3.48 | -3.48 | [6 empty] | 1 |  | -10.4588 | -10.4588 | -10.4588 | [6 empty] | 39
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 21 | 1 | 1 | TN | ug N/L | 0.0936 | 1.22 | -2.22 | TP | ug P/L | 0.0476 | 0.617 | -44.6 |  | 1 | 1 |  | -2.22 | -2.22 | -2.22 | [6 empty] | 1 |  | -11.387 | -11.387 | -11.387 | [6 empty] | 40
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 22 | 1 | 1 | TN | ug N/L | 0.0896 | 1.28 | 0.652 | TP | ug P/L | 0.0538 | 0.764 | -44.2 |  | 1 | 1 |  | 0.652 | 0.652 | 0.652 | [6 empty] | 1 |  | -2.014 | -2.014 | -2.014 | [6 empty] | 41
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 23 | 1 | 1 | TN | ug N/L | 0.095 | 1.3 | 1.41 | TP | ug P/L | 0.0539 | 0.721 | -44.2 |  | 1 | 1 |  | 1.41 | 1.41 | 1.41 | [6 empty] | 1 |  | -1.8668 | -1.8668 | -1.8668 | [6 empty] | 42
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 24 | 1 | 1 | TN | ug N/L | 0.0925 | 1.31 | 1.89 | TP | ug P/L | 0.0537 | 0.741 | -44.3 |  | 1 | 1 |  | 1.89 | 1.89 | 1.89 | [6 empty] | 1 |  | -2.1613 | -2.1613 | -2.1613 | [6 empty] | 43
[7 empty] | Known Conc |  |  | 0 |  | Known Conc |  |  | 0
[7 empty] | Mean |  | 1.262 | -0.3496 |  | Mean |  | 0.6806 | -44.38
[7 empty] | Std Dev |  | 0.049193 | 2.3671 |  | Std Dev |  | 0.087786 | 0.20494
[7 empty] | CV or %RSD |  | 3.8981 | -677.0934 |  | CV or %RSD |  | 12.8983 | -0.46178
[7 empty] | T value |  | 3.7469 | 3.7469 |  | T value |  | 3.7469 | 3.7469
[7 empty] | MDL |  | 0.18433 | 8.8695 |  | MDL |  | 0.32893 | 0.7679
[7 empty] | LOQ |  | 0.49193 | 23.6712 |  | LOQ |  | 0.87786 | 2.0494
[7 empty] | Error as % |  |  |  |  | Error as %
[7 empty] | Known/MDL |  |  |  |  | Known/MDL
[7 empty] | Mean PR |  |  |  |  | Mean PR
[7 empty] | S/N |  | 25.6538 | -0.14769 |  | S/N |  | 7.753 | -216.5522 | [5 empty] | MEANS | -0.3496 | [9 empty] | MEANS | -5.5776
|  |  | Different from 0? |  | one sample T-test |  |  |  | p value | 0.4149 |  |  |  | p value | 0.3504
[7 empty] | count |  | 5 |  |  |  |  | 5
[7 empty] | std error of mean |  | 1.0586 |  |  |  |  | 0.091652
[7 empty] | df |  | 4 |  |  |  |  | 4
[7 empty] | t |  | 0.33025 |  |  |  |  | 484.2255
[7 empty] | p value |  | 0.75779 |  |  |  |  | 1.0913e-10
DIGESTION CHECK - is organic material completely converted to inorganic?
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 25 | 1 | 1 | TN | ug N/L | 0.71 | 9.48 | 398 | TP | ug P/L | 0.0817 | 1.27 | -42.4 |  | 1 | 1 |  | 398 | 398 | 398 | [6 empty] | 1 |  | 34.1942 | 34.1942 | 34.1942 | [6 empty] | 44
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 26 | 1 | 1 | TN | ug N/L | 0.707 | 9.5 | 399 | TP | ug P/L | 0.0795 | 1.25 | -42.6 |  | 1 | 1 |  | 399 | 399 | 399 | [6 empty] | 1 |  | 31.6939 | 31.6939 | 31.6939 | [6 empty] | 45
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 27 | 1 | 1 | TN | ug N/L | 0.709 | 9.47 | 398 | TP | ug P/L | 0.0827 | 1.27 | -42.4 |  | 1 | 1 |  | 398 | 398 | 398 | [6 empty] | 1 |  | 35.3106 | 35.3106 | 35.3106 | [6 empty] | 46
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 28 | 1 | 1 | TN | ug N/L | 0.711 | 9.46 | 397 | TP | ug P/L | 0.0812 | 1.28 | -42.5 |  | 1 | 1 |  | 397 | 397 | 397 | [6 empty] | 1 |  | 33.6312 | 33.6312 | 33.6312 | [6 empty] | 47
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 29 | 1 | 1 | TN | ug N/L | 0.709 | 9.42 | 395 | TP | ug P/L | 0.0823 | 1.19 | -42.4 |  | 1 | 1 |  | 395 | 395 | 395 | [6 empty] | 1 |  | 34.8655 | 34.8655 | 34.8655 | [6 empty] | 48
[9 empty] | Known Conc | 400 |  |  |  | Known Conc | 40 | [5 empty] | Known Conc | 400 | [9 empty] | Known Conc | 40
[9 empty] | Mean | 397.4 |  |  |  | Mean | -42.46 | [5 empty] | Mean | 397.4 | [9 empty] | Mean | 33.9391
[9 empty] | Std Dev | 1.5166 |  |  |  | Std Dev | 0.089443 | [5 empty] | Std Dev | 1.5166 | [9 empty] | Std Dev | 1.4089
[9 empty] | CV or %RSD | 0.38162 |  |  |  | CV or %RSD | -0.21065 | [5 empty] | CV or %RSD | 0.38162 | [9 empty] | CV or %RSD | 4.1513
[9 empty] | T value | 3.7469 |  |  |  | T value | 3.7469 | [5 empty] | T value | 3.7469 | [9 empty] | T value | 3.7469
[9 empty] | MDL | 5.6825 |  |  |  | MDL | 0.33514 | [5 empty] | MDL | 5.6825 | [9 empty] | MDL | 5.2791
[9 empty] | LOQ | 15.1658 |  |  |  | LOQ | 0.89443 | [5 empty] | LOQ | 15.1658 | [9 empty] | LOQ | 14.0892
[9 empty] | Error as % | -0.65 |  |  |  | Error as % | -206.15 | [5 empty] | Error as % | -0.65 | [9 empty] | Error as % | -15.1523
[9 empty] | Known/MDL |  |  |  |  | Known/MDL | [6 empty] | Known/MDL | [10 empty] | Known/MDL
[9 empty] | Mean PR | 99.35 |  |  |  | Mean PR | -106.15 | [5 empty] | Mean PR | [10 empty] | Mean PR
[9 empty] | S/N | 262.0378 |  |  |  | S/N | -474.7172 | [5 empty] | S/N | 262.0378 | [9 empty] | S/N | 24.0887
SPIKED BLANKS- used for MDL determination (250 ul spiking soln 1000/10000 into 10 ml sample)
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 30 | 1 | 1 | TN | ug N/L | 0.433 | 5.75 | 217 | TP | ug P/L | 0.0724 | 1.02 | -43 |  | 1 | 1 |  | 217 | 217 | 217 | [6 empty] | 1 |  | 23.2103 | 23.2103 | 23.2103 | [6 empty] | 49
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | skp blank | 31 | 1 | 1 | TN | ug N/L | 0.436 | 5.89 | 224 | TP | ug P/L | 0.0746 | 1.06 | -42.9 |  | 1 | 1 |  | 224 | 224 | 224 | [6 empty] | 1 |  | 25.9067 | 25.9067 | 25.9067 | [6 empty] | 50
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 32 | 1 | 1 | TN | ug N/L | 0.441 | 5.9 | 224 | TP | ug P/L | 0.0761 | 0.935 | -42.8 |  | 1 | 1 |  | 224 | 224 | 224 | [6 empty] | 1 |  | 27.7103 | 27.7103 | 27.7103 | [6 empty] | 51
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 33 | 1 | 1 | TN | ug N/L | 0.437 | 5.85 | 222 | TP | ug P/L | 0.07 | 0.875 | -43.2 |  | 1 | 1 |  | 222 | 222 | 222 | [6 empty] | 1 |  | 20.1995 | 20.1995 | 20.1995 | [6 empty] | 52
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 34 | 1 | 1 | TN | ug N/L | 0.446 | 5.92 | 225 | TP | ug P/L | 0.0694 | 0.862 | -43.2 |  | 1 | 1 |  | 225 | 225 | 225 | [6 empty] | 1 |  | 19.4356 | 19.4356 | 19.4356 | [6 empty] | 53
[9 empty] | Known Conc | 243.9024 |  |  |  | Known Conc | 24.3902 | [5 empty] | Known Conc | 244 | [9 empty] | Known Conc | 24
[9 empty] | Mean | 222.4 |  |  |  | Mean | -43.02 | [5 empty] | Mean | 222.4 | [9 empty] | Mean | 23.2925
[9 empty] | Std Dev | 3.2094 |  |  |  | Std Dev | 0.17889 | [5 empty] | Std Dev | 3.2094 | [9 empty] | Std Dev | 3.5637
[9 empty] | CV or %RSD | 1.4431 |  |  |  | CV or %RSD | -0.41582 | [5 empty] | CV or %RSD | 1.4431 | [9 empty] | CV or %RSD | 15.2998
[9 empty] | T value | 3.7469 |  |  |  | T value | 3.7469 | [5 empty] | T value | 3.7469 | [9 empty] | T value | 3.7469
[9 empty] | MDL | 12.0253 |  |  |  | MDL | 0.67027 | [5 empty] | MDL | 12.0253 | [9 empty] | MDL | 13.353
[9 empty] | LOQ | 32.0936 |  |  |  | LOQ | 1.7889 | [5 empty] | LOQ | 32.0936 | [9 empty] | LOQ | 35.6369
[9 empty] | Error as % | -8.816 |  |  |  | Error as % | -276.382 | [5 empty] | Error as % | -8.8525 | [9 empty] | Error as % | -2.9481
[9 empty] | Known/MDL |  |  |  |  | Known/MDL | [6 empty] | Known/MDL | [10 empty] | Known/MDL
[9 empty] | Mean PR | 91.184 |  |  |  | Mean PR | -176.382 | [5 empty] | Mean PR | [10 empty] | Mean PR
[9 empty] | S/N | 69.2973 |  |  |  | S/N | -240.4891 | [5 empty] | S/N | 69.2973 | [9 empty] | S/N | 6.5361