## rolling spiked blank 250
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation # | PR P | PR N
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | spiked blank 250 | 10 | 1 | 1 | Total Phosphorus | ppb | 0.0585 | 1.33 | 23 | Total Nitrogen | ppb | 1.25 | 19.4 | 569 |  | 1.5 | 1 |  | 23 | 15.5 | 23.25 | [6 empty] | 1 |  | 569 | 228 | 342 | [7 empty] | 1
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | spiked blank + 250 uL | 10 | 1 | 1 | Total Phosphorus | ppb | 0.0743 | 1.32 | 21.9 | Total Nitrogen | ppb | 1.04 | 18 | 487 |  | 1.5 | 1 |  | 21.9 | 15.1 | 22.65 | [6 empty] | 1 |  | 487 | 108 | 162 | [7 empty] | 2
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 028 | 28 | 1 | 1 | TP | ug P/L | 0.0551 | 1.31 | 24 | TN | ug N/L | 0.422 | 7.07 | 8.16 |  | 1 | 1 |  | 24 | 24 | 24 | [6 empty] | 2 | using modeled decline in efficiency | 127.4681 | 127.4681 |  |  | 105.6869 | [5 empty] | 3 | 98.4
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 029 | 29 | 1 | 1 | TP | ug P/L | 0.0563 | 1.31 | 23.8 | TN | ug N/L | 0.512 | 8.69 | 22.1 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | [6 empty] | 2 | using modeled decline in efficiency | 292.5181 | 292.5181 | 251.6436 | [7 empty] | 4 | 97.58 | 103.1749
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 030 | 30 | 1 | 1 | TP | ug P/L | 0.0555 | 1.27 | 22.9 | TN | ug N/L | 0.495 | 8.41 | 19.7 |  | 1 | 1 |  | 22.9 | 22.9 | 22.9 | [6 empty] | 2 | using modeled decline in efficiency | 264.2396 | 264.2396 | 227.2878 | [7 empty] | 5 | 93.89 | 93.1889
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 031 | 31 | 1 | 1 | TP | ug P/L | 0.0557 | 1.28 | 23.1 | TN | ug N/L | 0.483 | 8.16 | 17.5 |  | 1 | 1 |  | 23.1 | 23.1 | 23.1 | [6 empty] | 2 | using modeled decline in efficiency | 238.9032 | 238.9032 | 205.5903 | [7 empty] | 6 | 94.71 | 84.2929
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 032 | 32 | 1 | 1 | TP | ug P/L | 0.0573 | 1.32 | 24.1 | TN | ug N/L | 0.517 | 8.7 | 22.2 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | [6 empty] | 2 | using modeled decline in efficiency | 293.5261 | 293.5261 | 254.9893 | [7 empty] | 7 | 98.81 | 104.5467
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 030 | 30 | 1 | 1 | TP | ug P/L | 0.0484 | 1.16 | 22.4 | TN | ug N/L | 0.309 | 5.28 | 230 |  | 1 | 1 |  | 22.4 | 22.4 | 22.4 | [6 empty] | 1 |  | 230 | 230 | 230 | [7 empty] | 8 | 91.84 | 94.3009
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 031 | 31 | 1 | 1 | TP | ug P/L | 0.0512 | 1.21 | 24.1 | TN | ug N/L | 0.331 | 5.72 | 262 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | [6 empty] | 1 |  | 262 | 262 | 262 | [7 empty] | 9 | 98.81 | 107.4211
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 032 | 32 | 1 | 1 | TP | ug P/L | 0.047 | 1.05 | 19.2 | TN | ug N/L | 0.324 | 5.55 | 249 |  | 1 | 1 |  | 19.2 | 19.2 | 19.2 | [6 empty] | 1 |  | 249 | 249 | 249 | [7 empty] | 10 | 78.72 | 102.091
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 033 | 33 | 1 | 1 | TP | ug P/L | 0.0514 | 1.21 | 24.1 | TN | ug N/L | 0.295 | 5.05 | 212 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | [6 empty] | 1 |  | 212 | 212 | 212 | [7 empty] | 11 | 98.81 | 86.9209
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 034 | 34 | 1 | 1 | TP | ug P/L | 0.0567 | 1.34 | 27.7 | TN | ug N/L | 0.32 | 5.51 | 246 |  | 1 | 1 |  | 27.7 | 27.7 | 27.7 | [6 empty] | 1 |  | 246 | 246 | 246 | [7 empty] | 12 | 113.57 | 100.861
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample30 | 30 | 1 | 1 | TP | ug P/L | 0.0578 | 1.29 | 18.7 | TN | ug N/L | 0.164 | 2.98 | 171 |  | 1 | 2 | reworked calibration | 25.4968 | 25.4968 | 25.4968 | [6 empty] | 1 |  | 181.0821 | 181.0821 | 181.0821 | [7 empty] | 13 | 104.5367 | 74.2444
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample31 | 31 | 1 | 1 | TP | ug P/L | 0.0588 | 1.33 | 19.6 | TN | ug N/L | 0.197 | 3.6 | 226 |  | 1 | 2 | reworked calibration | 26.6927 | 26.6927 | 26.6927 | [6 empty] | 1 |  | 239.415 | 239.415 | 239.415 | [7 empty] | 14 | 109.4403 | 98.1611
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample32 | 32 | 1 | 1 | TP | ug P/L | 0.0609 | 1.34 | 19.9 | TN | ug N/L | 0.225 | 4.04 | 266 |  | 1 | 2 | reworked calibration | 26.9919 | 26.9919 | 26.9919 | [6 empty] | 1 |  | 281.1202 | 281.1202 | 281.1202 | [7 empty] | 15 | 110.6669 | 115.2605
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample33 | 33 | 1 | 1 | TP | ug P/L | 0.0591 | 1.34 | 19.8 | TN | ug N/L | 0.218 | 3.9 | 253 |  | 1 | 2 | reworked calibration | 26.9919 | 26.9919 | 26.9919 | [6 empty] | 1 |  | 267.8227 | 267.8227 | 267.8227 | [7 empty] | 16 | 110.6669 | 109.8084
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample34 | 34 | 1 | 1 | TP | ug P/L | 0.0562 | 1.29 | 18.7 | TN | ug N/L | 0.188 | 3.43 | 212 |  | 1 | 2 | reworked calibration | 25.4968 | 25.4968 | 25.4968 | [6 empty] | 1 |  | 223.37 | 223.37 | 223.37 | [7 empty] | 17 | 104.5367 | 91.5826
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 30 | 1 | 1 | TP | ug P/L | 0.0448 | 1.04 | 23.5 | TN | ug N/L | 0.314 | 4.62 | 237 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 | [6 empty] | 1 |  | 237 | 237 | 237 | [7 empty] | 18 | 96.35 | 97.171
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 31 | 1 | 1 | TP | ug P/L | 0.0468 | 1.09 | 25.1 | TN | ug N/L | 0.32 | 4.67 | 241 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | [6 empty] | 1 |  | 241 | 241 | 241 | [7 empty] | 19 | 102.91 | 98.811
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 32 | 1 | 1 | TP | ug P/L | 0.0543 | 1.33 | 32.7 | TN | ug N/L | 0.32 | 4.67 | 240 |  | 1 | 1 |  | 32.7 | 32.7 | 32.7 | [6 empty] | 1 |  | 240 | 240 | 240 | [7 empty] | 20 | 134.07 | 98.401
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 33 | 1 | 1 | TP | ug P/L | 0.0498 | 1.12 | 26.2 | TN | ug N/L | 0.388 | 5.61 | 306 |  | 1 | 1 |  | 26.2 | 26.2 | 26.2 | [6 empty] | 1 |  | 306 | 306 | 306 | [7 empty] | 21 | 107.42 | 125.4613
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 34 | 1 | 1 | TP | ug P/L | 0.0446 | 1.05 | 23.8 | TN | ug N/L | 0.35 | 5.12 | 272 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | [6 empty] | 1 |  | 272 | 272 | 272 | [7 empty] | 22 | 97.58 | 111.5211
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 30 | 1 | 1 | TP | ug P/L | 0.0592 | 1.17 | 23.3 | TN | ug N/L | 0.252 | 3.83 | 178 |  | 1 | 1 |  | 23.3 | 23.3 | 23.3 | [6 empty] | 1 |  | 178 | 178 | 178 | [7 empty] | 23 | 95.53 | 72.9807
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 31 | 1 | 1 | TP | ug P/L | 0.0608 | 1.21 | 24.5 | TN | ug N/L | 0.26 | 3.9 | 183 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | [6 empty] | 1 |  | 183 | 183 | 183 | [7 empty] | 24 | 100.45 | 75.0308
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 32 | 1 | 1 | TP | ug P/L | 0.0593 | 1.19 | 23.9 | TN | ug N/L | 0.247 | 3.72 | 172 |  | 1 | 1 |  | 23.9 | 23.9 | 23.9 | [6 empty] | 1 |  | 172 | 172 | 172 | [7 empty] | 25 | 97.99 | 70.5207
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 33 | 1 | 1 | TP | ug P/L | 0.0597 | 1.23 | 25 | TN | ug N/L | 0.259 | 3.93 | 185 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 185 | 185 | 185 | [7 empty] | 26 | 102.5 | 75.8508
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 34 | 1 | 1 | TP | ug P/L | 0.0599 | 1.18 | 23.6 | TN | ug N/L | 0.274 | 4.1 | 195 |  | 1 | 1 |  | 23.6 | 23.6 | 23.6 | [6 empty] | 1 |  | 195 | 195 | 195 | [7 empty] | 27 | 96.76 | 79.9508
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 156 | 1 | 1 | TP | ug P/L | 0.0609 | 1.2 | 24 | TN | ug N/L | 0.293 | 4.33 | 210 |  | 1 | 1 |  | 24 | 24 | 24 | [6 empty] | 1 |  | 210 | 210 | 210 | [7 empty] | 28 | 98.4 | 86.1009
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 157 | 1 | 1 | TP | ug P/L | 0.0625 | 1.25 | 25.4 | TN | ug N/L | 0.307 | 4.54 | 224 |  | 1 | 1 |  | 25.4 | 25.4 | 25.4 | [6 empty] | 1 |  | 224 | 224 | 224 | [7 empty] | 29 | 104.14 | 91.8409
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 158 | 1 | 1 | TP | ug P/L | 0.0616 | 1.23 | 25 | TN | ug N/L | 0.301 | 4.45 | 218 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 218 | 218 | 218 | [7 empty] | 30 | 102.5 | 89.3809
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 159 | 1 | 1 | TP | ug P/L | 0.0621 | 1.28 | 26.4 | TN | ug N/L | 0.307 | 4.53 | 223 |  | 1 | 1 |  | 26.4 | 26.4 | 26.4 | [6 empty] | 1 |  | 223 | 223 | 223 | [7 empty] | 31 | 108.24 | 91.4309
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 160 | 1 | 1 | TP | ug P/L | 0.0619 | 1.24 | 25.1 | TN | ug N/L | 0.299 | 4.42 | 216 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | [6 empty] | 1 |  | 216 | 216 | 216 | [7 empty] | 32 | 102.91 | 88.5609
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 030 | 30 | 1 | 1 | TP | ug P/L | 0.0584 | 1.09 | 23.7 | TN | ug N/L | 0.274 | 3.79 | 174 |  | 1 | 1 |  | 23.7 | 23.7 | 23.7 | [6 empty] | 1 |  | 174 | 174 | 174 | [7 empty] | 33 | 97.17 | 71.3407
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 031 | 31 | 1 | 1 | TP | ug P/L | 0.0581 | 1.05 | 22.6 | TN | ug N/L | 0.342 | 4.65 | 230 |  | 1 | 1 |  | 22.6 | 22.6 | 22.6 | [6 empty] | 1 |  | 230 | 230 | 230 | [7 empty] | 34 | 92.66 | 94.3009
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 032 | 32 | 1 | 1 | TP | ug P/L | 0.0571 | 1.05 | 22.5 | TN | ug N/L | 0.341 | 4.65 | 230 |  | 1 | 1 |  | 22.5 | 22.5 | 22.5 | [6 empty] | 1 |  | 230 | 230 | 230 | [7 empty] | 35 | 92.25 | 94.3009
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 033 | 33 | 1 | 1 | TP | ug P/L | 0.058 | 1.14 | 25 | TN | ug N/L | 0.339 | 4.64 | 229 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 229 | 229 | 229 | [7 empty] | 36 | 102.5 | 93.8909
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 034 | 34 | 1 | 1 | TP | ug P/L | 0.0583 | 1.16 | 25.8 | TN | ug N/L | 0.364 | 5.06 | 257 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | [6 empty] | 1 |  | 257 | 257 | 257 | [7 empty] | 37 | 105.78 | 105.3711
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | spiked blank 1 | 30 | 1 | 1 | TP | ug P/L | 0.0665 | 1.2 | 21.4 | TN | ug N/L | 0.263 | 3.61 | 155 |  | 1 | 1 |  | 21.4 | 21.4 | 21.4 | [6 empty] | 1 |  | 155 | 155 | 155 | [7 empty] | 38 | 87.74 | 63.5506
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | spiked blank 2 | 31 | 1 | 1 | TP | ug P/L | 0.0559 | 1.03 | 17.4 | TN | ug N/L | 0.256 | 3.51 | 148 |  | 1 | 1 |  | 17.4 | 17.4 | 17.4 | [6 empty] | 1 |  | 148 | 148 | 148 | [7 empty] | 39 | 71.34 | 60.6806
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | spiked blank 3 | 32 | 1 | 1 | TP | ug P/L | 0.0555 | 1.03 | 17.2 | TN | ug N/L | 0.325 | 4.44 | 210 |  | 1 | 1 |  | 17.2 | 17.2 | 17.2 | [6 empty] | 1 |  | 210 | 210 | 210 | [7 empty] | 40 | 70.52 | 86.1009
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | spiked blank 4 | 33 | 1 | 1 | TP | ug P/L | 0.0562 | 1.06 | 18 | TN | ug N/L | 0.332 | 4.53 | 216 |  | 1 | 1 |  | 18 | 18 | 18 | [6 empty] | 1 |  | 216 | 216 | 216 | [7 empty] | 41 | 73.8 | 88.5609
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | spiked blank 5 | 34 | 1 | 1 | TP | ug P/L | 0.0564 | 1.08 | 18.4 | TN | ug N/L | 0.312 | 4.21 | 194 |  | 1 | 1 |  | 18.4 | 18.4 | 18.4 | [6 empty] | 1 |  | 194 | 194 | 194 | [7 empty] | 42 | 75.44 | 79.5408
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank1 | 30 | 1 | 1 | TP | ug P/L | 0.0821 | 1.39 | 30 | TN | ug N/L | 0.329 | 4.65 | 254 |  | 1 | 1 |  | 30 | 30 | 30 | [6 empty] | 1 |  | 254 | 254 | 254 | [7 empty] | 43 | 123 | 104.141
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank2 | 31 | 1 | 1 | TP | ug P/L | 0.0841 | 1.48 | 32.4 | TN | ug N/L | 0.335 | 4.69 | 256 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 | [6 empty] | 1 |  | 256 | 256 | 256 | [7 empty] | 44 | 132.84 | 104.961
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank3 | 32 | 1 | 1 | TP | ug P/L | 0.0827 | 1.49 | 32.7 | TN | ug N/L | 0.328 | 4.6 | 250 |  | 1 | 1 |  | 32.7 | 32.7 | 32.7 | [6 empty] | 1 |  | 250 | 250 | 250 | [7 empty] | 45 | 134.07 | 102.501
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank4 | 33 | 1 | 1 | TP | ug P/L | 0.0846 | 1.45 | 31.6 | TN | ug N/L | 0.328 | 4.6 | 250 |  | 1 | 1 |  | 31.6 | 31.6 | 31.6 | [6 empty] | 1 |  | 250 | 250 | 250 | [7 empty] | 46 | 129.56 | 102.501
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank 5 | 34 | 1 | 1 | TP | ug P/L | 0.0825 | 1.42 | 30.8 | TN | ug N/L | 0.331 | 4.65 | 254 |  | 1 | 1 |  | 30.8 | 30.8 | 30.8 | [6 empty] | 1 |  | 254 | 254 | 254 | [7 empty] | 47 | 126.28 | 104.141
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spiked blank BRN | 155 | 1 | 1 | TP | ug P/L | 0.0775 | 1.35 | 28.9 | TN | ug N/L | 0.267 | 3.76 | 191 |  | 1 | 1 |  | 28.9 | 28.9 | 28.9 | [6 empty] | 1 |  | 191 | 191 | 191 | [7 empty] | 48 | 118.49 | 78.3108
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Spike blank BRN | 156 | 1 | 1 | TP | ug P/L | 0.0817 | 1.42 | 30.8 | TN | ug N/L | 0.278 | 3.9 | 201 |  | 1 | 1 |  | 30.8 | 30.8 | 30.8 | [6 empty] | 1 |  | 201 | 201 | 201 | [7 empty] | 49 | 126.28 | 82.4108
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spiked blank BRN | 157 | 1 | 1 | TP | ug P/L | 0.0773 | 1.33 | 28.2 | TN | ug N/L | 0.27 | 3.83 | 195 |  | 1 | 1 |  | 28.2 | 28.2 | 28.2 | [6 empty] | 1 |  | 195 | 195 | 195 | [7 empty] | 50 | 115.62 | 79.9508
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | spikeblank1 | 30 | 1 | 1 | TP | ug P/L | 0.0791 | 1.31 | 26.5 | TN | ug N/L | 0.373 | 4.64 | 245 |  | 1 | 1 |  | 26.5 | 26.5 | 26.5 | [6 empty] | 1 |  | 245 | 245 | 245 | [7 empty] | 51 | 108.65 | 100.451
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | spikeblank2 | 31 | 1 | 1 | TP | ug P/L | 0.0788 | 1.29 | 26.2 | TN | ug N/L | 0.36 | 4.47 | 234 |  | 1 | 1 |  | 26.2 | 26.2 | 26.2 | [6 empty] | 1 |  | 234 | 234 | 234 | [7 empty] | 52 | 107.42 | 95.941
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | spikeblank3 | 32 | 1 | 1 | TP | ug P/L | 0.0799 | 1.35 | 27.7 | TN | ug N/L | 0.385 | 4.72 | 251 |  | 1 | 1 |  | 27.7 | 27.7 | 27.7 | [6 empty] | 1 |  | 251 | 251 | 251 | [7 empty] | 53 | 113.57 | 102.911
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | spikeblank4 | 33 | 1 | 1 | TP | ug P/L | 0.0784 | 1.3 | 26.4 | TN | ug N/L | 0.373 | 4.51 | 237 |  | 1 | 1 |  | 26.4 | 26.4 | 26.4 | [6 empty] | 1 |  | 237 | 237 | 237 | [7 empty] | 54 | 108.24 | 97.171
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | spikeblank5 | 34 | 1 | 1 | TP | ug P/L | 0.0769 | 1.27 | 25.4 | TN | ug N/L | 0.385 | 4.74 | 253 |  | 1 | 1 |  | 25.4 | 25.4 | 25.4 | [6 empty] | 1 |  | 253 | 253 | 253 | [7 empty] | 55 | 104.14 | 103.731
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 30 | 1 | 1 | TP | ug P/L | 0.0724 | 1.02 | -43 | TN | ug N/L | 0.433 | 5.75 | 217 |  | 1 | 1 |  | 23.2103 | 23.2103 | 23.2103 | [6 empty] | 1 |  | 217 | 217 | 217 | [7 empty] | 56 | 95.1622 | 88.9709
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | skp blank | 31 | 1 | 1 | TP | ug P/L | 0.0746 | 1.06 | -42.9 | TN | ug N/L | 0.436 | 5.89 | 224 |  | 1 | 1 |  | 25.9067 | 25.9067 | 25.9067 | [6 empty] | 1 |  | 224 | 224 | 224 | [7 empty] | 57 | 106.2173 | 91.8409
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 32 | 1 | 1 | TP | ug P/L | 0.0761 | 0.935 | -42.8 | TN | ug N/L | 0.441 | 5.9 | 224 |  | 1 | 1 |  | 27.7103 | 27.7103 | 27.7103 | [6 empty] | 1 |  | 224 | 224 | 224 | [7 empty] | 58 | 113.612 | 91.8409
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 33 | 1 | 1 | TP | ug P/L | 0.07 | 0.875 | -43.2 | TN | ug N/L | 0.437 | 5.85 | 222 |  | 1 | 1 |  | 20.1995 | 20.1995 | 20.1995 | [6 empty] | 1 |  | 222 | 222 | 222 | [7 empty] | 59 | 82.8181 | 91.0209
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk blank | 34 | 1 | 1 | TP | ug P/L | 0.0694 | 0.862 | -43.2 | TN | ug N/L | 0.446 | 5.92 | 225 |  | 1 | 1 |  | 19.4356 | 19.4356 | 19.4356 | [6 empty] | 1 |  | 225 | 225 | 225 | [7 empty] | 60 | 79.6858 | 92.2509
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkblank | 30 | 1 | 1 | TP | ug P/L | 0.139 | 2.38 | 25.7 | TN | ug N/L | 0.522 | 7.34 | 265 |  | 1 | 1 |  | 25.7 | 25.7 | 25.7 | [6 empty] | 1 |  | 265 | 265 | 265 | [7 empty] | 61 | 105.37 | 108.6511
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkblank | 31 | 1 | 1 | TP | ug P/L | 0.135 | 2.25 | 21.7 | TN | ug N/L | 0.492 | 7.08 | 252 |  | 1 | 1 |  | 21.7 | 21.7 | 21.7 | [6 empty] | 1 |  | 252 | 252 | 252 | [7 empty] | 62 | 88.97 | 103.321
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkblank | 32 | 1 | 1 | TP | ug P/L | 0.136 | 2.28 | 22.4 | TN | ug N/L | 0.52 | 7.14 | 255 |  | 1 | 1 |  | 22.4 | 22.4 | 22.4 | [6 empty] | 1 |  | 255 | 255 | 255 | [7 empty] | 63 | 91.84 | 104.551
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkblank | 33 | 1 | 1 | TP | ug P/L | 0.134 | 2.24 | 21.2 | TN | ug N/L | 0.488 | 6.62 | 230 |  | 1 | 1 |  | 21.2 | 21.2 | 21.2 | [6 empty] | 1 |  | 230 | 230 | 230 | [7 empty] | 64 | 86.92 | 94.3009
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkblank | 34 | 1 | 1 | TP | ug P/L | 0.135 | 2.31 | 23.5 | TN | ug N/L | 0.486 | 6.71 | 234 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 | [6 empty] | 1 |  | 234 | 234 | 234 | [7 empty] | 65 | 96.35 | 95.941
2023-02-22 | OM_2-22-2023_04-26-39PM BRN edit.omn | spkblank | 30 | 1 | 1 | TP | ug P/L | 0.137 | 2.33 | 24.1 | TN | ug N/L | 0.48 | 6.51 | 225 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | [6 empty] | 1 |  | 225 | 225 | 225 | [7 empty] | 66 | 98.81 | 92.2509
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | spk1 | 30 | 1 | 1 | TP | ug P/L | 0.063 | 1.11 | 22.8 | TN | ug N/L | 0.15 | 2.19 | 239 |  | 1 | 1 |  | 22.8 | 22.8 | 22.8 | [6 empty] | 1 |  | 239 | 239 | 239 | [7 empty] | 67 | 93.48 | 97.991
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | spk2 | 31 | 1 | 1 | TP | ug P/L | 0.0625 | 1.11 | 22.9 | TN | ug N/L | 0.146 | 2.08 | 222 |  | 1 | 1 |  | 22.9 | 22.9 | 22.9 | [6 empty] | 1 |  | 222 | 222 | 222 | [7 empty] | 68 | 93.89 | 91.0209
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | spk3 | 32 | 1 | 1 | TP | ug P/L | 0.0644 | 1.1 | 22.7 | TN | ug N/L | 0.14 | 2.09 | 223 |  | 1 | 1 |  | 22.7 | 22.7 | 22.7 | [6 empty] | 1 |  | 223 | 223 | 223 | [7 empty] | 69 | 93.07 | 91.4309
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | spk4 | 33 | 1 | 1 | TP | ug P/L | 0.062 | 1.1 | 22.5 | TN | ug N/L | 0.14 | 2.03 | 215 |  | 1 | 1 |  | 22.5 | 22.5 | 22.5 | [6 empty] | 1 |  | 215 | 215 | 215 | [7 empty] | 70 | 92.25 | 88.1509
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | spk5 | 34 | 1 | 1 | TP | ug P/L | 0.0644 | 1.17 | 24.7 | TN | ug N/L | 0.147 | 2.14 | 231 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | [6 empty] | 1 |  | 231 | 231 | 231 | [7 empty] | 71 | 101.27 | 94.7109
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | spik1 | 30 | 1 | 1 | TP | ug P/L | 0.0675 | 1.16 | 24.7 | TN | ug N/L | 0.132 | 1.89 | 193 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | [6 empty] | 1 |  | 193 | 193 | 193 | [7 empty] | 72 | 101.27 | 79.1308
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | spike2 | 31 | 1 | 1 | TP | ug P/L | 0.0684 | 1.18 | 25.5 | TN | ug N/L | 0.14 | 1.86 | 189 |  | 1 | 1 |  | 25.5 | 25.5 | 25.5 | [6 empty] | 1 |  | 189 | 189 | 189 | [7 empty] | 73 | 104.55 | 77.4908
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | spike3 | 32 | 1 | 1 | TP | ug P/L | 0.0677 | 1.19 | 25.6 | TN | ug N/L | 0.139 | 1.87 | 191 |  | 1 | 1 |  | 25.6 | 25.6 | 25.6 | [6 empty] | 1 |  | 191 | 191 | 191 | [7 empty] | 74 | 104.96 | 78.3108
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | spike4 | 33 | 1 | 1 | TP | ug P/L | 0.0672 | 1.16 | 24.8 | TN | ug N/L | 0.128 | 1.82 | 183 |  | 1 | 1 |  | 24.8 | 24.8 | 24.8 | [6 empty] | 1 |  | 183 | 183 | 183 | [7 empty] | 75 | 101.68 | 75.0308
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | spike5 | 34 | 1 | 1 | TP | ug P/L | 0.068 | 1.2 | 25.9 | TN | ug N/L | 0.133 | 1.92 | 197 |  | 1 | 1 |  | 25.9 | 25.9 | 25.9 | [6 empty] | 1 |  | 197 | 197 | 197 | [7 empty] | 76 | 106.19 | 80.7708
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | spk1 | 30 | 1 | 1 | TP | ug P/L | 0.0606 | 1.05 | 25.2 | TN | ug N/L | 0.127 | 1.83 | 206 |  | 1 | 1 |  | 25.2 | 25.2 | 25.2 | [6 empty] | 1 |  | 206 | 206 | 206
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | spk2 | 31 | 1 | 1 | TP | ug P/L | 0.0594 | 1 | 23.7 | TN | ug N/L | 0.138 | 1.88 | 214 |  | 1 | 1 |  | 23.7 | 23.7 | 23.7 | [6 empty] | 1 |  | 214 | 214 | 214
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | spk3 | 32 | 1 | 1 | TP | ug P/L | 0.0581 | 0.993 | 23.4 | TN | ug N/L | 0.138 | 1.9 | 217 |  | 1 | 1 |  | 23.4 | 23.4 | 23.4 | [6 empty] | 1 |  | 217 | 217 | 217
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | spk4 | 33 | 1 | 1 | TP | ug P/L | 0.0577 | 0.99 | 23.3 | TN | ug N/L | 0.131 | 1.92 | 220 |  | 1 | 1 |  | 23.3 | 23.3 | 23.3 | [6 empty] | 1 |  | 220 | 220 | 220
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | spk5 | 34 | 1 | 1 | TP | ug P/L | 0.0595 | 1.02 | 24.2 | TN | ug N/L | 0.135 | 1.88 | 214 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | [6 empty] | 1 |  | 214 | 214 | 214
TN
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: |  | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | [6 empty] | OBSERVATION #
| Mean | [5 empty] | Mean |  |  | 20.0582 |  |  |  |  | 209.5906 |  |  |  | Mean |  |  | 24.6979 | [7 empty] | Mean |  |  | 224.0041 | [6 empty] | Min | 3
| Std | [5 empty] | Std |  |  | 16.83 |  |  |  |  | 57.49 |  |  |  | Std |  |  | 3.196 | [7 empty] | Std |  |  | 29.1706 | [6 empty] | Max | 76
| CV | [5 empty] | CV |  |  | 83.9055 |  |  |  |  | 27.4297 |  |  |  | CV |  |  | 12.9402 | [7 empty] | CV |  |  | 13.0224
[7 empty] | S/N |  |  | 1.1918 |  |  |  |  | 3.6457 |  |  |  | S/N |  |  | 7.7279 | [7 empty] | S/N |  |  | 7.6791
Warning Limits | Xbar + 2SD |  |  |  |  | Warning Limits | Xbar + 2SD |  |  | 53.7181 |  |  |  |  | 324.5706 |  |  | Warning Limits | Xbar + 2SD |  |  | 31.0898 | 125.8804 | [5 empty] | Warning Limits | Xbar + 2SD |  |  | 282.3454
| Xbar - 2SD | [5 empty] | Xbar - 2SD |  |  | -13.6017 |  |  |  |  | 94.6106 |  |  |  | Xbar - 2SD |  |  | 18.306 | 74.1196 | [6 empty] | Xbar - 2SD |  |  | 165.6628
Control Limits | Xbar + 3SD |  |  |  |  | Control Limits | Xbar + 3SD |  |  | 70.5481 |  |  |  |  | 382.0606 |  |  | Control Limits | Xbar + 3SD |  |  | 34.2857 | [6 empty] | Control Limits | Xbar + 3SD |  |  | 311.5161
| Xbar - 3SD | [5 empty] | Xbar - 3SD |  |  | -30.4316 |  |  |  |  | 37.1207 |  |  |  | Xbar - 3SD |  |  | 15.11 | [7 empty] | Xbar - 3SD |  |  | 136.4922
Count | [5 empty] | Count |  |  |  | 79 |  |  |  |  | 79 |  |  | Count |  |  |  | 79 | [6 empty] | Count |  |  |  | 79
99th percentile.inc | [5 empty] | 99th percentile.inc |  |  |  | 32.7 |  |  |  |  | 279.48 |  |  | 99th percentile.inc |  |  |  | 32.7 | [6 empty] | 99th percentile.inc |  |  |  | 32.7
Maximum | [5 empty] | Maximum |  |  |  | 32.7 |  |  |  |  | 306 |  |  | Maximum |  |  |  | 32.7 | [6 empty] | Maximum |  |  |  | 32.7
[6 empty] | Known |  |  |  | 24.4 |  |  |  |  | 243.9 |  |  | Known |  |  |  | 24.4 | [6 empty] | Known |  |  |  | 243.9
Known | [5 empty] | MDL |  |  |  | 39.973 |  |  |  |  | 136.5451 |  |  | MDL |  |  |  | 7.5907 | [6 empty] | MDL |  |  |  | 69.2835
[6 empty] | LOQ |  |  |  | 168.2995 |  |  |  |  | 574.8999 |  |  | LOQ |  |  |  | 31.9595 | [6 empty] | LOQ |  |  |  | 291.7065
[6 empty] | ratio mean/mdl |  |  |  | 0.50179 |  |  |  |  | 1.535 |  |  | ratio mean/mdl |  |  |  | 3.2537 | [6 empty] | ratio mean/mdl |  |  |  | 3.2332
[21 empty] | Known Conc | 24.3902 | [9 empty] | Known Conc | 243.9024
[21 empty] | Mean | 24.6979 | [9 empty] | Mean | 224.0041
[21 empty] | Std Dev | 3.196 | [9 empty] | Std Dev | 29.1706
[21 empty] | CV or %RSD | 12.9402 | [9 empty] | CV or %RSD | 13.0224
[21 empty] | T value | 2.3751 | [9 empty] | T value | 2.3751
[21 empty] | MDL | 7.5907 | [9 empty] | MDL | 69.2835
[21 empty] | LOQ | 31.9595 | [9 empty] | LOQ | 291.7065
[21 empty] | Error as % | 1.2613 | [9 empty] | Error as % | -8.1583
[21 empty] | Known/MDL | 3.2132 | [9 empty] | Known/MDL | 3.5204
[21 empty] | Mean PR | 101.2613 | [9 empty] | Mean PR | 91.8417
[21 empty] | S/N | 7.7279 | [9 empty] | S/N | 7.6791
[21 empty] | STDEV.P | 3.1757
[21 empty] | STDEV.S | 3.196

## rolling digest check
As an MDLs surrogate… these aren't independently prepared but they ARE independently digested. And they are at concentrations within 10 x of the resulting MDLs
As a digestion check, TNs are running low. New Glycine? Raise temperature?
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 014 | 14 | 1 | 1 | Total Phosphorus | ppb | 0.0944 | 1.96 | 34.8 | Total Nitrogen | ppb | 1.49 | 23.2 | 684 |  | 1.5 | 1 |  | 44.2165 | 44.2165 | 44.2165 | [6 empty] | 1 |  | 401.9114 | 401.9114 | 401.9114 | [7 empty] | 1
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 015 | 15 | 1 | 1 | Total Phosphorus | ppb | 0.0787 | 1.78 | 31.5 | Total Nitrogen | ppb | 1.51 | 23.4 | 689 |  | 1.5 | 1 |  | 38.5562 | 38.5562 | 38.5562 | [6 empty] | 1 |  | 409.6283 | 409.6283 | 409.6283 | [7 empty] | 2
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 016 | 16 | 1 | 1 | Total Phosphorus | ppb | 0.0801 | 1.81 | 32 | Total Nitrogen | ppb | 1.47 | 23 | 676 |  | 1.5 | 1 |  | 39.5105 | 39.5105 | 39.5105 | [6 empty] | 1 |  | 394.1916 | 394.1916 | 394.1916 | [7 empty] | 3
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 017 | 17 | 1 | 1 | Total Phosphorus | ppb | 0.078 | 1.68 | 29.7 | Total Nitrogen | ppb | 1.47 | 22.9 | 675 |  | 1.5 | 1 |  | 35.344 | 35.344 | 35.344 | [6 empty] | 1 |  | 390.3306 | 390.3306 | 390.3306 | [7 empty] | 4
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 018 | 18 | 1 | 1 | Total Phosphorus | ppb | 0.0802 | 1.81 | 32.1 | Total Nitrogen | ppb | 1.47 | 23 | 677 |  | 1.5 | 1 |  | 39.5105 | 39.5105 | 39.5105 | [6 empty] | 1 |  | 394.1916 | 394.1916 | 394.1916 | [7 empty] | 5
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | digestion check | 14 | 1 | 1 | Total Phosphorus | ppb | 0.0961 | 1.98 | 33.2 | Total Nitrogen | ppb | 1.24 | 21.5 | 582 |  | 1.5 | 1 |  | 41.9776 | 41.9776 | 41.9776 | [6 empty] | 1 |  | 403.91 | 403.91 | 403.91 | [7 empty] | 6
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | digestion check | 15 | 1 | 1 | Total Phosphorus | ppb | 0.0947 | 2.04 | 34.2 | Total Nitrogen | ppb | 1.25 | 21.7 | 587 |  | 1.5 | 1 |  | 43.6042 | 43.6042 | 43.6042 | [6 empty] | 1 |  | 412.8095 | 412.8095 | 412.8095 | [7 empty] | 7
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | digestion check | 16 | 1 | 1 | Total Phosphorus | ppb | 0.0851 | 1.84 | 30.8 | Total Nitrogen | ppb | 1.22 | 21.2 | 575 |  | 1.5 | 1 |  | 38.175 | 38.175 | 38.175 | [6 empty] | 1 |  | 390.5386 | 390.5386 | 390.5386 | [7 empty] | 8
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | digestion check | 17 | 1 | 1 | Total Phosphorus | ppb | 0.0819 | 1.65 | 27.5 | Total Nitrogen | ppb | 1.23 | 21.3 | 577 |  | 1.5 | 1 |  | 32.9988 | 32.9988 | 32.9988 | [6 empty] | 1 |  | 394.9987 | 394.9987 | 394.9987 | [7 empty] | 9
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | digestion check | 18 | 1 | 1 | Total Phosphorus | ppb | 0.0871 | 1.88 | 31.4 | Total Nitrogen | ppb | 1.23 | 21.5 | 583 |  | 1.5 | 1 |  | 39.2625 | 39.2625 | 39.2625 | [6 empty] | 1 |  | 403.91 | 403.91 | 403.91 | [7 empty] | 10
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 023 | 23 | 1 | 1 | TP | ug P/L | 0.0836 | 1.83 | 38.9 | TN | ug N/L | 0.588 | 9.89 | 32.5 |  | 1 | 1 |  | 38.9 | 38.9 | 38.9 | [6 empty] | 1 |  | 306.698 | 306.698 | 352.1846 | [7 empty] | 11
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 024 | 24 | 1 | 1 | TP | ug P/L | 0.0827 | 1.88 | 40.4 | TN | ug N/L | 0.61 | 10.4 | 36.5 |  | 1 | 1 |  | 40.4 | 40.4 | 40.4 | [6 empty] | 1 |  | 350.5406 | 350.5406 | 398.68 | [7 empty] | 12
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 025 | 25 | 1 | 1 | TP | ug P/L | 0.0803 | 1.82 | 38.6 | TN | ug N/L | 0.602 | 10.2 | 34.9 |  | 1 | 1 |  | 38.6 | 38.6 | 38.6 | [6 empty] | 1 |  | 333.2991 | 333.2991 | 381.7425 | [7 empty] | 13
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 026 | 26 | 1 | 1 | TP | ug P/L | 0.0836 | 1.85 | 39.5 | TN | ug N/L | 0.643 | 10.8 | 40.6 |  | 1 | 1 |  | 39.5 | 39.5 | 39.5 | [6 empty] | 1 |  | 385.2102 | 385.2102 | 436.4607 | [7 empty] | 14
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 027 | 27 | 1 | 1 | TP | ug P/L | 0.0822 | 1.82 | 38.4 | TN | ug N/L | 0.608 | 10.3 | 35.6 |  | 1 | 1 |  | 38.4 | 38.4 | 38.4 | [6 empty] | 1 |  | 341.9121 | 341.9121 | 392.6024 | [7 empty] | 15
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 025 | 25 | 1 | 1 | TP | ug P/L | 0.0766 | 1.69 | 38.3 | TN | ug N/L | 0.368 | 6.26 | 302 |  | 1 | 1 |  | 38.3 | 38.3 | 38.3 | [6 empty] | 1 |  | 302 | 302 | 302 | [7 empty] | 16
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 026 | 26 | 1 | 1 | TP | ug P/L | 0.0732 | 1.67 | 37.5 | TN | ug N/L | 0.351 | 6.03 | 285 |  | 1 | 1 |  | 37.5 | 37.5 | 37.5 | [6 empty] | 1 |  | 285 | 285 | 285 | [7 empty] | 17
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 027 | 27 | 1 | 1 | TP | ug P/L | 0.0768 | 1.74 | 39.6 | TN | ug N/L | 0.428 | 7.26 | 376 |  | 1 | 1 |  | 39.6 | 39.6 | 39.6 | [6 empty] | 1 |  | 376 | 376 | 376 | [7 empty] | 18
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 028 | 28 | 1 | 1 | TP | ug P/L | 0.0743 | 1.69 | 38.2 | TN | ug N/L | 0.381 | 6.5 | 320 |  | 1 | 1 |  | 38.2 | 38.2 | 38.2 | [6 empty] | 1 |  | 320 | 320 | 320 | [7 empty] | 19
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 029 | 29 | 1 | 1 | TP | ug P/L | 0.0741 | 1.67 | 37.5 | TN | ug N/L | 0.435 | 7.39 | 386 |  | 1 | 1 |  | 37.5 | 37.5 | 37.5 | [6 empty] | 1 |  | 386 | 386 | 386 | [7 empty] | 20
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample25 | 25 | 1 | 1 | TP | ug P/L | 0.088 | 1.9 | 35.4 | TN | ug N/L | 0.25 | 4.4 | 297 |  | 1 | 2 | reworked calibration | 43.8644 | 43.8644 | 43.8644 | [6 empty] | 1 |  | 315.4328 | 315.4328 | 315.4328 | [7 empty] | 21
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample26 | 26 | 1 | 1 | TP | ug P/L | 0.0853 | 1.86 | 34.3 | TN | ug N/L | 0.194 | 3.51 | 219 |  | 1 | 2 | reworked calibration | 42.6515 | 42.6515 | 42.6515 | [6 empty] | 1 |  | 230.9158 | 230.9158 | 230.9158 | [7 empty] | 22
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample27 | 27 | 1 | 1 | TP | ug P/L | 0.0867 | 1.88 | 34.8 | TN | ug N/L | 0.28 | 5.09 | 358 |  | 1 | 2 | reworked calibration | 43.2578 | 43.2578 | 43.2578 | [6 empty] | 1 |  | 381.6766 | 381.6766 | 381.6766 | [7 empty] | 23
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample28 | 28 | 1 | 1 | TP | ug P/L | 0.0847 | 1.81 | 32.9 | TN | ug N/L | 0.232 | 4.2 | 280 |  | 1 | 2 | reworked calibration | 41.1371 | 41.1371 | 41.1371 | [6 empty] | 1 |  | 296.3492 | 296.3492 | 296.3492 | [7 empty] | 24
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample29 | 29 | 1 | 1 | TP | ug P/L | 0.0842 | 1.88 | 34.8 | TN | ug N/L | 0.208 | 3.79 | 243 |  | 1 | 2 | reworked calibration | 43.2578 | 43.2578 | 43.2578 | [6 empty] | 1 |  | 257.3927 | 257.3927 | 257.3927 | [7 empty] | 25
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample25 | 25 | 1 | 1 | TP | ug P/L | 0.088 | 1.9 | 35.4 | TN | ug N/L | 0.25 | 4.4 | 297 |  | 1 | 2 | reworked calibration | 43.8644 | 43.8644 | 43.8644 | [6 empty] | 1 |  | 315.4328 | 315.4328 | 315.4328 | [7 empty] | 26
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample26 | 26 | 1 | 1 | TP | ug P/L | 0.0853 | 1.86 | 34.3 | TN | ug N/L | 0.194 | 3.51 | 219 |  | 1 | 2 | reworked calibration | 42.6515 | 42.6515 | 42.6515 | [6 empty] | 1 |  | 230.9158 | 230.9158 | 230.9158 | [7 empty] | 27
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample27 | 27 | 1 | 1 | TP | ug P/L | 0.0867 | 1.88 | 34.8 | TN | ug N/L | 0.28 | 5.09 | 358 |  | 1 | 2 | reworked calibration | 43.2578 | 43.2578 | 43.2578 | [6 empty] | 1 |  | 381.6766 | 381.6766 | 381.6766 | [7 empty] | 28
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample28 | 28 | 1 | 1 | TP | ug P/L | 0.0847 | 1.81 | 32.9 | TN | ug N/L | 0.232 | 4.2 | 280 |  | 1 | 2 | reworked calibration | 41.1371 | 41.1371 | 41.1371 | [6 empty] | 1 |  | 296.3492 | 296.3492 | 296.3492 | [7 empty] | 29
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample29 | 29 | 1 | 1 | TP | ug P/L | 0.0842 | 1.88 | 34.8 | TN | ug N/L | 0.208 | 3.79 | 243 |  | 1 | 2 | reworked calibration | 43.2578 | 43.2578 | 43.2578 | [6 empty] | 1 |  | 257.3927 | 257.3927 | 257.3927 | [7 empty] | 30
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | digestion check | 25 | 1 | 1 | TP | ug P/L | 0.0788 | 1.63 | 42 | TN | ug N/L | 0.485 | 7.01 | 402 |  | 1 | 1 |  | 42 | 42 | 42 | [6 empty] | 1 |  | 402 | 402 | 402 | [7 empty] | 31
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | digestion check | 26 | 1 | 1 | TP | ug P/L | 0.0764 | 1.6 | 41.2 | TN | ug N/L | 0.474 | 6.82 | 389 |  | 1 | 1 |  | 41.2 | 41.2 | 41.2 | [6 empty] | 1 |  | 389 | 389 | 389 | [7 empty] | 32
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | digestion check | 27 | 1 | 1 | TP | ug P/L | 0.0775 | 1.61 | 41.6 | TN | ug N/L | 0.495 | 7.14 | 411 |  | 1 | 1 |  | 41.6 | 41.6 | 41.6 | [6 empty] | 1 |  | 411 | 411 | 411 | [7 empty] | 33
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | digestion check | 28 | 1 | 1 | TP | ug P/L | 0.0755 | 1.61 | 41.6 | TN | ug N/L | 0.488 | 7.07 | 407 |  | 1 | 1 |  | 41.6 | 41.6 | 41.6 | [6 empty] | 1 |  | 407 | 407 | 407 | [7 empty] | 34
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | digestion check | 29 | 1 | 1 | TP | ug P/L | 0.0759 | 1.6 | 41.1 | TN | ug N/L | 0.505 | 7.24 | 418 |  | 1 | 1 |  | 41.1 | 41.1 | 41.1 | [6 empty] | 1 |  | 418 | 418 | 418 | [7 empty] | 35
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | digest check | 25 | 1 | 1 | TP | ug P/L | 0.0857 | 1.69 | 37.8 | TN | ug N/L | 0.373 | 5.56 | 289 |  | 1 | 1 |  | 37.8 | 37.8 | 37.8 | [6 empty] | 1 |  | 289 | 289 | 289 | [7 empty] | 36
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | digest check | 26 | 1 | 1 | TP | ug P/L | 0.0869 | 1.71 | 38.3 | TN | ug N/L | 0.396 | 5.96 | 314 |  | 1 | 1 |  | 38.3 | 38.3 | 38.3 | [6 empty] | 1 |  | 314 | 314 | 314 | [7 empty] | 37
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | digest check | 27 | 1 | 1 | TP | ug P/L | 0.0868 | 1.69 | 37.7 | TN | ug N/L | 0.365 | 5.58 | 290 |  | 1 | 1 |  | 37.7 | 37.7 | 37.7 | [6 empty] | 1 |  | 290 | 290 | 290 | [7 empty] | 38
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | digest check | 28 | 1 | 1 | TP | ug P/L | 0.0866 | 1.69 | 37.6 | TN | ug N/L | 0.403 | 5.97 | 315 |  | 1 | 1 |  | 37.6 | 37.6 | 37.6 | [6 empty] | 1 |  | 315 | 315 | 315 | [7 empty] | 39
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | digest check | 29 | 1 | 1 | TP | ug P/L | 0.0868 | 1.69 | 37.7 | TN | ug N/L | 0.363 | 5.41 | 279 |  | 1 | 1 |  | 37.7 | 37.7 | 37.7 | [6 empty] | 1 |  | 279 | 279 | 279 | [7 empty] | 40
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 025 | 25 | 1 | 1 | TP | ug P/L | 0.0865 | 1.61 | 38.6 | TN | ug N/L | 0.439 | 5.98 | 317 |  | 1 | 1 |  | 38.6 | 38.6 | 38.6 | [6 empty] | 1 |  | 317 | 317 | 317 | [7 empty] | 41
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 026 | 26 | 1 | 1 | TP | ug P/L | 0.0869 | 1.66 | 39.8 | TN | ug N/L | 0.458 | 6.22 | 333 |  | 1 | 1 |  | 39.8 | 39.8 | 39.8 | [6 empty] | 1 |  | 333 | 333 | 333 | [7 empty] | 42
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 027 | 27 | 1 | 1 | TP | ug P/L | 0.087 | 1.61 | 38.4 | TN | ug N/L | 0.415 | 5.67 | 297 |  | 1 | 1 |  | 38.4 | 38.4 | 38.4 | [6 empty] | 1 |  | 297 | 297 | 297 | [7 empty] | 43
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 028 | 28 | 1 | 1 | TP | ug P/L | 0.0851 | 1.56 | 37.1 | TN | ug N/L | 0.489 | 6.6 | 357 |  | 1 | 1 |  | 37.1 | 37.1 | 37.1 | [6 empty] | 1 |  | 357 | 357 | 357 | [7 empty] | 44
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 029 | 29 | 1 | 1 | TP | ug P/L | 0.0848 | 1.52 | 35.9 | TN | ug N/L | 0.383 | 5.2 | 266 |  | 1 | 1 |  | 35.9 | 35.9 | 35.9 | [6 empty] | 1 |  | 266 | 266 | 266 | [7 empty] | 45
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | digest check 1 | 25 | 1 | 1 | TP | ug P/L | 0.095 | 1.71 | 34.1 | TN | ug N/L | 0.424 | 5.7 | 294 |  | 1 | 1 |  | 34.1 | 34.1 | 34.1 | [6 empty] | 1 |  | 294 | 294 | 294 | [7 empty] | 46
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | digest check 2 | 26 | 1 | 1 | TP | ug P/L | 0.097 | 1.73 | 34.5 | TN | ug N/L | 0.516 | 6.88 | 373 |  | 1 | 1 |  | 34.5 | 34.5 | 34.5 | [6 empty] | 1 |  | 373 | 373 | 373 | [7 empty] | 47
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | digest check 3 | 27 | 1 | 1 | TP | ug P/L | 0.0959 | 1.71 | 34.1 | TN | ug N/L | 0.433 | 5.81 | 302 |  | 1 | 1 |  | 34.1 | 34.1 | 34.1 | [6 empty] | 1 |  | 302 | 302 | 302 | [7 empty] | 48
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | digest check 4 | 28 | 1 | 1 | TP | ug P/L | 0.0969 | 1.74 | 34.7 | TN | ug N/L | 0.455 | 6.09 | 320 |  | 1 | 1 |  | 34.7 | 34.7 | 34.7 | [6 empty] | 1 |  | 320 | 320 | 320 | [7 empty] | 49
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | digest check 5 | 29 | 1 | 1 | TP | ug P/L | 0.0971 | 1.74 | 34.7 | TN | ug N/L | 0.437 | 5.87 | 306 |  | 1 | 1 |  | 34.7 | 34.7 | 34.7 | [6 empty] | 1 |  | 306 | 306 | 306 | [7 empty] | 50
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check1 | 25 | 1 | 1 | TP | ug P/L | 0.0941 | 1.67 | 37.7 | TN | ug N/L | 0.453 | 6.33 | 372 |  | 1 | 1 |  | 37.7 | 37.7 | 37.7 | [6 empty] | 1 |  | 372 | 372 | 372 | [7 empty] | 51
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check2 | 26 | 1 | 1 | TP | ug P/L | 0.0948 | 1.66 | 37.4 | TN | ug N/L | 0.436 | 6.12 | 357 |  | 1 | 1 |  | 37.4 | 37.4 | 37.4 | [6 empty] | 1 |  | 357 | 357 | 357 | [7 empty] | 52
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digestcheck3 | 27 | 1 | 1 | TP | ug P/L | 0.0908 | 1.58 | 35.2 | TN | ug N/L | 0.425 | 5.93 | 344 |  | 1 | 1 |  | 35.2 | 35.2 | 35.2 | [6 empty] | 1 |  | 344 | 344 | 344 | [7 empty] | 53
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digestcheck4 | 28 | 1 | 1 | TP | ug P/L | 0.0929 | 1.65 | 37.2 | TN | ug N/L | 0.448 | 6.21 | 363 |  | 1 | 1 |  | 37.2 | 37.2 | 37.2 | [6 empty] | 1 |  | 363 | 363 | 363 | [7 empty] | 54
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check5 | 29 | 1 | 1 | TP | ug P/L | 0.0941 | 1.65 | 37.3 | TN | ug N/L | 0.459 | 6.36 | 374 |  | 1 | 1 |  | 37.3 | 37.3 | 37.3 | [6 empty] | 1 |  | 374 | 374 | 374 | [7 empty] | 55
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | digestchk1 | 25 | 1 | 1 | TP | ug P/L | 0.104 | 1.74 | 39.1 | TN | ug N/L | 0.542 | 6.6 | 380 |  | 1 | 1 |  | 39.1 | 39.1 | 39.1 | [6 empty] | 1 |  | 380 | 380 | 380 | [7 empty] | 56
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | digestchk2 | 26 | 1 | 1 | TP | ug P/L | 0.101 | 1.71 | 38.1 | TN | ug N/L | 0.54 | 6.64 | 383 |  | 1 | 1 |  | 38.1 | 38.1 | 38.1 | [6 empty] | 1 |  | 383 | 383 | 383 | [7 empty] | 57
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | digestch3 | 27 | 1 | 1 | TP | ug P/L | 0.101 | 1.72 | 38.5 | TN | ug N/L | 0.552 | 6.7 | 387 |  | 1 | 1 |  | 38.5 | 38.5 | 38.5 | [6 empty] | 1 |  | 387 | 387 | 387 | [7 empty] | 58
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | digestchk4 | 28 | 1 | 1 | TP | ug P/L | 0.103 | 1.76 | 39.7 | TN | ug N/L | 0.554 | 6.71 | 387 |  | 1 | 1 |  | 39.7 | 39.7 | 39.7 | [6 empty] | 1 |  | 387 | 387 | 387 | [7 empty] | 59
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | digestchk5 | 29 | 1 | 1 | TP | ug P/L | 0.101 | 1.72 | 38.4 | TN | ug N/L | 0.536 | 6.53 | 375 |  | 1 | 1 |  | 38.4 | 38.4 | 38.4 | [6 empty] | 1 |  | 375 | 375 | 375 | [7 empty] | 60
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 25 | 1 | 1 | TP | ug P/L | 0.0817 | 1.27 | -42.4 | TN | ug N/L | 0.71 | 9.48 | 398 |  | 1 | 2 | reworked calibration | 34.1942 | 34.1942 | 34.1942 | [6 empty] | 1 |  | 398 | 398 | 398 | [7 empty] | 61
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 26 | 1 | 1 | TP | ug P/L | 0.0795 | 1.25 | -42.6 | TN | ug N/L | 0.707 | 9.5 | 399 |  | 1 | 2 | reworked calibration | 31.6939 | 31.6939 | 31.6939 | [6 empty] | 1 |  | 399 | 399 | 399 | [7 empty] | 62
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 27 | 1 | 1 | TP | ug P/L | 0.0827 | 1.27 | -42.4 | TN | ug N/L | 0.709 | 9.47 | 398 |  | 1 | 2 | reworked calibration | 35.3106 | 35.3106 | 35.3106 | [6 empty] | 1 |  | 398 | 398 | 398 | [7 empty] | 63
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 28 | 1 | 1 | TP | ug P/L | 0.0812 | 1.28 | -42.5 | TN | ug N/L | 0.711 | 9.46 | 397 |  | 1 | 2 | reworked calibration | 33.6312 | 33.6312 | 33.6312 | [6 empty] | 1 |  | 397 | 397 | 397 | [7 empty] | 64
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | digest | 29 | 1 | 1 | TP | ug P/L | 0.0823 | 1.19 | -42.4 | TN | ug N/L | 0.709 | 9.42 | 395 |  | 1 | 2 | reworked calibration | 34.8655 | 34.8655 | 34.8655 | [6 empty] | 1 |  | 395 | 395 | 395 | [7 empty] | 65
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | digestchk | 25 | 1 | 1 | TP | ug P/L | 0.157 | 2.66 | 35.1 | TN | ug N/L | 0.718 | 10.1 | 398 |  | 1 | 1 |  | 35.1 | 35.1 | 35.1 | [6 empty] | 1 |  | 398 | 398 | 398 | [7 empty] | 66
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | digestchk | 26 | 1 | 1 | TP | ug P/L | 0.158 | 2.69 | 36.1 | TN | ug N/L | 0.723 | 9.81 | 384 |  | 1 | 1 |  | 36.1 | 36.1 | 36.1 | [6 empty] | 1 |  | 384 | 384 | 384 | [7 empty] | 67
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | digestchk | 27 | 1 | 1 | TP | ug P/L | 0.16 | 2.76 | 38.4 | TN | ug N/L | 0.705 | 9.76 | 382 |  | 1 | 1 |  | 38.4 | 38.4 | 38.4 | [6 empty] | 1 |  | 382 | 382 | 382 | [7 empty] | 68
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | digsetchk | 28 | 1 | 1 | TP | ug P/L | 0.159 | 2.79 | 39.3 | TN | ug N/L | 0.74 | 10.5 | 417 |  | 1 | 1 |  | 39.3 | 39.3 | 39.3 | [6 empty] | 1 |  | 417 | 417 | 417 | [7 empty] | 69
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | digestchk | 29 | 1 | 1 | TP | ug P/L | 0.157 | 2.61 | 33.4 | TN | ug N/L | 0.716 | 9.9 | 389 |  | 1 | 1 |  | 33.4 | 33.4 | 33.4 | [6 empty] | 1 |  | 389 | 389 | 389 | [7 empty] | 70
2023-02-22 | OM_2-22-2023_04-26-39PM BRN edit.omn | digestchk | 25 | 1 | 1 | TP | ug P/L | 0.158 | 2.74 | 37.8 | TN | ug N/L | 0.724 | 9.75 | 382 |  | 1 | 1 |  | 37.8 | 37.8 | 37.8 | [6 empty] | 1 |  | 382 | 382 | 382 | [7 empty] | 71
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | digest1 | 25 | 1 | 1 | TP | ug P/L | 0.0942 | 1.63 | 38.8 | TN | ug N/L | 0.209 | 3.06 | 367 |  | 1 | 1 |  | 38.8 | 38.8 | 38.8 | [6 empty] | 1 |  | 367 | 367 | 367 | [7 empty] | 72
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | digest2 | 26 | 1 | 1 | TP | ug P/L | 0.0935 | 1.66 | 39.9 | TN | ug N/L | 0.229 | 3.19 | 387 |  | 1 | 1 |  | 39.9 | 39.9 | 39.9 | [6 empty] | 1 |  | 387 | 387 | 387 | [7 empty] | 73
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | digest3 | 27 | 1 | 1 | TP | ug P/L | 0.0945 | 1.66 | 39.7 | TN | ug N/L | 0.232 | 3.54 | 439 |  | 1 | 1 |  | 39.7 | 39.7 | 39.7 | [6 empty] | 1 |  | 439 | 439 | 439 | [7 empty] | 74
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | digest4 | 28 | 1 | 1 | TP | ug P/L | 0.0921 | 1.65 | 39.5 | TN | ug N/L | 0.221 | 3.14 | 379 |  | 1 | 1 |  | 39.5 | 39.5 | 39.5 | [6 empty] | 1 |  | 379 | 379 | 379 | [7 empty] | 75
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | digest5 | 29 | 1 | 1 | TP | ug P/L | 0.0946 | 1.69 | 40.6 | TN | ug N/L | 0.222 | 3.2 | 387 |  | 1 | 1 |  | 40.6 | 40.6 | 40.6 | [6 empty] | 1 |  | 387 | 387 | 387 | [7 empty] | 76
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | digest1 | 25 | 1 | 1 | TP | ug P/L | 0.0968 | 1.65 | 39.6 | TN | ug N/L | 0.204 | 2.79 | 327 |  | 1 | 1 |  | 39.6 | 39.6 | 39.6 | [6 empty] | 1 |  | 327 | 327 | 327 | [7 empty] | 77
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | digest2 | 26 | 1 | 1 | TP | ug P/L | 0.0957 | 1.67 | 40.3 | TN | ug N/L | 0.209 | 2.8 | 329 |  | 1 | 1 |  | 40.3 | 40.3 | 40.3 | [6 empty] | 1 |  | 329 | 329 | 329 | [7 empty] | 78
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | digest3 | 27 | 1 | 1 | TP | ug P/L | 0.0965 | 1.69 | 41 | TN | ug N/L | 0.21 | 2.8 | 329 |  | 1 | 1 |  | 41 | 41 | 41 | [6 empty] | 1 |  | 329 | 329 | 329 | [7 empty] | 79
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | digest4 | 28 | 1 | 1 | TP | ug P/L | 0.0957 | 1.66 | 40 | TN | ug N/L | 0.215 | 2.87 | 340 |  | 1 | 1 |  | 40 | 40 | 40 | [6 empty] | 1 |  | 340 | 340 | 340 | [7 empty] | 80
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | digest5 | 29 | 1 | 1 | TP | ug P/L | 0.0947 | 1.64 | 39.3 | TN | ug N/L | 0.214 | 2.84 | 334 |  | 1 | 1 |  | 39.3 | 39.3 | 39.3 | [6 empty] | 1 |  | 334 | 334 | 334 | [7 empty] | 81
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | digest1 | 25 | 1 | 1 | TP | ug P/L | 0.0875 | 1.52 | 39.5 | TN | ug N/L | 0.236 | 3.13 | 416 |  | 1 | 1 |  | 39.5 | 39.5 | 39.5 | [6 empty] | 1 |  | 416 | 416 | 416 | [7 empty] | 82
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | digest2 | 26 | 1 | 1 | TP | ug P/L | 0.0851 | 1.47 | 37.8 | TN | ug N/L | 0.221 | 2.93 | 383 |  | 1 | 1 |  | 37.8 | 37.8 | 37.8 | [6 empty] | 1 |  | 383 | 383 | 383 | [7 empty] | 83
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | digest3 | 27 | 1 | 1 | TP | ug P/L | 0.0873 | 1.5 | 38.9 | TN | ug N/L | 0.211 | 2.81 | 362 |  | 1 | 1 |  | 38.9 | 38.9 | 38.9 | [6 empty] | 1 |  | 362 | 362 | 362 | [7 empty] | 84
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | digest4 | 28 | 1 | 1 | TP | ug P/L | 0.0862 | 1.5 | 38.7 | TN | ug N/L | 0.21 | 2.79 | 360 |  | 1 | 1 |  | 38.7 | 38.7 | 38.7 | [6 empty] | 1 |  | 360 | 360 | 360 | [7 empty] | 85
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | digest5 | 29 | 1 | 1 | TP | ug P/L | 0.0865 | 1.47 | 38.1 | TN | ug N/L | 0.216 | 2.88 | 375 |  | 1 | 1 |  | 38.1 | 38.1 | 38.1 | [6 empty] | 1 |  | 375 | 375 | 375 | [7 empty] | 86
[20 empty] | TP | [10 empty] | TN
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | [6 empty] | OBSERVATION #
[19 empty] | Mean |  |  | 38.6859 | [7 empty] | Mean |  |  | 358.6584 | [6 empty] | Min | 1
[19 empty] | Std |  |  | 2.7785 | [7 empty] | Std |  |  | 48.0929 | [6 empty] | Max | 86
[19 empty] | CV |  |  | 7.1821 | [7 empty] | CV |  |  | 13.4091 | [6 empty] | Mean
[19 empty] | S/N |  |  | 13.9235 | [7 empty] | S/N |  |  | 7.4576 | [6 empty] | Whole series
[18 empty] | Warning Limits | Xbar + 2SD |  |  | 44.2428 | [6 empty] | Warning Limits | Xbar + 2SD |  |  | 454.8442
[19 empty] | Xbar - 2SD |  |  | 33.129 | [7 empty] | Xbar - 2SD |  |  | 262.4727
[18 empty] | Control Limits | Xbar + 3SD |  |  | 47.0213 | [6 empty] | Control Limits | Xbar + 3SD |  |  | 502.937
[19 empty] | Xbar - 3SD |  |  | 30.3505 | [7 empty] | Xbar - 3SD |  |  | 214.3798
[18 empty] | Count |  |  |  | 86 | [6 empty] | Count |  |  |  | 86
[18 empty] | 99th percentile.inc |  |  |  | 43.9172 | [6 empty] | 99th percentile.inc |  |  |  | 436.8416
[18 empty] | Maximum |  |  |  | 44.2165 | [6 empty] | Maximum |  |  |  | 439
[18 empty] | Known |  |  |  | 40 | [6 empty] | Known |  |  |  | 400
[18 empty] | MDL |  |  |  | 6.5878 | [6 empty] | MDL |  |  |  | 114.0292
[18 empty] | LOQ |  |  |  | 27.7846 | [6 empty] | LOQ |  |  |  | 480.9287
[18 empty] | ratio mean/mdl |  |  |  | 5.8724 | [6 empty] | ratio mean/mdl |  |  |  | 3.1453
[21 empty] | Known Conc | 40 | [9 empty] | Known Conc | 400
[21 empty] | Mean | 38.6859 | [9 empty] | Mean | 358.6584
[21 empty] | Std Dev | 2.7785 | [9 empty] | Std Dev | 48.0929
[21 empty] | CV or %RSD | 7.1821 | [9 empty] | CV or %RSD | 13.4091
[21 empty] | T value | 2.371 | [9 empty] | T value | 2.371
[21 empty] | MDL | 6.5878 | [9 empty] | MDL | 114.0292
[21 empty] | LOQ | 27.7846 | [9 empty] | LOQ | 480.9287
[21 empty] | Error as % | -3.2852 | [9 empty] | Error as % | -10.3354
[21 empty] | Known/MDL | 6.0718 | [9 empty] | Known/MDL | 3.5079
[21 empty] | Mean PR | 96.7148 | [9 empty] | Mean PR | 89.6646
[21 empty] | S/N | 13.9235 | [9 empty] | S/N | 7.4576

## rolling blanks
methods vary so interpretation could be a little tricky
Runs after summer 2020 use digested standards with bisulfate carrier for TP and water carrier for TN
THERE ARE KNOWN PROBLEMS ON 13JAN20 SO THEY MAY BE EXCLUDED.
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Carrier | Observation #
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 001 | 1 | 1 | 1 | Total Phosphorus | ppb | 0.0326 | 0.54 | 8.2 | Total Nitrogen | ppb | 0.781 | 11.9 | 348 |  | 1.5 | 1 |  | -4.694 | -4.694 | -4.694 | [6 empty] | 1 |  | -38.6703 | -38.6703 | -38.6703 | [6 empty] | bisulfate | 1
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 002 | 2 | 1 | 1 | Total Phosphorus | ppb | 0.0327 | 0.53 | 8.02 | Total Nitrogen | ppb | 0.784 | 11.9 | 348 |  | 1.5 | 1 |  | -5.073 | -5.073 | -5.073 | [6 empty] | 1 |  | -38.6703 | -38.6703 | -38.6703 | [6 empty] | bisulfate | 2
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 003 | 3 | 1 | 1 | Total Phosphorus | ppb | 0.0388 | 0.792 | 13 | Total Nitrogen | ppb | 0.817 | 12.5 | 365 |  | 1.5 | 1 |  | 4.6976 | 4.6976 | 4.6976 | [6 empty] | 1 |  | -15.0507 | -15.0507 | -15.0507 | [6 empty] | bisulfate | 3
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 004 | 4 | 1 | 1 | Total Phosphorus | ppb | 0.0429 | 0.694 | 11.1 | Total Nitrogen | ppb | 0.853 | 13 | 382 |  | 1.5 | 1 |  | 1.0818 | 1.0818 | 1.0818 | [6 empty] | 1 |  | 4.613 | 4.613 | 4.613 | [6 empty] | bisulfate | 4
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 005 | 5 | 1 | 1 | Total Phosphorus | ppb | 0.042 | 0.707 | 11.4 | Total Nitrogen | ppb | 0.855 | 13.1 | 385 |  | 1.5 | 1 |  | 1.5641 | 1.5641 | 1.5641 | [6 empty] | 1 |  | 8.5436 | 8.5436 | 8.5436 | [6 empty] | bisulfate | 5
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | water only blank | 1 | 1 | 1 | Total Phosphorus | ppb | 0.0294 | 0.476 | 7.92 | Total Nitrogen | ppb | 0.735 | 12.8 | 348 |  | 1.5 | 1 |  | 0.61457 | 0.61457 | 0.61457 | [6 empty] | 1 |  | 5.3012 | 5.3012 | 5.3012 | [6 empty] | bisulfate | 6
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | water only blank | 2 | 1 | 1 | Total Phosphorus | ppb | 0.0334 | 0.556 | 9.24 | Total Nitrogen | ppb | 0.75 | 13.1 | 356 |  | 1.5 | 1 |  | 2.8432 | 2.8432 | 2.8432 | [6 empty] | 1 |  | 19.42 | 19.42 | 19.42 | [6 empty] | bisulfate | 7
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | water only blank | 3 | 1 | 1 | Total Phosphorus | ppb | 0.0461 | 0.618 | 10.3 | Total Nitrogen | ppb | 0.734 | 12.8 | 349 |  | 1.5 | 1 |  | 4.5682 | 4.5682 | 4.5682 | [6 empty] | 1 |  | 5.3012 | 5.3012 | 5.3012 | [6 empty] | bisulfate | 8
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | water only blank | 4 | 1 | 1 | Total Phosphorus | ppb | 0.0298 | 0.465 | 7.74 | Total Nitrogen | ppb | 0.725 | 12.6 | 342 |  | 1.5 | 1 |  | 0.30788 | 0.30788 | 0.30788 | [6 empty] | 1 |  | -4.1262 | -4.1262 | -4.1262 | [6 empty] | bisulfate | 9
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | water only blank | 5 | 1 | 1 | Total Phosphorus | ppb | 0.0326 | 0.519 | 8.63 | Total Nitrogen | ppb | 0.737 | 12.8 | 348 |  | 1.5 | 1 |  | 1.8129 | 1.8129 | 1.8129 | [6 empty] | 1 |  | 5.3012 | 5.3012 | 5.3012 | [6 empty] | bisulfate | 10
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 018 | 18 | 1 | 1 | TP | ug P/L | 0.0219 | 0.482 | 0.588 | TN | ug N/L | 0.351 | 5.92 | -1.69 |  | 1 | 1 |  | 0.588 | 0.588 | 0.588 | [6 empty] | 1 |  | -20.74 | -20.74 | -3.6153 | [6 empty] | bisulfate for TP but water for TN | 11
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 019 | 19 | 1 | 1 | TP | ug P/L | 0.0221 | 0.492 | 0.857 | TN | ug N/L | 0.348 | 5.92 | -1.73 |  | 1 | 1 |  | 0.857 | 0.857 | 0.857 | [6 empty] | 1 |  | -20.74 | -20.74 | -3.0787 | [6 empty] | bisulfate for TP but water for TN | 12
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 020 | 20 | 1 | 1 | TP | ug P/L | 0.0441 | 0.433 | -0.783 | TN | ug N/L | 0.346 | 5.85 | -2.3 |  | 1 | 1 |  | -0.783 | -0.783 | -0.783 | [6 empty] | 1 |  | -26.2933 | -26.2933 | -8.726 | [6 empty] | bisulfate for TP but water for TN | 13
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 021 | 21 | 1 | 1 | TP | ug P/L | 0.0223 | 0.499 | 1.05 | TN | ug N/L | 0.406 | 6.83 | 6.16 |  | 1 | 1 |  | 1.05 | 1.05 | 1.05 | [6 empty] | 1 |  | 52.1473 | 52.1473 | [7 empty] | bisulfate for TP but water for TN | 14
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 022 | 22 | 1 | 1 | TP | ug P/L | 0.02 | 0.444 | -0.471 | TN | ug N/L | 0.343 | 5.8 | -2.72 |  | 1 | 1 |  | -0.471 | -0.471 | -0.471 | [6 empty] | 1 |  | -30.2553 | -30.2553 | -12.0868 | [6 empty] | bisulfate for TP but water for TN | 15
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 020 | 20 | 1 | 1 | TP | ug P/L | 0.0226 | 0.399 | -0.27 | TN | ug N/L | 0.103 | 1.86 | -20.5 |  | 1 | 1 |  | -0.27 | -0.27 | -0.27 | [6 empty] | 1 |  | -20.5 | -20.5 | -20.5 | [6 empty] | bisulfate for TP but water for TN | 16
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 021 | 21 | 1 | 1 | TP | ug P/L | 0.0238 | 0.447 | 1.17 | TN | ug N/L | 0.131 | 2.37 | 16.4 |  | 1 | 1 |  | 1.17 | 1.17 | 1.17 | [6 empty] | 1 |  | 16.4 | 16.4 | 16.4 | [6 empty] | bisulfate for TP but water for TN | 17
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 022 | 22 | 1 | 1 | TP | ug P/L | 0.0201 | 0.369 | -1.18 | TN | ug N/L | 0.109 | 1.97 | -12.5 |  | 1 | 1 |  | -1.18 | -1.18 | -1.18 | [6 empty] | 1 |  | -12.5 | -12.5 | -12.5 | [6 empty] | bisulfate for TP but water for TN | 18
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 023 | 23 | 1 | 1 | TP | ug P/L | 0.0222 | 0.417 | 0.273 | TN | ug N/L | 0.112 | 2.01 | -9.17 |  | 1 | 1 |  | 0.273 | 0.273 | 0.273 | [6 empty] | 1 |  | -9.17 | -9.17 | -9.17 | [6 empty] | bisulfate for TP but water for TN | 19
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 024 | 24 | 1 | 1 | TP | ug P/L | 0.0278 | 0.514 | 3.19 | TN | ug N/L | 0.126 | 2.23 | 6.67 |  | 1 | 1 |  | 3.19 | 3.19 | 3.19 | [6 empty] | 1 |  | 6.67 | 6.67 | 6.67 | [6 empty] | bisulfate for TP but water for TN | 20
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample20 | 20 | 1 | 1 | TP | ug P/L | 0.0521 | 1.18 |  | TN | ug N/L | 0.0404 | 0.784 | -25.6 |  | 1 | 3 | carry over | 22.2139 | 22.2139 | [7 empty] | 1 |  | -21.4483 | -21.4483 | -21.4483 | [6 empty] | bisulfate for TP but water for TN | 21
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample21 | 21 | 1 | 1 | TP | ug P/L | 0.0337 | 0.689 | 2.69 | TN | ug N/L | 0.0418 | 0.827 | -21.7 |  | 1 | 2 | reworked calibration | 7.6698 | 7.6698 | 7.6698 | [6 empty] | 1 |  | -17.5436 | -17.5436 | -17.5436 | [6 empty] | bisulfate for TP but water for TN | 22
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample22 | 22 | 1 | 1 | TP | ug P/L | 0.0326 | 0.653 | 1.75 | TN | ug N/L | 0.0457 | 1.03 | -3.08 |  | 1 | 2 | reworked calibration | 6.6104 | 6.6104 | 6.6104 | [6 empty] | 1 |  | 0.92298 | 0.92298 | 0.92298 | [6 empty] | bisulfate for TP but water for TN | 23
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample23 | 23 | 1 | 1 | TP | ug P/L | 0.0329 | 0.661 | 1.97 | TN | ug N/L | 0.0508 | 0.897 | -15.5 |  | 1 | 2 | reworked calibration | 6.8457 | 6.8457 | 6.8457 | [6 empty] | 1 |  | -11.182 | -11.182 | -11.182 | [6 empty] | bisulfate for TP but water for TN | 24
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample24 | 24 | 1 | 1 | TP | ug P/L | 0.0297 | 0.554 | -0.821 | TN | ug N/L | 0.0427 | 0.871 | -17.8 |  | 1 | 2 | reworked calibration | 3.7022 | 3.7022 | 3.7022 | [6 empty] | 1 |  | -13.5456 | -13.5456 | -13.5456 | [6 empty] | bisulfate for TP but water for TN | 25
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | water blank | 20 | 1 | 1 | TP | ug P/L | 0.0107 | 0.324 | -0.271 | TN | ug N/L | 0.0496 | 0.907 | -6.32 |  | 1 | 1 |  | -0.271 | -0.271 | -0.271 | [6 empty] | 1 |  | -6.32 | -6.32 | -6.32 | [6 empty] | bisulfate for TP but water for TN | 26
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | water blank | 21 | 1 | 1 | TP | ug P/L | 0.00974 | 0.321 | -0.334 | TN | ug N/L | 0.0459 | 0.839 | -10.6 |  | 1 | 1 |  | -0.334 | -0.334 | -0.334 | [6 empty] | 1 |  | -10.6 | -10.6 | -10.6 | [6 empty] | bisulfate for TP but water for TN | 27
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | water blank | 22 | 1 | 1 | TP | ug P/L | 0.011 | 0.333 | -0.0138 | TN | ug N/L | 0.0574 | 1.01 | -0.131 |  | 1 | 1 |  | -0.0138 | -0.0138 | -0.0138 | [6 empty] | 1 |  | -0.131 | -0.131 | -0.131 | [6 empty] | bisulfate for TP but water for TN | 28
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | water blank | 23 | 1 | 1 | TP | ug P/L | 0.011 | 0.33 | -0.0869 | TN | ug N/L | 0.0561 | 0.83 | -11.2 |  | 1 | 1 |  | -0.0869 | -0.0869 | -0.0869 | [6 empty] | 1 |  | -11.2 | -11.2 | -11.2 | [6 empty] | bisulfate for TP but water for TN | 29
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | water blank | 24 | 1 | 1 | TP | ug P/L | 0.0116 | 0.323 | -0.296 | TN | ug N/L | 0.064 | 1.11 | 6.52 |  | 1 | 1 |  | -0.296 | -0.296 | -0.296 | [6 empty] | 1 |  | 6.52 | 6.52 | 6.52 | [6 empty] | bisulfate for TP but water for TN | 30
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 020 | 20 | 1 | 1 | TP | ug P/L | 0.00961 | 0.267 | 0.354 | TN | ug N/L | 0.0596 | 0.968 | -5.93 |  | 1 | 1 |  | 0.354 | 0.354 | 0.354 | [6 empty] | 1 |  | -5.93 | -5.93 | -5.93 | [6 empty] | bisulfate for TP but water for TN | 31
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 021 | 21 | 1 | 1 | TP | ug P/L | 0.00787 | 0.223 | -0.867 | TN | ug N/L | 0.0645 | 1.06 | -0.157 |  | 1 | 1 |  | -0.867 | -0.867 | -0.867 | [6 empty] | 1 |  | -0.157 | -0.157 | -0.157 | [6 empty] | bisulfate for TP but water for TN | 32
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 022 | 22 | 1 | 1 | TP | ug P/L | 0.0104 | 0.241 | -0.38 | TN | ug N/L | 0.0677 | 1.08 | 0.862 |  | 1 | 1 |  | -0.38 | -0.38 | -0.38 | [6 empty] | 1 |  | 0.862 | 0.862 | 0.862 | [6 empty] | bisulfate for TP but water for TN | 33
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 023 | 23 | 1 | 1 | TP | ug P/L | 0.0118 | 0.264 | 0.287 | TN | ug N/L | 0.106 | 1.53 | 29.4 |  | 1 | 1 |  | 0.287 | 0.287 | 0.287 | [6 empty] | 1 |  | 29.4 | 29.4 | 29.4 | [6 empty] | bisulfate for TP but water for TN | 34
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 024 | 24 | 1 | 1 | TP | ug P/L | 0.011 | 0.238 | -0.452 | TN | ug N/L | 0.0782 | 1.19 | 8.35 |  | 1 | 1 |  | -0.452 | -0.452 | -0.452 | [6 empty] | 1 |  | 8.35 | 8.35 | 8.35 | [6 empty] | bisulfate for TP but water for TN | 35
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | water blank 1 | 20 | 1 | 1 | TP | ug P/L | 0.00803 | 0.226 | -2 | TN | ug N/L | 0.0517 | 0.852 | -29.7 |  | 1 | 1 |  | -2 | -2 | -2 | [6 empty] | 1 |  | -29.7 | -29.7 | -29.7 | [6 empty] | bisulfate for TP but water for TN | 36
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | water blank 2 | 21 | 1 | 1 | TP | ug P/L | 0.00822 | 0.228 | -1.95 | TN | ug N/L | 0.0498 | 0.832 | -31 |  | 1 | 1 |  | -1.95 | -1.95 | -1.95 | [6 empty] | 1 |  | -31 | -31 | -31 | [6 empty] | bisulfate for TP but water for TN | 37
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | water blank 3 | 22 | 1 | 1 | TP | ug P/L | 0.00828 | 0.216 | -2.24 | TN | ug N/L | 0.0521 | 0.846 | -30.1 |  | 1 | 1 |  | -2.24 | -2.24 | -2.24 | [6 empty] | 1 |  | -30.1 | -30.1 | -30.1 | [6 empty] | bisulfate for TP but water for TN | 38
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | water blank 4 | 23 | 1 | 1 | TP | ug P/L | 0.00855 | 0.207 | -2.46 | TN | ug N/L | 0.0438 | 0.687 | -40.7 |  | 1 | 1 |  | -2.46 | -2.46 | -2.46 | [6 empty] | 1 |  | -40.7 | -40.7 | -40.7 | [6 empty] | bisulfate for TP but water for TN | 39
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | water blank 5 | 24 | 1 | 1 | TP | ug P/L | 0.00832 | 0.235 | -1.78 | TN | ug N/L | 0.0502 | 0.802 | -33 |  | 1 | 1 |  | -1.78 | -1.78 | -1.78 | [6 empty] | 1 |  | -33 | -33 | -33 | [6 empty] | bisulfate for TP but water for TN | 40
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank1 | 20 | 1 | 1 | TP | ug P/L | 0.0136 | 0.308 | -0.453 | TN | ug N/L | 0.0585 | 0.989 | -5.79 |  | 1 | 1 |  | -0.453 | -0.453 | -0.453 | [6 empty] | 1 |  | -5.79 | -5.79 | -5.79 | [6 empty] | bisulfate for TP but water for TN | 41
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank2 | 21 | 1 | 1 | TP | ug P/L | 0.0125 | 0.345 | 0.576 | TN | ug N/L | 0.0608 | 1.01 | -4.53 |  | 1 | 1 |  | 0.576 | 0.576 | 0.576 | [6 empty] | 1 |  | -4.53 | -4.53 | -4.53 | [6 empty] | bisulfate for TP but water for TN | 42
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank3 | 22 | 1 | 1 | TP | ug P/L | 0.0134 | 0.348 | 0.675 | TN | ug N/L | 0.0589 | 0.973 | -6.93 |  | 1 | 1 |  | 0.675 | 0.675 | 0.675 | [6 empty] | 1 |  | -6.93 | -6.93 | -6.93 | [6 empty] | bisulfate for TP but water for TN | 43
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank4 | 23 | 1 | 1 | TP | ug P/L | 0.0126 | 0.337 | 0.353 | TN | ug N/L | 0.0623 | 0.972 | -7 |  | 1 | 1 |  | 0.353 | 0.353 | 0.353 | [6 empty] | 1 |  | -7 | -7 | -7 | [6 empty] | bisulfate for TP but water for TN | 44
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank5 | 24 | 1 | 1 | TP | ug P/L | 0.0115 | 0.301 | -0.66 | TN | ug N/L | 0.0674 | 1.09 | 1.59 |  | 1 | 1 |  | -0.66 | -0.66 | -0.66 | [6 empty] | 1 |  | 1.59 | 1.59 | 1.59 | [6 empty] | bisulfate for TP but water for TN | 45
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | waterblank1 | 20 | 1 | 1 | TP | ug P/L | 0.0218 | 0.366 | 0.0149 | TN | ug N/L | 0.063 | 0.948 | -6.52 |  | 1 | 1 |  | 0.0149 | 0.0149 | 0.0149 | [6 empty] | 1 |  | -6.52 | -6.52 | -6.52 | [6 empty] | bisulfate for TP but water for TN | 46
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | waterblank2 | 21 | 1 | 1 | TP | ug P/L | 0.0213 | 0.366 | 0.0252 | TN | ug N/L | 0.07 | 0.967 | -5.22 |  | 1 | 1 |  | 0.0252 | 0.0252 | 0.0252 | [6 empty] | 1 |  | -5.22 | -5.22 | -5.22 | [6 empty] | bisulfate for TP but water for TN | 47
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | waterblank3 | 22 | 1 | 1 | TP | ug P/L | 0.022 | 0.363 | -0.0623 | TN | ug N/L | 0.0629 | 0.961 | -5.61 |  | 1 | 1 |  | -0.0623 | -0.0623 | -0.0623 | [6 empty] | 1 |  | -5.61 | -5.61 | -5.61 | [6 empty] | bisulfate for TP but water for TN | 48
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | waterblank4 | 23 | 1 | 1 | TP | ug P/L | 0.0218 | 0.364 | -0.047 | TN | ug N/L | 0.071 | 0.975 | -4.68 |  | 1 | 1 |  | -0.047 | -0.047 | -0.047 | [6 empty] | 1 |  | -4.68 | -4.68 | -4.68 | [6 empty] | bisulfate for TP but water for TN | 49
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | waterblank5 | 24 | 1 | 1 | TP | ug P/L | 0.0209 | 0.365 | -7.23e-05 | TN | ug N/L | 0.0644 | 0.938 | -7.2 |  | 1 | 1 |  | -7.23e-05 | -7.23e-05 | -7.23e-05 | [6 empty] | 1 |  | -7.2 | -7.2 | -7.2 | [6 empty] | bisulfate for TP but water for TN | 50
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 20 | 1 | 1 | TP | ug P/L | 0.0482 | 0.56 | -44.6 | TN | ug N/L | 0.084 | 1.2 | -3.48 |  | 1 | 2 | reworked calibration | -10.4588 | -10.4588 | -10.4588 | [6 empty] | 1 |  | -3.48 | -3.48 | -3.48 | [6 empty] | reversed carriers | 51
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 21 | 1 | 1 | TP | ug P/L | 0.0476 | 0.617 | -44.6 | TN | ug N/L | 0.0936 | 1.22 | -2.22 |  | 1 | 2 | reworked calibration | -11.387 | -11.387 | -11.387 | [6 empty] | 1 |  | -2.22 | -2.22 | -2.22 | [6 empty] | reversed carriers | 52
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 22 | 1 | 1 | TP | ug P/L | 0.0538 | 0.764 | -44.2 | TN | ug N/L | 0.0896 | 1.28 | 0.652 |  | 1 | 2 | reworked calibration | -2.014 | -2.014 | -2.014 | [6 empty] | 1 |  | 0.652 | 0.652 | 0.652 | [6 empty] | reversed carriers | 53
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 23 | 1 | 1 | TP | ug P/L | 0.0539 | 0.721 | -44.2 | TN | ug N/L | 0.095 | 1.3 | 1.41 |  | 1 | 2 | reworked calibration | -1.8668 | -1.8668 | -1.8668 | [6 empty] | 1 |  | 1.41 | 1.41 | 1.41 | [6 empty] | reversed carriers | 54
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | water | 24 | 1 | 1 | TP | ug P/L | 0.0537 | 0.741 | -44.3 | TN | ug N/L | 0.0925 | 1.31 | 1.89 |  | 1 | 2 | reworked calibration | -2.1613 | -2.1613 | -2.1613 | [6 empty] | 1 |  | 1.89 | 1.89 | 1.89 | [6 empty] | reversed carriers | 55
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | waterblank | 20 | 1 | 1 | TP | ug P/L | 0.0912 | 1.63 | 1.27 | TN | ug N/L | 0.12 | 2.25 | 21.3 |  | 1 | 1 |  | 1.27 | 1.27 | 1.27 | [6 empty] | 1 |  | 21.3 | 21.3 | 21.3 | [6 empty] | bisulfate for TP but water for TN | 56
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | waterbnal | 21 | 1 | 1 | TP | ug P/L | 0.0902 | 1.57 | -0.434 | TN | ug N/L | 0.118 | 1.76 | -1.84 |  | 1 | 1 |  | -0.434 | -0.434 | -0.434 | [6 empty] | 1 |  | -1.84 | -1.84 | -1.84 | [6 empty] | bisulfate for TP but water for TN | 57
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | waterbnalk | 22 | 1 | 1 | TP | ug P/L | 0.0908 | 1.55 | -1.37 | TN | ug N/L | 0.117 | 1.79 | -0.506 |  | 1 | 1 |  | -1.37 | -1.37 | -1.37 | [6 empty] | 1 |  | -0.506 | -0.506 | -0.506 | [6 empty] | bisulfate for TP but water for TN | 58
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | waterblnak | 23 | 1 | 1 | TP | ug P/L | 0.0838 | 1.37 | -7.17 | TN | ug N/L | 0.136 | 2.5 | 33.1 |  | 1 | 1 |  | -7.17 | -7.17 | -7.17 | [6 empty] | 1 |  | 33.1 | 33.1 | 33.1 | [6 empty] | bisulfate for TP but water for TN | 59
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | waterblank | 24 | 1 | 1 | TP | ug P/L | 0.0904 | 1.56 | -0.962 | TN | ug N/L | 0.135 | 1.88 | 3.6 |  | 1 | 1 |  | -0.962 | -0.962 | -0.962 | [6 empty] | 1 |  | 3.6 | 3.6 | 3.6 | [6 empty] | bisulfate for TP but water for TN | 60
2023-02-22 | OM_2-22-2023_04-26-39PM BRN edit.omn | waterblank | 20 | 1 | 1 | TP | ug P/L | 0.0876 | 1.56 | -0.944 | TN | ug N/L | 0.123 | 1.74 | -2.84 |  | 1 | 1 |  | -0.944 | -0.944 | -0.944 | [6 empty] | 1 |  | -2.84 | -2.84 | -2.84 | [6 empty] | bisulfate for TP but water for TN | 61
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | waterblank1 | 20 | 1 | 1 | TP | ug P/L | 0.017 | 0.351 | -1.11 | TN | ug N/L | 0.0245 | 0.473 | -5.87 |  | 1 | 1 |  | -1.11 | -1.11 | -1.11 | [6 empty] | 1 |  | -5.87 | -5.87 | -5.87 | [6 empty] | bisulfate for TP but water for TN | 62
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | waterblan2 | 21 | 1 | 1 | TP | ug P/L | 0.0154 | 0.33 | -1.77 | TN | ug N/L | 0.0269 | 0.473 | -5.77 |  | 1 | 1 |  | -1.77 | -1.77 | -1.77 | [6 empty] | 1 |  | -5.77 | -5.77 | -5.77 | [6 empty] | bisulfate for TP but water for TN | 63
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | waterblank3 | 22 | 1 | 1 | TP | ug P/L | 0.0176 | 0.35 | -1.14 | TN | ug N/L | 0.0261 | 0.466 | -6.83 |  | 1 | 1 |  | -1.14 | -1.14 | -1.14 | [6 empty] | 1 |  | -6.83 | -6.83 | -6.83 | [6 empty] | bisulfate for TP but water for TN | 64
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | waterblank4 | 23 | 1 | 1 | TP | ug P/L | 0.0164 | 0.333 | -1.69 | TN | ug N/L | 0.0282 | 0.407 | -15 |  | 1 | 1 |  | -1.69 | -1.69 | -1.69 | [6 empty] | 1 |  | -15 | -15 | -15 | [6 empty] | bisulfate for TP but water for TN | 65
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | waterblank5 | 24 | 1 | 1 | TP | ug P/L | 0.0165 | 0.379 | -0.225 | TN | ug N/L | 0.0283 | 0.479 | -5.05 |  | 1 | 1 |  | -0.225 | -0.225 | -0.225 | [6 empty] | 1 |  | -5.05 | -5.05 | -5.05 | [6 empty] | bisulfate for TP but water for TN | 66
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | waterblank1 | 20 | 1 | 1 | TP | ug P/L | 0.0178 | 0.33 | -0.907 | TN | ug N/L | 0.0231 | 0.356 | -18.7 |  | 1 | 1 |  | -0.907 | -0.907 | -0.907 | [6 empty] | 1 |  | -18.7 | -18.7 | -18.7 | [6 empty] | bisulfate for TP but water for TN | 67
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | waterblank2 | 21 | 1 | 1 | TP | ug P/L | 0.0214 | 0.357 | -0.0674 | TN | ug N/L | 0.0241 | 0.364 | -17.6 |  | 1 | 1 |  | -0.0674 | -0.0674 | -0.0674 | [6 empty] | 1 |  | -17.6 | -17.6 | -17.6 | [6 empty] | bisulfate for TP but water for TN | 68
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | waterblank3 | 22 | 1 | 1 | TP | ug P/L | 0.0189 | 0.362 | 0.1 | TN | ug N/L | 0.0294 | 0.447 | -6.58 |  | 1 | 1 |  | 0.1 | 0.1 | 0.1 | [6 empty] | 1 |  | -6.58 | -6.58 | -6.58 | [6 empty] | bisulfate for TP but water for TN | 69
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | waterblank4 | 23 | 1 | 1 | TP | ug P/L | 0.0178 | 0.32 | -1.21 | TN | ug N/L | 0.0227 | 0.361 | -17.9 |  | 1 | 1 |  | -1.21 | -1.21 | -1.21 | [6 empty] | 1 |  | -17.9 | -17.9 | -17.9 | [6 empty] | bisulfate for TP but water for TN | 70
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | waterblank5 | 24 | 1 | 1 | TP | ug P/L | 0.0201 | 0.374 | 0.457 | TN | ug N/L | 0.0249 | 0.407 | -11.9 |  | 1 | 1 |  | 0.457 | 0.457 | 0.457 | [6 empty] | 1 |  | -11.9 | -11.9 | -11.9 | [6 empty] | bisulfate for TP but water for TN | 71
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | waterblank1 | 20 | 1 | 1 | TP | ug P/L | 0.0112 | 0.222 | -0.0471 | TN | ug N/L | 0.0204 | 0.292 | -27.4 |  | 1 | 1 |  | -0.0471 | -0.0471 | -0.0471 | [6 empty] | 1 |  | -27.4 | -27.4 | -27.4 | [6 empty] | bisulfate for TP but water for TN | 72
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | waterblank2 | 21 | 1 | 1 | TP | ug P/L | 0.0107 | 0.212 | -0.337 | TN | ug N/L | 0.022 | 0.324 | -22.8 |  | 1 | 1 |  | -0.337 | -0.337 | -0.337 | [6 empty] | 1 |  | -22.8 | -22.8 | -22.8 | [6 empty] | bisulfate for TP but water for TN | 73
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | waterblank3 | 22 | 1 | 1 | TP | ug P/L | 0.0112 | 0.179 | -1.34 | TN | ug N/L | 0.0306 | 0.346 | -19.6 |  | 1 | 1 |  | -1.34 | -1.34 | -1.34 | [6 empty] | 1 |  | -19.6 | -19.6 | -19.6 | [6 empty] | bisulfate for TP but water for TN | 74
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | waterblank4 | 23 | 1 | 1 | TP | ug P/L | 0.0109 | 0.18 | -1.31 | TN | ug N/L | 0.0212 | 0.314 | -24.3 |  | 1 | 1 |  | -1.31 | -1.31 | -1.31 | [6 empty] | 1 |  | -24.3 | -24.3 | -24.3 | [6 empty] | bisulfate for TP but water for TN | 75
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | waterblank5 | 24 | 1 | 1 | TP | ug P/L | 0.00907 | 0.185 | -1.16 | TN | ug N/L | 0.0215 | 0.327 | -22.4 |  | 1 | 1 |  | -1.16 | -1.16 | -1.16 | [6 empty] | 1 |  | -22.4 | -22.4 | -22.4 | [6 empty] | bisulfate for TP but water for TN | 76
[20 empty] | TP | [10 empty] | TN
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Carrier | Observation #
[19 empty] | Mean |  |  | -0.32463 | [7 empty] | Mean |  |  | -7.1956 | [6 empty] | Min | 1 | Mean
[19 empty] | Std |  |  | 2.9105 | [7 empty] | Std |  |  | 14.3272 | [6 empty] | Max | 76
[19 empty] | CV |  |  | -896.5695 | [7 empty] | CV |  |  | -199.1105
[19 empty] | S/N |  |  | -0.11154 | [7 empty] | S/N |  |  | -0.50223
[18 empty] | Warning Limits | Xbar + 2SD |  |  | 5.4964 | [6 empty] | Warning Limits | Xbar + 2SD |  |  | 21.4588
[19 empty] | Xbar - 2SD |  |  | -6.1457 | [7 empty] | Xbar - 2SD |  |  | -35.85
[18 empty] | Control Limits | Xbar + 3SD |  |  | 8.407 | [6 empty] | Control Limits | Xbar + 3SD |  |  | 35.786
[19 empty] | Xbar - 3SD |  |  | -9.0562 | [7 empty] | Xbar - 3SD |  |  | -50.1773
[18 empty] | Count |  |  |  | 75 | [6 empty] | Count |  |  |  | 76
[15 empty] | MDLb if N >100 |  |  | 99th percentile.inc |  |  |  | 7.06 |  |  |  | MDLb if N >100 |  |  | 99th percentile.inc |  |  |  | 11.3058
[15 empty] | MDLb if N <100 |  |  | Maximum |  |  |  | 7.6698 |  |  |  | MDLb if N <100 |  |  | Maximum |  |  |  | 22.2139
[15 empty] | MDLb if all are >0 |  |  | MDL |  |  |  | 6.9207 | [6 empty] | MDL |  |  |  | 34.0572
[18 empty] | LOQ |  |  |  | 29.1053 | [6 empty] | LOQ |  |  |  | 143.2721

## rolling spike PR
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 030 | 30 | 1 | 1 | Total Phosphorus | ppb | 0.0949 | 2.08 | 37.1 | Total Nitrogen | ppb | 2.39 | 36.8 | 1090 |  | 1.5 | 1 |  | 47.9029 | 48.3829 | 47.9029 |  |  |  |  | 116.6233 | PASS | 1 |  | 920.2523 | 936.1023 | 920.2523 |  |  |  |  | 90.904 | PASS |  | 1
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 042 | 42 | 1 | 1 | Total Phosphorus | ppb | 0.0698 | 1.6 | 28.2 | Total Nitrogen | ppb | 1.45 | 22.8 | 671 |  | 1.5 | 1 |  | 32.7393 | 33.2193 | 32.7393 |  |  |  |  | 124.4217 | FAIL | 1 |  | 386.469 | 402.319 | 386.469 |  |  |  |  | 84.575 | PASS |  | 2
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 054 | 54 | 1 | 1 | Total Phosphorus | ppb | 0.0798 | 1.24 | 21.4 | Total Nitrogen | ppb | 1.65 | 26 | 766 |  | 1.5 | 1 |  | 20.6354 | 21.1154 | 20.6354 |  |  |  |  | -6.1798 | FAIL | 1 |  | 509.6928 | 525.5428 | 509.6928 |  |  |  |  | 6.6329 | FAIL |  | 3
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 066 | 66 | 1 | 1 | Total Phosphorus | ppb | 0.117 | 2.5 | 45 | Total Nitrogen | ppb | 2.07 | 32.5 | 962 |  | 1.5 | 1 |  | 60.257 | 60.737 | 60.257 |  |  |  |  | 131.9564 | FAIL | 1 |  | 757.7723 | 773.6223 | 757.7723 |  |  |  |  | 94.5322 | PASS |  | 4
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 090 | 90 | 1 | 1 | Total Phosphorus | ppb | 0.112 | 2.27 | 40.7 | Total Nitrogen | ppb | 2.08 | 32.7 | 967 |  | 1.5 | 1 |  | 53.5973 | 54.0773 | 53.5973 |  |  |  |  | 103.3194 | PASS | 1 |  | 765.3584 | 781.2084 | 765.3584 |  |  |  |  | 85.3832 | PASS |  | 5
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 102 | 102 | 1 | 1 | Total Phosphorus | ppb | 0.0902 | 2.02 | 36 | Total Nitrogen | ppb | 1.65 | 22.4 | 660 |  | 1.5 | 1 |  | 46.0684 | 46.5484 | 46.0684 |  |  |  |  | 137.16 | FAIL | 3 | chk std failed | 371.0153 | 386.8653 | 371.0153 | [7 empty] | 6
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 114 | 114 | 1 | 1 | Total Phosphorus | ppb | 0.101 | 2.2 | 39.4 | Total Nitrogen | ppb | 1.5 | 17.8 | 523 |  | 1.5 | 1 |  | 51.5197 | 51.9997 | 51.5197 |  |  |  |  | 64.692 | FAIL | 3 | chk std failed | 192.4891 | 208.3391 | 192.4891 | [7 empty] | 7
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 126 | 126 | 1 | 1 | Total Phosphorus | ppb | 0.116 | 2.44 | 43.8 | Total Nitrogen | ppb | 1.7 | 18.2 | 535 |  | 1.5 | 1 |  | 58.5444 | 59.0244 | 58.5444 |  |  |  |  | 111.6986 | PASS | 3 | chk std failed | 208.0722 | 223.9222 | 208.0722 | [7 empty] | 8
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 138 | 138 | 1 | 1 | Total Phosphorus | ppb | 0.115 | 2.43 | 43.7 | Total Nitrogen | ppb | 1.5 | 14.9 | 436 |  | 1.5 | 1 |  | 58.2572 | 58.7372 | 58.2572 |  |  |  |  | 106.5882 | PASS | 3 | chk std failed | 79.1745 | 95.0245 | 79.1745 | [7 empty] | 9
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | Sample 102 | 102 | 1 | 1 | Total Phosphorus | ppb | 0.0832 | 1.74 | 36.5 | Total Nitrogen | ppb | 1.48 | 25.6 | 668 |  | 1.5 | 1 |  | 37.2771 | 37.7571 | 37.2771 |  |  |  |  | 136.3946 | FAIL | 1 |  | 494.3292 | 510.1792 | 494.3292 |  |  |  |  | 93.005 | PASS |  | 10
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | Sample 114 | 114 | 1 | 1 | Total Phosphorus | ppb | 0.0925 | 1.91 | 39.4 | Total Nitrogen | ppb | 1.5 | 25.9 | 674 |  | 1.5 | 1 |  | 42.6599 | 43.1399 | 42.6599 |  |  |  |  | 70.5362 | FAIL | 1 |  | 505.8529 | 521.7029 | 505.8529 |  |  |  |  | 69.822 | FAIL |  | 11
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | Sample 126 | 126 | 1 | 1 | Total Phosphorus | ppb | 0.11 | 2.2 | 44.5 | Total Nitrogen | ppb | 1.84 | 31.9 | 829 |  | 1.5 | 1 |  | 51.5197 | 51.9997 | 51.5197 |  |  |  |  | 82.8974 | PASS | 1 |  | 734.9972 | 750.8472 | 734.9972 |  |  |  |  | 80.5951 | PASS |  | 12
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | Sample 138 | 138 | 1 | 1 | Total Phosphorus | ppb | 0.109 | 2.19 | 44.3 | Total Nitrogen | ppb | 1.86 | 32.2 | 836 |  | 1.5 | 1 |  | 51.2209 | 51.7009 | 51.2209 |  |  |  |  | 113.7639 | PASS | 1 |  | 746.3879 | 762.2379 | 746.3879 |  |  |  |  | 83.7327 | PASS |  | 13
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | F 04oct19 50 t r1 SPK | 30 | 1 | 1 | Total Phosphorus | ppb | 0.125 | 2.6 | 43.8 | Total Nitrogen | ppb | 2.29 | 40.1 | 1080 |  | 1.5 | 1 |  | 58.6996 | 56.6696 | 58.6996 |  |  |  |  | 117.4128 | PASS | 1 |  | 1180.8045 | 1174.5645 | 1180.8045 |  |  |  |  | 105.9369 | PASS |  | 14
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | B 22aug19 100 t r2 SPK | 42 | 1 | 1 | Total Phosphorus | ppb | 0.222 | 4.28 | 73.9 | Total Nitrogen | ppb | 1.29 | 22.6 | 613 |  | 1.5 | 1 |  | 103.0449 | 101.0149 | 103.0449 |  |  |  |  | 108.8578 | PASS | 1 |  | 452.7102 | 446.4702 | 452.7102 |  |  |  |  | 94.4114 | PASS |  | 15
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | F 19aug19 t 3.8 SPK | 54 | 1 | 1 | Total Phosphorus | ppb | 0.154 | 2.64 | 44.6 | Total Nitrogen | ppb | 1.36 | 24 | 650 |  | 1.5 | 1 |  | 59.7718 | 57.7418 | 59.7718 |  |  |  |  | 115.2531 | PASS | 1 |  | 514.3004 | 508.0604 | 514.3004 |  |  |  |  | 99.6344 | PASS |  | 16
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | F 19aug19 t 9 SPK | 66 | 1 | 1 | Total Phosphorus | ppb | 0.124 | 2.52 | 42.6 | Total Nitrogen | ppb | 1.81 | 31.9 | 863 |  | 1.5 | 1 |  | 56.5527 | 54.5227 | 56.5527 |  |  |  |  | 100.9628 | PASS | 1 |  | 850.9497 | 844.7097 | 850.9497 |  |  |  |  | 94.5946 | PASS |  | 17
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | B 04oct19 50 t r1 SPK | 78 | 1 | 1 | Total Phosphorus | ppb | 0.0971 | 2.02 | 33.8 | Total Nitrogen | ppb | 1.32 | 23.1 | 625 |  | 1.5 | 1 |  | 43.0622 | 41.0322 | 43.0622 |  |  |  |  | 92.7858 | PASS | 1 |  | 474.7734 | 468.5334 | 474.7734 |  |  |  |  | 96.1574 | PASS |  | 18
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | F 23oct19 t isco r2 SPK | 90 | 1 | 1 | Total Phosphorus | ppb | 0.3 | 5.69 | 100 | Total Nitrogen | ppb | 1.71 | 29.9 | 808 |  | 1.5 | 1 |  | 139.174 | 137.144 | 139.174 |  |  |  |  | 97.4787 | PASS | 1 |  | 767.4719 | 761.2319 | 767.4719 |  |  |  |  | 89.7495 | PASS |  | 19
2020-10-22 | OM_10-22-2020_10-45-05AM.OMN | Sample 044 | 44 | 1 | 1 | TP | ug P/L | 0.1 | 2.04 | 55.3 | TN | ug N/L | 0.736 | 12.4 | 341 |  | 1 | 1 |  | 55.3 | 55.3 | 55.3 |  |  |  |  | 111.93 | PASS | 1 |  | 341 | 341 | 341 |  |  |  |  | 171.65 | FAIL |  | 20
2020-10-22 | OM_10-22-2020_10-45-05AM.OMN | Sample 056 | 56 | 1 | 1 | TP | ug P/L | 0.0721 | 1.55 | 40.1 | TN | ug N/L | 0.502 | 8.56 | 72.4 |  | 1 | 1 |  | 40.1 | 40.1 | 40.1 |  |  |  |  | 114.81 | PASS | 1 |  | 72.4 | 72.4 | 72.4 |  |  |  |  | 105.684 | PASS |  | 21
2020-10-22 | OM_10-22-2020_11-52-24AM.OMN | Sample 092 | 92 | 1 | 1 | TP | ug P/L | 0.219 | 4.21 | 115 | TN | ug N/L | 1.02 | 17.4 | 743 |  | 1 | 1 |  | 115 | 115 | 115 |  |  |  |  | 83.5 | PASS | 1 |  | 743 | 743 | 743 |  |  |  |  | 146.23 | FAIL |  | 22
2020-10-22 | OM_10-22-2020_11-52-24AM.OMN | Sample 104 | 104 | 1 | 1 | TP | ug P/L | 0.108 | 2.18 | 57.9 | TN | ug N/L | 1.05 | 18 | 786 |  | 1 | 1 |  | 57.9 | 57.9 | 57.9 |  |  |  |  | 110.99 | PASS | 1 |  | 786 | 786 | 786 |  |  |  |  | 187.46 | FAIL |  | 23
2020-10-22 | OM_10-22-2020_11-52-24AM.OMN | Sample 116 | 116 | 1 | 1 | TP | ug P/L | 0.124 | 2.55 | 68.6 | TN | ug N/L | 0.748 | 12.8 | 394 |  | 1 | 1 |  | 68.6 | 68.6 | 68.6 |  |  |  |  | 119.26 | PASS | 1 |  | 394 | 394 | 394 |  |  |  |  | 95.54 | PASS |  | 24
2020-10-22 | OM_10-22-2020_11-52-24AM.OMN | Sample 140 | 140 | 1 | 1 | TP | ug P/L | 0.0542 | 1.2 | 28.1 | TN | ug N/L | 0.563 | 9.51 | 141 |  | 1 | 1 |  | 28.1 | 28.1 | 28.1 |  |  |  |  | -17.19 | FAIL | 1 |  | 141 | 141 | 141 |  |  |  |  | -39.39 | FAIL |  | 25
2020-10-22 | OM_10-22-2020_11-52-24AM.OMN | Sample 152 | 152 | 1 | 1 | TP | ug P/L | 0.057 | 1.32 | 31.9 | TN | ug N/L | 0.869 | 14.9 | 553 |  | 1 | 1 |  | 31.9 | 31.9 | 31.9 |  |  |  |  | 6.79 | FAIL | 1 |  | 553 | 553 | 553 |  |  |  |  | 13.93 | FAIL |  | 26
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 046 | 46 | 1 | 1 | TP | ug P/L | 0.124 | 2.55 | 63.5 | TN | ug N/L | 0.642 | 10.9 | 654 |  | 1 | 1 |  | 63.5 | 63.5 | 63.5 |  |  |  |  | 78.35 | FAIL | 1 |  | 654 | 654 | 654 |  |  |  |  | 86.14 | PASS |  | 27
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 058 | 58 | 1 | 1 | TP | ug P/L | 0.0889 | 1.95 | 45.9 | TN | ug N/L | 0.473 | 8.02 | 433 |  | 1 | 1 |  | 45.9 | 45.9 | 45.9 |  |  |  |  | 98.99 | PASS | 1 |  | 433 | 433 | 433 |  |  |  |  | 90.33 | PASS |  | 28
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 070 | 70 | 1 | 1 | TP | ug P/L | 0.119 | 2.46 | 60.9 | TN | ug N/L | 0.801 | 13.6 | 856 |  | 1 | 1 |  | 60.9 | 60.9 | 60.9 |  |  |  |  | 100.49 | PASS | 1 |  | 856 | 856 | 856 |  |  |  |  | 78.96 | FAIL |  | 29
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 082 | 82 | 1 | 1 | TP | ug P/L | 0.174 | 3.44 | 89.1 | TN | ug N/L | 1.61 | 27.2 | 1950 |  | 1 | 1 |  | 89.1 | 89.1 | 89.1 |  |  |  |  | 100.51 | PASS | 1 |  | 1950 | 1950 | 1950 |  |  |  |  | 91.5 | PASS |  | 30
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 094 | 94 | 1 | 1 | TP | ug P/L | 0.103 | 2.17 | 52.4 | TN | ug N/L | 0.906 | 15.3 | 994 |  | 1 | 1 |  | 52.4 | 52.4 | 52.4 |  |  |  |  | 101.24 | PASS | 1 |  | 994 | 994 | 994 |  |  |  |  | 82.34 | PASS |  | 31
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 106 | 106 | 1 | 1 | TP | ug P/L | 0.134 | 2.7 | 67.7 | TN | ug N/L | 0.669 | 11.4 | 691 |  | 1 | 1 |  | 67.7 | 67.7 | 67.7 |  |  |  |  | 95.97 | PASS | 1 |  | 691 | 691 | 691 |  |  |  |  | 89.31 | PASS |  | 32
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 118 | 118 | 1 | 1 | TP | ug P/L | 0.0994 | 2.01 | 47.5 | TN | ug N/L | 1.26 | 21.5 | 1480 |  | 1 | 1 |  | 47.5 | 47.5 | 47.5 |  |  |  |  | 96.35 | PASS | 1 |  | 1480 | 1480 | 1480 |  |  |  |  | 86.8 | PASS |  | 33
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 130 | 130 | 1 | 1 | TP | ug P/L | 0.101 | 2.1 | 50.3 | TN | ug N/L | 0.618 | 10.5 | 621 |  | 1 | 1 |  | 50.3 | 50.3 | 50.3 |  |  |  |  | 96.23 | PASS | 1 |  | 621 | 621 | 621 |  |  |  |  | 94.61 | PASS |  | 34
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 142 | 142 | 1 | 1 | TP | ug P/L | 0.112 | 2.36 | 57.7 | TN | ug N/L | 0.701 | 12 | 734 |  | 1 | 1 |  | 57.7 | 57.7 | 57.7 |  |  |  |  | 100.57 | PASS | 1 |  | 734 | 734 | 734 |  |  |  |  | 91.74 | PASS |  | 35
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 154 | 154 | 1 | 1 | TP | ug P/L | 0.103 | 2.11 | 50.6 | TN | ug N/L | 0.348 | 5.99 | 282 |  | 1 | 1 |  | 50.6 | 50.6 | 50.6 |  |  |  |  | 111.06 | PASS | 1 |  | 282 | 282 | 282 |  |  |  |  | 97.14 | PASS |  | 36
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 46 | 1 | 1 | TP | ug P/L | 0.153 | 3.11 | 69.8 | TN | ug N/L | 0.29 | 5.25 | 372 |  | 1 | 2 | reworked calibration | 81.1145 | 81.1145 | 81.1145 |  |  |  |  | 96.0675 | PASS | 1 |  | 397.1273 | 397.1273 | 397.1273 |  |  |  |  | 94.882 | PASS |  | 37
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 58 | 1 | 1 | TP | ug P/L | 0.0973 | 2.05 | 39.5 | TN | ug N/L | 0.355 | 6.39 | 472 |  | 1 | 2 | reworked calibration | 48.4233 | 48.4233 | 48.4233 |  |  |  |  | 98.9386 | PASS | 1 |  | 508.191 | 508.191 | 508.191 |  |  |  |  | 104.264 | PASS |  | 38
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 70 | 1 | 1 | TP | ug P/L | 0.0924 | 1.95 | 36.8 | TN | ug N/L | 0.345 | 6.23 | 458 |  | 1 | 2 | reworked calibration | 45.3822 | 45.3822 | 45.3822 |  |  |  |  | 98.4033 | PASS | 1 |  | 492.4996 | 492.4996 | 492.4996 |  |  |  |  | 118.9796 | PASS |  | 39
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 82 | 1 | 1 | TP | ug P/L | 0.118 | 2.4 | 49.3 | TN | ug N/L | 0.564 | 10.2 | 801 |  | 1 | 2 | reworked calibration | 59.1255 | 59.1255 | 59.1255 |  |  |  |  | 100.818 | PASS | 1 |  | 891.8231 | 891.8231 | 891.8231 |  |  |  |  | 106.4747 | PASS |  | 40
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 94 | 1 | 1 | TP | ug P/L | 0.117 | 2.42 | 49.9 | TN | ug N/L | 0.525 | 9.4 | 734 |  | 1 | 2 | reworked calibration | 59.7398 | 59.7398 | 59.7398 |  |  |  |  | 102.1313 | PASS | 1 |  | 809.6814 | 809.6814 | 809.6814 |  |  |  |  | 107.4098 | PASS |  | 41
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 106 | 1 | 1 | TP | ug P/L | 0.0941 | 2.02 | 38.6 | TN | ug N/L | 0.374 | 6.83 | 511 |  | 1 | 2 | reworked calibration | 47.5102 | 47.5102 | 47.5102 |  |  |  |  | 110.7022 | PASS | 1 |  | 551.5165 | 551.5165 | 551.5165 |  |  |  |  | 104.1497 | PASS |  | 42
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 118 | 1 | 1 | TP | ug P/L | 0.0801 | 1.73 | 30.6 | TN | ug N/L | 0.362 | 6.67 | 497 |  | 1 | 2 | reworked calibration | 38.7178 | 38.7178 | 38.7178 |  |  |  |  | 103.743 | PASS | 1 |  | 535.7322 | 535.7322 | 535.7322 |  |  |  |  | 101.4917 | PASS |  | 43
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 130 | 1 | 1 | TP | ug P/L | 0.0842 | 1.79 | 32.4 | TN | ug N/L | 0.267 | 4.92 | 343 |  | 1 | 2 | reworked calibration | 40.5318 | 40.5318 | 40.5318 |  |  |  |  | 96.3611 | PASS | 1 |  | 365.2974 | 365.2974 | 365.2974 |  |  |  |  | 78.8375 | FAIL |  | 44
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 142 | 1 | 1 | TP | ug P/L | 0.108 | 2.25 | 45.2 | TN | ug N/L | 0.462 | 8.49 | 656 |  | 1 | 2 | reworked calibration | 54.5277 | 54.5277 | 54.5277 |  |  |  |  | 117.989 | PASS | 1 |  | 717.2721 | 717.2721 | 717.2721 |  |  |  |  | 98.6487 | PASS |  | 45
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 154 | 1 | 1 | TP | ug P/L | 0.109 | 2.25 | 45.1 | TN | ug N/L | 0.482 | 8.87 | 688 |  | 1 | 2 | reworked calibration | 54.5277 | 54.5277 | 54.5277 |  |  |  |  | 102.4123 | PASS | 1 |  | 755.7277 | 755.7277 | 755.7277 |  |  |  |  | 90.0317 | PASS |  | 46
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 46 | 1 | 1 | TP | ug P/L | 0.0959 | 1.95 | 52.1 | TN | ug N/L | 0.645 | 9.21 | 552 |  | 1 | 1 |  | 52.1 | 52.1 | 52.1 |  |  |  |  | 96.01 | PASS | 1 |  | 552 | 552 | 552 |  |  |  |  | 92.72 | PASS |  | 47
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 58 | 1 | 1 | TP | ug P/L | 0.105 | 2.13 | 57.7 | TN | ug N/L | 0.682 | 9.74 | 587 |  | 1 | 1 |  | 57.7 | 57.7 | 57.7 |  |  |  |  | 113.37 | PASS | 1 |  | 587 | 587 | 587 |  |  |  |  | 85.07 | PASS |  | 48
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 70 | 1 | 1 | TP | ug P/L | 0.0757 | 1.61 | 41.6 | TN | ug N/L | 0.603 | 8.63 | 512 |  | 1 | 1 |  | 41.6 | 41.6 | 41.6 |  |  |  |  | 92.16 | PASS | 1 |  | 512 | 512 | 512 |  |  |  |  | 88.72 | PASS |  | 49
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 82 | 1 | 1 | TP | ug P/L | 0.0611 | 1.3 | 31.8 | TN | ug N/L | 0.424 | 6.27 | 351 |  | 1 | 1 |  | 31.8 | 31.8 | 31.8 |  |  |  |  | 97.3 | PASS | 1 |  | 351 | 351 | 351 |  |  |  |  | 108.19 | PASS |  | 50
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 94 | 1 | 1 | TP | ug P/L | 0.178 | 3.38 | 95.1 | TN | ug N/L | 0.849 | 12 | 738 |  | 1 | 1 |  | 95.1 | 95.1 | 95.1 |  |  |  |  | 91.11 | PASS | 1 |  | 738 | 738 | 738 |  |  |  |  | 81.38 | PASS |  | 51
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 106 | 1 | 1 | TP | ug P/L | 0.0709 | 1.56 | 39.8 | TN | ug N/L | 0.525 | 7.51 | 437 |  | 1 | 1 |  | 39.8 | 39.8 | 39.8 |  |  |  |  | 105.58 | PASS | 1 |  | 437 | 437 | 437 |  |  |  |  | 87.97 | PASS |  | 52
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 118 | 1 | 1 | TP | ug P/L | 0.0988 | 2.01 | 53.8 | TN | ug N/L | 1.1 | 15.5 | 963 |  | 1 | 1 |  | 53.8 | 53.8 | 53.8 |  |  |  |  | 69.78 | FAIL | 1 |  | 963 | 963 | 963 |  |  |  |  | 90.03 | PASS |  | 53
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 130 | 1 | 1 | TP | ug P/L | 0.0784 | 1.71 | 44.7 | TN | ug N/L | 0.499 | 7.17 | 413 |  | 1 | 1 |  | 44.7 | 44.7 | 44.7 |  |  |  |  | 120.07 | FAIL | 1 |  | 413 | 413 | 413 |  |  |  |  | 83.73 | PASS |  | 54
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 142 | 1 | 1 | TP | ug P/L | 0.0629 | 1.37 | 34 | TN | ug N/L | 0.572 | 8.2 | 483 |  | 1 | 1 |  | 34 | 34 | 34 |  |  |  |  | 103.32 | PASS | 1 |  | 483 | 483 | 483 |  |  |  |  | 106.83 | PASS |  | 55
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 154 | 1 | 1 | TP | ug P/L | 0.0766 | 1.59 | 40.8 | TN | ug N/L | 0.612 | 8.7 | 517 |  | 1 | 1 |  | 40.8 | 40.8 | 40.8 |  |  |  |  | 108.88 | PASS | 1 |  | 517 | 517 | 517 |  |  |  |  | 112.77 | PASS |  | 56
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 166 | 1 | 1 | TP | ug P/L | 0.0773 | 1.49 | 37.8 | TN | ug N/L | 0.582 | 8.3 | 491 |  | 1 | 1 |  | 37.8 | 37.8 | 37.8 |  |  |  |  | 111.78 | PASS | 1 |  | 491 | 491 | 491 |  |  |  |  | 95.71 | PASS |  | 57
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK F 17aug20 9.0 | 178 | 1 | 1 | TP | ug P/L | 0.119 | 2.29 | 62.5 | TN | ug N/L | 0.833 | 11.8 | 724 |  | 1 | 1 |  | 62.5 | 62.5 | 62.5 |  |  |  |  | 99.85 | PASS | 1 |  | 724 | 724 | 724 |  |  |  |  | 86.44 | PASS |  | 58
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 46 | 1 | 1 | TP | ug P/L | 0.082 | 1.62 | 35.8 | TN | ug N/L | 0.491 | 7.24 | 396 |  | 1 | 1 |  | 35.8 | 35.8 | 35.8 |  |  |  |  | 103.18 | PASS | 1 |  | 396 | 396 | 396 |  |  |  |  | 97.56 | PASS |  | 59
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 58 | 1 | 1 | TP | ug P/L | 0.104 | 2.04 | 47.3 | TN | ug N/L | 0.519 | 7.63 | 422 |  | 1 | 1 |  | 47.3 | 47.3 | 47.3 |  |  |  |  | 98.33 | PASS | 1 |  | 422 | 422 | 422 |  |  |  |  | 90.22 | PASS |  | 60
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 82 | 1 | 1 | TP | ug P/L | 0.11 | 2.16 | 50.7 | TN | ug N/L | 0.64 | 9.34 | 532 |  | 1 | 1 |  | 50.7 | 50.7 | 50.7 |  |  |  |  | 109.87 | PASS | 1 |  | 532 | 532 | 532 |  |  |  |  | 120.92 | FAIL |  | 61
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 94 | 1 | 1 | TP | ug P/L | 0.123 | 2.34 | 55.7 | TN | ug N/L | 0.709 | 10.3 | 594 |  | 1 | 1 |  | 55.7 | 55.7 | 55.7 |  |  |  |  | 101.17 | PASS | 1 |  | 594 | 594 | 594 |  |  |  |  | 103.54 | PASS |  | 62
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 106 | 1 | 1 | TP | ug P/L | 0.107 | 2.05 | 47.7 | TN | ug N/L | 0.55 | 8.05 | 448 |  | 1 | 1 |  | 47.7 | 47.7 | 47.7 |  |  |  |  | 98.77 | PASS | 1 |  | 448 | 448 | 448 |  |  |  |  | 88.88 | PASS |  | 63
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 118 | 1 | 1 | TP | ug P/L | 0.123 | 2.33 | 55.3 | TN | ug N/L | 0.528 | 7.66 | 424 |  | 1 | 1 |  | 55.3 | 55.3 | 55.3 |  |  |  |  | 104.73 | PASS | 1 |  | 424 | 424 | 424 |  |  |  |  | 89.84 | PASS |  | 64
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 130 | 1 | 1 | TP | ug P/L | 0.133 | 2.53 | 61 | TN | ug N/L | 0.563 | 8.19 | 458 |  | 1 | 1 |  | 61 | 61 | 61 |  |  |  |  | 105.3 | PASS | 1 |  | 458 | 458 | 458 |  |  |  |  | 69.38 | FAIL |  | 65
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 142 | 1 | 1 | TP | ug P/L | 0.0964 | 1.86 | 42.5 | TN | ug N/L | 0.628 | 9.15 | 520 |  | 1 | 1 |  | 42.5 | 42.5 | 42.5 |  |  |  |  | 98.65 | PASS | 1 |  | 520 | 520 | 520 |  |  |  |  | 90.4 | PASS |  | 66
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 154 | 1 | 1 | TP | ug P/L | 0.0777 | 1.53 | 33.3 | TN | ug N/L | 0.482 | 7.04 | 383 |  | 1 | 1 |  | 33.3 | 33.3 | 33.3 |  |  |  |  | 100.93 | PASS | 1 |  | 383 | 383 | 383 |  |  |  |  | 97.83 | PASS |  | 67
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 046 | 46 | 1 | 1 | TP | ug P/L | 0.115 | 2.08 | 51.9 | TN | ug N/L | 1.03 | 13.9 | 853 |  | 1 | 1 |  | 51.9 | 51.9 | 51.9 |  |  |  |  | 103.59 | PASS | 1 |  | 853 | 853 | 853 |  |  |  |  | 76.93 | FAIL |  | 68
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 058 | 58 | 1 | 1 | TP | ug P/L | 0.105 | 1.89 | 46.4 | TN | ug N/L | 0.63 | 8.52 | 484 |  | 1 | 1 |  | 46.4 | 46.4 | 46.4 |  |  |  |  | 97.84 | PASS | 1 |  | 484 | 484 | 484 |  |  |  |  | 89.24 | PASS |  | 69
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 070 | 70 | 1 | 1 | TP | ug P/L | 0.121 | 2.1 | 52.5 | TN | ug N/L | 0.824 | 11.1 | 658 |  | 1 | 1 |  | 52.5 | 52.5 | 52.5 |  |  |  |  | 104.05 | PASS | 1 |  | 658 | 658 | 658 |  |  |  |  | 104.58 | PASS |  | 70
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 082 | 82 | 1 | 1 | TP | ug P/L | 0.195 | 3.15 | 82.2 | TN | ug N/L | 0.995 | 13.4 | 818 |  | 1 | 1 |  | 82.2 | 82.2 | 82.2 |  |  |  |  | 110.22 | PASS | 1 |  | 818 | 818 | 818 |  |  |  |  | 25.78 | FAIL |  | 71
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 094 | 94 | 1 | 1 | TP | ug P/L | 0.118 | 2.06 | 51.3 | TN | ug N/L | 0.631 | 8.54 | 486 |  | 1 | 1 |  | 51.3 | 51.3 | 51.3 |  |  |  |  | 101.53 | PASS | 1 |  | 486 | 486 | 486 |  |  |  |  | 110.46 | PASS |  | 72
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 106 | 106 | 1 | 1 | TP | ug P/L | 0.0987 | 1.75 | 42.4 | TN | ug N/L | 0.664 | 9.01 | 517 |  | 1 | 1 |  | 42.4 | 42.4 | 42.4 |  |  |  |  | 104.64 | PASS | 1 |  | 517 | 517 | 517 |  |  |  |  | 77.97 | FAIL |  | 73
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 118 | 118 | 1 | 1 | TP | ug P/L | 0.116 | 2.06 | 51.4 | TN | ug N/L | 0.844 | 11.4 | 678 |  | 1 | 1 |  | 51.4 | 51.4 | 51.4 |  |  |  |  | 107.54 | PASS | 1 |  | 678 | 678 | 678 |  |  |  |  | 89.18 | PASS |  | 74
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 130 | 130 | 1 | 1 | TP | ug P/L | 0.0743 | 1.35 | 30.9 | TN | ug N/L | 0.466 | 6.35 | 341 |  | 1 | 1 |  | 30.9 | 30.9 | 30.9 |  |  |  |  | 98.33 | PASS | 1 |  | 341 | 341 | 341 |  |  |  |  | 87.41 | PASS |  | 75
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 142 | 142 | 1 | 1 | TP | ug P/L | 0.0944 | 1.74 | 42.1 | TN | ug N/L | 1.23 | 16.6 | 1040 |  | 1 | 1 |  | 42.1 | 42.1 | 42.1 |  |  |  |  | 108.21 | PASS | 1 |  | 1040 | 1040 | 1040 |  |  |  |  | 75.2 | FAIL |  | 76
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 154 | 154 | 1 | 1 | TP | ug P/L | 0.13 | 2.26 | 57 | TN | ug N/L | 1.02 | 13.8 | 844 |  | 1 | 1 |  | 57 | 57 | 57 |  |  |  |  | 91.7 | PASS | 1 |  | 844 | 844 | 844 |  |  |  |  | 78.84 | FAIL |  | 77
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample46 | 46 | 1 | 1 | TP | ug P/L | 0.0996 | 1.75 | 35.1 | TN | ug N/L | 0.513 | 6.96 | 379 |  | 1 | 1 |  | 35.1 | 35.1 | 35.1 |  |  |  |  | 81.11 | PASS | 1 |  | 379 | 379 | 379 |  |  |  |  | 43.79 | FAIL |  | 78
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample58 | 58 | 1 | 1 | TP | ug P/L | 0.105 | 1.93 | 39.4 | TN | ug N/L | 0.513 | 6.74 | 364 |  | 1 | 1 |  | 39.4 | 39.4 | 39.4 |  |  |  |  | 88.74 | PASS | 1 |  | 364 | 364 | 364 |  |  |  |  | 91.24 | PASS |  | 79
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample70 | 70 | 1 | 1 | TP | ug P/L | 0.106 | 1.89 | 38.4 | TN | ug N/L | 0.672 | 9.02 | 517 |  | 1 | 1 |  | 38.4 | 38.4 | 38.4 |  |  |  |  | 67.84 | FAIL | 1 |  | 517 | 517 | 517 |  |  |  |  | 77.97 | FAIL |  | 80
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample82 | 82 | 1 | 1 | TP | ug P/L | 0.123 | 2.16 | 45.3 | TN | ug N/L | 0.623 | 8.29 | 468 |  | 1 | 1 |  | 45.3 | 45.3 | 45.3 |  |  |  |  | 104.13 | PASS | 1 |  | 468 | 468 | 468 |  |  |  |  | 82.28 | PASS |  | 81
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample94 | 94 | 1 | 1 | TP | ug P/L | 0.201 | 3.48 | 79.3 | TN | ug N/L | 1.38 | 18.4 | 1150 |  | 1 | 1 |  | 79.3 | 79.3 | 79.3 |  |  |  |  | 62.33 | FAIL | 1 |  | 1150 | 1150 | 1150 |  |  |  |  | -16.5 | FAIL |  | 82
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample106 | 106 | 1 | 1 | TP | ug P/L | 0.132 | 2.33 | 49.6 | TN | ug N/L | 1.13 | 15.3 | 939 |  | 1 | 1 |  | 49.6 | 49.6 | 49.6 |  |  |  |  | 80.16 | PASS | 1 |  | 939 | 939 | 939 |  |  |  |  | 45.39 | FAIL |  | 83
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample118 | 118 | 1 | 1 | TP | ug P/L | 0.108 | 1.92 | 39.3 | TN | ug N/L | 0.745 | 10.1 | 588 |  | 1 | 1 |  | 39.3 | 39.3 | 39.3 |  |  |  |  | 67.93 | FAIL | 1 |  | 588 | 588 | 588 |  |  |  |  | 79.88 | FAIL |  | 84
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample130 | 130 | 1 | 1 | TP | ug P/L | 0.0823 | 1.48 | 28.3 | TN | ug N/L | 0.448 | 6.06 | 318 |  | 1 | 1 |  | 28.3 | 28.3 | 28.3 |  |  |  |  | 80.19 | PASS | 1 |  | 318 | 318 | 318 |  |  |  |  | 90.46 | PASS |  | 85
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample142 | 142 | 1 | 1 | TP | ug P/L | 0.0821 | 1.47 | 28 | TN | ug N/L | 0.486 | 6.56 | 352 |  | 1 | 1 |  | 28 | 28 | 28 |  |  |  |  | 88.88 | PASS | 1 |  | 352 | 352 | 352 |  |  |  |  | 105.84 | PASS |  | 86
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample154 | 154 | 1 | 1 | TP | ug P/L | 0.0985 | 1.75 | 35 | TN | ug N/L | 0.639 | 8.51 | 482 |  | 1 | 1 |  | 35 | 35 | 35 |  |  |  |  | 87.9 | PASS | 1 |  | 482 | 482 | 482 |  |  |  |  | 95.22 | PASS |  | 87
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 046 | 46 | 1 | 1 | TP | ug P/L | 0.12 | 2.07 | 48.9 | TN | ug N/L | 1.91 | 26.2 | 1730 |  | 1 | 1 |  | 48.9 | 48.9 | 48.9 |  |  |  |  | 92.89 | PASS | 1 |  | 1730 | 1730 | 1730 |  |  |  |  | 61.3 | FAIL |  | 88
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 058 | 58 | 1 | 1 | TP | ug P/L | 0.12 | 2.06 | 48.7 | TN | ug N/L | 0.587 | 8.02 | 491 |  | 1 | 1 |  | 48.7 | 48.7 | 48.7 |  |  |  |  | 130.87 | FAIL | 1 |  | 491 | 491 | 491 |  |  |  |  | 91.71 | PASS |  | 89
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 070 | 70 | 1 | 1 | TP | ug P/L | 0.0988 | 1.71 | 38.9 | TN | ug N/L | 0.538 | 7.37 | 445 |  | 1 | 1 |  | 38.9 | 38.9 | 38.9 |  |  |  |  | 129.09 | FAIL | 1 |  | 445 | 445 | 445 |  |  |  |  | 102.45 | PASS |  | 90
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 082 | 82 | 1 | 1 | TP | ug P/L | 0.16 | 2.76 | 67.9 | TN | ug N/L | 1.03 | 14.1 | 913 |  | 1 | 1 |  | 67.9 | 67.9 | 67.9 |  |  |  |  | 83.59 | PASS | 1 |  | 913 | 913 | 913 |  |  |  |  | 177.93 | FAIL |  | 91
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 094 | 94 | 1 | 1 | TP | ug P/L | 0.119 | 2.04 | 48.1 | TN | ug N/L | 0.565 | 7.83 | 477 |  | 1 | 1 |  | 48.1 | 48.1 | 48.1 |  |  |  |  | 132.41 | FAIL | 1 |  | 477 | 477 | 477 |  |  |  |  | 104.37 | PASS |  | 92
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 106 | 106 | 1 | 1 | TP | ug P/L | 0.108 | 1.86 | 43.2 | TN | ug N/L | 0.5 | 6.86 | 409 |  | 1 | 1 |  | 43.2 | 43.2 | 43.2 |  |  |  |  | 135.52 | FAIL | 1 |  | 409 | 409 | 409 |  |  |  |  | 97.29 | PASS |  | 93
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 118 | 118 | 1 | 1 | TP | ug P/L | 0.16 | 2.74 | 67.6 | TN | ug N/L | 0.503 | 6.95 | 415 |  | 1 | 1 |  | 67.6 | 67.6 | 67.6 |  |  |  |  | 123.16 | FAIL | 1 |  | 415 | 415 | 415 |  |  |  |  | 125.35 | FAIL |  | 94
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 130 | 130 | 1 | 1 | TP | ug P/L | 0.15 | 2.58 | 63.1 | TN | ug N/L | 2.02 | 27.9 | 1840 |  | 1 | 1 |  | 63.1 | 63.1 | 63.1 |  |  |  |  | 153.91 | FAIL | 1 |  | 1840 | 1840 | 1840 |  |  |  |  | 66.4 | FAIL |  | 95
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 142 | 142 | 1 | 1 | TP | ug P/L | 0.142 | 2.46 | 59.6 | TN | ug N/L | 0.768 | 10.6 | 667 |  | 1 | 1 |  | 59.6 | 59.6 | 59.6 |  |  |  |  | 137.16 | FAIL | 1 |  | 667 | 667 | 667 |  |  |  |  | 109.07 | PASS |  | 96
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 154 | 154 | 1 | 1 | TP | ug P/L | 0.153 | 2.59 | 63.4 | TN | ug N/L | 1.09 | 14.9 | 969 |  | 1 | 1 |  | 63.4 | 63.4 | 63.4 |  |  |  |  | 133.14 | FAIL | 1 |  | 969 | 969 | 969 |  |  |  |  | 90.49 | PASS |  | 97
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 046 | 46 | 1 | 1 | TP | ug P/L | 0.134 | 2.26 | 54.1 | TN | ug N/L | 0.748 | 9.05 | 548 |  | 1 | 1 |  | 54.1 | 54.1 | 54.1 |  |  |  |  | 116.21 | PASS | 1 |  | 548 | 548 | 548 |  |  |  |  | 98.68 | PASS |  | 98
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 058 | 58 | 1 | 1 | TP | ug P/L | 0.124 | 2.03 | 47.2 | TN | ug N/L | 0.817 | 9.88 | 605 |  | 1 | 1 |  | 47.2 | 47.2 | 47.2 |  |  |  |  | 107.52 | PASS | 1 |  | 605 | 605 | 605 |  |  |  |  | 113.25 | PASS |  | 99
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 070 | 70 | 1 | 1 | TP | ug P/L | 0.128 | 2 | 46.5 | TN | ug N/L | 0.715 | 8.73 | 526 |  | 1 | 1 |  | 46.5 | 46.5 | 46.5 |  |  |  |  | 113.05 | PASS | 1 |  | 526 | 526 | 526 |  |  |  |  | 96.46 | PASS |  | 100
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 082 | 82 | 1 | 1 | TP | ug P/L | 0.101 | 1.68 | 37.3 | TN | ug N/L | 0.644 | 7.97 | 474 |  | 1 | 1 |  | 37.3 | 37.3 | 37.3 |  |  |  |  | 108.53 | PASS | 1 |  | 474 | 474 | 474 |  |  |  |  | 97.94 | PASS |  | 101
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 094 | 94 | 1 | 1 | TP | ug P/L | 0.161 | 2.66 | 65.5 | TN | ug N/L | 1.76 | 21.6 | 1420 |  | 1 | 1 |  | 65.5 | 65.5 | 65.5 |  |  |  |  | 115.75 | PASS | 1 |  | 1420 | 1420 | 1420 |  |  |  |  | 106.2 | PASS |  | 102
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 106 | 106 | 1 | 1 | TP | ug P/L | 0.119 | 1.97 | 45.6 | TN | ug N/L | 0.958 | 11.7 | 734 |  | 1 | 1 |  | 45.6 | 45.6 | 45.6 |  |  |  |  | 113.76 | PASS | 1 |  | 734 | 734 | 734 |  |  |  |  | 149.34 | FAIL |  | 103
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 118 | 118 | 1 | 1 | TP | ug P/L | 0.144 | 2.4 | 58.2 | TN | ug N/L | 0.775 | 9.68 | 591 |  | 1 | 1 |  | 58.2 | 58.2 | 58.2 |  |  |  |  | 115.02 | PASS | 1 |  | 591 | 591 | 591 |  |  |  |  | 106.71 | PASS |  | 104
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 130 | 130 | 1 | 1 | TP | ug P/L | 0.124 | 2.03 | 47.4 | TN | ug N/L | 0.707 | 8.89 | 537 |  | 1 | 1 |  | 47.4 | 47.4 | 47.4 |  |  |  |  | 105.94 | PASS | 1 |  | 537 | 537 | 537 |  |  |  |  | 104.17 | PASS |  | 105
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 142 | 142 | 1 | 1 | TP | ug P/L | 0.165 | 2.73 | 67.6 | TN | ug N/L | 1.31 | 16.2 | 1050 |  | 1 | 1 |  | 67.6 | 67.6 | 67.6 |  |  |  |  | 105.16 | PASS | 1 |  | 1050 | 1050 | 1050 |  |  |  |  | 104.5 | PASS |  | 106
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 154 | 154 | 1 | 1 | TP | ug P/L | 0.18 | 3 | 75.6 | TN | ug N/L | 0.452 | 5.78 | 323 |  | 1 | 1 |  | 75.6 | 75.6 | 75.6 |  |  |  |  | 132.36 | FAIL | 1 |  | 323 | 323 | 323 |  |  |  |  | 69.63 | FAIL |  | 107
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk10 | 46 | 1 | 1 | TP | ug P/L | 0.0958 | 1.64 | -41.5 | TN | ug N/L | 1.07 | 14.4 | 639 |  | 1 | 1 |  | 48.7766 | 48.7766 | 48.7766 |  |  |  |  | 100.2514 | PASS | 2 | reworked calibration | 639 | 639 | 639 |  |  |  |  | 87.99 | PASS |  | 108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk20 | 58 | 1 | 1 | TP | ug P/L | 0.105 | 1.9 | -41 | TN | ug N/L | 2.81 | 37.7 | 1800 |  | 1 | 1 |  | 56.9462 | 56.9462 | 56.9462 |  |  |  |  | 109.9372 | PASS | 2 | reworked calibration | 1800 | 1800 | 1800 |  |  |  |  | 106 | PASS |  | 109
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup30 | 70 | 1 | 1 | TP | ug P/L | 0.098 | 1.72 | -41.4 | TN | ug N/L | 1.46 | 19.7 | 900 |  | 1 | 1 |  | 50.8269 | 50.8269 | 50.8269 |  |  |  |  | 99.9414 | PASS | 2 | reworked calibration | 900 | 900 | 900 |  |  |  |  | 96.2 | PASS |  | 110
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk40 | 82 | 1 | 1 | TP | ug P/L | 0.1 | 1.74 | -41.3 | TN | ug N/L | 2.09 | 28.1 | 1310 |  | 1 | 1 |  | 52.638 | 52.638 | 52.638 |  |  |  |  | 122.4791 | FAIL | 2 | reworked calibration | 1310 | 1310 | 1310 |  |  |  |  | 97.1 | PASS |  | 111
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk50 | 94 | 1 | 1 | TP | ug P/L | 0.101 | 1.85 | -41.2 | TN | ug N/L | 1.2 | 16.2 | 726 |  | 1 | 1 |  | 53.5247 | 53.5247 | 53.5247 |  |  |  |  | 132.0941 | FAIL | 2 | reworked calibration | 726 | 726 | 726 |  |  |  |  | 106.46 | PASS |  | 112
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk60 | 106 | 1 | 1 | TP | ug P/L | 0.136 | 2.71 | -39 | TN | ug N/L | 3.47 | 47.2 | 2280 |  | 1 | 1 |  | 76.6538 | 76.6538 | 76.6538 |  |  |  |  | 104.8025 | PASS | 2 | reworked calibration | 2280 | 2280 | 2280 |  |  |  |  | 122.8 | FAIL |  | 113
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk70 | 118 | 1 | 1 | TP | ug P/L | 0.107 | 1.95 | -40.8 | TN | ug N/L | 2.14 | 29 | 1360 |  | 1 | 1 |  | 58.5816 | 58.5816 | 58.5816 |  |  |  |  | 91.9017 | PASS | 2 | reworked calibration | 1360 | 1360 | 1360 |  |  |  |  | 81.6 | PASS |  | 114
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk80 | 130 | 1 | 1 | TP | ug P/L | 0.107 | 1.99 | -40.9 | TN | ug N/L | 3.25 | 44.1 | 2120 |  | 1 | 1 |  | 58.5816 | 58.5816 | 58.5816 |  |  |  |  | 81.1579 | PASS | 2 | reworked calibration | 2120 | 2120 | 2120 |  |  |  |  | 89.2 | PASS |  | 115
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk90 | 142 | 1 | 1 | TP | ug P/L | 0.0892 | 1.38 | -42 | TN | ug N/L | 0.961 | 13 | 571 |  | 1 | 1 |  | 42.2614 | 42.2614 | 42.2614 |  |  |  |  | 112.6603 | PASS | 2 | reworked calibration | 571 | 571 | 571 |  |  |  |  | 97.71 | PASS |  | 116
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | spk100 | 154 | 1 | 1 | TP | ug P/L | 0.102 | 1.67 | -41.1 | TN | ug N/L | 1.66 | 22.4 | 1030 |  | 1 | 1 |  | 54.3989 | 54.3989 | 54.3989 |  |  |  |  | 126.2444 | FAIL | 2 | reworked calibration | 1030 | 1030 | 1030 |  |  |  |  | 99.1 | PASS |  | 117
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkS10 | 46 | 1 | 1 | TP | ug P/L | 0.178 | 2.95 | 44.5 | TN | ug N/L | 0.943 | 12.8 | 529 |  | 1 | 1 |  | 44.5 | 44.5 | 44.5 |  |  |  |  | 96.05 | PASS | 1 |  | 529 | 529 | 529 |  |  |  |  | 96.49 | PASS |  | 118
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkS20 | 58 | 1 | 1 | TP | ug P/L | 0.161 | 2.68 | 35.6 | TN | ug N/L | 0.862 | 11.8 | 483 |  | 1 | 1 |  | 35.6 | 35.6 | 35.6 |  |  |  |  | 108.64 | PASS | 1 |  | 483 | 483 | 483 |  |  |  |  | 82.83 | PASS |  | 119
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkS30 | 70 | 1 | 1 | TP | ug P/L | 0.274 | 4.44 | 94.3 | TN | ug N/L | 0.913 | 12.9 | 537 |  | 1 | 1 |  | 94.3 | 94.3 | 94.3 |  |  |  |  | 47.43 | FAIL | 1 |  | 537 | 537 | 537 |  |  |  |  | 125.37 | FAIL |  | 120
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkS40 | 82 | 1 | 1 | TP | ug P/L | 0.163 | 2.74 | 37.8 | TN | ug N/L | 0.847 | 11.6 | 474 |  | 1 | 1 |  | 37.8 | 37.8 | 37.8 |  |  |  |  | 112.58 | PASS | 1 |  | 474 | 474 | 474 |  |  |  |  | 97.54 | PASS |  | 121
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkS50 | 94 | 1 | 1 | TP | ug P/L | 0.143 | 2.43 | 27.6 | TN | ug N/L | 0.736 | 10.1 | 398 |  | 1 | 1 |  | 27.6 | 27.6 | 27.6 |  |  |  |  | 84.52 | PASS | 1 |  | 398 | 398 | 398 |  |  |  |  | 107.58 | PASS |  | 122
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spikS60 | 106 | 1 | 1 | TP | ug P/L | 0.188 | 3.18 | 52.2 | TN | ug N/L | 0.971 | 13.5 | 564 |  | 1 | 1 |  | 52.2 | 52.2 | 52.2 |  |  |  |  | 100.02 | PASS | 1 |  | 564 | 564 | 564 |  |  |  |  | 113.24 | PASS |  | 123
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkS70 | 118 | 1 | 1 | TP | ug P/L | 0.179 | 3.09 | 49.4 | TN | ug N/L | 1 | 13.8 | 581 |  | 1 | 1 |  | 49.4 | 49.4 | 49.4 |  |  |  |  | 103.74 | PASS | 1 |  | 581 | 581 | 581 |  |  |  |  | 168.21 | FAIL |  | 124
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkS80 | 130 | 1 | 1 | TP | ug P/L | 0.216 | 3.66 | 68.2 | TN | ug N/L | 0.757 | 10.4 | 412 |  | 1 | 1 |  | 68.2 | 68.2 | 68.2 |  |  |  |  | 90.82 | PASS | 1 |  | 412 | 412 | 412 |  |  |  |  | 106.12 | PASS |  | 125
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkS90 | 142 | 1 | 1 | TP | ug P/L | 0.182 | 3.15 | 51.2 | TN | ug N/L | 0.95 | 13.2 | 551 |  | 1 | 1 |  | 51.2 | 51.2 | 51.2 |  |  |  |  | 112.32 | PASS | 1 |  | 551 | 551 | 551 |  |  |  |  | 107.91 | PASS |  | 126
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | spkS100 | 154 | 1 | 1 | TP | ug P/L | 0.184 | 3.21 | 53.3 | TN | ug N/L | 1.24 | 16.5 | 714 |  | 1 | 1 |  | 53.3 | 53.3 | 53.3 |  |  |  |  | 121.33 | FAIL | 1 |  | 714 | 714 | 714 |  |  |  |  | 98.34 | PASS |  | 127
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | S10 SPK | 46 | 1 | 1 | TP | ug P/L | 0.22 | 3.79 | 103 | TN | ug N/L | 0.661 | 9.95 | 1490 |  | 1 | 1 |  | 103 | 103 | 103 |  |  |  |  | 94.3 | PASS | 1 |  | 1490 | 1490 | 1490 |  |  |  |  | 102.9 | PASS |  | 128
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | S20 SPK | 58 | 1 | 1 | TP | ug P/L | 0.104 | 1.9 | 47.1 | TN | ug N/L | 0.28 | 3.99 | 508 |  | 1 | 1 |  | 47.1 | 47.1 | 47.1 |  |  |  |  | 96.71 | PASS | 1 |  | 508 | 508 | 508 |  |  |  |  | 73.88 | FAIL |  | 129
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | s30 SPK | 70 | 1 | 1 | TP | ug P/L | 0.227 | 3.89 | 105 | TN | ug N/L | 0.566 | 8.33 | 1210 |  | 1 | 1 |  | 105 | 105 | 105 |  |  |  |  | 71.7 | FAIL | 1 |  | 1210 | 1210 | 1210 |  |  |  |  | 64.1 | FAIL |  | 130
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | s40 SPK | 82 | 1 | 1 | TP | ug P/L | 0.107 | 1.95 | 48.6 | TN | ug N/L | 0.284 | 4.04 | 515 |  | 1 | 1 |  | 48.6 | 48.6 | 48.6 |  |  |  |  | 123.26 | FAIL | 1 |  | 515 | 515 | 515 |  |  |  |  | 95.95 | PASS |  | 131
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | s50 SPK | 94 | 1 | 1 | TP | ug P/L | 0.0999 | 1.74 | 42.2 | TN | ug N/L | 0.419 | 6.16 | 849 |  | 1 | 1 |  | 42.2 | 42.2 | 42.2 |  |  |  |  | 10.582 | FAIL | 1 |  | 849 | 849 | 849 |  |  |  |  | 102.49 | PASS |  | 132
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | s60 SPK | 106 | 1 | 1 | TP | ug P/L | 0.0994 | 1.71 | 41.4 | TN | ug N/L | 0.249 | 3.64 | 454 |  | 1 | 1 |  | 41.4 | 41.4 | 41.4 |  |  |  |  | 98.54 | PASS | 1 |  | 454 | 454 | 454 |  |  |  |  | 88.54 | PASS |  | 133
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | s70 SPK | 118 | 1 | 1 | TP | ug P/L | 0.103 | 1.8 | 44 | TN | ug N/L | 0.412 | 5.98 | 820 |  | 1 | 1 |  | 44 | 44 | 44 |  |  |  |  | 96.4 | PASS | 1 |  | 820 | 820 | 820 |  |  |  |  | 82.2 | PASS |  | 134
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | s80 SPK | 130 | 1 | 1 | TP | ug P/L | 0.114 | 2.01 | 50.5 | TN | ug N/L | 0.547 | 8.08 | 1170 |  | 1 | 1 |  | 50.5 | 50.5 | 50.5 |  |  |  |  | 75.85 | FAIL | 1 |  | 1170 | 1170 | 1170 |  |  |  |  | 71.7 | FAIL |  | 135
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | s90 SPK | 142 | 1 | 1 | TP | ug P/L | 0.106 | 1.87 | 46.3 | TN | ug N/L | 0.274 | 3.96 | 503 |  | 1 | 1 |  | 46.3 | 46.3 | 46.3 |  |  |  |  | 93.43 | PASS | 1 |  | 503 | 503 | 503 |  |  |  |  | 74.63 | FAIL |  | 136
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | s100 SPK | 154 | 1 | 1 | TP | ug P/L | 0.097 | 1.7 | 41.2 | TN | ug N/L | 0.252 | 3.73 | 467 |  | 1 | 1 |  | 41.2 | 41.2 | 41.2 |  |  |  |  | 101.32 | PASS | 1 |  | 467 | 467 | 467 |  |  |  |  | 97.87 | PASS |  | 137
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | s108 SPK | 164 | 1 | 1 | TP | ug P/L | 0.0971 | 1.72 | 41.7 | TN | ug N/L | 0.271 | 3.72 | 466 |  | 1 | 1 |  | 41.7 | 41.7 | 41.7 |  |  |  |  | 103.77 | PASS | 1 |  | 466 | 466 | 466 |  |  |  |  | 69.46 | FAIL |  | 138
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | s10 SPK | 46 | 1 | 1 | TP | ug P/L | 0.136 | 2.31 | 59.9 | TN | ug N/L | 0.499 | 6.76 | 998 |  | 1 | 1 |  | 59.9 | 59.9 | 59.9 |  |  |  |  | 102.79 | PASS | 1 |  | 998 | 998 | 998 |  |  |  |  | 76.78 | FAIL |  | 139
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK S20 | 58 | 1 | 1 | TP | ug P/L | 0.153 | 2.59 | 68.2 | TN | ug N/L | 0.632 | 8.8 | 1390 |  | 1 | 1 |  | 68.2 | 68.2 | 68.2 |  |  |  |  | 98.42 | PASS | 1 |  | 1390 | 1390 | 1390 |  |  |  |  | 85.9 | PASS |  | 140
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s30 | 70 | 1 | 1 | TP | ug P/L | 0.111 | 1.9 | 47.3 | TN | ug N/L | 0.257 | 3.53 | 442 |  | 1 | 1 |  | 47.3 | 47.3 | 47.3 |  |  |  |  | 106.73 | PASS | 1 |  | 442 | 442 | 442 |  |  |  |  | 85.22 | PASS |  | 141
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s40 | 82 | 1 | 1 | TP | ug P/L | 0.131 | 2.15 | 55 | TN | ug N/L | 0.374 | 5.06 | 694 |  | 1 | 1 |  | 55 | 55 | 55 |  |  |  |  | 51.9 | FAIL | 1 |  | 694 | 694 | 694 |  |  |  |  | 69.74 | FAIL |  | 142
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s50 | 94 | 1 | 1 | TP | ug P/L | 0.165 | 2.77 | 73.5 | TN | ug N/L | 0.379 | 5.49 | 769 |  | 1 | 1 |  | 73.5 | 73.5 | 73.5 |  |  |  |  | 106.95 | PASS | 1 |  | 769 | 769 | 769 |  |  |  |  | 138.49 | FAIL |  | 143
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s60 | 106 | 1 | 1 | TP | ug P/L | 0.13 | 2.22 | 56.9 | TN | ug N/L | 0.428 | 5.77 | 819 |  | 1 | 1 |  | 56.9 | 56.9 | 56.9 |  |  |  |  | 94.49 | PASS | 1 |  | 819 | 819 | 819 |  |  |  |  | 71.79 | FAIL |  | 144
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s70 | 118 | 1 | 1 | TP | ug P/L | 0.125 | 2.13 | 54.4 | TN | ug N/L | 0.489 | 6.66 | 980 |  | 1 | 1 |  | 54.4 | 54.4 | 54.4 |  |  |  |  | 118.64 | PASS | 1 |  | 980 | 980 | 980 |  |  |  |  | 75.4 | FAIL |  | 145
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s80 | 130 | 1 | 1 | TP | ug P/L | 0.0859 | 1.5 | 35 | TN | ug N/L | 0.228 | 3.1 | 375 |  | 1 | 1 |  | 35 | 35 | 35 |  |  |  |  | 108.14 | PASS | 1 |  | 375 | 375 | 375 |  |  |  |  | 74.55 | FAIL |  | 146
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s90 | 142 | 1 | 1 | TP | ug P/L | 0.117 | 2.02 | 50.9 | TN | ug N/L | 0.277 | 3.73 | 474 |  | 1 | 1 |  | 50.9 | 50.9 | 50.9 |  |  |  |  | 105.49 | PASS | 1 |  | 474 | 474 | 474 |  |  |  |  | 82.34 | PASS |  | 147
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s100 | 154 | 1 | 1 | TP | ug P/L | 0.13 | 2.26 | 58.1 | TN | ug N/L | 0.334 | 4.49 | 598 |  | 1 | 1 |  | 58.1 | 58.1 | 58.1 |  |  |  |  | 112.61 | PASS | 1 |  | 598 | 598 | 598 |  |  |  |  | 98.78 | PASS |  | 148
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s110 | 166 | 1 | 1 | TP | ug P/L | 0.132 | 2.3 | 59.5 | TN | ug N/L | 0.48 | 6.61 | 971 |  | 1 | 1 |  | 59.5 | 59.5 | 59.5 |  |  |  |  | 101.15 | PASS | 1 |  | 971 | 971 | 971 |  |  |  |  | 88.51 | PASS |  | 149
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s120 | 178 | 1 | 1 | TP | ug P/L | 0.117 | 1.99 | 49.9 | TN | ug N/L | 0.343 | 4.66 | 627 |  | 1 | 1 |  | 49.9 | 49.9 | 49.9 |  |  |  |  | 108.59 | PASS | 1 |  | 627 | 627 | 627 |  |  |  |  | 96.27 | PASS |  | 150
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | SPK s130 | 190 | 1 | 1 | TP | ug P/L | 0.0922 | 1.6 | 38.3 | TN | ug N/L | 0.236 | 3.17 | 386 |  | 1 | 1 |  | 38.3 | 38.3 | 38.3 |  |  |  |  | 101.03 | PASS | 1 |  | 386 | 386 | 386 |  |  |  |  | 74.26 | FAIL |  | 151
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | s10 SPK | 46 | 1 | 1 | TP | ug P/L | 0.125 | 2.13 | 57.9 | TN | ug N/L | 0.333 | 4.44 | 641 |  | 1 | 1 |  | 57.9 | 57.9 | 57.9 |  |  |  |  | 106.19 | PASS | 1 |  | 641 | 641 | 641 |  |  |  |  | 102.41 | PASS |  | 152
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | s20 SPK | 58 | 1 | 1 | TP | ug P/L | 0.131 | 2.22 | 60.8 | TN | ug N/L | 0.54 | 7.44 | 1200 |  | 1 | 1 |  | 60.8 | 60.8 | 60.8 |  |  |  |  | 100.08 | PASS | 1 |  | 1200 | 1200 | 1200 |  |  |  |  | 48 | FAIL |  | 153
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | s30 SPK | 70 | 1 | 1 | TP | ug P/L | 0.0706 | 1.22 | 30.2 | TN | ug N/L | 0.219 | 2.96 | 388 |  | 1 | 1 |  | 30.2 | 30.2 | 30.2 |  |  |  |  | 106.7 | PASS | 1 |  | 388 | 388 | 388 |  |  |  |  | 103.88 | PASS |  | 154
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | s40 SPK | 82 | 1 | 1 | TP | ug P/L | 0.134 | 2.28 | 62.5 | TN | ug N/L | 0.58 | 8.07 | 1330 |  | 1 | 1 |  | 62.5 | 62.5 | 62.5 |  |  |  |  | 96.25 | PASS | 1 |  | 1330 | 1330 | 1330 |  |  |  |  | 97.3 | PASS |  | 155
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | SPK s50 | 94 | 1 | 1 | TP | ug P/L | 0.158 | 2.66 | 74 | TN | ug N/L | 0.504 | 6.99 | 1110 |  | 1 | 1 |  | 74 | 74 | 74 |  |  |  |  | 109.8 | PASS | 1 |  | 1110 | 1110 | 1110 |  |  |  |  | 93.9 | PASS |  | 156
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | SPK s60 | 106 | 1 | 1 | TP | ug P/L | 0.0776 | 1.33 | 33.8 | TN | ug N/L | 0.175 | 2.37 | 292 |  | 1 | 1 |  | 33.8 | 33.8 | 33.8 |  |  |  |  | 94.18 | PASS | 1 |  | 292 | 292 | 292 |  |  |  |  | 46.52 | FAIL |  | 157
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | SPK s70 | 118 | 1 | 1 | TP | ug P/L | 0.0751 | 1.34 | 34.1 | TN | ug N/L | 0.203 | 2.85 | 370 |  | 1 | 1 |  | 34.1 | 34.1 | 34.1 |  |  |  |  | 109.09 | PASS | 1 |  | 370 | 370 | 370 |  |  |  |  | 103.7 | PASS |  | 158
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | SPK s80 | 130 | 1 | 1 | TP | ug P/L | 0.149 | 2.54 | 70.5 | TN | ug N/L | 0.511 | 7.08 | 1130 |  | 1 | 1 |  | 70.5 | 70.5 | 70.5 |  |  |  |  | 115.05 | PASS | 1 |  | 1130 | 1130 | 1130 |  |  |  |  | 84.9 | PASS |  | 159
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | SPK s90 | 142 | 1 | 1 | TP | ug P/L | 0.0667 | 1.17 | 28.8 | TN | ug N/L | 0.208 | 2.89 | 376 |  | 1 | 1 |  | 28.8 | 28.8 | 28.8 |  |  |  |  | 106.24 | PASS | 1 |  | 376 | 376 | 376 |  |  |  |  | 103.36 | PASS |  | 160
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | SPK s100 | 154 | 1 | 1 | TP | ug P/L | 0.103 | 1.78 | 47.3 | TN | ug N/L | 0.281 | 3.91 | 548 |  | 1 | 1 |  | 47.3 | 47.3 | 47.3 |  |  |  |  | 117.13 | PASS | 1 |  | 548 | 548 | 548 |  |  |  |  | 105.08 | PASS |  | 161
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | SPK s110 | 166 | 1 | 1 | TP | ug P/L | 0.132 | 2.25 | 61.5 | TN | ug N/L | 0.651 | 9.44 | 1610 |  | 1 | 1 |  | 61.5 | 61.5 | 61.5 |  |  |  |  | 107.35 | PASS | 1 |  | 1610 | 1610 | 1610 |  |  |  |  | 132.1 | FAIL |  | 162
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | SPK 120 | 178 | 1 | 1 | TP | ug P/L | 0.11 | 1.91 | 51.3 | TN | ug N/L | 0.372 | 5.18 | 772 |  | 1 | 1 |  | 51.3 | 51.3 | 51.3 |  |  |  |  | 105.93 | PASS | 1 |  | 772 | 772 | 772 |  |  |  |  | 107.72 | PASS |  | 163
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | SPK s130 | 190 | 1 | 1 | TP | ug P/L | 0.121 | 2.07 | 56.2 | TN | ug N/L | 0.388 | 5.38 | 808 |  | 1 | 1 |  | 56.2 | 56.2 | 56.2 |  |  |  |  | 110.02 | PASS | 1 |  | 808 | 808 | 808 |  |  |  |  | 132.08 | FAIL |  | 164
[27 empty] | PO4-P- PR of spikes | [10 empty] | NO3-N- PR of spikes
[25 empty] | Mean |  | 100.8646 | [10 empty] | 92.4239 |  | Min | 1
[25 empty] | Std |  | 22.8585 | [10 empty] | 27.3514 |  | Max | 164
[25 empty] | CV |  | 22.6626 | [10 empty] | 29.5934 |  | Mean
[25 empty] | S/N |  | 4.4126 | [10 empty] | 3.3791 |  | Whole series
[24 empty] | Warning Limits | Xbar + 2SD |  | 146.5817 | [10 empty] | 147.1268
[25 empty] | Xbar - 2SD |  | 55.1476 | [10 empty] | 37.7211
[24 empty] | Control Limits | Xbar + 3SD |  | 169.4402 | [10 empty] | 174.4782
[25 empty] | Xbar - 3SD |  | 32.2891 | [10 empty] | 10.3697
Count
99th percentile.inc
Maximum
Known
MDL
LOQ
ratio mean/mdl

## rolling dup RPD
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 043 | 43 | 1 | 1 | TP | ug P/L | 0.0646 | 1.46 | 28.1 | TN | ug N/L | 0.565 | 9.58 | 29.8 |  | 1 | 1 |  | 28.1 | 28.1 | 28.1 |  |  | 1.7953 | PASS |  |  | 2 | using modeled decline in efficiency | 381.7128 | 381.7128 | 346.237 |  |  | 8.5056 | PASS |  |  |  | 1
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 055 | 55 | 1 | 1 | TP | ug P/L | 0.0666 | 1.46 | 28.1 | TN | ug N/L | 0.466 | 7.85 | 14.9 |  | 1 | 1 |  | 28.1 | 28.1 | 28.1 |  |  | 16.6395 | PASS |  |  | 2 | using modeled decline in efficiency | 207.3709 | 207.3709 | 198.9491 |  |  | 19.969 | PASS |  |  |  | 2
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 067 | 67 | 1 | 1 | TP | ug P/L | 0.0722 | 1.61 | 32.6 | TN | ug N/L | 0.856 | 14.5 | 72.5 |  | 1 | 1 |  | 32.6 | 32.6 | 32.6 |  |  | 4.2042 | PASS |  |  | 2 | using modeled decline in efficiency | 855.8403 | 855.8403 | 839.2601 |  |  | 2.2851 | PASS |  |  |  | 3
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 079 | 79 | 1 | 1 | TP | ug P/L | 0.0495 | 1.17 | 20 | TN | ug N/L | 0.527 | 8.93 | 24.2 |  | 1 | 1 |  | 20 | 20 | 20 |  |  | 5.1282 | PASS |  |  | 2 | using modeled decline in efficiency | 316.674 | 316.674 | 333.3864 |  |  | 13.7071 | PASS |  |  |  | 4
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 091 | 91 | 1 | 1 | TP | ug P/L | 0.0628 | 1.38 | 25.9 | TN | ug N/L | 0.624 | 10.6 | 38.7 |  | 1 | 1 |  | 25.9 | 25.9 | 25.9 |  |  | 4.1588 | PASS |  |  | 2 | using modeled decline in efficiency | 482.6446 | 482.6446 | 511.4898 |  |  | 7.8651 | PASS |  |  |  | 5
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 103 | 103 | 1 | 1 | TP | ug P/L | 0.0488 | 1.12 | 18.7 | TN | ug N/L | 0.503 | 8.52 | 20.7 |  | 1 | 1 |  | 18.7 | 18.7 | 18.7 |  |  | 0.53619 | PASS |  |  | 2 | using modeled decline in efficiency | 275.3614 | 275.3614 | 320.3854 |  |  | 3.9719 | PASS |  |  |  | 6
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 115 | 115 | 1 | 1 | TP | ug P/L | 0.0566 | 1.29 | 23.5 | TN | ug N/L | 0.465 | 8 | 16.2 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 |  |  | 3.0238 | PASS |  |  | 2 | using modeled decline in efficiency | 222.6444 | 222.6444 | 281.6251 |  |  | 1.6705 | PASS |  |  |  | 7
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 127 | 127 | 1 | 1 | TP | ug P/L | 0.0733 | 1.59 | 31.9 | TN | ug N/L | 0.82 | 14.1 | 69 |  | 1 | 1 |  | 31.9 | 31.9 | 31.9 |  |  | 0 | PASS |  |  | 2 | using modeled decline in efficiency | 818.4919 | 818.4919 | 927.8125 |  |  | 1.1337 | PASS |  |  |  | 8
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 139 | 139 | 1 | 1 | TP | ug P/L | 0.0471 | 1.1 | 18 | TN | ug N/L | 0.405 | 6.93 | 6.99 |  | 1 | 1 |  | 18 | 18 | 18 |  |  | 1.6529 | PASS |  |  | 2 | using modeled decline in efficiency | 113.0412 | 113.0412 | 198.8718 |  |  | 9.7193 | PASS |  |  |  | 9
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 45 | 1 | 1 | TP | ug P/L | 0.0632 | 1.48 | 32.1 | TN | ug N/L | 0.858 | 14.7 | 943 |  | 1 | 1 |  | 32.1 | 32.1 | 32.1 |  |  | 3.1646 | PASS |  |  | 1 |  | 943 | 943 | 943 |  |  | 5.7673 | PASS |  |  |  | 10
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 57 | 1 | 1 | TP | ug P/L | 0.0738 | 1.65 | 36.9 | TN | ug N/L | 0.881 | 15 | 967 |  | 1 | 1 |  | 36.9 | 36.9 | 36.9 |  |  | 5.0198 | PASS |  |  | 1 |  | 967 | 967 | 967 |  |  | 5.6353 | PASS |  |  |  | 11
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 69 | 1 | 1 | TP | ug P/L | 0.0712 | 1.6 | 35.5 | TN | ug N/L | 1.22 | 20.6 | 1410 |  | 1 | 1 |  | 35.5 | 35.5 | 35.5 |  |  | 1.9915 | PASS |  |  | 1 |  | 1410 | 1410 | 1410 |  |  | 12.0301 | PASS |  |  |  | 12
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 81 | 1 | 1 | TP | ug P/L | 0.0674 | 1.49 | 32.4 | TN | ug N/L | 1.24 | 20.9 | 1440 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 |  |  | 3.0395 | PASS |  |  | 1 |  | 1440 | 1440 | 1440 |  |  | 11.7647 | PASS |  |  |  | 13
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 93 | 1 | 1 | TP | ug P/L | 0.065 | 1.5 | 32.6 | TN | ug N/L | 0.597 | 10.2 | 597 |  | 1 | 1 |  | 32.6 | 32.6 | 32.6 |  |  | 7.6433 | PASS |  |  | 1 |  | 597 | 597 | 597 |  |  | 2.8887 | PASS |  |  |  | 14
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 105 | 1 | 1 | TP | ug P/L | 0.0967 | 2.08 | 49.5 | TN | ug N/L | 0.408 | 7 | 357 |  | 1 | 1 |  | 49.5 | 49.5 | 49.5 |  |  | 11.9914 | PASS |  |  | 1 |  | 357 | 357 | 357 |  |  | 10.3093 | PASS |  |  |  | 15
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 117 | 1 | 1 | TP | ug P/L | 0.0505 | 1.12 | 21.2 | TN | ug N/L | 0.275 | 4.71 | 188 |  | 1 | 1 |  | 21.2 | 21.2 | 21.2 |  |  | 2.7907 | PASS |  |  | 1 |  | 188 | 188 | 188 |  |  | 11.0553 | PASS |  |  |  | 16
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 129 | 1 | 1 | TP | ug P/L | 0.0779 | 1.66 | 37.4 | TN | ug N/L | 0.766 | 13 | 812 |  | 1 | 1 |  | 37.4 | 37.4 | 37.4 |  |  | 0.80537 | PASS |  |  | 1 |  | 812 | 812 | 812 |  |  | 0.74166 | PASS |  |  |  | 17
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 141 | 1 | 1 | TP | ug P/L | 0.0453 | 1.1 | 20.7 | TN | ug N/L | 0.336 | 5.79 | 267 |  | 1 | 1 |  | 20.7 | 20.7 | 20.7 |  |  | 17.2185 | PASS |  |  | 1 |  | 267 | 267 | 267 |  |  | 1.8553 | PASS |  |  |  | 18
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 153 | 1 | 1 | TP | ug P/L | 0.0722 | 1.67 | 37.6 | TN | ug N/L | 0.514 | 8.77 | 489 |  | 1 | 1 |  | 37.6 | 37.6 | 37.6 |  |  | 7.7348 | PASS |  |  | 1 |  | 489 | 489 | 489 |  |  | 5.7597 | PASS |  |  |  | 19
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 45 | 1 | 1 | TP | ug P/L | 0.0485 | 1.09 | 13.3 | TN | ug N/L | 0.167 | 3.05 | 178 |  | 1 | 2 | reworked calibration | 19.5346 | 19.5346 | 19.5346 |  |  | 0 | PASS |  |  | 1 |  | 187.6427 | 187.6427 | 187.6427 |  |  | 2.0185 | PASS |  |  |  | 20
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 57 | 1 | 1 | TP | ug P/L | 0.0487 | 1.12 | 13.9 | TN | ug N/L | 0.176 | 3.22 | 192 |  | 1 | 2 | reworked calibration | 20.427 | 20.427 | 20.427 |  |  | 0 | PASS |  |  | 1 |  | 203.6024 | 203.6024 | 203.6024 |  |  | 1.3943 | PASS |  |  |  | 21
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 69 | 1 | 1 | TP | ug P/L | 0.0334 | 0.713 | 3.3 | TN | ug N/L | 0.154 | 2.86 | 160 |  | 1 | 2 | reworked calibration | 8.3765 | 8.3765 | 8.3765 |  |  | 4.3103 | PASS |  |  | 1 |  | 169.8505 | 169.8505 | 169.8505 |  |  | 17.3138 | PASS |  |  |  | 22
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 81 | 1 | 1 | TP | ug P/L | 0.0648 | 1.41 | 21.9 | TN | ug N/L | 0.25 | 4.51 | 307 |  | 1 | 2 | reworked calibration | 29.0883 | 29.0883 | 29.0883 |  |  | 8.5891 | PASS |  |  | 1 |  | 325.9513 | 325.9513 | 325.9513 |  |  | 15.1065 | PASS |  |  |  | 23
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 93 | 1 | 1 | TP | ug P/L | 0.0341 | 0.763 | 4.63 | TN | ug N/L | 0.158 | 2.86 | 160 |  | 1 | 2 | reworked calibration | 9.8503 | 9.8503 | 9.8503 |  |  | 3.6583 | PASS |  |  | 1 |  | 169.8505 | 169.8505 | 169.8505 |  |  | 1.6652 | PASS |  |  |  | 24
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 105 | 1 | 1 | TP | ug P/L | 0.0733 | 1.53 | 25.1 | TN | ug N/L | 1.21 | 22.2 | 1800 |  | 1 | 2 | reworked calibration | 32.6905 | 32.6905 | 32.6905 |  |  | 18.4182 | PASS |  |  | 1 |  | 2225.3094 | 2225.3094 | 2225.3094 |  |  | 0.53609 | PASS |  |  |  | 25
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 117 | 1 | 1 | TP | ug P/L | 0.0651 | 1.46 | 23.3 | TN | ug N/L | 0.22 | 4.08 | 269 |  | 1 | 2 | reworked calibration | 30.5879 | 30.5879 | 30.5879 |  |  | 2.9862 | PASS |  |  | 1 |  | 284.9243 | 284.9243 | 284.9243 |  |  | 3.393 | PASS |  |  |  | 26
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 129 | 1 | 1 | TP | ug P/L | 0.0687 | 1.52 | 24.9 | TN | ug N/L | 0.459 | 8.44 | 651 |  | 1 | 2 | reworked calibration | 32.3899 | 32.3899 | 32.3899 |  |  | 10.5626 | PASS |  |  | 1 |  | 712.2263 | 712.2263 | 712.2263 |  |  | 15.4123 | PASS |  |  |  | 27
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 141 | 1 | 1 | TP | ug P/L | 0.0364 | 0.804 | 5.69 | TN | ug N/L | 0.188 | 3.46 | 214 |  | 1 | 2 | reworked calibration | 11.0602 | 11.0602 | 11.0602 |  |  | 4.9222 | PASS |  |  | 1 |  | 226.1987 | 226.1987 | 226.1987 |  |  | 45.8326 | FAIL |  |  |  | 28
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 153 | 1 | 1 | TP | ug P/L | 0.0784 | 1.77 | 31.7 | TN | ug N/L | 0.293 | 5.47 | 391 |  | 1 | 2 | reworked calibration | 39.9269 | 39.9269 | 39.9269 |  |  | 37.5961 | FAIL |  |  | 1 |  | 418.4271 | 418.4271 | 418.4271 |  |  | 48.3522 | FAIL |  |  |  | 29
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 45 | 1 | 1 | TP | ug P/L | 0.0538 | 1.2 | 28.7 | TN | ug N/L | 0.804 | 11.4 | 700 |  | 1 | 1 |  | 28.7 | 28.7 | 28.7 |  |  | 2.4096 | PASS |  |  | 1 |  | 700 | 700 | 700 |  |  | 4.0816 | PASS |  |  |  | 30
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 57 | 1 | 1 | TP | ug P/L | 0.0342 | 0.745 | 14.2 | TN | ug N/L | 0.239 | 3.56 | 162 |  | 1 | 1 |  | 14.2 | 14.2 | 14.2 |  |  | 6.1433 | PASS |  |  | 1 |  | 162 | 162 | 162 |  |  | 8.8496 | PASS |  |  |  | 31
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 69 | 1 | 1 | TP | ug P/L | 0.0518 | 1.12 | 26.1 | TN | ug N/L | 0.464 | 6.67 | 379 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 |  |  | 1.5209 | PASS |  |  | 1 |  | 379 | 379 | 379 |  |  | 0.7947 | PASS |  |  |  | 32
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 81 | 1 | 1 | TP | ug P/L | 0.073 | 1.6 | 41.1 | TN | ug N/L | 0.908 | 12.9 | 796 |  | 1 | 1 |  | 41.1 | 41.1 | 41.1 |  |  | 6.015 | PASS |  |  | 1 |  | 796 | 796 | 796 |  |  | 2.6735 | PASS |  |  |  | 33
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 93 | 1 | 1 | TP | ug P/L | 0.0705 | 1.5 | 38 | TN | ug N/L | 1.05 | 14.8 | 918 |  | 1 | 1 |  | 38 | 38 | 38 |  |  | 3.871 | PASS |  |  | 1 |  | 918 | 918 | 918 |  |  | 11.0345 | PASS |  |  |  | 34
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 105 | 1 | 1 | TP | ug P/L | 0.103 | 2.13 | 57.5 | TN | ug N/L | 0.933 | 13.3 | 820 |  | 1 | 1 |  | 57.5 | 57.5 | 57.5 |  |  | 0.5231 | PASS |  |  | 1 |  | 820 | 820 | 820 |  |  | 15.3644 | PASS |  |  |  | 35
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 117 | 1 | 1 | TP | ug P/L | 0.0869 | 1.83 | 48.3 | TN | ug N/L | 0.36 | 5.24 | 280 |  | 1 | 1 |  | 48.3 | 48.3 | 48.3 |  |  | 5.0454 | PASS |  |  | 1 |  | 280 | 280 | 280 |  |  | 0.35778 | PASS |  |  |  | 36
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 129 | 1 | 1 | TP | ug P/L | 0.0277 | 0.588 | 9.21 | TN | ug N/L | 0.283 | 4.13 | 203 |  | 1 | 1 |  | 9.21 | 9.21 | 9.21 |  |  | 0.32626 | PASS |  |  | 1 |  | 203 | 203 | 203 |  |  | 6.6158 | PASS |  |  |  | 37
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 141 | 1 | 1 | TP | ug P/L | 0.0399 | 0.944 | 20.6 | TN | ug N/L | 0.378 | 5.5 | 298 |  | 1 | 1 |  | 20.6 | 20.6 | 20.6 |  |  | 61.5873 | FAIL |  |  | 1 |  | 298 | 298 | 298 |  |  | 4.4597 | PASS |  |  |  | 38
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 153 | 1 | 1 | TP | ug P/L | 0.0457 | 1.04 | 23.8 | TN | ug N/L | 0.368 | 5.31 | 285 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 |  |  | 7.4074 | PASS |  |  | 1 |  | 285 | 285 | 285 |  |  | 15.235 | PASS |  |  |  | 39
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 165 | 1 | 1 | TP | ug P/L | 0.104 | 1.97 | 52.7 | TN | ug N/L | 0.388 | 5.59 | 305 |  | 1 | 1 |  | 52.7 | 52.7 | 52.7 |  |  | 2.8874 | PASS |  |  | 1 |  | 305 | 305 | 305 |  |  | 25.0923 | FAIL |  |  |  | 40
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP F 17aug20 0.1 | 177 | 1 | 1 | TP | ug P/L | 0.0642 | 1.32 | 32.5 | TN | ug N/L | 0.752 | 10.7 | 649 |  | 1 | 1 |  | 32.5 | 32.5 | 32.5 |  |  | 3.4429 | PASS |  |  | 1 |  | 649 | 649 | 649 |  |  | 8.2718 | PASS |  |  |  | 41
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | DUP | 45 | 1 | 1 | TP | ug P/L | 0.0542 | 1.12 | 21.8 | TN | ug N/L | 0.345 | 5.15 | 262 |  | 1 | 1 |  | 21.8 | 21.8 | 21.8 |  |  | 0.45977 | PASS |  |  | 1 |  | 262 | 262 | 262 |  |  | 39.8169 | FAIL |  |  |  | 42
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | DUP | 57 | 1 | 1 | TP | ug P/L | 0.0448 | 0.957 | 17.4 | TN | ug N/L | 0.248 | 3.7 | 170 |  | 1 | 1 |  | 17.4 | 17.4 | 17.4 |  |  | 3.3898 | PASS |  |  | 1 |  | 170 | 170 | 170 |  |  | 2.3256 | PASS |  |  |  | 43
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | DUP | 69 | 1 | 1 | TP | ug P/L | 0.0424 | 0.912 | 16.1 | TN | ug N/L | 0.256 | 3.81 | 177 |  | 1 | 1 |  | 16.1 | 16.1 | 16.1 |  |  | 1.25 | PASS |  |  | 1 |  | 177 | 177 | 177 |  |  | 0.56338 | PASS |  |  |  | 44
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 81 | 1 | 1 | TP | ug P/L | 0.0484 | 1.05 | 19.9 | TN | ug N/L | 0.28 | 4.17 | 200 |  | 1 | 1 |  | 19.9 | 19.9 | 19.9 |  |  | 1.0101 | PASS |  |  | 1 |  | 200 | 200 | 200 |  |  | 3.9216 | PASS |  |  |  | 45
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 93 | 1 | 1 | TP | ug P/L | 0.0397 | 0.868 | 14.9 | TN | ug N/L | 0.248 | 3.72 | 172 |  | 1 | 1 |  | 14.9 | 14.9 | 14.9 |  |  | 5.8632 | PASS |  |  | 1 |  | 172 | 172 | 172 |  |  | 3.9886 | PASS |  |  |  | 46
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 105 | 1 | 1 | TP | ug P/L | 0.0508 | 1.08 | 20.8 | TN | ug N/L | 0.273 | 4.08 | 194 |  | 1 | 1 |  | 20.8 | 20.8 | 20.8 |  |  | 2.3753 | PASS |  |  | 1 |  | 194 | 194 | 194 |  |  | 3.1414 | PASS |  |  |  | 47
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 117 | 1 | 1 | TP | ug P/L | 0.154 | 2.91 | 71.5 | TN | ug N/L | 0.631 | 9.12 | 517 |  | 1 | 1 |  | 71.5 | 71.5 | 71.5 |  |  | 2.4062 | PASS |  |  | 1 |  | 517 | 517 | 517 |  |  | 15.6413 | PASS |  |  |  | 48
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 129 | 1 | 1 | TP | ug P/L | 0.152 | 2.88 | 70.6 | TN | ug N/L | 0.531 | 7.74 | 428 |  | 1 | 1 |  | 70.6 | 70.6 | 70.6 |  |  | 10.9772 | PASS |  |  | 1 |  | 428 | 428 | 428 |  |  | 2.5375 | PASS |  |  |  | 49
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 141 | 1 | 1 | TP | ug P/L | 0.0638 | 1.34 | 27.9 | TN | ug N/L | 0.324 | 4.8 | 240 |  | 1 | 1 |  | 27.9 | 27.9 | 27.9 |  |  | 13.3779 | PASS |  |  | 1 |  | 240 | 240 | 240 |  |  | 17.4905 | PASS |  |  |  | 50
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 153 | 1 | 1 | TP | ug P/L | 0.0316 | 0.73 | 11 | TN | ug N/L | 0.315 | 4.67 | 232 |  | 1 | 1 |  | 11 | 11 | 11 |  |  | 11.6402 | PASS |  |  | 1 |  | 232 | 232 | 232 |  |  | 0.8658 | PASS |  |  |  | 51
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 045 | 45 | 1 | 1 | TP | ug P/L | 0.0778 | 1.53 | 36.3 | TN | ug N/L | 1.35 | 18.1 | 1150 |  | 1 | 1 |  | 36.3 | 36.3 | 36.3 |  |  | 17.0404 | PASS |  |  | 1 |  | 1150 | 1150 | 1150 |  |  | 10.0457 | PASS |  |  |  | 52
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 057 | 57 | 1 | 1 | TP | ug P/L | 0.0314 | 0.654 | 11.3 | TN | ug N/L | 0.308 | 4.28 | 206 |  | 1 | 1 |  | 11.3 | 11.3 | 11.3 |  |  | 4.329 | PASS |  |  | 1 |  | 206 | 206 | 206 |  |  | 2.3981 | PASS |  |  |  | 53
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 069 | 69 | 1 | 1 | TP | ug P/L | 0.0802 | 1.46 | 34.3 | TN | ug N/L | 0.633 | 8.63 | 492 |  | 1 | 1 |  | 34.3 | 34.3 | 34.3 |  |  | 1.173 | PASS |  |  | 1 |  | 492 | 492 | 492 |  |  | 1.8462 | PASS |  |  |  | 54
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 081 | 81 | 1 | 1 | TP | ug P/L | 0.0764 | 1.4 | 32.5 | TN | ug N/L | 0.675 | 9.14 | 526 |  | 1 | 1 |  | 32.5 | 32.5 | 32.5 |  |  | 9.0032 | PASS |  |  | 1 |  | 526 | 526 | 526 |  |  | 11.245 | PASS |  |  |  | 55
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 105 | 105 | 1 | 1 | TP | ug P/L | 0.101 | 1.79 | 43.7 | TN | ug N/L | 0.792 | 10.8 | 639 |  | 1 | 1 |  | 43.7 | 43.7 | 43.7 |  |  | 3.7296 | PASS |  |  | 1 |  | 639 | 639 | 639 |  |  | 18.275 | PASS |  |  |  | 56
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 117 | 117 | 1 | 1 | TP | ug P/L | 0.0687 | 1.27 | 28.9 | TN | ug N/L | 0.515 | 7.07 | 388 |  | 1 | 1 |  | 28.9 | 28.9 | 28.9 |  |  | 2.3932 | PASS |  |  | 1 |  | 388 | 388 | 388 |  |  | 7.9208 | PASS |  |  |  | 57
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 129 | 129 | 1 | 1 | TP | ug P/L | 0.0473 | 0.944 | 19.5 | TN | ug N/L | 0.301 | 4.19 | 200 |  | 1 | 1 |  | 19.5 | 19.5 | 19.5 |  |  | 5.8047 | PASS |  |  | 1 |  | 200 | 200 | 200 |  |  | 9.9738 | PASS |  |  |  | 58
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 141 | 141 | 1 | 1 | TP | ug P/L | 0.0295 | 0.607 | 9.98 | TN | ug N/L | 0.228 | 3.2 | 136 |  | 1 | 1 |  | 9.98 | 9.98 | 9.98 |  |  | 0.90589 | PASS |  |  | 1 |  | 136 | 136 | 136 |  |  | 1.4599 | PASS |  |  |  | 59
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 153 | 153 | 1 | 1 | TP | ug P/L | 0.0587 | 1.11 | 24.3 | TN | ug N/L | 0.362 | 5 | 253 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 |  |  | 3.3473 | PASS |  |  | 1 |  | 253 | 253 | 253 |  |  | 0.79365 | PASS |  |  |  | 60
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample45 | 45 | 1 | 1 | TP | ug P/L | 0.056 | 1.07 | 18.3 | TN | ug N/L | 0.477 | 6.48 | 347 |  | 1 | 1 |  | 18.3 | 18.3 | 18.3 |  |  | 15.6171 | PASS |  |  | 1 |  | 347 | 347 | 347 |  |  | 3.3994 | PASS |  |  |  | 61
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample57 | 57 | 1 | 1 | TP | ug P/L | 0.0864 | 1.61 | 31.5 | TN | ug N/L | 1.21 | 16 | 984 |  | 1 | 1 |  | 31.5 | 31.5 | 31.5 |  |  | 3.2258 | PASS |  |  | 1 |  | 984 | 984 | 984 |  |  | 1.4127 | PASS |  |  |  | 62
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample69 | 69 | 1 | 1 | TP | ug P/L | 0.112 | 2.03 | 42 | TN | ug N/L | 0.248 | 3.45 | 143 |  | 1 | 1 |  | 42 | 42 | 42 |  |  | 15.1088 | PASS |  |  | 1 |  | 143 | 143 | 143 |  |  | 40.5383 | FAIL |  |  |  | 63
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample81 | 81 | 1 | 1 | TP | ug P/L | 0.0882 | 1.57 | 30.6 | TN | ug N/L | 0.793 | 10.5 | 619 |  | 1 | 1 |  | 30.6 | 30.6 | 30.6 |  |  | 11.0345 | PASS |  |  | 1 |  | 619 | 619 | 619 |  |  | 1.2841 | PASS |  |  |  | 64
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample93 | 93 | 1 | 1 | TP | ug P/L | 0.0883 | 1.58 | 30.7 | TN | ug N/L | 1.41 | 18.7 | 1170 |  | 1 | 1 |  | 30.7 | 30.7 | 30.7 |  |  | 3.8339 | PASS |  |  | 1 |  | 1170 | 1170 | 1170 |  |  | 13.6986 | PASS |  |  |  | 65
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample105 | 105 | 1 | 1 | TP | ug P/L | 0.0459 | 0.9 | 14.1 | TN | ug N/L | 0.252 | 3.48 | 146 |  | 1 | 1 |  | 14.1 | 14.1 | 14.1 |  |  | 10.101 | PASS |  |  | 1 |  | 146 | 146 | 146 |  |  | 0.68259 | PASS |  |  |  | 66
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample117 | 117 | 1 | 1 | TP | ug P/L | 0.0803 | 1.5 | 28.9 | TN | ug N/L | 0.329 | 4.5 | 214 |  | 1 | 1 |  | 28.9 | 28.9 | 28.9 |  |  | 3.4014 | PASS |  |  | 1 |  | 214 | 214 | 214 |  |  | 3.2184 | PASS |  |  |  | 67
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample129 | 129 | 1 | 1 | TP | ug P/L | 0.127 | 2.26 | 47.7 | TN | ug N/L | 0.592 | 7.97 | 447 |  | 1 | 1 |  | 47.7 | 47.7 | 47.7 |  |  | 1.0428 | PASS |  |  | 1 |  | 447 | 447 | 447 |  |  | 6.2839 | PASS |  |  |  | 68
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample141 | 141 | 1 | 1 | TP | ug P/L | 0.101 | 1.81 | 36.6 | TN | ug N/L | 1.01 | 13.7 | 829 |  | 1 | 1 |  | 36.6 | 36.6 | 36.6 |  |  | 2.1622 | PASS |  |  | 1 |  | 829 | 829 | 829 |  |  | 11.4796 | PASS |  |  |  | 69
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample153 | 153 | 1 | 1 | TP | ug P/L | 0.0525 | 0.975 | 16 | TN | ug N/L | 0.662 | 9.02 | 517 |  | 1 | 1 |  | 16 | 16 | 16 |  |  | 25.6131 | FAIL |  |  | 1 |  | 517 | 517 | 517 |  |  | 4.3478 | PASS |  |  |  | 70
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 045 | 45 | 1 | 1 | TP | ug P/L | 0.0553 | 1.05 | 20.3 | TN | ug N/L | 0.551 | 7.61 | 462 |  | 1 | 1 |  | 20.3 | 20.3 | 20.3 |  |  | 4.534 | PASS |  |  | 1 |  | 462 | 462 | 462 |  |  | 3.5242 | PASS |  |  |  | 71
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 057 | 57 | 1 | 1 | TP | ug P/L | 0.0639 | 1.17 | 23.9 | TN | ug N/L | 2.16 | 29.6 | 1960 |  | 1 | 1 |  | 23.9 | 23.9 | 23.9 |  |  | 17.9048 | PASS |  |  | 2 | oor | 1960 | 1960 | 1960 |  |  | 0.51151 | PASS |  |  |  | 72
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 069 | 69 | 1 | 1 | TP | ug P/L | 0.0561 | 1.01 | 19.4 | TN | ug N/L | 0.408 | 5.64 | 324 |  | 1 | 1 |  | 19.4 | 19.4 | 19.4 |  |  | 1.5584 | PASS |  |  | 1 |  | 324 | 324 | 324 |  |  | 7.7151 | PASS |  |  |  | 73
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 081 | 81 | 1 | 1 | TP | ug P/L | 0.0225 | 0.466 | 3.99 | TN | ug N/L | 0.247 | 3.48 | 171 |  | 1 | 1 |  | 3.99 | 3.99 | 3.99 |  |  | 2.5381 | PASS |  |  | 1 |  | 171 | 171 | 171 |  |  | 4.7904 | PASS |  |  |  | 74
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 093 | 93 | 1 | 1 | TP | ug P/L | 0.0567 | 1.04 | 20.3 | TN | ug N/L | 0.411 | 5.74 | 330 |  | 1 | 1 |  | 20.3 | 20.3 | 20.3 |  |  | 2.4331 | PASS |  |  | 1 |  | 330 | 330 | 330 |  |  | 0.30349 | PASS |  |  |  | 75
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 105 | 105 | 1 | 1 | TP | ug P/L | 0.066 | 1.22 | 25.2 | TN | ug N/L | 1.55 | 21.3 | 1400 |  | 1 | 1 |  | 25.2 | 25.2 | 25.2 |  |  | 4.6512 | PASS |  |  | 1 |  | 1400 | 1400 | 1400 |  |  | 4.1958 | PASS |  |  |  | 76
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 117 | 117 | 1 | 1 | TP | ug P/L | 0.199 | 3.45 | 86.9 | TN | ug N/L | 0.433 | 6.04 | 352 |  | 1 | 1 |  | 86.9 | 86.9 | 86.9 |  |  | 42.6437 | FAIL |  |  | 1 |  | 352 | 352 | 352 |  |  | 59.3407 | FAIL |  |  |  | 77
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 129 | 129 | 1 | 1 | TP | ug P/L | 0.0668 | 1.2 | 24.6 | TN | ug N/L | 0.103 | 1.47 | 28.8 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 |  |  | 9.3023 | PASS |  |  | 1 |  | 28.8 | 28.8 | 28.8 |  |  | 75.4595 | FAIL |  |  |  | 78
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 141 | 141 | 1 | 1 | TP | ug P/L | 0.0547 | 0.998 | 19 | TN | ug N/L | 0.32 | 4.53 | 245 |  | 1 | 1 |  | 19 | 19 | 19 |  |  | 3.6176 | PASS |  |  | 1 |  | 245 | 245 | 245 |  |  | 12.1212 | PASS |  |  |  | 79
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 153 | 153 | 1 | 1 | TP | ug P/L | 0.0509 | 0.979 | 18.4 | TN | ug N/L | 0.519 | 7.17 | 431 |  | 1 | 1 |  | 18.4 | 18.4 | 18.4 |  |  | 1.6438 | PASS |  |  | 1 |  | 431 | 431 | 431 |  |  | 0.69849 | PASS |  |  |  | 80
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 045 | 45 | 1 | 1 | TP | ug P/L | 0.0775 | 1.32 | 26.9 | TN | ug N/L | 0.479 | 5.92 | 333 |  | 1 | 1 |  | 26.9 | 26.9 | 26.9 |  |  | 0 | PASS |  |  | 1 |  | 333 | 333 | 333 |  |  | 1.7857 | PASS |  |  |  | 81
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 057 | 57 | 1 | 1 | TP | ug P/L | 0.0728 | 1.25 | 24.9 | TN | ug N/L | 0.448 | 5.55 | 308 |  | 1 | 1 |  | 24.9 | 24.9 | 24.9 |  |  | 9.2437 | PASS |  |  | 1 |  | 308 | 308 | 308 |  |  | 2.2989 | PASS |  |  |  | 82
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 069 | 69 | 1 | 1 | TP | ug P/L | 0.0646 | 1.11 | 21 | TN | ug N/L | 0.485 | 6.29 | 358 |  | 1 | 1 |  | 21 | 21 | 21 |  |  | 5.8824 | PASS |  |  | 1 |  | 358 | 358 | 358 |  |  | 3.6984 | PASS |  |  |  | 83
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 081 | 81 | 1 | 1 | TP | ug P/L | 0.0898 | 1.54 | 33.2 | TN | ug N/L | 1.25 | 15.3 | 977 |  | 1 | 1 |  | 33.2 | 33.2 | 33.2 |  |  | 6.5319 | PASS |  |  | 1 |  | 977 | 977 | 977 |  |  | 1.1323 | PASS |  |  |  | 84
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 093 | 93 | 1 | 1 | TP | ug P/L | 0.0882 | 1.47 | 31.2 | TN | ug N/L | 0.455 | 5.72 | 319 |  | 1 | 1 |  | 31.2 | 31.2 | 31.2 |  |  | 4.0816 | PASS |  |  | 1 |  | 319 | 319 | 319 |  |  | 1.1323 | PASS |  |  |  | 85
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 105 | 105 | 1 | 1 | TP | ug P/L | 0.063 | 1.06 | 19.4 | TN | ug N/L | 0.448 | 5.73 | 320 |  | 1 | 1 |  | 19.4 | 19.4 | 19.4 |  |  | 11.6505 | PASS |  |  | 1 |  | 320 | 320 | 320 |  |  | 3.0769 | PASS |  |  |  | 86
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 117 | 117 | 1 | 1 | TP | ug P/L | 0.0853 | 1.42 | 29.9 | TN | ug N/L | 0.457 | 5.76 | 322 |  | 1 | 1 |  | 29.9 | 29.9 | 29.9 |  |  | 1.6584 | PASS |  |  | 1 |  | 322 | 322 | 322 |  |  | 0.31008 | PASS |  |  |  | 87
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 129 | 129 | 1 | 1 | TP | ug P/L | 0.0695 | 1.17 | 22.6 | TN | ug N/L | 0.427 | 5.45 | 301 |  | 1 | 1 |  | 22.6 | 22.6 | 22.6 |  |  | 5.1724 | PASS |  |  | 1 |  | 301 | 301 | 301 |  |  | 11.9718 | PASS |  |  |  | 88
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 141 | 141 | 1 | 1 | TP | ug P/L | 0.075 | 1.29 | 26.1 | TN | ug N/L | 0.482 | 6.11 | 346 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 |  |  | 1.5444 | PASS |  |  | 1 |  | 346 | 346 | 346 |  |  | 0.57971 | PASS |  |  |  | 89
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 153 | 153 | 1 | 1 | TP | ug P/L | 0.0873 | 1.48 | 31.6 | TN | ug N/L | 3.14 | 38.9 | 2650 |  | 1 | 1 |  | 31.6 | 31.6 | 31.6 |  |  | 6.7278 | PASS |  |  | 3 | oor | 2650 | 2650 | 2650 |  |  | 5.8252 | PASS |  |  |  | 90
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup5 | 45 | 1 | 1 | TP | ug P/L | 0.0882 | 1.48 | -42 | TN | ug N/L | 3.39 | 45.6 | 2190 |  | 1 | 1 |  | 41.2266 | 41.2266 | 41.2266 |  |  | 2.836 | PASS |  |  | 2 | reworked calibration | 2190 | 2190 | 2190 |  |  | 0.91743 | PASS |  |  |  | 91
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup15 | 57 | 1 | 1 | TP | ug P/L | 0.0682 | 0.797 | -43.3 | TN | ug N/L | 0.54 | 7.19 | 287 |  | 1 | 1 |  | 17.894 | 17.894 | 17.894 |  |  | 7.6005 | PASS |  |  | 2 | reworked calibration | 287 | 287 | 287 |  |  | 0.69444 | PASS |  |  |  | 92
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup25 | 69 | 1 | 1 | TP | ug P/L | 0.0993 | 1.71 | -41.3 | TN | ug N/L | 3 | 40.5 | 1930 |  | 1 | 1 |  | 52.0098 | 52.0098 | 52.0098 |  |  | 1.2006 | PASS |  |  | 2 | reworked calibration | 1930 | 1930 | 1930 |  |  | 1.0309 | PASS |  |  |  | 93
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup35 | 81 | 1 | 1 | TP | ug P/L | 0.0691 | 0.712 | -43.3 | TN | ug N/L | 0.746 | 10.1 | 430 |  | 1 | 1 |  | 19.0519 | 19.0519 | 19.0519 |  |  | 8.9301 | PASS |  |  | 2 | reworked calibration | 430 | 430 | 430 |  |  | 1.1696 | PASS |  |  |  | 94
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup45 | 93 | 1 | 1 | TP | ug P/L | 0.0702 | 0.778 | -43.2 | TN | ug N/L | 1.45 | 19.7 | 899 |  | 1 | 1 |  | 20.4532 | 20.4532 | 20.4532 |  |  | 2.514 | PASS |  |  | 2 | reworked calibration | 899 | 899 | 899 |  |  | 0.22222 | PASS |  |  |  | 95
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup55 | 105 | 1 | 1 | TP | ug P/L | 0.0701 | 0.797 | -43.2 | TN | ug N/L | 0.729 | 9.84 | 416 |  | 1 | 1 |  | 20.3264 | 20.3264 | 20.3264 |  |  | 6.0218 | PASS |  |  | 2 | reworked calibration | 416 | 416 | 416 |  |  | 0.47962 | PASS |  |  |  | 96
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup65 | 117 | 1 | 1 | TP | ug P/L | 0.0893 | 1.44 | -42 | TN | ug N/L | 4.38 | 59.6 | 2910 |  | 1 | 1 |  | 42.3642 | 42.3642 | 42.3642 |  |  | 4.7721 | PASS |  |  | 2 | reworked calibration | 2910 | 2910 | 2910 |  |  | 0.68966 | PASS |  |  |  | 97
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup75 | 129 | 1 | 1 | TP | ug P/L | 0.121 | 2.28 | -39.9 | TN | ug N/L | 1.93 | 26.2 | 1220 |  | 1 | 1 |  | 68.624 | 68.624 | 68.624 |  |  | 35.336 | FAIL |  |  | 2 | reworked calibration | 1220 | 1220 | 1220 |  |  | 15.9292 | PASS |  |  |  | 98
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup85 | 141 | 1 | 1 | TP | ug P/L | 0.0709 | 0.675 | -43.1 | TN | ug N/L | 0.503 | 6.8 | 268 |  | 1 | 1 |  | 21.337 | 21.337 | 21.337 |  |  | 35.5995 | FAIL |  |  | 2 | reworked calibration | 268 | 268 | 268 |  |  | 1.8832 | PASS |  |  |  | 99
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | dup95 | 153 | 1 | 1 | TP | ug P/L | 0.0787 | 1.1 | -42.6 | TN | ug N/L | 1.47 | 19.8 | 905 |  | 1 | 1 |  | 30.7696 | 30.7696 | 30.7696 |  |  | 7.5344 | PASS |  |  | 2 | reworked calibration | 905 | 905 | 905 |  |  | 2.6876 | PASS |  |  |  | 100
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS5 | 45 | 1 | 1 | TP | ug P/L | 0.123 | 2.12 | 17.3 | TN | ug N/L | 0.47 | 6.24 | 212 |  | 1 | 1 |  | 17.3 | 17.3 | 17.3 |  |  | 9.8901 | PASS |  |  | 1 |  | 212 | 212 | 212 |  |  | 16.0521 | PASS |  |  |  | 101
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS15 | 57 | 1 | 1 | TP | ug P/L | 0.127 | 2.2 | 20 | TN | ug N/L | 1.34 | 18.5 | 811 |  | 1 | 1 |  | 20 | 20 | 20 |  |  | 23.4637 | FAIL |  |  | 1 |  | 811 | 811 | 811 |  |  | 0.3706 | PASS |  |  |  | 102
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS25 | 69 | 1 | 1 | TP | ug P/L | 0.115 | 1.91 | 10.4 | TN | ug N/L | 0.547 | 7.63 | 279 |  | 1 | 1 |  | 10.4 | 10.4 | 10.4 |  |  | 0.95694 | PASS |  |  | 1 |  | 279 | 279 | 279 |  |  | 16.145 | PASS |  |  |  | 103
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS35 | 81 | 1 | 1 | TP | ug P/L | 0.174 | 2.97 | 45.2 | TN | ug N/L | 0.917 | 11.8 | 481 |  | 1 | 1 |  | 45.2 | 45.2 | 45.2 |  |  | 5.4545 | PASS |  |  | 1 |  | 481 | 481 | 481 |  |  | 9.6934 | PASS |  |  |  | 104
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS45 | 93 | 1 | 1 | TP | ug P/L | 0.132 | 2.28 | 22.7 | TN | ug N/L | 1.03 | 14.3 | 605 |  | 1 | 1 |  | 22.7 | 22.7 | 22.7 |  |  | 1.3304 | PASS |  |  | 1 |  | 605 | 605 | 605 |  |  | 2.0033 | PASS |  |  |  | 105
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS55 | 105 | 1 | 1 | TP | ug P/L | 0.132 | 2.25 | 21.6 | TN | ug N/L | 0.562 | 7.9 | 292 |  | 1 | 1 |  | 21.6 | 21.6 | 21.6 |  |  | 6.7114 | PASS |  |  | 1 |  | 292 | 292 | 292 |  |  | 0.68729 | PASS |  |  |  | 106
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS65 | 117 | 1 | 1 | TP | ug P/L | 0.141 | 2.41 | 26.9 | TN | ug N/L | 0.883 | 12.4 | 509 |  | 1 | 1 |  | 26.9 | 26.9 | 26.9 |  |  | 0.37106 | PASS |  |  | 1 |  | 509 | 509 | 509 |  |  | 0.78278 | PASS |  |  |  | 107
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS75 | 129 | 1 | 1 | TP | ug P/L | 0.217 | 3.65 | 68.1 | TN | ug N/L | 0.499 | 7.05 | 251 |  | 1 | 1 |  | 68.1 | 68.1 | 68.1 |  |  | 20.6715 | FAIL |  |  | 1 |  | 251 | 251 | 251 |  |  | 8.2988 | PASS |  |  |  | 108
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS85 | 141 | 1 | 1 | TP | ug P/L | 0.137 | 2.38 | 25.8 | TN | ug N/L | 0.661 | 9.13 | 351 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 |  |  | 4.5455 | PASS |  |  | 1 |  | 351 | 351 | 351 |  |  | 2.809 | PASS |  |  |  | 109
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS95 | 153 | 1 | 1 | TP | ug P/L | 0.114 | 2.01 | 13.9 | TN | ug N/L | 0.445 | 6.19 | 209 |  | 1 | 1 |  | 13.9 | 13.9 | 13.9 |  |  | 13.0268 | PASS |  |  | 1 |  | 209 | 209 | 209 |  |  | 2.4213 | PASS |  |  |  | 110
44979 | OM_2-22-2023_11-17-20AM BRN edit.omn | dupS105 | 163 | 1 | 1 | TP | ug P/L | 0.106 | 1.84 | 8.11 | TN | ug N/L | 0.387 | 5.42 | 172 |  | 1 | 1 |  | 8.11 | 8.11 | 8.11 |  |  | 9.5126 | PASS |  |  | 1 |  | 172 | 172 | 172 |  |  | 2.8653 | PASS |  |  |  | 111
45000 | OM_3-15-2023_12-40-40PM.OMN | s5 DUP | 45 | 1 | 1 | TP | ug P/L | 0.0501 | 0.935 | 17.4 | TN | ug N/L | 0.334 | 4.83 | 637 |  | 1 | 1 |  | 17.4 | 17.4 | 17.4 |  |  | 5.9172 | PASS |  |  | 1 |  | 637 | 637 | 637 |  |  | 4.6586 | PASS |  |  |  | 112
45000 | OM_3-15-2023_12-40-40PM.OMN | s15 DUP | 57 | 1 | 1 | TP | ug P/L | 0.048 | 0.931 | 17.2 | TN | ug N/L | 0.151 | 2.22 | 243 |  | 1 | 1 |  | 17.2 | 17.2 | 17.2 |  |  | 22.6537 | FAIL |  |  | 1 |  | 243 | 243 | 243 |  |  | 8.5837 | PASS |  |  |  | 113
45000 | OM_3-15-2023_12-40-40PM.OMN | s25 DUP | 69 | 1 | 1 | TP | ug P/L | 0.0591 | 1.13 | 23.5 | TN | ug N/L | 0.164 | 2.55 | 292 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 |  |  | 13.1519 | PASS |  |  | 1 |  | 292 | 292 | 292 |  |  | 14.7059 | PASS |  |  |  | 114
45000 | OM_3-15-2023_12-40-40PM.OMN | s35 DUP | 81 | 1 | 1 | TP | ug P/L | 0.0468 | 0.904 | 16.4 | TN | ug N/L | 0.318 | 4.63 | 605 |  | 1 | 1 |  | 16.4 | 16.4 | 16.4 |  |  | 4.1791 | PASS |  |  | 1 |  | 605 | 605 | 605 |  |  | 4.8387 | PASS |  |  |  | 115
45000 | OM_3-15-2023_12-40-40PM.OMN | s45 DUP | 93 | 1 | 1 | TP | ug P/L | 0.069 | 1.25 | 27.2 | TN | ug N/L | 0.488 | 7.08 | 999 |  | 1 | 1 |  | 27.2 | 27.2 | 27.2 |  |  | 0 | PASS |  |  | 1 |  | 999 | 999 | 999 |  |  | 3.9816 | PASS |  |  |  | 116
45000 | OM_3-15-2023_12-40-40PM.OMN | s55 DUP | 105 | 1 | 1 | TP | ug P/L | 0.0494 | 0.919 | 16.9 | TN | ug N/L | 0.211 | 3.15 | 381 |  | 1 | 1 |  | 16.9 | 16.9 | 16.9 |  |  | 1.1905 | PASS |  |  | 1 |  | 381 | 381 | 381 |  |  | 1.321 | PASS |  |  |  | 117
45000 | OM_3-15-2023_12-40-40PM.OMN | s65 DUP | 117 | 1 | 1 | TP | ug P/L | 0.0482 | 0.896 | 16.2 | TN | ug N/L | 0.148 | 2.27 | 250 |  | 1 | 1 |  | 16.2 | 16.2 | 16.2 |  |  | 12.459 | PASS |  |  | 1 |  | 250 | 250 | 250 |  |  | 1.9802 | PASS |  |  |  | 118
45000 | OM_3-15-2023_12-40-40PM.OMN | s75 DUP | 129 | 1 | 1 | TP | ug P/L | 0.0773 | 1.4 | 31.7 | TN | ug N/L | 0.19 | 2.74 | 319 |  | 1 | 1 |  | 31.7 | 31.7 | 31.7 |  |  | 9.9338 | PASS |  |  | 1 |  | 319 | 319 | 319 |  |  | 13.3779 | PASS |  |  |  | 119
45000 | OM_3-15-2023_12-40-40PM.OMN | s85 DUP | 141 | 1 | 1 | TP | ug P/L | 0.0622 | 1.14 | 23.7 | TN | ug N/L | 0.686 | 10.3 | 1560 |  | 1 | 1 |  | 23.7 | 23.7 | 23.7 |  |  | 24.4444 | FAIL |  |  | 1 |  | 1560 | 1560 | 1560 |  |  | 1.2903 | PASS |  |  |  | 120
45000 | OM_3-15-2023_12-40-40PM.OMN | s95 DUP | 153 | 1 | 1 | TP | ug P/L | 0.0665 | 1.22 | 26.2 | TN | ug N/L | 0.449 | 6.48 | 900 |  | 1 | 1 |  | 26.2 | 26.2 | 26.2 |  |  | 6.6421 | PASS |  |  | 1 |  | 900 | 900 | 900 |  |  | 0.99502 | PASS |  |  |  | 121
45000 | OM_3-15-2023_12-40-40PM.OMN | s105 DUP | 163 | 1 | 1 | TP | ug P/L | 0.0859 | 1.55 | 36.5 | TN | ug N/L | 0.892 | 14 | 2250 |  | 1 | 1 |  | 36.5 | 36.5 | 36.5 |  |  | 1.6304 | PASS |  |  | 1 |  | 2250 | 2250 | 2250 |  |  | 20.5882 | FAIL |  |  |  | 122
45005 | OM_3-20-2023_10-53-22AM.OMN | s5 DUP | 45 | 1 | 1 | TP | ug P/L | 0.103 | 1.74 | 42.3 | TN | ug N/L | 0.134 | 1.87 | 191 |  | 1 | 1 |  | 42.3 | 42.3 | 42.3 |  |  | 0 | PASS |  |  | 1 |  | 191 | 191 | 191 |  |  | 19.5402 | PASS |  |  |  | 123
45005 | OM_3-20-2023_10-53-22AM.OMN | dup S15 | 57 | 1 | 1 | TP | ug P/L | 0.112 | 1.93 | 48.1 | TN | ug N/L | 0.438 | 5.97 | 854 |  | 1 | 1 |  | 48.1 | 48.1 | 48.1 |  |  | 9.6934 | PASS |  |  | 1 |  | 854 | 854 | 854 |  |  | 4.3528 | PASS |  |  |  | 124
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP s25 | 69 | 1 | 1 | TP | ug P/L | 0.0312 | 0.568 | 6.48 | TN | ug N/L | 0.107 | 1.54 | 144 |  | 1 | 1 |  | 6.48 | 6.48 | 6.48 |  |  | 1.0861 | PASS |  |  | 1 |  | 144 | 144 | 144 |  |  | 4.7458 | PASS |  |  |  | 125
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP s35 | 81 | 1 | 1 | TP | ug P/L | 0.0272 | 0.472 | 3.49 | TN | ug N/L | 0.109 | 1.55 | 145 |  | 1 | 1 |  | 3.49 | 3.49 | 3.49 |  |  | 83.5696 | FAIL |  |  | 1 |  | 145 | 145 | 145 |  |  | 15.2866 | PASS |  |  |  | 126
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP s45 | 93 | 1 | 1 | TP | ug P/L | 0.0191 | 0.357 | -0.0728 | TN | ug N/L | 0.0565 | 0.833 | 45.2 |  | 1 | 1 |  | -0.0728 | -0.0728 | -0.0728 |  |  | 161.3691 | FAIL |  |  | 1 |  | 45.2 | 45.2 | 45.2 |  |  | 0 | PASS |  |  |  | 127
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP S55 | 105 | 1 | 1 | TP | ug P/L | 0.0945 | 1.65 | 39.6 | TN | ug N/L | 0.321 | 4.37 | 579 |  | 1 | 1 |  | 39.6 | 39.6 | 39.6 |  |  | 8.4643 | PASS |  |  | 1 |  | 579 | 579 | 579 |  |  | 2.0513 | PASS |  |  |  | 128
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP s65 | 117 | 1 | 1 | TP | ug P/L | 0.0889 | 1.51 | 35.3 | TN | ug N/L | 0.549 | 7.58 | 1150 |  | 1 | 1 |  | 35.3 | 35.3 | 35.3 |  |  | 4.0462 | PASS |  |  | 1 |  | 1150 | 1150 | 1150 |  |  | 2.6432 | PASS |  |  |  | 129
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP s75 | 129 | 1 | 1 | TP | ug P/L | 0.0663 | 1.2 | 26 | TN | ug N/L | 0.357 | 4.89 | 665 |  | 1 | 1 |  | 26 | 26 | 26 |  |  | 0 | PASS |  |  | 1 |  | 665 | 665 | 665 |  |  | 3.2086 | PASS |  |  |  | 130
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP s85 | 141 | 1 | 1 | TP | ug P/L | 0.174 | 2.96 | 79.5 | TN | ug N/L | 0.795 | 11.6 | 1990 |  | 1 | 1 |  | 79.5 | 79.5 | 79.5 |  |  | 1.3933 | PASS |  |  | 1 |  | 1990 | 1990 | 1990 |  |  | 21.727 | FAIL |  |  |  | 131
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP s95 | 153 | 1 | 1 | TP | ug P/L | 0.0763 | 1.28 | 28.3 | TN | ug N/L | 0.434 | 5.9 | 842 |  | 1 | 1 |  | 28.3 | 28.3 | 28.3 |  |  | 0.35398 | PASS |  |  | 1 |  | 842 | 842 | 842 |  |  | 4.1212 | PASS |  |  |  | 132
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP s105 | 165 | 1 | 1 | TP | ug P/L | 0.0262 | 0.486 | 3.92 | TN | ug N/L | 0.11 | 1.54 | 144 |  | 1 | 1 |  | 3.92 | 3.92 | 3.92 |  |  | 13.1108 | PASS |  |  | 1 |  | 144 | 144 | 144 |  |  | 3.413 | PASS |  |  |  | 133
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP s115 | 177 | 1 | 1 | TP | ug P/L | 0.116 | 2.02 | 51 | TN | ug N/L | 0.431 | 5.9 | 842 |  | 1 | 1 |  | 51 | 51 | 51 |  |  | 0.19627 | PASS |  |  | 1 |  | 842 | 842 | 842 |  |  | 2.7694 | PASS |  |  |  | 134
45005 | OM_3-20-2023_10-53-22AM.OMN | DUP s125 | 189 | 1 | 1 | TP | ug P/L | 0.061 | 1.09 | 22.6 | TN | ug N/L | 0.165 | 2.35 | 261 |  | 1 | 1 |  | 22.6 | 22.6 | 22.6 |  |  | 4.7516 | PASS |  |  | 1 |  | 261 | 261 | 261 |  |  | 3.0189 | PASS |  |  |  | 135
45007 | OM_3-22-2023_10-55-57AM.OMN | s5 DUP | 45 | 1 | 1 | TP | ug P/L | 0.0492 | 0.861 | 19.4 | TN | ug N/L | 0.252 | 3.37 | 457 |  | 1 | 1 |  | 19.4 | 19.4 | 19.4 |  |  | 5.5138 | PASS |  |  | 1 |  | 457 | 457 | 457 |  |  | 2.2124 | PASS |  |  |  | 136
45007 | OM_3-22-2023_10-55-57AM.OMN | s15 DUP | 57 | 1 | 1 | TP | ug P/L | 0.021 | 0.403 | 5.47 | TN | ug N/L | 0.108 | 1.52 | 157 |  | 1 | 1 |  | 5.47 | 5.47 | 5.47 |  |  | 20.9493 | PASS |  |  | 1 |  | 157 | 157 | 157 |  |  | 13.0952 | PASS |  |  |  | 137
45007 | OM_3-22-2023_10-55-57AM.OMN | s25 DUP | 69 | 1 | 1 | TP | ug P/L | 0.0757 | 1.31 | 33 | TN | ug N/L | 0.229 | 3.14 | 418 |  | 1 | 1 |  | 33 | 33 | 33 |  |  | 5.9282 | PASS |  |  | 1 |  | 418 | 418 | 418 |  |  | 0.96154 | PASS |  |  |  | 138
45007 | OM_3-22-2023_10-55-57AM.OMN | s35 DUP | 81 | 1 | 1 | TP | ug P/L | 0.0152 | 0.307 | 2.54 | TN | ug N/L | 0.0956 | 1.38 | 135 |  | 1 | 1 |  | 2.54 | 2.54 | 2.54 |  |  | 17.5589 | PASS |  |  | 1 |  | 135 | 135 | 135 |  |  | 0.74349 | PASS |  |  |  | 139
45007 | OM_3-22-2023_10-55-57AM.OMN | DUP s45 | 93 | 1 | 1 | TP | ug P/L | 0.0554 | 1.02 | 24.3 | TN | ug N/L | 0.204 | 2.82 | 365 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 |  |  | 4.6316 | PASS |  |  | 1 |  | 365 | 365 | 365 |  |  | 3.0598 | PASS |  |  |  | 140
45007 | OM_3-22-2023_10-55-57AM.OMN | DUP s55 | 105 | 1 | 1 | TP | ug P/L | 0.0657 | 1.13 | 27.4 | TN | ug N/L | 0.208 | 2.84 | 368 |  | 1 | 1 |  | 27.4 | 27.4 | 27.4 |  |  | 2.1661 | PASS |  |  | 1 |  | 368 | 368 | 368 |  |  | 0.27211 | PASS |  |  |  | 141
45007 | OM_3-22-2023_10-55-57AM.OMN | DUP s65 | 117 | 1 | 1 | TP | ug P/L | 0.139 | 2.4 | 66.1 | TN | ug N/L | 0.319 | 4.37 | 628 |  | 1 | 1 |  | 66.1 | 66.1 | 66.1 |  |  | 9.0119 | PASS |  |  | 1 |  | 628 | 628 | 628 |  |  | 2.0917 | PASS |  |  |  | 142
45007 | OM_3-22-2023_10-55-57AM.OMN | DUP s75 | 129 | 1 | 1 | TP | ug P/L | 0.0949 | 1.62 | 42.4 | TN | ug N/L | 0.303 | 4.13 | 586 |  | 1 | 1 |  | 42.4 | 42.4 | 42.4 |  |  | 1.6647 | PASS |  |  | 1 |  | 586 | 586 | 586 |  |  | 4.8293 | PASS |  |  |  | 143
45007 | OM_3-22-2023_10-55-57AM.OMN | DUP s85 | 141 | 1 | 1 | TP | ug P/L | 0.066 | 1.17 | 28.9 | TN | ug N/L | 0.217 | 2.99 | 393 |  | 1 | 1 |  | 28.9 | 28.9 | 28.9 |  |  | 3.7351 | PASS |  |  | 1 |  | 393 | 393 | 393 |  |  | 6.1652 | PASS |  |  |  | 144
45007 | OM_3-22-2023_10-55-57AM.OMN | DUP s95 | 153 | 1 | 1 | TP | ug P/L | 0.0539 | 0.957 | 22.3 | TN | ug N/L | 0.152 | 2.19 | 263 |  | 1 | 1 |  | 22.3 | 22.3 | 22.3 |  |  | 0 | PASS |  |  | 1 |  | 263 | 263 | 263 |  |  | 0.75758 | PASS |  |  |  | 145
45007 | OM_3-22-2023_10-55-57AM.OMN | DUP s105 | 165 | 1 | 1 | TP | ug P/L | 0.0597 | 1.05 | 25 | TN | ug N/L | 0.18 | 2.48 | 310 |  | 1 | 1 |  | 25 | 25 | 25 |  |  | 2.3715 | PASS |  |  | 1 |  | 310 | 310 | 310 |  |  | 6.5523 | PASS |  |  |  | 146
45007 | OM_3-22-2023_10-55-57AM.OMN | DUP 115 | 177 | 1 | 1 | TP | ug P/L | 0.0714 | 1.22 | 30.4 | TN | ug N/L | 0.194 | 2.74 | 352 |  | 1 | 1 |  | 30.4 | 30.4 | 30.4 |  |  | 1.6584 | PASS |  |  | 1 |  | 352 | 352 | 352 |  |  | 28.201 | FAIL |  |  |  | 147
45007 | OM_3-22-2023_10-55-57AM.OMN | DUP s125 | 189 | 1 | 1 | TP | ug P/L | 0.194 | 3.32 | 93.8 | TN | ug N/L | 0.235 | 3.27 | 439 |  | 1 | 1 |  | 93.8 | 93.8 | 93.8 |  |  | 9.2582 | PASS |  |  | 1 |  | 439 | 439 | 439 |  |  | 5.3215 | PASS |  |  |  | 148
[25 empty] | PO4-P- RPD of duplicates | [10 empty] | NO3-N- RPD of duplicates |  |  |  | OBSERVATION #
[19 empty] | Mean | [5 empty] | 8.8423 | [10 empty] | 7.8166 |  |  |  | Min | 1
[19 empty] | Std | [5 empty] | 16.6436 | [10 empty] | 11.0702 |  |  |  | Max | 148
[19 empty] | CV | [5 empty] | 188.2268 | [10 empty] | 141.6255
[19 empty] | S/N | [5 empty] | 0.53127 | [10 empty] | 0.70609
[18 empty] | Warning Limits | Xbar + 2SD | [5 empty] | 42.1295 | [10 empty] | 29.9571
[19 empty] | Xbar - 2SD | [5 empty] | -24.4449 | [10 empty] | -14.3239
[18 empty] | Control Limits | Xbar + 3SD | [5 empty] | 58.7731 | [10 empty] | 41.0273
[19 empty] | Xbar - 3SD | [5 empty] | -41.0885 | [10 empty] | -25.3942
[18 empty] | Count | [6 empty] | 148 | [10 empty] | 148
[18 empty] | 99th percentile.inc | [6 empty] | 73.2379 | [10 empty] | 54.1761
[18 empty] | Maximum | [6 empty] | 161.3691 | [10 empty] | 75.4595

## rolling CCC
check standard
continuing calibration check
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 15 | 1 | 1 | TP | ug P/L | 0.0571 | 1.3 | 23.8 | TN | ug N/L | 0.493 | 8.3 | 18.8 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | -4.8 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 253.1018 | 253.1018 | 214.3084 | -14.2767 | PASS | [5 empty] | 1
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 16 | 1 | 1 | TP | ug P/L | 0.0732 | 1.32 | 24.3 | TN | ug N/L | 0.499 | 8.41 | 19.7 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 264.2396 | 264.2396 | 233.6407 | -6.5437 | PASS | [5 empty] | 2
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 17 | 1 | 1 | TP | ug P/L | 0.0569 | 1.31 | 24 | TN | ug N/L | 0.497 | 8.46 | 20.1 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 269.297 | 269.297 | 247.9649 | -0.81405 | PASS | [5 empty] | 3
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 15 | 1 | 1 | TP | ug P/L | 0.057 | 1.29 | 23.5 | TN | ug N/L | 0.485 | 8.35 | 19.2 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 | -6 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 258.1665 | 258.1665 | 247.781 | -0.8876 | PASS | [5 empty] | 4
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 16 | 1 | 1 | TP | ug P/L | 0.0577 | 1.35 | 25 | TN | ug N/L | 0.502 | 8.46 | 20.2 |  | 1 | 1 |  | 25 | 25 | 25 | 0 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 269.297 | 269.297 | 268.3253 | 7.3301 | PASS | [5 empty] | 5
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0561 | 1.3 | 23.6 | TN | ug N/L | 0.48 | 8.13 | 17.3 |  | 1 | 1 |  | 23.6 | 23.6 | 23.6 | [6 empty] | 2 | using modeled decline in efficiency | 235.8573 | 235.8573 | 249.9837 | [7 empty] | 6
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 17 | 1 | 1 | TP | ug P/L | 0.0564 | 1.32 | 24.2 | TN | ug N/L | 0.487 | 8.34 | 19.1 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 257.1538 | 257.1538 | 284.5594 | 13.8238 | PASS | [5 empty] | 7
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 15 | 1 | 1 | TP | ug P/L | 0.0588 | 1.38 | 25.8 | TN | ug N/L | 0.461 | 7.84 | 14.8 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | 3.2 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 206.3517 | 206.3517 | 246.2969 | -1.4812 | PASS | [5 empty] | 8
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 16 | 1 | 1 | TP | ug P/L | 0.059 | 1.34 | 24.7 | TN | ug N/L | 0.466 | 7.97 | 15.9 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | -1.2 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 219.5921 | 219.5921 | 270.1384 | 8.0554 | PASS | [5 empty] | 9
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 17 | 1 | 1 | TP | ug P/L | 0.056 | 1.32 | 24.2 | TN | ug N/L | 0.461 | 7.83 | 14.8 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 205.3322 | 205.3322 | 267.1999 | 6.88 | PASS | [5 empty] | 10
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 15 | 1 | 1 | TP | ug P/L | 0.0563 | 1.23 | 21.6 | TN | ug N/L | 0.438 | 7.5 | 11.9 |  | 1 | 1 |  | 21.6 | 21.6 | 21.6 | -13.6 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 171.6169 | 171.6169 | 244.5824 | -2.167 | PASS | [5 empty] | 11
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 16 | 1 | 1 | TP | ug P/L | 0.0568 | 1.32 | 24.1 | TN | ug N/L | 0.431 | 7.37 | 10.7 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | -3.6 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 158.2954 | 158.2954 | 242.2342 | -3.1063 | PASS | [5 empty] | 12
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 17 | 1 | 1 | TP | ug P/L | 0.0568 | 1.32 | 24.2 | TN | ug N/L | 0.381 | 6.54 | 3.63 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 72.7149 | 72.7149 | 165.5676 | -33.7729 | FAIL | [5 empty] | 13
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 8 | 1 | 1 | TP | ug P/L | 0.0489 | 1.15 | 19.3 | TN | ug N/L | 0.451 | 7.79 | 14.3 |  | 1 | 1 |  | 19.3 | 19.3 | 19.3 | [6 empty] | 2 | using modeled decline in efficiency | 201.2532 | 201.2532 | 309.588 | [7 empty] | 14
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.056 | 1.25 | 25.1 | TN | ug N/L | 0.289 | 4.97 | 206 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | 0.4 | PASS |  |  |  |  | 1 |  | 206 | 206 | 206 | -17.6 | PASS | [5 empty] | 15
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.067 | 1.24 | 24.8 | TN | ug N/L | 0.335 | 5.74 | 263 |  | 1 | 1 |  | 24.8 | 24.8 | 24.8 | -0.8 | PASS |  |  |  |  | 1 |  | 263 | 263 | 263 | 5.2 | PASS | [5 empty] | 16
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0542 | 1.24 | 24.9 | TN | ug N/L | 0.31 | 5.3 | 231 |  | 1 | 1 |  | 24.9 | 24.9 | 24.9 | -0.4 | PASS |  |  |  |  | 1 |  | 231 | 231 | 231 | -7.6 | PASS | [5 empty] | 17
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0519 | 1.23 | 24.5 | TN | ug N/L | 0.316 | 5.42 | 239 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | -2 | PASS |  |  |  |  | 1 |  | 239 | 239 | 239 | -4.4 | PASS | [5 empty] | 18
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0539 | 1.25 | 25.3 | TN | ug N/L | 0.32 | 5.47 | 244 |  | 1 | 1 |  | 25.3 | 25.3 | 25.3 | 1.2 | PASS |  |  |  |  | 1 |  | 244 | 244 | 244 | -2.4 | PASS | [5 empty] | 19
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0547 | 1.23 | 24.7 | TN | ug N/L | 0.307 | 5.27 | 229 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | -1.2 | PASS |  |  |  |  | 1 |  | 229 | 229 | 229 | -8.4 | PASS | [5 empty] | 20
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0556 | 1.31 | 27 | TN | ug N/L | 0.315 | 5.44 | 241 |  | 1 | 1 |  | 27 | 27 | 27 | 8 | PASS |  |  |  |  | 1 |  | 241 | 241 | 241 | -3.6 | PASS | [5 empty] | 21
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0515 | 1.22 | 24.2 | TN | ug N/L | 0.353 | 5.97 | 280 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 280 | 280 | 280 | 12 | PASS | [5 empty] | 22
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.057 | 1.31 | 27.1 | TN | ug N/L | 0.301 | 5.16 | 220 |  | 1 | 1 |  | 27.1 | 27.1 | 27.1 | 8.4 | PASS |  |  |  |  | 1 |  | 220 | 220 | 220 | -12 | PASS | [5 empty] | 23
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0524 | 1.23 | 24.6 | TN | ug N/L | 0.312 | 5.38 | 237 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 237 | 237 | 237 | -5.2 | PASS | [5 empty] | 24
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0562 | 1.28 | 26.1 | TN | ug N/L | 0.303 | 5.26 | 228 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | 4.4 | PASS |  |  |  |  | 1 |  | 228 | 228 | 228 | -8.8 | PASS | [5 empty] | 25
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.052 | 1.19 | 23.3 | TN | ug N/L | 0.306 | 5.22 | 225 |  | 1 | 1 |  | 23.3 | 23.3 | 23.3 | -6.8 | PASS |  |  |  |  | 1 |  | 225 | 225 | 225 | -10 | PASS | [5 empty] | 26
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0551 | 1.28 | 18.3 | TN | ug N/L | 0.196 | 3.57 | 224 |  | 1 | 2 | reworked calibration | 25.1979 | 25.1979 | 25.1979 | 0.79179 | PASS |  |  |  |  | 1 |  | 236.5807 | 236.5807 | 236.5807 | -5.3677 | PASS | [5 empty] | 27
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0552 | 1.26 | 17.8 | TN | ug N/L | 0.191 | 3.49 | 216 |  | 1 | 2 | reworked calibration | 24.6005 | 24.6005 | 24.6005 | -1.5978 | PASS |  |  |  |  | 1 |  | 229.0286 | 229.0286 | 229.0286 | -8.3886 | PASS | [5 empty] | 28
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.052 | 1.19 | 16 | TN | ug N/L | 0.188 | 3.4 | 209 |  | 1 | 2 | reworked calibration | 22.512 | 22.512 | 22.512 | -9.9521 | PASS |  |  |  |  | 1 |  | 220.5425 | 220.5425 | 220.5425 | -11.783 | PASS | [5 empty] | 29
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0521 | 1.22 | 16.6 | TN | ug N/L | 0.177 | 3.25 | 195 |  | 1 | 2 | reworked calibration | 23.4066 | 23.4066 | 23.4066 | -6.3735 | PASS |  |  |  |  | 1 |  | 206.4228 | 206.4228 | 206.4228 | -17.4309 | PASS | [5 empty] | 30
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0518 | 1.18 | 15.6 | TN | ug N/L | 0.19 | 3.45 | 213 |  | 1 | 2 | reworked calibration | 22.2139 | 22.2139 | 22.2139 | -11.1444 | PASS |  |  |  |  | 1 |  | 225.2557 | 225.2557 | 225.2557 | -9.8977 | PASS | [5 empty] | 31
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0515 | 1.18 | 15.7 | TN | ug N/L | 0.188 | 3.41 | 210 |  | 1 | 2 | reworked calibration | 22.2139 | 22.2139 | 22.2139 | -11.1444 | PASS |  |  |  |  | 1 |  | 221.4849 | 221.4849 | 221.4849 | -11.4061 | PASS | [5 empty] | 32
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0495 | 1.14 | 14.5 | TN | ug N/L | 0.188 | 3.4 | 209 |  | 1 | 2 | reworked calibration | 21.0224 | 21.0224 | 21.0224 | -15.9106 | PASS |  |  |  |  | 1 |  | 220.5425 | 220.5425 | 220.5425 | -11.783 | PASS | [5 empty] | 33
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0506 | 1.19 | 15.9 | TN | ug N/L | 0.176 | 3.27 | 197 |  | 1 | 2 | reworked calibration | 22.512 | 22.512 | 22.512 | -9.9521 | PASS |  |  |  |  | 1 |  | 208.3037 | 208.3037 | 208.3037 | -16.6785 | PASS | [5 empty] | 34
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0498 | 1.17 | 15.3 | TN | ug N/L | 0.175 | 3.11 | 183 |  | 1 | 2 | reworked calibration | 21.9159 | 21.9159 | 21.9159 | -12.3364 | PASS |  |  |  |  | 1 |  | 193.2712 | 193.2712 | 193.2712 | -22.6915 | FAIL | [5 empty] | 35
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0513 | 1.19 | 16 | TN | ug N/L | 0.18 | 3.32 | 202 |  | 1 | 2 | reworked calibration | 22.512 | 22.512 | 22.512 | -9.9521 | PASS |  |  |  |  | 1 |  | 213.0083 | 213.0083 | 213.0083 | -14.7967 | PASS | [5 empty] | 36
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0495 | 1.08 | 13 | TN | ug N/L | 0.173 | 3.2 | 191 |  | 1 | 2 | reworked calibration | 19.2373 | 19.2373 | 19.2373 | -23.0509 | FAIL |  |  |  |  | 1 |  | 201.7228 | 201.7228 | 201.7228 | -19.3109 | PASS | [5 empty] | 37
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0511 | 1.17 | 15.3 | TN | ug N/L | 0.172 | 3.19 | 189 |  | 1 | 2 | reworked calibration | 21.9159 | 21.9159 | 21.9159 | -12.3364 | PASS |  |  |  |  | 1 |  | 200.7832 | 200.7832 | 200.7832 | -19.6867 | PASS | [5 empty] | 38
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0531 | 1.12 | 26 | TN | ug N/L | 0.337 | 4.92 | 258 |  | 1 | 1 |  | 26 | 26 | 26 | 4 | PASS |  |  |  |  | 1 |  | 258 | 258 | 258 | 3.2 | PASS | [5 empty] | 39
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0466 | 1.06 | 24.3 | TN | ug N/L | 0.269 | 3.94 | 189 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 1 |  | 189 | 189 | 189 | -24.4 | FAIL | [5 empty] | 40
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0495 | 1.12 | 26.1 | TN | ug N/L | 0.281 | 4.13 | 202 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | [6 empty] | 1 |  | 202 | 202 | 202 | [7 empty] | 41
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0511 | 1.16 | 27.5 | TN | ug N/L | 0.323 | 4.71 | 243 |  | 1 | 1 |  | 27.5 | 27.5 | 27.5 | 10 | PASS |  |  |  |  | 1 |  | 243 | 243 | 243 | -2.8 | PASS | [5 empty] | 42
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0506 | 1.11 | 26 | TN | ug N/L | 0.367 | 5.33 | 286 |  | 1 | 1 |  | 26 | 26 | 26 | 4 | PASS |  |  |  |  | 1 |  | 286 | 286 | 286 | 14.4 | PASS | [5 empty] | 43
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0503 | 1.13 | 26.3 | TN | ug N/L | 0.335 | 4.85 | 253 |  | 1 | 1 |  | 26.3 | 26.3 | 26.3 | [6 empty] | 1 |  | 253 | 253 | 253 | [7 empty] | 44
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0489 | 1.11 | 25.8 | TN | ug N/L | 0.33 | 4.8 | 250 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | 3.2 | PASS |  |  |  |  | 1 |  | 250 | 250 | 250 | 0 | PASS | [5 empty] | 45
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0492 | 1.12 | 26.2 | TN | ug N/L | 0.309 | 4.52 | 230 |  | 1 | 1 |  | 26.2 | 26.2 | 26.2 | 4.8 | PASS |  |  |  |  | 1 |  | 230 | 230 | 230 | -8 | PASS | [5 empty] | 46
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0477 | 0.998 | 22.3 | TN | ug N/L | 0.32 | 4.66 | 240 |  | 1 | 1 |  | 22.3 | 22.3 | 22.3 | [6 empty] | 1 |  | 240 | 240 | 240 | [7 empty] | 47
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0475 | 1.07 | 24.5 | TN | ug N/L | 0.31 | 4.49 | 228 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | -2 | PASS |  |  |  |  | 1 |  | 228 | 228 | 228 | -8.8 | PASS | [5 empty] | 48
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0583 | 1.17 | 27.7 | TN | ug N/L | 0.353 | 5.09 | 270 |  | 1 | 1 |  | 27.7 | 27.7 | 27.7 | 10.8 | PASS |  |  |  |  | 1 |  | 270 | 270 | 270 | 8 | PASS | [5 empty] | 49
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.049 | 1.12 | 26.1 | TN | ug N/L | 0.356 | 5.11 | 271 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | 4.4 | PASS |  |  |  |  | 1 |  | 271 | 271 | 271 | 8.4 | PASS | [5 empty] | 50
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0514 | 1.15 | 26.9 | TN | ug N/L | 0.351 | 5.1 | 271 |  | 1 | 1 |  | 26.9 | 26.9 | 26.9 | 7.6 | PASS |  |  |  |  | 1 |  | 271 | 271 | 271 | 8.4 | PASS | [5 empty] | 51
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 18 | 1 | 1 | TP | ug P/L | 0.0585 | 1.2 | 24.2 | TN | ug N/L | 0.264 | 4.01 | 190 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | [6 empty] | 1 |  | 190 | 190 | 190 | [7 empty] | 52
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 19 | 1 | 1 | TP | ug P/L | 0.0602 | 1.22 | 24.8 | TN | ug N/L | 0.317 | 4.73 | 235 |  | 1 | 1 |  | 24.8 | 24.8 | 24.8 | -0.8 | PASS |  |  |  |  | 1 |  | 235 | 235 | 235 | -6 | PASS | [5 empty] | 53
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 17 | 1 | 1 | TP | ug P/L | 0.0596 | 1.22 | 24.6 | TN | ug N/L | 0.318 | 4.7 | 234 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 234 | 234 | 234 | -6.4 | PASS | [5 empty] | 54
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 18 | 1 | 1 | TP | ug P/L | 0.0594 | 1.24 | 25.2 | TN | ug N/L | 0.313 | 4.63 | 230 |  | 1 | 1 |  | 25.2 | 25.2 | 25.2 | 0.8 | PASS |  |  |  |  | 1 |  | 230 | 230 | 230 | -8 | PASS | [5 empty] | 55
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 19 | 1 | 1 | TP | ug P/L | 0.0596 | 1.2 | 24.2 | TN | ug N/L | 0.301 | 4.49 | 221 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 221 | 221 | 221 | -11.6 | PASS | [5 empty] | 56
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 7 | 1 | 1 | TP | ug P/L | 0.0597 | 1.23 | 24.9 | TN | ug N/L | 0.321 | 4.77 | 238 |  | 1 | 1 |  | 24.9 | 24.9 | 24.9 | -0.4 | PASS |  |  |  |  | 1 |  | 238 | 238 | 238 | -4.8 | PASS | [5 empty] | 57
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 17 | 1 | 1 | TP | ug P/L | 0.0612 | 1.22 | 24.6 | TN | ug N/L | 0.314 | 4.69 | 233 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 233 | 233 | 233 | -6.8 | PASS | [5 empty] | 58
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 18 | 1 | 1 | TP | ug P/L | 0.0602 | 1.23 | 24.9 | TN | ug N/L | 0.301 | 4.44 | 217 |  | 1 | 1 |  | 24.9 | 24.9 | 24.9 | -0.4 | PASS |  |  |  |  | 1 |  | 217 | 217 | 217 | -13.2 | PASS | [5 empty] | 59
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 19 | 1 | 1 | TP | ug P/L | 0.0605 | 1.23 | 25.1 | TN | ug N/L | 0.308 | 4.55 | 224 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | 0.4 | PASS |  |  |  |  | 1 |  | 224 | 224 | 224 | -10.4 | PASS | [5 empty] | 60
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 17 | 1 | 1 | TP | ug P/L | 0.0604 | 1.21 | 24.3 | TN | ug N/L | 0.319 | 4.74 | 236 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 1 |  | 236 | 236 | 236 | -5.6 | PASS | [5 empty] | 61
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 7 | 1 | 1 | TP | ug P/L | 0.0606 | 1.22 | 24.6 | TN | ug N/L | 0.313 | 4.65 | 230 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 230 | 230 | 230 | -8 | PASS | [5 empty] | 62
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.058 | 1.1 | 24 | TN | ug N/L | 0.333 | 4.59 | 226 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 1 |  | 226 | 226 | 226 | -9.6 | PASS | [5 empty] | 63
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0591 | 1.12 | 24.4 | TN | ug N/L | 0.329 | 4.53 | 222 |  | 1 | 1 |  | 24.4 | 24.4 | 24.4 | -2.4 | PASS |  |  |  |  | 1 |  | 222 | 222 | 222 | -11.2 | PASS | [5 empty] | 64
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0579 | 1.11 | 24.3 | TN | ug N/L | 0.323 | 4.49 | 220 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | [6 empty] | 1 |  | 220 | 220 | 220 | [7 empty] | 65
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0591 | 1.11 | 24.3 | TN | ug N/L | 0.328 | 4.52 | 221 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 1 |  | 221 | 221 | 221 | -11.6 | PASS | [5 empty] | 66
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.058 | 1.09 | 23.8 | TN | ug N/L | 0.32 | 4.39 | 213 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | -4.8 | PASS |  |  |  |  | 1 |  | 213 | 213 | 213 | -14.8 | PASS | [5 empty] | 67
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0588 | 1.11 | 24.2 | TN | ug N/L | 0.311 | 4.31 | 208 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 208 | 208 | 208 | -16.8 | PASS | [5 empty] | 68
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0582 | 1.07 | 23.2 | TN | ug N/L | 0.313 | 4.31 | 208 |  | 1 | 1 |  | 23.2 | 23.2 | 23.2 | -7.2 | PASS |  |  |  |  | 1 |  | 208 | 208 | 208 | -16.8 | PASS | [5 empty] | 69
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0597 | 1.12 | 24.6 | TN | ug N/L | 0.319 | 4.43 | 216 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 216 | 216 | 216 | -13.6 | PASS | [5 empty] | 70
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0592 | 1.1 | 24 | TN | ug N/L | 0.301 | 4.19 | 200 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 1 |  | 200 | 200 | 200 | -20 | PASS | [5 empty] | 71
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0591 | 1.14 | 25.1 | TN | ug N/L | 0.314 | 4.32 | 209 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | [6 empty] | 1 |  | 209 | 209 | 209 | [7 empty] | 72
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0662 | 1.26 | 22.9 | TN | ug N/L | 0.378 | 5.09 | 253 |  | 1 | 1 |  | 22.9 | 22.9 | 22.9 | [6 empty] | 1 |  | 253 | 253 | 253 | [7 empty] | 73
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0646 | 1.2 | 21.5 | TN | ug N/L | 0.304 | 4.13 | 190 |  | 1 | 1 |  | 21.5 | 21.5 | 21.5 | -14 | PASS |  |  |  |  | 1 |  | 190 | 190 | 190 | -24 | FAIL | [5 empty] | 74
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0649 | 1.22 | 21.8 | TN | ug N/L | 0.355 | 4.8 | 234 |  | 1 | 1 |  | 21.8 | 21.8 | 21.8 | -12.8 | PASS |  |  |  |  | 1 |  | 234 | 234 | 234 | -6.4 | PASS | [5 empty] | 75
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0633 | 1.19 | 21.2 | TN | ug N/L | 0.338 | 4.56 | 218 |  | 1 | 1 |  | 21.2 | 21.2 | 21.2 | -15.2 | PASS |  |  |  |  | 1 |  | 218 | 218 | 218 | -12.8 | PASS | [5 empty] | 76
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0663 | 1.26 | 22.8 | TN | ug N/L | 0.339 | 4.61 | 221 |  | 1 | 1 |  | 22.8 | 22.8 | 22.8 | -8.8 | PASS |  |  |  |  | 1 |  | 221 | 221 | 221 | -11.6 | PASS | [5 empty] | 77
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0641 | 1.2 | 21.4 | TN | ug N/L | 0.343 | 4.65 | 224 |  | 1 | 1 |  | 21.4 | 21.4 | 21.4 | -14.4 | PASS |  |  |  |  | 1 |  | 224 | 224 | 224 | -10.4 | PASS | [5 empty] | 78
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0643 | 1.19 | 21.1 | TN | ug N/L | 0.323 | 4.38 | 206 |  | 1 | 1 |  | 21.1 | 21.1 | 21.1 | -15.6 | PASS |  |  |  |  | 1 |  | 206 | 206 | 206 | -17.6 | PASS | [5 empty] | 79
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.066 | 1.19 | 21.3 | TN | ug N/L | 0.265 | 3.64 | 157 |  | 1 | 1 |  | 21.3 | 21.3 | 21.3 | -14.8 | PASS |  |  |  |  | 1 |  | 157 | 157 | 157 | -37.2 | FAIL | [5 empty] | 80
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.068 | 1.28 | 23.5 | TN | ug N/L | 0.331 | 4.58 | 219 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 | -6 | PASS |  |  |  |  | 1 |  | 219 | 219 | 219 | -12.4 | PASS | [5 empty] | 81
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0698 | 1.3 | 23.9 | TN | ug N/L | 0.327 | 4.46 | 211 |  | 1 | 1 |  | 23.9 | 23.9 | 23.9 | -4.4 | PASS |  |  |  |  | 1 |  | 211 | 211 | 211 | -15.6 | PASS | [5 empty] | 82
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0726 | 1.33 | 24.6 | TN | ug N/L | 0.361 | 4.9 | 240 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 240 | 240 | 240 | -4 | PASS | [5 empty] | 83
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0738 | 1.36 | 25.4 | TN | ug N/L | 0.357 | 4.86 | 238 |  | 1 | 1 |  | 25.4 | 25.4 | 25.4 | 1.6 | PASS |  |  |  |  | 1 |  | 238 | 238 | 238 | -4.8 | PASS | [5 empty] | 84
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0702 | 1.25 | 26.1 | TN | ug N/L | 0.327 | 4.56 | 247 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | 4.4 | PASS |  |  |  |  | 1 |  | 247 | 247 | 247 | -1.2 | PASS | [5 empty] | 85
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0687 | 1.19 | 24.2 | TN | ug N/L | 0.32 | 4.47 | 241 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 241 | 241 | 241 | -3.6 | PASS | [5 empty] | 86
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0692 | 1.22 | 25.3 | TN | ug N/L | 0.321 | 4.46 | 240 |  | 1 | 1 |  | 25.3 | 25.3 | 25.3 | 1.2 | PASS |  |  |  |  | 1 |  | 240 | 240 | 240 | -4 | PASS | [5 empty] | 87
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0717 | 1.28 | 27 | TN | ug N/L | 0.313 | 4.36 | 233 |  | 1 | 1 |  | 27 | 27 | 27 | 8 | PASS |  |  |  |  | 1 |  | 233 | 233 | 233 | -6.8 | PASS | [5 empty] | 88
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk250 | 7 | 1 | 1 | TP | ug P/L | 0.0695 | 1.25 | 26 | TN | ug N/L | 0.303 | 4.24 | 225 |  | 1 | 1 |  | 26 | 26 | 26 | 4 | PASS |  |  |  |  | 1 |  | 225 | 225 | 225 | -10 | PASS | [5 empty] | 89
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0688 | 1.23 | 25.5 | TN | ug N/L | 0.301 | 4.23 | 224 |  | 1 | 1 |  | 25.5 | 25.5 | 25.5 | 2 | PASS |  |  |  |  | 1 |  | 224 | 224 | 224 | -10.4 | PASS | [5 empty] | 90
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0672 | 1.19 | 24.2 | TN | ug N/L | 0.301 | 4.2 | 222 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 222 | 222 | 222 | -11.2 | PASS | [5 empty] | 91
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.069 | 1.2 | 24.6 | TN | ug N/L | 0.289 | 4.06 | 212 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 212 | 212 | 212 | -15.2 | PASS | [5 empty] | 92
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0695 | 1.27 | 26.5 | TN | ug N/L | 0.3 | 4.24 | 224 |  | 1 | 1 |  | 26.5 | 26.5 | 26.5 | 6 | PASS |  |  |  |  | 1 |  | 224 | 224 | 224 | -10.4 | PASS | [5 empty] | 93
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0679 | 1.2 | 24.5 | TN | ug N/L | 0.3 | 4.2 | 222 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | -2 | PASS |  |  |  |  | 1 |  | 222 | 222 | 222 | -11.2 | PASS | [5 empty] | 94
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk250 | 7 | 1 | 1 | TP | ug P/L | 0.0689 | 1.24 | 25.8 | TN | ug N/L | 0.341 | 4.77 | 262 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | 3.2 | PASS |  |  |  |  | 1 |  | 262 | 262 | 262 | 4.8 | PASS | [5 empty] | 95
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0742 | 1.28 | 25.8 | TN | ug N/L | 0.39 | 4.8 | 257 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | 3.2 | PASS |  |  |  |  | 1 |  | 257 | 257 | 257 | 2.8 | PASS | [5 empty] | 96
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0738 | 1.24 | 24.6 | TN | ug N/L | 0.387 | 4.78 | 255 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 255 | 255 | 255 | 2 | PASS | [5 empty] | 97
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0765 | 1.29 | 26 | TN | ug N/L | 0.374 | 4.61 | 243 |  | 1 | 1 |  | 26 | 26 | 26 | 4 | PASS |  |  |  |  | 1 |  | 243 | 243 | 243 | -2.8 | PASS | [5 empty] | 98
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0774 | 1.35 | 27.7 | TN | ug N/L | 0.371 | 4.65 | 246 |  | 1 | 1 |  | 27.7 | 27.7 | 27.7 | 10.8 | PASS |  |  |  |  | 1 |  | 246 | 246 | 246 | -1.6 | PASS | [5 empty] | 99
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0764 | 1.33 | 27.4 | TN | ug N/L | 0.405 | 5.08 | 275 |  | 1 | 1 |  | 27.4 | 27.4 | 27.4 | 9.6 | PASS |  |  |  |  | 1 |  | 275 | 275 | 275 | 10 | PASS | [5 empty] | 100
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 250 | 7 | 1 | 1 | TP | ug P/L | 0.0732 | 1.22 | 24 | TN | ug N/L | 0.348 | 4.43 | 231 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 1 |  | 231 | 231 | 231 | -7.6 | PASS | [5 empty] | 101
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0751 | 1.26 | 25.3 | TN | ug N/L | 0.356 | 4.51 | 237 |  | 1 | 1 |  | 25.3 | 25.3 | 25.3 | 1.2 | PASS |  |  |  |  | 1 |  | 237 | 237 | 237 | -5.2 | PASS | [5 empty] | 102
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.075 | 1.28 | 25.7 | TN | ug N/L | 0.37 | 4.76 | 254 |  | 1 | 1 |  | 25.7 | 25.7 | 25.7 | 2.8 | PASS |  |  |  |  | 1 |  | 254 | 254 | 254 | 1.6 | PASS | [5 empty] | 103
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0734 | 1.25 | 24.8 | TN | ug N/L | 0.361 | 4.65 | 246 |  | 1 | 1 |  | 24.8 | 24.8 | 24.8 | -0.8 | PASS |  |  |  |  | 1 |  | 246 | 246 | 246 | -1.6 | PASS | [5 empty] | 104
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0751 | 1.29 | 26.1 | TN | ug N/L | 0.361 | 4.65 | 246 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | 4.4 | PASS |  |  |  |  | 1 |  | 246 | 246 | 246 | -1.6 | PASS | [5 empty] | 105
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 250 | 7 | 1 | 1 | TP | ug P/L | 0.0762 | 1.28 | 25.9 | TN | ug N/L | 0.29 | 3.78 | 187 |  | 1 | 1 |  | 25.9 | 25.9 | 25.9 | 3.6 | PASS |  |  |  |  | 1 |  | 187 | 187 | 187 | -25.2 | FAIL | [5 empty] | 106
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.071 | 0.863 | -43.1 | TN | ug N/L | 0.47 | 6.26 | 242 |  | 1 | 2 | reworked calibration | 21.4628 | 21.4628 | 21.4628 | -14.1488 | PASS |  |  |  |  | 1 |  | 242 | 242 | 242 | -3.2 | PASS | [5 empty] | 107
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.071 | 0.861 | -43.1 | TN | ug N/L | 0.47 | 6.28 | 243 |  | 1 | 2 | reworked calibration | 21.4628 | 21.4628 | 21.4628 | -14.1488 | PASS |  |  |  |  | 1 |  | 243 | 243 | 243 | -2.8 | PASS | [5 empty] | 108
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0721 | 0.877 | -43.1 | TN | ug N/L | 0.47 | 6.35 | 246 |  | 1 | 2 | reworked calibration | 22.8379 | 22.8379 | 22.8379 | -8.6484 | PASS |  |  |  |  | 1 |  | 246 | 246 | 246 | -1.6 | PASS | [5 empty] | 109
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0715 | 0.884 | -43.1 | TN | ug N/L | 0.469 | 6.28 | 243 |  | 1 | 2 | reworked calibration | 22.0897 | 22.0897 | 22.0897 | -11.6411 | PASS |  |  |  |  | 1 |  | 243 | 243 | 243 | -2.8 | PASS | [5 empty] | 110
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0725 | 0.92 | -43 | TN | ug N/L | 0.47 | 6.3 | 244 |  | 1 | 2 | reworked calibration | 23.3342 | 23.3342 | 23.3342 | -6.6633 | PASS |  |  |  |  | 1 |  | 244 | 244 | 244 | -2.4 | PASS | [5 empty] | 111
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0719 | 0.911 | -43.1 | TN | ug N/L | 0.471 | 6.35 | 246 |  | 1 | 2 | reworked calibration | 22.589 | 22.589 | 22.589 | -9.6439 | PASS |  |  |  |  | 1 |  | 246 | 246 | 246 | -1.6 | PASS | [5 empty] | 112
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | call250 | 7 | 1 | 1 | TP | ug P/L | 0.0697 | 0.881 | -43.2 | TN | ug N/L | 0.471 | 6.35 | 246 |  | 1 | 2 | reworked calibration | 19.8181 | 19.8181 | 19.8181 | -20.7276 | FAIL |  |  |  |  | 1 |  | 246 | 246 | 246 | [7 empty] | 113
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0707 | 0.884 | -43.2 | TN | ug N/L | 0.462 | 6.23 | 240 |  | 1 | 2 | reworked calibration | 21.0851 | 21.0851 | 21.0851 | -15.6594 | PASS |  |  |  |  | 1 |  | 240 | 240 | 240 | -4 | PASS | [5 empty] | 114
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0699 | 0.891 | -43.2 | TN | ug N/L | 0.459 | 6.21 | 239 |  | 1 | 2 | reworked calibration | 20.0725 | 20.0725 | 20.0725 | -19.7099 | PASS |  |  |  |  | 1 |  | 239 | 239 | 239 | -4.4 | PASS | [5 empty] | 115
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0729 | 0.883 | -43 | TN | ug N/L | 0.472 | 6.36 | 247 |  | 1 | 2 | reworked calibration | 23.8284 | 23.8284 | 23.8284 | -4.6862 | PASS |  |  |  |  | 1 |  | 247 | 247 | 247 | -1.2 | PASS | [5 empty] | 116
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.072 | 0.878 | -43.1 | TN | ug N/L | 0.466 | 6.36 | 247 |  | 1 | 2 | reworked calibration | 22.7135 | 22.7135 | 22.7135 | -9.1459 | PASS |  |  |  |  | 1 |  | 247 | 247 | 247 | -1.2 | PASS | [5 empty] | 117
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | ch25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0701 | 0.872 | -43.2 | TN | ug N/L | 0.473 | 6.38 | 248 |  | 1 | 2 | reworked calibration | 20.3264 | 20.3264 | 20.3264 | -18.6943 | PASS |  |  |  |  | 1 |  | 248 | 248 | 248 | -0.8 | PASS | [5 empty] | 118
2023-02-08 | OM_2-8-2023_12-05-24PM BRN edit flat.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0695 | 0.878 | -43.2 | TN | ug N/L | 0.451 | 6.34 | 246 |  | 1 | 2 | reworked calibration | 15.3339 | 15.3339 | 15.3339 | -38.6646 | FAIL |  |  |  |  | 1 |  | 246 | 246 | 246 | [7 empty] | 119
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.137 | 2.3 | 23.3 | TN | ug N/L | 0.515 | 7.07 | 252 |  | 1 | 1 |  | 23.3 | 23.3 | 23.3 | 0.8 | PASS |  |  |  |  | 1 |  | 252 | 252 | 252 | -6.8 | PASS | [5 empty] | 120
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.136 | 2.31 | 23.4 | TN | ug N/L | 0.524 | 7.35 | 265 |  | 1 | 1 |  | 23.4 | 23.4 | 23.4 | 6 | PASS |  |  |  |  | 1 |  | 265 | 265 | 265 | -6.4 | PASS | [5 empty] | 121
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.141 | 2.32 | 23.9 | TN | ug N/L | 0.646 | 8.75 | 333 |  | 1 | 1 |  | 23.9 | 23.9 | 23.9 | 33.2 | FAIL |  |  |  |  | 1 |  | 333 | 333 | 333 | -4.4 | PASS | [5 empty] | 122
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.123 | 2.12 | 17.2 | TN | ug N/L | 0.478 | 6.64 | 231 |  | 1 | 1 |  | 17.2 | 17.2 | 17.2 | -7.6 | PASS |  |  |  |  | 1 |  | 231 | 231 | 231 | -31.2 | FAIL | [5 empty] | 123
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.135 | 2.29 | 23 | TN | ug N/L | 0.581 | 7.96 | 295 |  | 1 | 1 |  | 23 | 23 | 23 | 18 | PASS |  |  |  |  | 1 |  | 295 | 295 | 295 | -8 | PASS | [5 empty] | 124
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.134 | 2.29 | 22.8 | TN | ug N/L | 0.515 | 7.19 | 257 |  | 1 | 1 |  | 22.8 | 22.8 | 22.8 | 2.8 | PASS |  |  |  |  | 1 |  | 257 | 257 | 257 | -8.8 | PASS | [5 empty] | 125
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.134 | 2.3 | 23.2 | TN | ug N/L | 0.515 | 7.73 | 284 |  | 1 | 1 |  | 23.2 | 23.2 | 23.2 | 13.6 | PASS |  |  |  |  | 1 |  | 284 | 284 | 284 | -7.2 | PASS | [5 empty] | 126
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 17 | 1 | 1 | TP | ug P/L | 0.136 | 2.35 | 25 | TN | ug N/L | 0.526 | 7.14 | 255 |  | 1 | 1 |  | 25 | 25 | 25 | 2 | PASS |  |  |  |  | 1 |  | 255 | 255 | 255 | 0 | PASS | [5 empty] | 127
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 17 | 1 | 1 | TP | ug P/L | 0.132 | 2.28 | 22.7 | TN | ug N/L | 0.525 | 7.47 | 271 |  | 1 | 1 |  | 22.7 | 22.7 | 22.7 | 8.4 | PASS |  |  |  |  | 1 |  | 271 | 271 | 271 | -9.2 | PASS | [5 empty] | 128
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.136 | 2.34 | 24.7 | TN | ug N/L | 0.511 | 7.11 | 254 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | 1.6 | PASS |  |  |  |  | 1 |  | 254 | 254 | 254 | -1.2 | PASS | [5 empty] | 129
2023-02-22 | OM_2-22-2023_11-17-20AM BRN edit.omn | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.135 | 2.31 | 23.6 | TN | ug N/L | 0.514 | 7.26 | 261 |  | 1 | 1 |  | 23.6 | 23.6 | 23.6 | 4.4 | PASS |  |  |  |  | 1 |  | 261 | 261 | 261 | -5.6 | PASS | [5 empty] | 130
2023-02-22 | OM_2-22-2023_04-26-39PM BRN edit.omn | chk 25/260 | 17 | 1 | 1 | TP | ug P/L | 0.142 | 2.46 | 28.3 | TN | ug N/L | 0.562 | 8.47 | 319 |  | 1 | 1 |  | 28.3 | 28.3 | 28.3 | 27.6 | FAIL |  |  |  |  | 1 |  | 319 | 319 | 319 | [7 empty] | 131
2023-02-22 | OM_2-22-2023_04-26-39PM BRN edit.omn | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.137 | 2.38 | 26 | TN | ug N/L | 0.515 | 7.12 | 254 |  | 1 | 1 |  | 26 | 26 | 26 | 1.6 | PASS |  |  |  |  | 1 |  | 254 | 254 | 254 | 4 | PASS | [5 empty] | 132
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0635 | 1.14 | 23.8 | TN | ug N/L | 0.146 | 2.2 | 239 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | -4.8 | PASS |  |  |  |  | 1 |  | 239 | 239 | 239 | -4.4 | PASS | [5 empty] | 66
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0635 | 1.14 | 23.8 | TN | ug N/L | 0.149 | 2.17 | 236 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | -4.8 | PASS |  |  |  |  | 1 |  | 236 | 236 | 236 | -5.6 | PASS | [5 empty] | 81
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0628 | 1.15 | 24 | TN | ug N/L | 0.152 | 2.17 | 235 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 1 |  | 235 | 235 | 235 | -6 | PASS | [5 empty] | 96
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0647 | 1.19 | 25.2 | TN | ug N/L | 0.149 | 2.14 | 231 |  | 1 | 1 |  | 25.2 | 25.2 | 25.2 | 0.8 | PASS |  |  |  |  | 1 |  | 231 | 231 | 231 | -7.6 | PASS | [5 empty] | 111
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0668 | 1.18 | 25.1 | TN | ug N/L | 0.149 | 2.08 | 222 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | 0.4 | PASS |  |  |  |  | 1 |  | 222 | 222 | 222 | -11.2 | PASS | [5 empty] | 131
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0653 | 1.15 | 24.1 | TN | ug N/L | 0.153 | 2.27 | 250 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | -3.6 | PASS |  |  |  |  | 1 |  | 250 | 250 | 250 | 0 | PASS | [5 empty] | 150
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0624 | 1.15 | 24 | TN | ug N/L | 0.147 | 2.1 | 226 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 1 |  | 226 | 226 | 226 | -9.6 | PASS | [5 empty] | 165
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.064 | 1.16 | 24.5 | TN | ug N/L | 0.148 | 2.14 | 232 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | -2 | PASS |  |  |  |  | 1 |  | 232 | 232 | 232 | -7.2 | PASS | [5 empty] | 180
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0646 | 1.15 | 24.1 | TN | ug N/L | 0.133 | 1.94 | 202 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | -3.6 | PASS |  |  |  |  | 1 |  | 202 | 202 | 202 | -19.2 | PASS | [5 empty] | 195
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 2.9 | 9.73 | 254 | TN | ug N/L | -0.024 | -0.2 | -98.4 |  | 1 | 1 |  | 254 | 254 | 254 | 916 | FAIL |  |  |  |  | 1 |  | -98.4 | -98.4 | -98.4 | -139.36 | FAIL | [5 empty] | 210
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0635 | 1.15 | 24.2 | TN | ug N/L | 0.144 | 2.13 | 230 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 230 | 230 | 230 | -8 | PASS | [5 empty] | 228
2023-03-15 | OM_3-15-2023_12-40-40PM.OMN | chk 250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0641 | 1.11 | 22.9 | TN | ug N/L | 0.139 | 2.1 | 225 |  | 1 | 1 |  | 22.9 | 22.9 | 22.9 | [6 empty] | 1 |  | 225 | 225 | 225 | [7 empty] | 246
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0638 | 1.13 | 23.7 | TN | ug N/L | 0.124 | 1.73 | 170 |  | 1 | 1 |  | 23.7 | 23.7 | 23.7 | [6 empty] | 1 |  | 170 | 170 | 170 | [7 empty] | 36
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0647 | 1.11 | 23 | TN | ug N/L | 0.151 | 2 | 210 |  | 1 | 1 |  | 23 | 23 | 23 | -8 | PASS |  |  |  |  | 1 |  | 210 | 210 | 210 | -16 | PASS | [5 empty] | 66
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0653 | 1.13 | 23.8 | TN | ug N/L | 0.141 | 1.94 | 202 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | -4.8 | PASS |  |  |  |  | 1 |  | 202 | 202 | 202 | -19.2 | PASS | [5 empty] | 81
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.066 | 1.15 | 24.3 | TN | ug N/L | 0.14 | 1.92 | 198 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 1 |  | 198 | 198 | 198 | -20.8 | FAIL | [5 empty] | 96
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.065 | 1.15 | 24.4 | TN | ug N/L | 0.144 | 1.88 | 191 |  | 1 | 1 |  | 24.4 | 24.4 | 24.4 | -2.4 | PASS |  |  |  |  | 1 |  | 191 | 191 | 191 | -23.6 | FAIL | [5 empty] | 111
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 2 | 18.2 | 492 | TN | ug N/L | -0.0281 | -0.221 | -93.4 |  | 1 | 1 |  | 492 | 492 | 492 | 1868 | FAIL |  |  |  |  | 1 |  | -93.4 | -93.4 | -93.4 | -137.36 | FAIL | [5 empty] | 126
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 3.4 | 23 | 603 | TN | ug N/L | -0.026 | -0.213 | -92.3 |  | 1 | 1 |  | 603 | 603 | 603 | 2312 | FAIL |  |  |  |  | 1 |  | -92.3 | -92.3 | -92.3 | -136.92 | FAIL | [5 empty] | 146
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0641 | 1.1 | 23 | TN | ug N/L | 0.148 | 2.03 | 213 |  | 1 | 1 |  | 23 | 23 | 23 | -8 | PASS |  |  |  |  | 1 |  | 213 | 213 | 213 | -14.8 | PASS | [5 empty] | 166
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0677 | 1.18 | 25.5 | TN | ug N/L | 0.141 | 1.96 | 203 |  | 1 | 1 |  | 25.5 | 25.5 | 25.5 | 2 | PASS |  |  |  |  | 1 |  | 203 | 203 | 203 | -18.8 | PASS | [5 empty] | 181
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0654 | 1.17 | 24.9 | TN | ug N/L | 0.136 | 1.93 | 199 |  | 1 | 1 |  | 24.9 | 24.9 | 24.9 | -0.4 | PASS |  |  |  |  | 1 |  | 199 | 199 | 199 | -20.4 | FAIL | [5 empty] | 196
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0658 | 1.15 | 24.4 | TN | ug N/L | 0.128 | 1.83 | 185 |  | 1 | 1 |  | 24.4 | 24.4 | 24.4 | -2.4 | PASS |  |  |  |  | 1 |  | 185 | 185 | 185 | -26 | FAIL | [5 empty] | 211
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0652 | 1.16 | 24.7 | TN | ug N/L | 0.141 | 2 | 210 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | -1.2 | PASS |  |  |  |  | 1 |  | 210 | 210 | 210 | -16 | PASS | [5 empty] | 226
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0645 | 1.12 | 23.5 | TN | ug N/L | 0.137 | 1.92 | 198 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 | -6 | PASS |  |  |  |  | 1 |  | 198 | 198 | 198 | -20.8 | FAIL | [5 empty] | 241
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0653 | 1.11 | 23.2 | TN | ug N/L | 0.11 | 1.55 | 145 |  | 1 | 1 |  | 23.2 | 23.2 | 23.2 | -7.2 | PASS |  |  |  |  | 1 |  | 145 | 145 | 145 | -42 | FAIL
2023-03-20 | OM_3-20-2023_10-53-22AM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0687 | 1.15 | 24.4 | TN | ug N/L | 0.138 | 1.93 | 200 |  | 1 | 1 |  | 24.4 | 24.4 | 24.4 | -2.4 | PASS |  |  |  |  | 1 |  | 200 | 200 | 200 | -20 | PASS
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0595 | 1.03 | 24.6 | TN | ug N/L | 0.13 | 1.78 | 198 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | [6 empty] | 1 |  | 198 | 198 | 198 | [7 empty] | 36
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0565 | 0.964 | 22.5 | TN | ug N/L | 0.152 | 2.07 | 244 |  | 1 | 1 |  | 22.5 | 22.5 | 22.5 | -10 | PASS |  |  |  |  | 1 |  | 244 | 244 | 244 | -2.4 | PASS | [5 empty] | 66
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0585 | 1.05 | 25.2 | TN | ug N/L | 0.144 | 1.96 | 227 |  | 1 | 1 |  | 25.2 | 25.2 | 25.2 | 0.8 | PASS |  |  |  |  | 1 |  | 227 | 227 | 227 | -9.2 | PASS | [5 empty] | 81
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0577 | 1.01 | 24 | TN | ug N/L | 0.143 | 2.03 | 238 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 1 |  | 238 | 238 | 238 | -4.8 | PASS | [5 empty] | 96
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0604 | 1.04 | 24.8 | TN | ug N/L | 0.138 | 1.95 | 224 |  | 1 | 1 |  | 24.8 | 24.8 | 24.8 | -0.8 | PASS |  |  |  |  | 1 |  | 224 | 224 | 224 | -10.4 | PASS | [5 empty] | 111
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0591 | 1.06 | 25.5 | TN | ug N/L | 0.138 | 2 | 233 |  | 1 | 1 |  | 25.5 | 25.5 | 25.5 | 2 | PASS |  |  |  |  | 1 |  | 233 | 233 | 233 | -6.8 | PASS | [5 empty] | 126
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0572 | 0.988 | 23.3 | TN | ug N/L | 0.138 | 1.97 | 228 |  | 1 | 1 |  | 23.3 | 23.3 | 23.3 | -6.8 | PASS |  |  |  |  | 1 |  | 228 | 228 | 228 | -8.8 | PASS | [5 empty] | 146
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0582 | 1.03 | 24.5 | TN | ug N/L | 0.137 | 1.97 | 228 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | -2 | PASS |  |  |  |  | 1 |  | 228 | 228 | 228 | -8.8 | PASS | [5 empty] | 165
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0582 | 1.02 | 24.1 | TN | ug N/L | 0.138 | 2.05 | 241 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | -3.6 | PASS |  |  |  |  | 1 |  | 241 | 241 | 241 | -3.6 | PASS | [5 empty] | 180
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0586 | 0.994 | 23.4 | TN | ug N/L | 0.146 | 1.99 | 231 |  | 1 | 1 |  | 23.4 | 23.4 | 23.4 | -6.4 | PASS |  |  |  |  | 1 |  | 231 | 231 | 231 | -7.6 | PASS | [5 empty] | 195
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0588 | 1.03 | 24.7 | TN | ug N/L | 0.141 | 1.97 | 228 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | -1.2 | PASS |  |  |  |  | 1 |  | 228 | 228 | 228 | -8.8 | PASS | [5 empty] | 210
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.059 | 1.05 | 25.1 | TN | ug N/L | 0.141 | 1.96 | 227 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | 0.4 | PASS |  |  |  |  | 1 |  | 227 | 227 | 227 | -9.2 | PASS | [5 empty] | 225
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0584 | 1.02 | 24.3 | TN | ug N/L | 0.132 | 1.86 | 210 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 1 |  | 210 | 210 | 210 | -16 | PASS | [5 empty] | 240
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0603 | 1.03 | 24.6 | TN | ug N/L | 0.136 | 1.96 | 226 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 226 | 226 | 226 | -9.6 | PASS
2023-03-22 | OM_3-22-2023_10-55-57AM.OMN | chk250/25 | 7 | 1 | 1 | TP | ug P/L | 0.0601 | 1.02 | 24.3 | TN | ug N/L | 0.142 | 1.99 | 231 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 1 |  | 231 | 231 | 231 | -7.6 | PASS
[20 empty] | TP | [10 empty] | TN
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | [6 empty] | OBSERVATION #
[19 empty] | Mean |  |  | 31.3966 | [7 empty] | Mean |  |  | 226.8047 | [6 empty] | Min | 1
[19 empty] | Std |  |  | 58.8038 | [7 empty] | Std |  |  | 50.3614 | [6 empty] | Max | 246
[19 empty] | CV |  |  | 187.2932 | [7 empty] | CV |  |  | 22.2048 | [6 empty] | Mean
[19 empty] | S/N |  |  | 0.53392 | limits as percent of known | [6 empty] | S/N |  |  | 4.5035 | limits as percent of known | [5 empty] | Whole series
[18 empty] | Warning Limits | Xbar + 2SD |  |  | 149.0042 | 596.0166 | [5 empty] | Warning Limits | Xbar + 2SD |  |  | 327.5276 | 131.011
[19 empty] | Xbar - 2SD |  |  | -86.2109 | -344.8435 | [6 empty] | Xbar - 2SD |  |  | 126.0818 | 50.4327
[18 empty] | Control Limits | Xbar + 3SD |  |  | 207.8079 | 831.2316 | [5 empty] | Control Limits | Xbar + 3SD |  |  | 377.889 | 151.1556
[19 empty] | Xbar - 3SD |  |  | -145.0146 | -580.0585 | [6 empty] | Xbar - 3SD |  |  | 75.7204 | 30.2882
[18 empty] | Count |  |  |  | 174 | [6 empty] | Count |  |  |  | 174
[18 empty] | 99th percentile.inc |  |  |  | 318.26 | [6 empty] | 99th percentile.inc |  |  |  | 312.1292
[18 empty] | Maximum |  |  |  | 603 | [6 empty] | Maximum |  |  |  | 333
[18 empty] | Known |  |  |  | 40 | [6 empty] | Known |  |  |  | 250
[18 empty] | MDL |  |  |  | 138.0768 | [6 empty] | MDL |  |  |  | 118.2535
[18 empty] | LOQ |  |  |  | 588.0376 | [6 empty] | LOQ |  |  |  | 503.6144
[18 empty] | ratio mean/mdl |  |  |  | 0.22739 | [6 empty] | ratio mean/mdl |  |  |  | 1.918
[21 empty] | Known Conc | 25 | [9 empty] | Known Conc | 250
[21 empty] | Mean | 31.3966 | [9 empty] | Mean | 226.8047
[21 empty] | Std Dev | 58.8038 | [9 empty] | Std Dev | 50.3614
[21 empty] | CV or %RSD | 187.2932 | [9 empty] | CV or %RSD | 22.2048
[21 empty] | T value | 2.3481 | [9 empty] | T value | 2.3481
[21 empty] | MDL | 138.0768 | [9 empty] | MDL | 118.2535
[21 empty] | LOQ | 588.0376 | [9 empty] | LOQ | 503.6144
[21 empty] | Error as % | 25.5865 | [9 empty] | Error as % | -9.2781
[21 empty] | Known/MDL | 0.18106 | [9 empty] | Known/MDL | 2.1141
[21 empty] | Mean PR | 125.5865 | [9 empty] | Mean PR | 90.7219
[21 empty] | S/N | 0.53392 | [9 empty] | S/N | 4.5035

## more recent old style MDL calcs
2019 Field Season
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Observation # | PR P | PR N
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | MDL at +150 | 139 | 1 | 1 | Total Phosphorus | ppb | 0.027 | 0.591 | 17.2 | Total Nitrogen | ppb | 0.846 | 14.7 | 384 |  | 1.5 | 1 |  | 17.2 | 9.7 | 14.55 | [6 empty] | 1 |  | 384 | 43 | 64.5
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | MDL at +150 | 140 | 1 | 1 | Total Phosphorus | ppb | 0.0288 | 0.624 | 17.8 | Total Nitrogen | ppb | 0.839 | 14.5 | 379 |  | 1.5 | 1 |  | 17.8 | 10.3 | 15.45 | [6 empty] | 1 |  | 379 | 38 | 57
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | MDL at +150 | 141 | 1 | 1 | Total Phosphorus | ppb | 0.0323 | 0.716 | 19.3 | Total Nitrogen | ppb | 0.89 | 15.3 | 400 |  | 1.5 | 1 |  | 19.3 | 11.8 | 17.7 | [6 empty] | 1 |  | 400 | 59 | 88.5
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | MDL at +150 | 142 | 1 | 1 | Total Phosphorus | ppb | 0.028 | 0.594 | 17.3 | Total Nitrogen | ppb | 0.853 | 14.8 | 388 |  | 1.5 | 1 |  | 17.3 | 9.8 | 14.7 | [6 empty] | 1 |  | 388 | 47 | 70.5
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | MDL at +150 | 143 | 1 | 1 | Total Phosphorus | ppb | 0.0293 | 0.603 | 17.4 | Total Nitrogen | ppb | 0.863 | 14.9 | 390 |  | 1.5 | 1 |  | 17.4 | 9.9 | 14.85 | [6 empty] | 1 |  | 390 | 49 | 73.5
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | MDL at +150 | 144 | 1 | 1 | Total Phosphorus | ppb | 0.0293 | 0.631 | 17.9 | Total Nitrogen | ppb | 0.856 | 14.8 | 389 |  | 1.5 | 1 |  | 17.9 | 10.4 | 15.6 | [6 empty] | 1 |  | 389 | 48 | 72
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | MDL at +150 | 145 | 1 | 1 | Total Phosphorus | ppb | 0.0306 | 0.666 | 18.5 | Total Nitrogen | ppb | 0.863 | 14.9 | 392 |  | 1.5 | 1 |  | 18.5 | 11 | 16.5 | [6 empty] | 1 |  | 392 | 51 | 76.5
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | MDL at +150 | 146 | 1 | 1 | Total Phosphorus | ppb | 0.0334 | 0.752 | 19.9 | Total Nitrogen | ppb | 0.78 | 13.5 | 353 |  | 1.5 | 1 |  | 19.9 | 12.4 | 18.6 | [6 empty] | 1 |  | 353 | 12 | 18
[21 empty] | Known Conc | 14.7783 | [9 empty] | Known Conc | 147.7833
[21 empty] | Mean | 15.9937 | [9 empty] | Mean | 65.0625
[21 empty] | Std Dev | 1.487 | [9 empty] | Std Dev | 21.0754
[21 empty] | CV or %RSD | 9.2971 | [9 empty] | CV or %RSD | 32.3926
[21 empty] | T value | 3.1427 | [9 empty] | T value | 3.1427
[21 empty] | MDL | 4.673 | [9 empty] | MDL | 66.2331
[21 empty] | LOQ | 14.8695 | [9 empty] | LOQ | 210.7544
[21 empty] | Error as % | 8.2244 | [9 empty] | Error as % | -55.9744
[21 empty] | Known/MDL | 3.1625 | [9 empty] | Known/MDL | 2.2313
[21 empty] | Mean PR | 108.2244 | [9 empty] | Mean PR | 44.0256
[21 empty] | S/N | 10.7561 | [9 empty] | S/N | 3.0871
2019 FIELD SEASON
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Observation # | PR P | PR N
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | MDL spiked at 150 uL | 94 | 1 | 1 | Total Phosphorus | ppb | 0.0454 | 1 | 16.7 | Total Nitrogen | ppb | 0.791 | 13.7 | 372 |  | 1.5 | 1 |  | 16.7 | 9.9 | 14.85 | [6 empty] | 2 | don't know what's going on and can't check the original record | 372 | -7 | -10.5
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | MDL spiked at 150 uL | 95 | 1 | 1 | Total Phosphorus | ppb | 0.0437 | 0.949 | 15.7 | Total Nitrogen | ppb | 0.802 | 14 | 379 |  | 1.5 | 1 |  | 15.7 | 8.9 | 13.35 | [6 empty] | 2 | don't know what's going on and can't check the original record | 379 | 0 | 0
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | MDL spiked at 150 uL | 96 | 1 | 1 | Total Phosphorus | ppb | 0.057 | 1.48 | 24.7 | Total Nitrogen | ppb | 0.795 | 13.9 | 376 |  | 1.5 | 1 |  | 24.7 | 17.9 | 26.85 | [6 empty] | 2 | don't know what's going on and can't check the original record | 376 | -3 | -4.5
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | MDL spiked at 150 uL | 97 | 1 | 1 | Total Phosphorus | ppb | 0.0442 | 0.98 | 16.2 | Total Nitrogen | ppb | 0.802 | 14 | 381 |  | 1.5 | 1 |  | 16.2 | 9.4 | 14.1 | [6 empty] | 2 | don't know what's going on and can't check the original record | 381 | 2 | 3
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | MDL spiked at 150 uL | 98 | 1 | 1 | Total Phosphorus | ppb | 0.0435 | 0.934 | 15.5 | Total Nitrogen | ppb | 0.812 | 14.5 | 395 |  | 1.5 | 1 |  | 15.5 | 8.7 | 13.05 | [6 empty] | 2 | don't know what's going on and can't check the original record | 395 | 16 | 24
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | MDL spiked at 150 uL | 99 | 1 | 1 | Total Phosphorus | ppb | 0.0448 | 0.98 | 16.2 | Total Nitrogen | ppb | 0.827 | 14.4 | 391 |  | 1.5 | 1 |  | 16.2 | 9.4 | 14.1 | [6 empty] | 2 | don't know what's going on and can't check the original record | 391 | 12 | 18
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | MDL spiked at 150 uL | 100 | 1 | 1 | Total Phosphorus | ppb | 0.0436 | 1.01 | 16.8 | Total Nitrogen | ppb | 0.799 | 14 | 380 |  | 1.5 | 1 |  | 16.8 | 10 | 15 | [6 empty] | 2 | don't know what's going on and can't check the original record | 380 | 1 | 1.5
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | MDL spiked at 150 uL | 101 | 1 | 1 | Total Phosphorus | ppb | 0.0476 | 0.986 | 16.4 | Total Nitrogen | ppb | 0.83 | 16 | 434 |  | 1.5 | 1 |  | 16.4 | 9.6 | 14.4 |  |  | 7.5676 | PASS |  |  | 2 | don't know what's going on and can't check the original record | 434 | 55 | 82.5
[21 empty] | Known Conc | 14.7783 | [9 empty] | Known Conc | 147.7833
[21 empty] | Mean | 15.7125 | [9 empty] | Mean | 14.25
[21 empty] | Std Dev | 4.5495 | [9 empty] | Std Dev | 29.8173
[21 empty] | CV or %RSD | 28.955 | [9 empty] | CV or %RSD | 209.2442
[21 empty] | T value | 3.1427 | [9 empty] | T value | 3.1427
[21 empty] | MDL | 14.2977 | [9 empty] | MDL | 93.7059
[21 empty] | LOQ | 45.4955 | [9 empty] | LOQ | 298.173
[21 empty] | Error as % | 6.3212 | [9 empty] | Error as % | -90.3575
[21 empty] | Known/MDL | 1.0336 | [9 empty] | Known/MDL | 1.5771
[21 empty] | Mean PR | 106.3212 | [9 empty] | Mean PR | 9.6425
[21 empty] | S/N | 3.4536 | [9 empty] | S/N | 0.47791
2020 FIELD SEASON
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Observation # | PR P | PR N
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 30 | 1 | 1 | TP | ug P/L | 0.0448 | 1.04 | 23.5 | TN | ug N/L | 0.314 | 4.62 | 237 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 | [6 empty] | 1 |  | 237 | 237 | 237 | [6 empty] | 16 | #REF! | #REF!
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 31 | 1 | 1 | TP | ug P/L | 0.0468 | 1.09 | 25.1 | TN | ug N/L | 0.32 | 4.67 | 241 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | [6 empty] | 1 |  | 241 | 241 | 241 | [6 empty] | 17 | #REF! | #REF!
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 32 | 1 | 1 | TP | ug P/L | 0.0543 | 1.33 | 32.7 | TN | ug N/L | 0.32 | 4.67 | 240 |  | 1 | 1 |  | 32.7 | 32.7 | 32.7 | [6 empty] | 1 |  | 240 | 240 | 240 | [6 empty] | 18 | #REF! | #REF!
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 33 | 1 | 1 | TP | ug P/L | 0.0498 | 1.12 | 26.2 | TN | ug N/L | 0.388 | 5.61 | 306 |  | 1 | 1 |  | 26.2 | 26.2 | 26.2 | [6 empty] | 1 |  | 306 | 306 | 306 | [6 empty] | 19 | #REF! | #REF!
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 34 | 1 | 1 | TP | ug P/L | 0.0446 | 1.05 | 23.8 | TN | ug N/L | 0.35 | 5.12 | 272 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | [6 empty] | 1 |  | 272 | 272 | 272 | [6 empty] | 20 | #REF! | #REF!
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 179 | 1 | 1 | TP | ug P/L | 0.0558 | 1.23 | 29.6 | TN | ug N/L | 0.382 | 5.54 | 301 |  | 1 | 1 |  | 29.6 | 29.6 | 29.6 | [6 empty] | 1 |  | 301 | 301 | 301
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 180 | 1 | 1 | TP | ug P/L | 0.0538 | 1.14 | 26.9 | TN | ug N/L | 0.298 | 4.35 | 218 |  | 1 | 1 |  | 26.9 | 26.9 | 26.9 | [6 empty] | 1 |  | 218 | 218 | 218
[21 empty] | Known Conc | 24.3902 | [9 empty] | Known Conc | 243.9024
[21 empty] | Mean | 26.8286 | [9 empty] | Mean | 259.2857
[21 empty] | Std Dev | 3.3115 | [9 empty] | Std Dev | 34.1356
[21 empty] | CV or %RSD | 12.343 | [9 empty] | CV or %RSD | 13.1652
[21 empty] | T value | 3.1427 | [9 empty] | T value | 3.1427
[21 empty] | MDL | 10.4068 | [9 empty] | MDL | 107.2768
[21 empty] | LOQ | 33.1145 | [9 empty] | LOQ | 341.3558
[21 empty] | Error as % | 9.9971 | [9 empty] | Error as % | 6.3071
[21 empty] | Known/MDL | 2.3437 | [9 empty] | Known/MDL | 2.2736
[21 empty] | Mean PR | 109.9971 | [9 empty] | Mean PR | 106.3071
[21 empty] | S/N | 8.1018 | [9 empty] | S/N | 7.5958
2021 Field Season Batch 1
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Observation # | PR P | PR N
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 30 | 1 | 1 | TP | ug P/L | 0.0592 | 1.17 | 23.3 | TN | ug N/L | 0.252 | 3.83 | 178 |  | 1 | 1 |  | 23.3 | 23.3 | 23.3 | [6 empty] | 1 |  | 178 | 178 | 178 | [6 empty] | 21 | #REF! | #REF!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 31 | 1 | 1 | TP | ug P/L | 0.0608 | 1.21 | 24.5 | TN | ug N/L | 0.26 | 3.9 | 183 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | [6 empty] | 1 |  | 183 | 183 | 183 | [6 empty] | 22 | #REF! | #REF!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 32 | 1 | 1 | TP | ug P/L | 0.0593 | 1.19 | 23.9 | TN | ug N/L | 0.247 | 3.72 | 172 |  | 1 | 1 |  | 23.9 | 23.9 | 23.9 | [6 empty] | 1 |  | 172 | 172 | 172 | [6 empty] | 23 | #REF! | #REF!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 33 | 1 | 1 | TP | ug P/L | 0.0597 | 1.23 | 25 | TN | ug N/L | 0.259 | 3.93 | 185 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 185 | 185 | 185 | [6 empty] | 24 | #REF! | #REF!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 34 | 1 | 1 | TP | ug P/L | 0.0599 | 1.18 | 23.6 | TN | ug N/L | 0.274 | 4.1 | 195 |  | 1 | 1 |  | 23.6 | 23.6 | 23.6 | [6 empty] | 1 |  | 195 | 195 | 195 | [6 empty] | 25 | #REF! | #REF!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 156 | 1 | 1 | TP | ug P/L | 0.0609 | 1.2 | 24 | TN | ug N/L | 0.293 | 4.33 | 210 |  | 1 | 1 |  | 24 | 24 | 24 | [6 empty] | 1 |  | 210 | 210 | 210 | [6 empty] | 26 | #REF! | #REF!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 157 | 1 | 1 | TP | ug P/L | 0.0625 | 1.25 | 25.4 | TN | ug N/L | 0.307 | 4.54 | 224 |  | 1 | 1 |  | 25.4 | 25.4 | 25.4 | [6 empty] | 1 |  | 224 | 224 | 224 | [6 empty] | 27 | #REF! | #REF!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 158 | 1 | 1 | TP | ug P/L | 0.0616 | 1.23 | 25 | TN | ug N/L | 0.301 | 4.45 | 218 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 218 | 218 | 218 | [6 empty] | 28 | #REF! | #REF!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 159 | 1 | 1 | TP | ug P/L | 0.0621 | 1.28 | 26.4 | TN | ug N/L | 0.307 | 4.53 | 223 |  | 1 | 1 |  | 26.4 | 26.4 | 26.4 | [6 empty] | 1 |  | 223 | 223 | 223 | [6 empty] | 29 | #REF! | #REF!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 160 | 1 | 1 | TP | ug P/L | 0.0619 | 1.24 | 25.1 | TN | ug N/L | 0.299 | 4.42 | 216 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | [6 empty] | 1 |  | 216 | 216 | 216 | [6 empty] | 30 | #REF! | #REF!
[21 empty] | Known Conc | 24.3902 | [9 empty] | Known Conc | 243.9024
[21 empty] | Mean | 24.62 | [9 empty] | Mean | 200.4
[21 empty] | Std Dev | 0.94257 | [9 empty] | Std Dev | 19.9733
[21 empty] | CV or %RSD | 3.8285 | [9 empty] | CV or %RSD | 9.9667
[21 empty] | T value | 3.1427 | [9 empty] | T value | 3.1427
[21 empty] | MDL | 2.9622 | [9 empty] | MDL | 62.7695
[21 empty] | LOQ | 9.4257 | [9 empty] | LOQ | 199.7332
[21 empty] | Error as % | -1.9124 | [9 empty] | Error as % | -7.2222
[21 empty] | Known/MDL | 8.4734 | [9 empty] | Known/MDL | 3.4412
[21 empty] | Mean PR | 98.0876 | [9 empty] | Mean PR | 92.7778
[21 empty] | S/N | 26.12 | [9 empty] | S/N | 10.0334
2021 Field Season Batch 4
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Observation # | PR P | PR N
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank1 | 30 | 1 | 1 | TP | ug P/L | 0.0821 | 1.39 | 30 | TN | ug N/L | 0.329 | 4.65 | 254 |  | 1 | 1 |  | 30 | 30 | 30 | [6 empty] | 1 |  | 254 | 254 | 254 | [6 empty] | 41 | #REF! | #REF!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank2 | 31 | 1 | 1 | TP | ug P/L | 0.0841 | 1.48 | 32.4 | TN | ug N/L | 0.335 | 4.69 | 256 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 | [6 empty] | 1 |  | 256 | 256 | 256 | [6 empty] | 42 | #REF! | #REF!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank3 | 32 | 1 | 1 | TP | ug P/L | 0.0827 | 1.49 | 32.7 | TN | ug N/L | 0.328 | 4.6 | 250 |  | 1 | 1 |  | 32.7 | 32.7 | 32.7 | [6 empty] | 1 |  | 250 | 250 | 250 | [6 empty] | 43 | #REF! | #REF!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank4 | 33 | 1 | 1 | TP | ug P/L | 0.0846 | 1.45 | 31.6 | TN | ug N/L | 0.328 | 4.6 | 250 |  | 1 | 1 |  | 31.6 | 31.6 | 31.6 | [6 empty] | 1 |  | 250 | 250 | 250 | [6 empty] | 44 | #REF! | #REF!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank 5 | 34 | 1 | 1 | TP | ug P/L | 0.0825 | 1.42 | 30.8 | TN | ug N/L | 0.331 | 4.65 | 254 |  | 1 | 1 |  | 30.8 | 30.8 | 30.8 | [6 empty] | 1 |  | 254 | 254 | 254 | [6 empty] | 45 | #REF! | #REF!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spiked blank BRN | 155 | 1 | 1 | TP | ug P/L | 0.0775 | 1.35 | 28.9 | TN | ug N/L | 0.267 | 3.76 | 191 |  | 1 | 1 |  | 28.9 | 28.9 | 28.9 | [6 empty] | 1 |  | 191 | 191 | 191 | [6 empty] | 46 | #REF! | #REF!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Spike blank BRN | 156 | 1 | 1 | TP | ug P/L | 0.0817 | 1.42 | 30.8 | TN | ug N/L | 0.278 | 3.9 | 201 |  | 1 | 1 |  | 30.8 | 30.8 | 30.8 | [6 empty] | 1 |  | 201 | 201 | 201 | [6 empty] | 47 | #REF! | #REF!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spiked blank BRN | 157 | 1 | 1 | TP | ug P/L | 0.0773 | 1.33 | 28.2 | TN | ug N/L | 0.27 | 3.83 | 195 |  | 1 | 1 |  | 28.2 | 28.2 | 28.2 | [6 empty] | 1 |  | 195 | 195 | 195 | [6 empty] | 48 | #REF! | #REF!
[21 empty] | Known Conc | 24.3902 | [9 empty] | Known Conc | 243.9024
[21 empty] | Mean | 30.675 | [9 empty] | Mean | 231.375
[21 empty] | Std Dev | 1.5899 | [9 empty] | Std Dev | 29.7607
[21 empty] | CV or %RSD | 5.1831 | [9 empty] | CV or %RSD | 12.8625
[21 empty] | T value | 3.1427 | [9 empty] | T value | 3.1427
[21 empty] | MDL | 4.9966 | [9 empty] | MDL | 93.5279
[21 empty] | LOQ | 15.8992 | [9 empty] | LOQ | 297.6065
[21 empty] | Error as % | 25.7675 | [9 empty] | Error as % | -5.1363
[21 empty] | Known/MDL | 4.8814 | [9 empty] | Known/MDL | 2.6078
[21 empty] | Mean PR | 125.7675 | [9 empty] | Mean PR | 94.8637
[21 empty] | S/N | 19.2934 | [9 empty] | S/N | 7.7745