## rolling spiked blank 250
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation # | PR P | PR N
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 028 | 28 | 1 | 1 | TP | ug P/L | 0.0551 | 1.31 | 24 | TN | ug N/L | 0.422 | 7.07 | 8.16 |  | 1 | 1 |  | 24 | 24 | 24 | [6 empty] | 2 | using modeled decline in efficiency | 127.4681 | 127.4681 |  |  | 105.6869 | [5 empty] | 1 | 98.4
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 029 | 29 | 1 | 1 | TP | ug P/L | 0.0563 | 1.31 | 23.8 | TN | ug N/L | 0.512 | 8.69 | 22.1 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | [6 empty] | 2 | using modeled decline in efficiency | 292.5181 | 292.5181 | 251.6436 | [7 empty] | 2 | 97.58 | 103.1749
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 030 | 30 | 1 | 1 | TP | ug P/L | 0.0555 | 1.27 | 22.9 | TN | ug N/L | 0.495 | 8.41 | 19.7 |  | 1 | 1 |  | 22.9 | 22.9 | 22.9 | [6 empty] | 2 | using modeled decline in efficiency | 264.2396 | 264.2396 | 227.2878 | [7 empty] | 3 | 93.89 | 93.1889
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 031 | 31 | 1 | 1 | TP | ug P/L | 0.0557 | 1.28 | 23.1 | TN | ug N/L | 0.483 | 8.16 | 17.5 |  | 1 | 1 |  | 23.1 | 23.1 | 23.1 | [6 empty] | 2 | using modeled decline in efficiency | 238.9032 | 238.9032 | 205.5903 | [7 empty] | 4 | 94.71 | 84.2929
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 032 | 32 | 1 | 1 | TP | ug P/L | 0.0573 | 1.32 | 24.1 | TN | ug N/L | 0.517 | 8.7 | 22.2 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | [6 empty] | 2 | using modeled decline in efficiency | 293.5261 | 293.5261 | 254.9893 | [7 empty] | 5 | 98.81 | 104.5467
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 030 | 30 | 1 | 1 | TP | ug P/L | 0.0484 | 1.16 | 22.4 | TN | ug N/L | 0.309 | 5.28 | 230 |  | 1 | 1 |  | 22.4 | 22.4 | 22.4 | [6 empty] | 1 |  | 230 | 230 | 230 | [7 empty] | 6 | 91.84 | 94.3009
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 031 | 31 | 1 | 1 | TP | ug P/L | 0.0512 | 1.21 | 24.1 | TN | ug N/L | 0.331 | 5.72 | 262 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | [6 empty] | 1 |  | 262 | 262 | 262 | [7 empty] | 7 | 98.81 | 107.4211
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 032 | 32 | 1 | 1 | TP | ug P/L | 0.047 | 1.05 | 19.2 | TN | ug N/L | 0.324 | 5.55 | 249 |  | 1 | 1 |  | 19.2 | 19.2 | 19.2 | [6 empty] | 1 |  | 249 | 249 | 249 | [7 empty] | 8 | 78.72 | 102.091
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 033 | 33 | 1 | 1 | TP | ug P/L | 0.0514 | 1.21 | 24.1 | TN | ug N/L | 0.295 | 5.05 | 212 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | [6 empty] | 1 |  | 212 | 212 | 212 | [7 empty] | 9 | 98.81 | 86.9209
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 034 | 34 | 1 | 1 | TP | ug P/L | 0.0567 | 1.34 | 27.7 | TN | ug N/L | 0.32 | 5.51 | 246 |  | 1 | 1 |  | 27.7 | 27.7 | 27.7 | [6 empty] | 1 |  | 246 | 246 | 246 | [7 empty] | 10 | 113.57 | 100.861
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample30 | 30 | 1 | 1 | TP | ug P/L | 0.0578 | 1.29 | 18.7 | TN | ug N/L | 0.164 | 2.98 | 171 |  | 1 | 2 | reworked calibration | 25.4968 | 25.4968 | 25.4968 | [6 empty] | 1 |  | 181.0821 | 181.0821 | 181.0821 | [7 empty] | 11 | 104.5367 | 74.2444
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample31 | 31 | 1 | 1 | TP | ug P/L | 0.0588 | 1.33 | 19.6 | TN | ug N/L | 0.197 | 3.6 | 226 |  | 1 | 2 | reworked calibration | 26.6927 | 26.6927 | 26.6927 | [6 empty] | 1 |  | 239.415 | 239.415 | 239.415 | [7 empty] | 12 | 109.4403 | 98.1611
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample32 | 32 | 1 | 1 | TP | ug P/L | 0.0609 | 1.34 | 19.9 | TN | ug N/L | 0.225 | 4.04 | 266 |  | 1 | 2 | reworked calibration | 26.9919 | 26.9919 | 26.9919 | [6 empty] | 1 |  | 281.1202 | 281.1202 | 281.1202 | [7 empty] | 13 | 110.6669 | 115.2605
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample33 | 33 | 1 | 1 | TP | ug P/L | 0.0591 | 1.34 | 19.8 | TN | ug N/L | 0.218 | 3.9 | 253 |  | 1 | 2 | reworked calibration | 26.9919 | 26.9919 | 26.9919 | [6 empty] | 1 |  | 267.8227 | 267.8227 | 267.8227 | [7 empty] | 14 | 110.6669 | 109.8084
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample34 | 34 | 1 | 1 | TP | ug P/L | 0.0562 | 1.29 | 18.7 | TN | ug N/L | 0.188 | 3.43 | 212 |  | 1 | 2 | reworked calibration | 25.4968 | 25.4968 | 25.4968 | [6 empty] | 1 |  | 223.37 | 223.37 | 223.37 | [7 empty] | 15 | 104.5367 | 91.5826
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 30 | 1 | 1 | TP | ug P/L | 0.0448 | 1.04 | 23.5 | TN | ug N/L | 0.314 | 4.62 | 237 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 | [6 empty] | 1 |  | 237 | 237 | 237 | [7 empty] | 16 | 96.35 | 97.171
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 31 | 1 | 1 | TP | ug P/L | 0.0468 | 1.09 | 25.1 | TN | ug N/L | 0.32 | 4.67 | 241 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | [6 empty] | 1 |  | 241 | 241 | 241 | [7 empty] | 17 | 102.91 | 98.811
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 32 | 1 | 1 | TP | ug P/L | 0.0543 | 1.33 | 32.7 | TN | ug N/L | 0.32 | 4.67 | 240 |  | 1 | 1 |  | 32.7 | 32.7 | 32.7 | [6 empty] | 1 |  | 240 | 240 | 240 | [7 empty] | 18 | 134.07 | 98.401
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 33 | 1 | 1 | TP | ug P/L | 0.0498 | 1.12 | 26.2 | TN | ug N/L | 0.388 | 5.61 | 306 |  | 1 | 1 |  | 26.2 | 26.2 | 26.2 | [6 empty] | 1 |  | 306 | 306 | 306 | [7 empty] | 19 | 107.42 | 125.4613
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 34 | 1 | 1 | TP | ug P/L | 0.0446 | 1.05 | 23.8 | TN | ug N/L | 0.35 | 5.12 | 272 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | [6 empty] | 1 |  | 272 | 272 | 272 | [7 empty] | 20 | 97.58 | 111.5211
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 30 | 1 | 1 | TP | ug P/L | 0.0592 | 1.17 | 23.3 | TN | ug N/L | 0.252 | 3.83 | 178 |  | 1 | 1 |  | 23.3 | 23.3 | 23.3 | [6 empty] | 1 |  | 178 | 178 | 178 | [7 empty] | 21 | 95.53 | 72.9807
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 31 | 1 | 1 | TP | ug P/L | 0.0608 | 1.21 | 24.5 | TN | ug N/L | 0.26 | 3.9 | 183 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | [6 empty] | 1 |  | 183 | 183 | 183 | [7 empty] | 22 | 100.45 | 75.0308
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 32 | 1 | 1 | TP | ug P/L | 0.0593 | 1.19 | 23.9 | TN | ug N/L | 0.247 | 3.72 | 172 |  | 1 | 1 |  | 23.9 | 23.9 | 23.9 | [6 empty] | 1 |  | 172 | 172 | 172 | [7 empty] | 23 | 97.99 | 70.5207
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 33 | 1 | 1 | TP | ug P/L | 0.0597 | 1.23 | 25 | TN | ug N/L | 0.259 | 3.93 | 185 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 185 | 185 | 185 | [7 empty] | 24 | 102.5 | 75.8508
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 34 | 1 | 1 | TP | ug P/L | 0.0599 | 1.18 | 23.6 | TN | ug N/L | 0.274 | 4.1 | 195 |  | 1 | 1 |  | 23.6 | 23.6 | 23.6 | [6 empty] | 1 |  | 195 | 195 | 195 | [7 empty] | 25 | 96.76 | 79.9508
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 156 | 1 | 1 | TP | ug P/L | 0.0609 | 1.2 | 24 | TN | ug N/L | 0.293 | 4.33 | 210 |  | 1 | 1 |  | 24 | 24 | 24 | [6 empty] | 1 |  | 210 | 210 | 210 | [7 empty] | 26 | 98.4 | 86.1009
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 157 | 1 | 1 | TP | ug P/L | 0.0625 | 1.25 | 25.4 | TN | ug N/L | 0.307 | 4.54 | 224 |  | 1 | 1 |  | 25.4 | 25.4 | 25.4 | [6 empty] | 1 |  | 224 | 224 | 224 | [7 empty] | 27 | 104.14 | 91.8409
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 158 | 1 | 1 | TP | ug P/L | 0.0616 | 1.23 | 25 | TN | ug N/L | 0.301 | 4.45 | 218 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 218 | 218 | 218 | [7 empty] | 28 | 102.5 | 89.3809
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 159 | 1 | 1 | TP | ug P/L | 0.0621 | 1.28 | 26.4 | TN | ug N/L | 0.307 | 4.53 | 223 |  | 1 | 1 |  | 26.4 | 26.4 | 26.4 | [6 empty] | 1 |  | 223 | 223 | 223 | [7 empty] | 29 | 108.24 | 91.4309
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 160 | 1 | 1 | TP | ug P/L | 0.0619 | 1.24 | 25.1 | TN | ug N/L | 0.299 | 4.42 | 216 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | [6 empty] | 1 |  | 216 | 216 | 216 | [7 empty] | 30 | 102.91 | 88.5609
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 030 | 30 | 1 | 1 | TP | ug P/L | 0.0584 | 1.09 | 23.7 | TN | ug N/L | 0.274 | 3.79 | 174 |  | 1 | 1 |  | 23.7 | 23.7 | 23.7 | [6 empty] | 1 |  | 174 | 174 | 174 | [7 empty] | 31 | 97.17 | 71.3407
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 031 | 31 | 1 | 1 | TP | ug P/L | 0.0581 | 1.05 | 22.6 | TN | ug N/L | 0.342 | 4.65 | 230 |  | 1 | 1 |  | 22.6 | 22.6 | 22.6 | [6 empty] | 1 |  | 230 | 230 | 230 | [7 empty] | 32 | 92.66 | 94.3009
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 032 | 32 | 1 | 1 | TP | ug P/L | 0.0571 | 1.05 | 22.5 | TN | ug N/L | 0.341 | 4.65 | 230 |  | 1 | 1 |  | 22.5 | 22.5 | 22.5 | [6 empty] | 1 |  | 230 | 230 | 230 | [7 empty] | 33 | 92.25 | 94.3009
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 033 | 33 | 1 | 1 | TP | ug P/L | 0.058 | 1.14 | 25 | TN | ug N/L | 0.339 | 4.64 | 229 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 229 | 229 | 229 | [7 empty] | 34 | 102.5 | 93.8909
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 034 | 34 | 1 | 1 | TP | ug P/L | 0.0583 | 1.16 | 25.8 | TN | ug N/L | 0.364 | 5.06 | 257 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | [6 empty] | 1 |  | 257 | 257 | 257 | [7 empty] | 35 | 105.78 | 105.3711
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | spiked blank 1 | 30 | 1 | 1 | TP | ug P/L | 0.0665 | 1.2 | 21.4 | TN | ug N/L | 0.263 | 3.61 | 155 |  | 1 | 1 |  | 21.4 | 21.4 | 21.4 | [6 empty] | 1 |  | 155 | 155 | 155 | [7 empty] | 36 | 87.74 | 63.5506
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | spiked blank 2 | 31 | 1 | 1 | TP | ug P/L | 0.0559 | 1.03 | 17.4 | TN | ug N/L | 0.256 | 3.51 | 148 |  | 1 | 1 |  | 17.4 | 17.4 | 17.4 | [6 empty] | 1 |  | 148 | 148 | 148 | [7 empty] | 37 | 71.34 | 60.6806
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | spiked blank 3 | 32 | 1 | 1 | TP | ug P/L | 0.0555 | 1.03 | 17.2 | TN | ug N/L | 0.325 | 4.44 | 210 |  | 1 | 1 |  | 17.2 | 17.2 | 17.2 | [6 empty] | 1 |  | 210 | 210 | 210 | [7 empty] | 38 | 70.52 | 86.1009
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | spiked blank 4 | 33 | 1 | 1 | TP | ug P/L | 0.0562 | 1.06 | 18 | TN | ug N/L | 0.332 | 4.53 | 216 |  | 1 | 1 |  | 18 | 18 | 18 | [6 empty] | 1 |  | 216 | 216 | 216 | [7 empty] | 39 | 73.8 | 88.5609
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | spiked blank 5 | 34 | 1 | 1 | TP | ug P/L | 0.0564 | 1.08 | 18.4 | TN | ug N/L | 0.312 | 4.21 | 194 |  | 1 | 1 |  | 18.4 | 18.4 | 18.4 | [6 empty] | 1 |  | 194 | 194 | 194 | [7 empty] | 40 | 75.44 | 79.5408
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank1 | 30 | 1 | 1 | TP | ug P/L | 0.0821 | 1.39 | 30 | TN | ug N/L | 0.329 | 4.65 | 254 |  | 1 | 1 |  | 30 | 30 | 30 | [6 empty] | 1 |  | 254 | 254 | 254 | [7 empty] | 41 | 123 | 104.141
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank2 | 31 | 1 | 1 | TP | ug P/L | 0.0841 | 1.48 | 32.4 | TN | ug N/L | 0.335 | 4.69 | 256 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 | [6 empty] | 1 |  | 256 | 256 | 256 | [7 empty] | 42 | 132.84 | 104.961
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank3 | 32 | 1 | 1 | TP | ug P/L | 0.0827 | 1.49 | 32.7 | TN | ug N/L | 0.328 | 4.6 | 250 |  | 1 | 1 |  | 32.7 | 32.7 | 32.7 | [6 empty] | 1 |  | 250 | 250 | 250 | [7 empty] | 43 | 134.07 | 102.501
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank4 | 33 | 1 | 1 | TP | ug P/L | 0.0846 | 1.45 | 31.6 | TN | ug N/L | 0.328 | 4.6 | 250 |  | 1 | 1 |  | 31.6 | 31.6 | 31.6 | [6 empty] | 1 |  | 250 | 250 | 250 | [7 empty] | 44 | 129.56 | 102.501
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank 5 | 34 | 1 | 1 | TP | ug P/L | 0.0825 | 1.42 | 30.8 | TN | ug N/L | 0.331 | 4.65 | 254 |  | 1 | 1 |  | 30.8 | 30.8 | 30.8 | [6 empty] | 1 |  | 254 | 254 | 254 | [7 empty] | 45 | 126.28 | 104.141
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spiked blank BRN | 155 | 1 | 1 | TP | ug P/L | 0.0775 | 1.35 | 28.9 | TN | ug N/L | 0.267 | 3.76 | 191 |  | 1 | 1 |  | 28.9 | 28.9 | 28.9 | [6 empty] | 1 |  | 191 | 191 | 191 | [7 empty] | 46 | 118.49 | 78.3108
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Spike blank BRN | 156 | 1 | 1 | TP | ug P/L | 0.0817 | 1.42 | 30.8 | TN | ug N/L | 0.278 | 3.9 | 201 |  | 1 | 1 |  | 30.8 | 30.8 | 30.8 | [6 empty] | 1 |  | 201 | 201 | 201 | [7 empty] | 47 | 126.28 | 82.4108
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spiked blank BRN | 157 | 1 | 1 | TP | ug P/L | 0.0773 | 1.33 | 28.2 | TN | ug N/L | 0.27 | 3.83 | 195 |  | 1 | 1 |  | 28.2 | 28.2 | 28.2 | [6 empty] | 1 |  | 195 | 195 | 195 | [7 empty] | 48 | 115.62 | 79.9508
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | spikeblank1 | 30 | 1 | 1 | TP | ug P/L | 0.0791 | 1.31 | 26.5 | TN | ug N/L | 0.373 | 4.64 | 245 |  | 1 | 1 |  | 26.5 | 26.5 | 26.5 | [6 empty] | 1 |  | 245 | 245 | 245 | [7 empty] | 49 | 108.65 | 100.451
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | spikeblank2 | 31 | 1 | 1 | TP | ug P/L | 0.0788 | 1.29 | 26.2 | TN | ug N/L | 0.36 | 4.47 | 234 |  | 1 | 1 |  | 26.2 | 26.2 | 26.2 | [6 empty] | 1 |  | 234 | 234 | 234 | [7 empty] | 50 | 107.42 | 95.941
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | spikeblank3 | 32 | 1 | 1 | TP | ug P/L | 0.0799 | 1.35 | 27.7 | TN | ug N/L | 0.385 | 4.72 | 251 |  | 1 | 1 |  | 27.7 | 27.7 | 27.7 | [6 empty] | 1 |  | 251 | 251 | 251 | [7 empty] | 51 | 113.57 | 102.911
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | spikeblank4 | 33 | 1 | 1 | TP | ug P/L | 0.0784 | 1.3 | 26.4 | TN | ug N/L | 0.373 | 4.51 | 237 |  | 1 | 1 |  | 26.4 | 26.4 | 26.4 | [6 empty] | 1 |  | 237 | 237 | 237 | [7 empty] | 52 | 108.24 | 97.171
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | spikeblank5 | 34 | 1 | 1 | TP | ug P/L | 0.0769 | 1.27 | 25.4 | TN | ug N/L | 0.385 | 4.74 | 253 |  | 1 | 1 |  | 25.4 | 25.4 | 25.4 | [6 empty] | 1 |  | 253 | 253 | 253 | [7 empty] | 53 | 104.14 | 103.731
TN
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: |  | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | [6 empty] | OBSERVATION #
| Mean | [5 empty] | Mean |  |  | 24.5132 |  |  |  |  | 204.3521 |  |  |  | Mean |  |  | 25.173 | [7 empty] | Mean |  |  | 225.8716 | [6 empty] | Min | 1
| Std | [5 empty] | Std |  |  | 3.9477 |  |  |  |  | 68.4287 |  |  |  | Std |  |  | 3.5923 | [7 empty] | Std |  |  | 32.8658 | [6 empty] | Max | 53
| CV | [5 empty] | CV |  |  | 16.1043 |  |  |  |  | 33.4857 |  |  |  | CV |  |  | 14.2706 | [7 empty] | CV |  |  | 14.5507
[7 empty] | S/N |  |  | 6.2095 |  |  |  |  | 2.9864 |  |  |  | S/N |  |  | 7.0074 | [7 empty] | S/N |  |  | 6.8725
Warning Limits | Xbar + 2SD |  |  |  |  | Warning Limits | Xbar + 2SD |  |  | 32.4086 |  |  |  |  | 341.2095 |  |  | Warning Limits | Xbar + 2SD |  |  | 32.3577 | 128.5412 | [5 empty] | Warning Limits | Xbar + 2SD |  |  | 291.6032
| Xbar - 2SD | [5 empty] | Xbar - 2SD |  |  | 16.6178 |  |  |  |  | 67.4947 |  |  |  | Xbar - 2SD |  |  | 17.9883 | 71.4588 | [6 empty] | Xbar - 2SD |  |  | 160.14
Control Limits | Xbar + 3SD |  |  |  |  | Control Limits | Xbar + 3SD |  |  | 36.3563 |  |  |  |  | 409.6382 |  |  | Control Limits | Xbar + 3SD |  |  | 35.9501 | [6 empty] | Control Limits | Xbar + 3SD |  |  | 324.469
| Xbar - 3SD | [5 empty] | Xbar - 3SD |  |  | 12.6701 |  |  |  |  | -0.93403 |  |  |  | Xbar - 3SD |  |  | 14.396 | [7 empty] | Xbar - 3SD |  |  | 127.2742
Count | [5 empty] | Count |  |  |  | 53 |  |  |  |  | 53 |  |  | Count |  |  |  | 33 | [6 empty] | Count |  |  |  | 33
99th percentile.inc | [5 empty] | 99th percentile.inc |  |  |  | 32.7 |  |  |  |  | 288.32 |  |  | 99th percentile.inc |  |  |  | 31.75 | [6 empty] | 99th percentile.inc |  |  |  | 256.68
Maximum | [5 empty] | Maximum |  |  |  | 32.7 |  |  |  |  | 306 |  |  | Maximum |  |  |  | 32.7 | [6 empty] | Maximum |  |  |  | 32.7
[6 empty] | Known |  |  |  | 24.4 |  |  |  |  | 243.9 |  |  | Known |  |  |  | 24.4 | [6 empty] | Known |  |  |  | 243.9
Known | [5 empty] | MDL |  |  |  | 9.4753 |  |  |  |  | 164.2443 |  |  | MDL |  |  |  | 8.7965 | [6 empty] | MDL |  |  |  | 80.4778
[6 empty] | LOQ |  |  |  | 39.4769 |  |  |  |  | 684.287 |  |  | LOQ |  |  |  | 35.9235 | [6 empty] | LOQ |  |  |  | 328.658
[6 empty] | ratio mean/mdl |  |  |  | 2.5871 |  |  |  |  | 1.2442 |  |  | ratio mean/mdl |  |  |  | 2.8617 | [6 empty] | ratio mean/mdl |  |  |  | 2.8066
[21 empty] | Known Conc | 24.3902 | [9 empty] | Known Conc | 243.9024
[21 empty] | Mean | 25.3273 | [9 empty] | Mean | 215.697
[21 empty] | Std Dev | 4.0805 | [9 empty] | Std Dev | 31.2165
[21 empty] | CV or %RSD | 16.1109 | [9 empty] | CV or %RSD | 14.4724
[21 empty] | T value | 2.4487 | [9 empty] | T value | 2.4487
[21 empty] | MDL | 9.9917 | [9 empty] | MDL | 76.4391
[21 empty] | LOQ | 40.8046 | [9 empty] | LOQ | 312.1647
[21 empty] | Error as % | 3.8418 | [9 empty] | Error as % | -11.5642
[21 empty] | Known/MDL | 2.441 | [9 empty] | Known/MDL | 3.1908
[21 empty] | Mean PR | 103.8418 | [9 empty] | Mean PR | 88.4358
[21 empty] | S/N | 6.207 | [9 empty] | S/N | 6.9097
[21 empty] | STDEV.P | 3.5583
[21 empty] | STDEV.S | 3.5923

## rolling digest check
As an MDLs surrogate… these aren't independently prepared but they ARE independently digested. And they are at concentrations within 10 x of the resulting MDLs
As a digestion check, TNs are running low. New Glycine? Raise temperature?
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 014 | 14 | 1 | 1 | Total Phosphorus | ppb | 0.0944 | 1.96 | 34.8 | Total Nitrogen | ppb | 1.49 | 23.2 | 684 |  | 1.5 | 1 |  | 44.2165 | 44.2165 | 44.2165 | [6 empty] | 1 |  | 401.9114 | 401.9114 | 401.9114 | [7 empty] | 1
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 015 | 15 | 1 | 1 | Total Phosphorus | ppb | 0.0787 | 1.78 | 31.5 | Total Nitrogen | ppb | 1.51 | 23.4 | 689 |  | 1.5 | 1 |  | 38.5562 | 38.5562 | 38.5562 | [6 empty] | 1 |  | 409.6283 | 409.6283 | 409.6283 | [7 empty] | 2
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 016 | 16 | 1 | 1 | Total Phosphorus | ppb | 0.0801 | 1.81 | 32 | Total Nitrogen | ppb | 1.47 | 23 | 676 |  | 1.5 | 1 |  | 39.5105 | 39.5105 | 39.5105 | [6 empty] | 1 |  | 394.1916 | 394.1916 | 394.1916 | [7 empty] | 3
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 017 | 17 | 1 | 1 | Total Phosphorus | ppb | 0.078 | 1.68 | 29.7 | Total Nitrogen | ppb | 1.47 | 22.9 | 675 |  | 1.5 | 1 |  | 35.344 | 35.344 | 35.344 | [6 empty] | 1 |  | 390.3306 | 390.3306 | 390.3306 | [7 empty] | 4
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 018 | 18 | 1 | 1 | Total Phosphorus | ppb | 0.0802 | 1.81 | 32.1 | Total Nitrogen | ppb | 1.47 | 23 | 677 |  | 1.5 | 1 |  | 39.5105 | 39.5105 | 39.5105 | [6 empty] | 1 |  | 394.1916 | 394.1916 | 394.1916 | [7 empty] | 5
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | digestion check | 14 | 1 | 1 | Total Phosphorus | ppb | 0.0961 | 1.98 | 33.2 | Total Nitrogen | ppb | 1.24 | 21.5 | 582 |  | 1.5 | 1 |  | 41.9776 | 41.9776 | 41.9776 | [6 empty] | 1 |  | 403.91 | 403.91 | 403.91 | [7 empty] | 6
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | digestion check | 15 | 1 | 1 | Total Phosphorus | ppb | 0.0947 | 2.04 | 34.2 | Total Nitrogen | ppb | 1.25 | 21.7 | 587 |  | 1.5 | 1 |  | 43.6042 | 43.6042 | 43.6042 | [6 empty] | 1 |  | 412.8095 | 412.8095 | 412.8095 | [7 empty] | 7
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | digestion check | 16 | 1 | 1 | Total Phosphorus | ppb | 0.0851 | 1.84 | 30.8 | Total Nitrogen | ppb | 1.22 | 21.2 | 575 |  | 1.5 | 1 |  | 38.175 | 38.175 | 38.175 | [6 empty] | 1 |  | 390.5386 | 390.5386 | 390.5386 | [7 empty] | 8
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | digestion check | 17 | 1 | 1 | Total Phosphorus | ppb | 0.0819 | 1.65 | 27.5 | Total Nitrogen | ppb | 1.23 | 21.3 | 577 |  | 1.5 | 1 |  | 32.9988 | 32.9988 | 32.9988 | [6 empty] | 1 |  | 394.9987 | 394.9987 | 394.9987 | [7 empty] | 9
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | digestion check | 18 | 1 | 1 | Total Phosphorus | ppb | 0.0871 | 1.88 | 31.4 | Total Nitrogen | ppb | 1.23 | 21.5 | 583 |  | 1.5 | 1 |  | 39.2625 | 39.2625 | 39.2625 | [6 empty] | 1 |  | 403.91 | 403.91 | 403.91 | [7 empty] | 10
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 023 | 23 | 1 | 1 | TP | ug P/L | 0.0836 | 1.83 | 38.9 | TN | ug N/L | 0.588 | 9.89 | 32.5 |  | 1 | 1 |  | 38.9 | 38.9 | 38.9 | [6 empty] | 1 |  | 306.698 | 306.698 | 352.1846 | [7 empty] | 11
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 024 | 24 | 1 | 1 | TP | ug P/L | 0.0827 | 1.88 | 40.4 | TN | ug N/L | 0.61 | 10.4 | 36.5 |  | 1 | 1 |  | 40.4 | 40.4 | 40.4 | [6 empty] | 1 |  | 350.5406 | 350.5406 | 398.68 | [7 empty] | 12
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 025 | 25 | 1 | 1 | TP | ug P/L | 0.0803 | 1.82 | 38.6 | TN | ug N/L | 0.602 | 10.2 | 34.9 |  | 1 | 1 |  | 38.6 | 38.6 | 38.6 | [6 empty] | 1 |  | 333.2991 | 333.2991 | 381.7425 | [7 empty] | 13
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 026 | 26 | 1 | 1 | TP | ug P/L | 0.0836 | 1.85 | 39.5 | TN | ug N/L | 0.643 | 10.8 | 40.6 |  | 1 | 1 |  | 39.5 | 39.5 | 39.5 | [6 empty] | 1 |  | 385.2102 | 385.2102 | 436.4607 | [7 empty] | 14
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 027 | 27 | 1 | 1 | TP | ug P/L | 0.0822 | 1.82 | 38.4 | TN | ug N/L | 0.608 | 10.3 | 35.6 |  | 1 | 1 |  | 38.4 | 38.4 | 38.4 | [6 empty] | 1 |  | 341.9121 | 341.9121 | 392.6024 | [7 empty] | 15
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 025 | 25 | 1 | 1 | TP | ug P/L | 0.0766 | 1.69 | 38.3 | TN | ug N/L | 0.368 | 6.26 | 302 |  | 1 | 1 |  | 38.3 | 38.3 | 38.3 | [6 empty] | 1 |  | 302 | 302 | 302 | [7 empty] | 16
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 026 | 26 | 1 | 1 | TP | ug P/L | 0.0732 | 1.67 | 37.5 | TN | ug N/L | 0.351 | 6.03 | 285 |  | 1 | 1 |  | 37.5 | 37.5 | 37.5 | [6 empty] | 1 |  | 285 | 285 | 285 | [7 empty] | 17
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 027 | 27 | 1 | 1 | TP | ug P/L | 0.0768 | 1.74 | 39.6 | TN | ug N/L | 0.428 | 7.26 | 376 |  | 1 | 1 |  | 39.6 | 39.6 | 39.6 | [6 empty] | 1 |  | 376 | 376 | 376 | [7 empty] | 18
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 028 | 28 | 1 | 1 | TP | ug P/L | 0.0743 | 1.69 | 38.2 | TN | ug N/L | 0.381 | 6.5 | 320 |  | 1 | 1 |  | 38.2 | 38.2 | 38.2 | [6 empty] | 1 |  | 320 | 320 | 320 | [7 empty] | 19
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 029 | 29 | 1 | 1 | TP | ug P/L | 0.0741 | 1.67 | 37.5 | TN | ug N/L | 0.435 | 7.39 | 386 |  | 1 | 1 |  | 37.5 | 37.5 | 37.5 | [6 empty] | 1 |  | 386 | 386 | 386 | [7 empty] | 20
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample25 | 25 | 1 | 1 | TP | ug P/L | 0.088 | 1.9 | 35.4 | TN | ug N/L | 0.25 | 4.4 | 297 |  | 1 | 2 | reworked calibration | 43.8644 | 43.8644 | 43.8644 | [6 empty] | 1 |  | 315.4328 | 315.4328 | 315.4328 | [7 empty] | 21
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample26 | 26 | 1 | 1 | TP | ug P/L | 0.0853 | 1.86 | 34.3 | TN | ug N/L | 0.194 | 3.51 | 219 |  | 1 | 2 | reworked calibration | 42.6515 | 42.6515 | 42.6515 | [6 empty] | 1 |  | 230.9158 | 230.9158 | 230.9158 | [7 empty] | 22
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample27 | 27 | 1 | 1 | TP | ug P/L | 0.0867 | 1.88 | 34.8 | TN | ug N/L | 0.28 | 5.09 | 358 |  | 1 | 2 | reworked calibration | 43.2578 | 43.2578 | 43.2578 | [6 empty] | 1 |  | 381.6766 | 381.6766 | 381.6766 | [7 empty] | 23
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample28 | 28 | 1 | 1 | TP | ug P/L | 0.0847 | 1.81 | 32.9 | TN | ug N/L | 0.232 | 4.2 | 280 |  | 1 | 2 | reworked calibration | 41.1371 | 41.1371 | 41.1371 | [6 empty] | 1 |  | 296.3492 | 296.3492 | 296.3492 | [7 empty] | 24
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample29 | 29 | 1 | 1 | TP | ug P/L | 0.0842 | 1.88 | 34.8 | TN | ug N/L | 0.208 | 3.79 | 243 |  | 1 | 2 | reworked calibration | 43.2578 | 43.2578 | 43.2578 | [6 empty] | 1 |  | 257.3927 | 257.3927 | 257.3927 | [7 empty] | 25
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample25 | 25 | 1 | 1 | TP | ug P/L | 0.088 | 1.9 | 35.4 | TN | ug N/L | 0.25 | 4.4 | 297 |  | 1 | 2 | reworked calibration | 43.8644 | 43.8644 | 43.8644 | [6 empty] | 1 |  | 315.4328 | 315.4328 | 315.4328 | [7 empty] | 26
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample26 | 26 | 1 | 1 | TP | ug P/L | 0.0853 | 1.86 | 34.3 | TN | ug N/L | 0.194 | 3.51 | 219 |  | 1 | 2 | reworked calibration | 42.6515 | 42.6515 | 42.6515 | [6 empty] | 1 |  | 230.9158 | 230.9158 | 230.9158 | [7 empty] | 27
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample27 | 27 | 1 | 1 | TP | ug P/L | 0.0867 | 1.88 | 34.8 | TN | ug N/L | 0.28 | 5.09 | 358 |  | 1 | 2 | reworked calibration | 43.2578 | 43.2578 | 43.2578 | [6 empty] | 1 |  | 381.6766 | 381.6766 | 381.6766 | [7 empty] | 28
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample28 | 28 | 1 | 1 | TP | ug P/L | 0.0847 | 1.81 | 32.9 | TN | ug N/L | 0.232 | 4.2 | 280 |  | 1 | 2 | reworked calibration | 41.1371 | 41.1371 | 41.1371 | [6 empty] | 1 |  | 296.3492 | 296.3492 | 296.3492 | [7 empty] | 29
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample29 | 29 | 1 | 1 | TP | ug P/L | 0.0842 | 1.88 | 34.8 | TN | ug N/L | 0.208 | 3.79 | 243 |  | 1 | 2 | reworked calibration | 43.2578 | 43.2578 | 43.2578 | [6 empty] | 1 |  | 257.3927 | 257.3927 | 257.3927 | [7 empty] | 30
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | digestion check | 25 | 1 | 1 | TP | ug P/L | 0.0788 | 1.63 | 42 | TN | ug N/L | 0.485 | 7.01 | 402 |  | 1 | 1 |  | 42 | 42 | 42 | [6 empty] | 1 |  | 402 | 402 | 402 | [7 empty] | 31
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | digestion check | 26 | 1 | 1 | TP | ug P/L | 0.0764 | 1.6 | 41.2 | TN | ug N/L | 0.474 | 6.82 | 389 |  | 1 | 1 |  | 41.2 | 41.2 | 41.2 | [6 empty] | 1 |  | 389 | 389 | 389 | [7 empty] | 32
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | digestion check | 27 | 1 | 1 | TP | ug P/L | 0.0775 | 1.61 | 41.6 | TN | ug N/L | 0.495 | 7.14 | 411 |  | 1 | 1 |  | 41.6 | 41.6 | 41.6 | [6 empty] | 1 |  | 411 | 411 | 411 | [7 empty] | 33
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | digestion check | 28 | 1 | 1 | TP | ug P/L | 0.0755 | 1.61 | 41.6 | TN | ug N/L | 0.488 | 7.07 | 407 |  | 1 | 1 |  | 41.6 | 41.6 | 41.6 | [6 empty] | 1 |  | 407 | 407 | 407 | [7 empty] | 34
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | digestion check | 29 | 1 | 1 | TP | ug P/L | 0.0759 | 1.6 | 41.1 | TN | ug N/L | 0.505 | 7.24 | 418 |  | 1 | 1 |  | 41.1 | 41.1 | 41.1 | [6 empty] | 1 |  | 418 | 418 | 418 | [7 empty] | 35
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | digest check | 25 | 1 | 1 | TP | ug P/L | 0.0857 | 1.69 | 37.8 | TN | ug N/L | 0.373 | 5.56 | 289 |  | 1 | 1 |  | 37.8 | 37.8 | 37.8 | [6 empty] | 1 |  | 289 | 289 | 289 | [7 empty] | 36
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | digest check | 26 | 1 | 1 | TP | ug P/L | 0.0869 | 1.71 | 38.3 | TN | ug N/L | 0.396 | 5.96 | 314 |  | 1 | 1 |  | 38.3 | 38.3 | 38.3 | [6 empty] | 1 |  | 314 | 314 | 314 | [7 empty] | 37
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | digest check | 27 | 1 | 1 | TP | ug P/L | 0.0868 | 1.69 | 37.7 | TN | ug N/L | 0.365 | 5.58 | 290 |  | 1 | 1 |  | 37.7 | 37.7 | 37.7 | [6 empty] | 1 |  | 290 | 290 | 290 | [7 empty] | 38
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | digest check | 28 | 1 | 1 | TP | ug P/L | 0.0866 | 1.69 | 37.6 | TN | ug N/L | 0.403 | 5.97 | 315 |  | 1 | 1 |  | 37.6 | 37.6 | 37.6 | [6 empty] | 1 |  | 315 | 315 | 315 | [7 empty] | 39
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | digest check | 29 | 1 | 1 | TP | ug P/L | 0.0868 | 1.69 | 37.7 | TN | ug N/L | 0.363 | 5.41 | 279 |  | 1 | 1 |  | 37.7 | 37.7 | 37.7 | [6 empty] | 1 |  | 279 | 279 | 279 | [7 empty] | 40
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 025 | 25 | 1 | 1 | TP | ug P/L | 0.0865 | 1.61 | 38.6 | TN | ug N/L | 0.439 | 5.98 | 317 |  | 1 | 1 |  | 38.6 | 38.6 | 38.6 | [6 empty] | 1 |  | 317 | 317 | 317 | [7 empty] | 41
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 026 | 26 | 1 | 1 | TP | ug P/L | 0.0869 | 1.66 | 39.8 | TN | ug N/L | 0.458 | 6.22 | 333 |  | 1 | 1 |  | 39.8 | 39.8 | 39.8 | [6 empty] | 1 |  | 333 | 333 | 333 | [7 empty] | 42
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 027 | 27 | 1 | 1 | TP | ug P/L | 0.087 | 1.61 | 38.4 | TN | ug N/L | 0.415 | 5.67 | 297 |  | 1 | 1 |  | 38.4 | 38.4 | 38.4 | [6 empty] | 1 |  | 297 | 297 | 297 | [7 empty] | 43
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 028 | 28 | 1 | 1 | TP | ug P/L | 0.0851 | 1.56 | 37.1 | TN | ug N/L | 0.489 | 6.6 | 357 |  | 1 | 1 |  | 37.1 | 37.1 | 37.1 | [6 empty] | 1 |  | 357 | 357 | 357 | [7 empty] | 44
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 029 | 29 | 1 | 1 | TP | ug P/L | 0.0848 | 1.52 | 35.9 | TN | ug N/L | 0.383 | 5.2 | 266 |  | 1 | 1 |  | 35.9 | 35.9 | 35.9 | [6 empty] | 1 |  | 266 | 266 | 266 | [7 empty] | 45
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | digest check 1 | 25 | 1 | 1 | TP | ug P/L | 0.095 | 1.71 | 34.1 | TN | ug N/L | 0.424 | 5.7 | 294 |  | 1 | 1 |  | 34.1 | 34.1 | 34.1 | [6 empty] | 1 |  | 294 | 294 | 294 | [7 empty] | 46
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | digest check 2 | 26 | 1 | 1 | TP | ug P/L | 0.097 | 1.73 | 34.5 | TN | ug N/L | 0.516 | 6.88 | 373 |  | 1 | 1 |  | 34.5 | 34.5 | 34.5 | [6 empty] | 1 |  | 373 | 373 | 373 | [7 empty] | 47
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | digest check 3 | 27 | 1 | 1 | TP | ug P/L | 0.0959 | 1.71 | 34.1 | TN | ug N/L | 0.433 | 5.81 | 302 |  | 1 | 1 |  | 34.1 | 34.1 | 34.1 | [6 empty] | 1 |  | 302 | 302 | 302 | [7 empty] | 48
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | digest check 4 | 28 | 1 | 1 | TP | ug P/L | 0.0969 | 1.74 | 34.7 | TN | ug N/L | 0.455 | 6.09 | 320 |  | 1 | 1 |  | 34.7 | 34.7 | 34.7 | [6 empty] | 1 |  | 320 | 320 | 320 | [7 empty] | 49
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | digest check 5 | 29 | 1 | 1 | TP | ug P/L | 0.0971 | 1.74 | 34.7 | TN | ug N/L | 0.437 | 5.87 | 306 |  | 1 | 1 |  | 34.7 | 34.7 | 34.7 | [6 empty] | 1 |  | 306 | 306 | 306 | [7 empty] | 50
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check1 | 25 | 1 | 1 | TP | ug P/L | 0.0941 | 1.67 | 37.7 | TN | ug N/L | 0.453 | 6.33 | 372 |  | 1 | 1 |  | 37.7 | 37.7 | 37.7 | [6 empty] | 1 |  | 372 | 372 | 372 | [7 empty] | 51
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check2 | 26 | 1 | 1 | TP | ug P/L | 0.0948 | 1.66 | 37.4 | TN | ug N/L | 0.436 | 6.12 | 357 |  | 1 | 1 |  | 37.4 | 37.4 | 37.4 | [6 empty] | 1 |  | 357 | 357 | 357 | [7 empty] | 52
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digestcheck3 | 27 | 1 | 1 | TP | ug P/L | 0.0908 | 1.58 | 35.2 | TN | ug N/L | 0.425 | 5.93 | 344 |  | 1 | 1 |  | 35.2 | 35.2 | 35.2 | [6 empty] | 1 |  | 344 | 344 | 344 | [7 empty] | 53
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digestcheck4 | 28 | 1 | 1 | TP | ug P/L | 0.0929 | 1.65 | 37.2 | TN | ug N/L | 0.448 | 6.21 | 363 |  | 1 | 1 |  | 37.2 | 37.2 | 37.2 | [6 empty] | 1 |  | 363 | 363 | 363 | [7 empty] | 54
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check5 | 29 | 1 | 1 | TP | ug P/L | 0.0941 | 1.65 | 37.3 | TN | ug N/L | 0.459 | 6.36 | 374 |  | 1 | 1 |  | 37.3 | 37.3 | 37.3 | [6 empty] | 1 |  | 374 | 374 | 374 | [7 empty] | 55
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | digestchk1 | 25 | 1 | 1 | TP | ug P/L | 0.104 | 1.74 | 39.1 | TN | ug N/L | 0.542 | 6.6 | 380 |  | 1 | 1 |  | 39.1 | 39.1 | 39.1 | [6 empty] | 1 |  | 380 | 380 | 380 | [7 empty] | 56
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | digestchk2 | 26 | 1 | 1 | TP | ug P/L | 0.101 | 1.71 | 38.1 | TN | ug N/L | 0.54 | 6.64 | 383 |  | 1 | 1 |  | 38.1 | 38.1 | 38.1 | [6 empty] | 1 |  | 383 | 383 | 383 | [7 empty] | 57
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | digestch3 | 27 | 1 | 1 | TP | ug P/L | 0.101 | 1.72 | 38.5 | TN | ug N/L | 0.552 | 6.7 | 387 |  | 1 | 1 |  | 38.5 | 38.5 | 38.5 | [6 empty] | 1 |  | 387 | 387 | 387 | [7 empty] | 58
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | digestchk4 | 28 | 1 | 1 | TP | ug P/L | 0.103 | 1.76 | 39.7 | TN | ug N/L | 0.554 | 6.71 | 387 |  | 1 | 1 |  | 39.7 | 39.7 | 39.7 | [6 empty] | 1 |  | 387 | 387 | 387 | [7 empty] | 59
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | digestchk5 | 29 | 1 | 1 | TP | ug P/L | 0.101 | 1.72 | 38.4 | TN | ug N/L | 0.536 | 6.53 | 375 |  | 1 | 1 |  | 38.4 | 38.4 | 38.4 | [6 empty] | 1 |  | 375 | 375 | 375 | [7 empty] | 60
[20 empty] | TP | [10 empty] | TN
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | [6 empty] | OBSERVATION #
[19 empty] | Mean |  |  | 39.0915 | [7 empty] | Mean |  |  | 349.8604 | [6 empty] | Min | 1
[19 empty] | Std |  |  | 2.7844 | [7 empty] | Std |  |  | 52.1428 | [6 empty] | Max | 60
[19 empty] | CV |  |  | 7.1226 | [7 empty] | CV |  |  | 14.9039 | [6 empty] | Mean
[19 empty] | S/N |  |  | 14.0397 | [7 empty] | S/N |  |  | 6.7097 | [6 empty] | Whole series
[18 empty] | Warning Limits | Xbar + 2SD |  |  | 44.6603 | [6 empty] | Warning Limits | Xbar + 2SD |  |  | 454.1461
[19 empty] | Xbar - 2SD |  |  | 33.5228 | [7 empty] | Xbar - 2SD |  |  | 245.5747
[18 empty] | Control Limits | Xbar + 3SD |  |  | 47.4446 | [6 empty] | Control Limits | Xbar + 3SD |  |  | 506.2889
[19 empty] | Xbar - 3SD |  |  | 30.7385 | [7 empty] | Xbar - 3SD |  |  | 193.4319
[18 empty] | Count |  |  |  | 60 | [6 empty] | Count |  |  |  | 60
[18 empty] | 99th percentile.inc |  |  |  | 44.0087 | [6 empty] | 99th percentile.inc |  |  |  | 425.5689
[18 empty] | Maximum |  |  |  | 44.2165 | [6 empty] | Maximum |  |  |  | 436.4607
[18 empty] | Known |  |  |  | 40 | [6 empty] | Known |  |  |  | 400
[18 empty] | MDL |  |  |  | 6.658 | [6 empty] | MDL |  |  |  | 124.6855
[18 empty] | LOQ |  |  |  | 27.8435 | [6 empty] | LOQ |  |  |  | 521.4284
[18 empty] | ratio mean/mdl |  |  |  | 5.8713 | [6 empty] | ratio mean/mdl |  |  |  | 2.8059
[21 empty] | Known Conc | 40 | [9 empty] | Known Conc | 400
[21 empty] | Mean | 39.0915 | [9 empty] | Mean | 349.8604
[21 empty] | Std Dev | 2.7844 | [9 empty] | Std Dev | 52.1428
[21 empty] | CV or %RSD | 7.1226 | [9 empty] | CV or %RSD | 14.9039
[21 empty] | T value | 2.3912 | [9 empty] | T value | 2.3912
[21 empty] | MDL | 6.658 | [9 empty] | MDL | 124.6855
[21 empty] | LOQ | 27.8435 | [9 empty] | LOQ | 521.4284
[21 empty] | Error as % | -2.2711 | [9 empty] | Error as % | -12.5349
[21 empty] | Known/MDL | 6.0078 | [9 empty] | Known/MDL | 3.2081
[21 empty] | Mean PR | 97.7289 | [9 empty] | Mean PR | 87.4651
[21 empty] | S/N | 14.0397 | [9 empty] | S/N | 6.7097

## rolling blanks
methods vary so interpretation could be a little tricky
Runs after summer 2020 use digested standards with bisulfate carrier for TP and water carrier for TN
THERE ARE KNOWN PROBLEMS ON 13JAN20 SO THEY MAY BE EXCLUDED.
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Carrier | Observation #
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 001 | 1 | 1 | 1 | Total Phosphorus | ppb | 0.0326 | 0.54 | 8.2 | Total Nitrogen | ppb | 0.781 | 11.9 | 348 |  | 1.5 | 1 |  | -4.694 | -4.694 | -4.694 | [6 empty] | 1 |  | -38.6703 | -38.6703 | -38.6703 | [6 empty] | bisulfate | 1
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 002 | 2 | 1 | 1 | Total Phosphorus | ppb | 0.0327 | 0.53 | 8.02 | Total Nitrogen | ppb | 0.784 | 11.9 | 348 |  | 1.5 | 1 |  | -5.073 | -5.073 | -5.073 | [6 empty] | 1 |  | -38.6703 | -38.6703 | -38.6703 | [6 empty] | bisulfate | 2
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 003 | 3 | 1 | 1 | Total Phosphorus | ppb | 0.0388 | 0.792 | 13 | Total Nitrogen | ppb | 0.817 | 12.5 | 365 |  | 1.5 | 1 |  | 4.6976 | 4.6976 | 4.6976 | [6 empty] | 1 |  | -15.0507 | -15.0507 | -15.0507 | [6 empty] | bisulfate | 3
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 004 | 4 | 1 | 1 | Total Phosphorus | ppb | 0.0429 | 0.694 | 11.1 | Total Nitrogen | ppb | 0.853 | 13 | 382 |  | 1.5 | 1 |  | 1.0818 | 1.0818 | 1.0818 | [6 empty] | 1 |  | 4.613 | 4.613 | 4.613 | [6 empty] | bisulfate | 4
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 005 | 5 | 1 | 1 | Total Phosphorus | ppb | 0.042 | 0.707 | 11.4 | Total Nitrogen | ppb | 0.855 | 13.1 | 385 |  | 1.5 | 1 |  | 1.5641 | 1.5641 | 1.5641 | [6 empty] | 1 |  | 8.5436 | 8.5436 | 8.5436 | [6 empty] | bisulfate | 5
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | water only blank | 1 | 1 | 1 | Total Phosphorus | ppb | 0.0294 | 0.476 | 7.92 | Total Nitrogen | ppb | 0.735 | 12.8 | 348 |  | 1.5 | 1 |  | 0.61457 | 0.61457 | 0.61457 | [6 empty] | 1 |  | 5.3012 | 5.3012 | 5.3012 | [6 empty] | bisulfate | 6
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | water only blank | 2 | 1 | 1 | Total Phosphorus | ppb | 0.0334 | 0.556 | 9.24 | Total Nitrogen | ppb | 0.75 | 13.1 | 356 |  | 1.5 | 1 |  | 2.8432 | 2.8432 | 2.8432 | [6 empty] | 1 |  | 19.42 | 19.42 | 19.42 | [6 empty] | bisulfate | 7
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | water only blank | 3 | 1 | 1 | Total Phosphorus | ppb | 0.0461 | 0.618 | 10.3 | Total Nitrogen | ppb | 0.734 | 12.8 | 349 |  | 1.5 | 1 |  | 4.5682 | 4.5682 | 4.5682 | [6 empty] | 1 |  | 5.3012 | 5.3012 | 5.3012 | [6 empty] | bisulfate | 8
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | water only blank | 4 | 1 | 1 | Total Phosphorus | ppb | 0.0298 | 0.465 | 7.74 | Total Nitrogen | ppb | 0.725 | 12.6 | 342 |  | 1.5 | 1 |  | 0.30788 | 0.30788 | 0.30788 | [6 empty] | 1 |  | -4.1262 | -4.1262 | -4.1262 | [6 empty] | bisulfate | 9
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | water only blank | 5 | 1 | 1 | Total Phosphorus | ppb | 0.0326 | 0.519 | 8.63 | Total Nitrogen | ppb | 0.737 | 12.8 | 348 |  | 1.5 | 1 |  | 1.8129 | 1.8129 | 1.8129 | [6 empty] | 1 |  | 5.3012 | 5.3012 | 5.3012 | [6 empty] | bisulfate | 10
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 018 | 18 | 1 | 1 | TP | ug P/L | 0.0219 | 0.482 | 0.588 | TN | ug N/L | 0.351 | 5.92 | -1.69 |  | 1 | 1 |  | 0.588 | 0.588 | 0.588 | [6 empty] | 1 |  | -20.74 | -20.74 | -3.6153 | [6 empty] | bisulfate for TP but water for TN | 11
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 019 | 19 | 1 | 1 | TP | ug P/L | 0.0221 | 0.492 | 0.857 | TN | ug N/L | 0.348 | 5.92 | -1.73 |  | 1 | 1 |  | 0.857 | 0.857 | 0.857 | [6 empty] | 1 |  | -20.74 | -20.74 | -3.0787 | [6 empty] | bisulfate for TP but water for TN | 12
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 020 | 20 | 1 | 1 | TP | ug P/L | 0.0441 | 0.433 | -0.783 | TN | ug N/L | 0.346 | 5.85 | -2.3 |  | 1 | 1 |  | -0.783 | -0.783 | -0.783 | [6 empty] | 1 |  | -26.2933 | -26.2933 | -8.726 | [6 empty] | bisulfate for TP but water for TN | 13
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 021 | 21 | 1 | 1 | TP | ug P/L | 0.0223 | 0.499 | 1.05 | TN | ug N/L | 0.406 | 6.83 | 6.16 |  | 1 | 1 |  | 1.05 | 1.05 | 1.05 | [6 empty] | 1 |  | 52.1473 | 52.1473 | [7 empty] | bisulfate for TP but water for TN | 14
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 022 | 22 | 1 | 1 | TP | ug P/L | 0.02 | 0.444 | -0.471 | TN | ug N/L | 0.343 | 5.8 | -2.72 |  | 1 | 1 |  | -0.471 | -0.471 | -0.471 | [6 empty] | 1 |  | -30.2553 | -30.2553 | -12.0868 | [6 empty] | bisulfate for TP but water for TN | 15
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 020 | 20 | 1 | 1 | TP | ug P/L | 0.0226 | 0.399 | -0.27 | TN | ug N/L | 0.103 | 1.86 | -20.5 |  | 1 | 1 |  | -0.27 | -0.27 | -0.27 | [6 empty] | 1 |  | -20.5 | -20.5 | -20.5 | [6 empty] | bisulfate for TP but water for TN | 16
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 021 | 21 | 1 | 1 | TP | ug P/L | 0.0238 | 0.447 | 1.17 | TN | ug N/L | 0.131 | 2.37 | 16.4 |  | 1 | 1 |  | 1.17 | 1.17 | 1.17 | [6 empty] | 1 |  | 16.4 | 16.4 | 16.4 | [6 empty] | bisulfate for TP but water for TN | 17
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 022 | 22 | 1 | 1 | TP | ug P/L | 0.0201 | 0.369 | -1.18 | TN | ug N/L | 0.109 | 1.97 | -12.5 |  | 1 | 1 |  | -1.18 | -1.18 | -1.18 | [6 empty] | 1 |  | -12.5 | -12.5 | -12.5 | [6 empty] | bisulfate for TP but water for TN | 18
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 023 | 23 | 1 | 1 | TP | ug P/L | 0.0222 | 0.417 | 0.273 | TN | ug N/L | 0.112 | 2.01 | -9.17 |  | 1 | 1 |  | 0.273 | 0.273 | 0.273 | [6 empty] | 1 |  | -9.17 | -9.17 | -9.17 | [6 empty] | bisulfate for TP but water for TN | 19
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 024 | 24 | 1 | 1 | TP | ug P/L | 0.0278 | 0.514 | 3.19 | TN | ug N/L | 0.126 | 2.23 | 6.67 |  | 1 | 1 |  | 3.19 | 3.19 | 3.19 | [6 empty] | 1 |  | 6.67 | 6.67 | 6.67 | [6 empty] | bisulfate for TP but water for TN | 20
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample20 | 20 | 1 | 1 | TP | ug P/L | 0.0521 | 1.18 |  | TN | ug N/L | 0.0404 | 0.784 | -25.6 |  | 1 | 3 | carry over | 22.2139 | 22.2139 | [7 empty] | 1 |  | -21.4483 | -21.4483 | -21.4483 | [6 empty] | bisulfate for TP but water for TN | 21
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample21 | 21 | 1 | 1 | TP | ug P/L | 0.0337 | 0.689 | 2.69 | TN | ug N/L | 0.0418 | 0.827 | -21.7 |  | 1 | 2 | reworked calibration | 7.6698 | 7.6698 | 7.6698 | [6 empty] | 1 |  | -17.5436 | -17.5436 | -17.5436 | [6 empty] | bisulfate for TP but water for TN | 22
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample22 | 22 | 1 | 1 | TP | ug P/L | 0.0326 | 0.653 | 1.75 | TN | ug N/L | 0.0457 | 1.03 | -3.08 |  | 1 | 2 | reworked calibration | 6.6104 | 6.6104 | 6.6104 | [6 empty] | 1 |  | 0.92298 | 0.92298 | 0.92298 | [6 empty] | bisulfate for TP but water for TN | 23
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample23 | 23 | 1 | 1 | TP | ug P/L | 0.0329 | 0.661 | 1.97 | TN | ug N/L | 0.0508 | 0.897 | -15.5 |  | 1 | 2 | reworked calibration | 6.8457 | 6.8457 | 6.8457 | [6 empty] | 1 |  | -11.182 | -11.182 | -11.182 | [6 empty] | bisulfate for TP but water for TN | 24
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | Sample24 | 24 | 1 | 1 | TP | ug P/L | 0.0297 | 0.554 | -0.821 | TN | ug N/L | 0.0427 | 0.871 | -17.8 |  | 1 | 2 | reworked calibration | 3.7022 | 3.7022 | 3.7022 | [6 empty] | 1 |  | -13.5456 | -13.5456 | -13.5456 | [6 empty] | bisulfate for TP but water for TN | 25
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | water blank | 20 | 1 | 1 | TP | ug P/L | 0.0107 | 0.324 | -0.271 | TN | ug N/L | 0.0496 | 0.907 | -6.32 |  | 1 | 1 |  | -0.271 | -0.271 | -0.271 | [6 empty] | 1 |  | -6.32 | -6.32 | -6.32 | [6 empty] | bisulfate for TP but water for TN | 26
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | water blank | 21 | 1 | 1 | TP | ug P/L | 0.00974 | 0.321 | -0.334 | TN | ug N/L | 0.0459 | 0.839 | -10.6 |  | 1 | 1 |  | -0.334 | -0.334 | -0.334 | [6 empty] | 1 |  | -10.6 | -10.6 | -10.6 | [6 empty] | bisulfate for TP but water for TN | 27
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | water blank | 22 | 1 | 1 | TP | ug P/L | 0.011 | 0.333 | -0.0138 | TN | ug N/L | 0.0574 | 1.01 | -0.131 |  | 1 | 1 |  | -0.0138 | -0.0138 | -0.0138 | [6 empty] | 1 |  | -0.131 | -0.131 | -0.131 | [6 empty] | bisulfate for TP but water for TN | 28
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | water blank | 23 | 1 | 1 | TP | ug P/L | 0.011 | 0.33 | -0.0869 | TN | ug N/L | 0.0561 | 0.83 | -11.2 |  | 1 | 1 |  | -0.0869 | -0.0869 | -0.0869 | [6 empty] | 1 |  | -11.2 | -11.2 | -11.2 | [6 empty] | bisulfate for TP but water for TN | 29
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | water blank | 24 | 1 | 1 | TP | ug P/L | 0.0116 | 0.323 | -0.296 | TN | ug N/L | 0.064 | 1.11 | 6.52 |  | 1 | 1 |  | -0.296 | -0.296 | -0.296 | [6 empty] | 1 |  | 6.52 | 6.52 | 6.52 | [6 empty] | bisulfate for TP but water for TN | 30
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 020 | 20 | 1 | 1 | TP | ug P/L | 0.00961 | 0.267 | 0.354 | TN | ug N/L | 0.0596 | 0.968 | -5.93 |  | 1 | 1 |  | 0.354 | 0.354 | 0.354 | [6 empty] | 1 |  | -5.93 | -5.93 | -5.93 | [6 empty] | bisulfate for TP but water for TN | 31
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 021 | 21 | 1 | 1 | TP | ug P/L | 0.00787 | 0.223 | -0.867 | TN | ug N/L | 0.0645 | 1.06 | -0.157 |  | 1 | 1 |  | -0.867 | -0.867 | -0.867 | [6 empty] | 1 |  | -0.157 | -0.157 | -0.157 | [6 empty] | bisulfate for TP but water for TN | 32
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 022 | 22 | 1 | 1 | TP | ug P/L | 0.0104 | 0.241 | -0.38 | TN | ug N/L | 0.0677 | 1.08 | 0.862 |  | 1 | 1 |  | -0.38 | -0.38 | -0.38 | [6 empty] | 1 |  | 0.862 | 0.862 | 0.862 | [6 empty] | bisulfate for TP but water for TN | 33
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 023 | 23 | 1 | 1 | TP | ug P/L | 0.0118 | 0.264 | 0.287 | TN | ug N/L | 0.106 | 1.53 | 29.4 |  | 1 | 1 |  | 0.287 | 0.287 | 0.287 | [6 empty] | 1 |  | 29.4 | 29.4 | 29.4 | [6 empty] | bisulfate for TP but water for TN | 34
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 024 | 24 | 1 | 1 | TP | ug P/L | 0.011 | 0.238 | -0.452 | TN | ug N/L | 0.0782 | 1.19 | 8.35 |  | 1 | 1 |  | -0.452 | -0.452 | -0.452 | [6 empty] | 1 |  | 8.35 | 8.35 | 8.35 | [6 empty] | bisulfate for TP but water for TN | 35
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | water blank 1 | 20 | 1 | 1 | TP | ug P/L | 0.00803 | 0.226 | -2 | TN | ug N/L | 0.0517 | 0.852 | -29.7 |  | 1 | 1 |  | -2 | -2 | -2 | [6 empty] | 1 |  | -29.7 | -29.7 | -29.7 | [6 empty] | bisulfate for TP but water for TN | 36
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | water blank 2 | 21 | 1 | 1 | TP | ug P/L | 0.00822 | 0.228 | -1.95 | TN | ug N/L | 0.0498 | 0.832 | -31 |  | 1 | 1 |  | -1.95 | -1.95 | -1.95 | [6 empty] | 1 |  | -31 | -31 | -31 | [6 empty] | bisulfate for TP but water for TN | 37
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | water blank 3 | 22 | 1 | 1 | TP | ug P/L | 0.00828 | 0.216 | -2.24 | TN | ug N/L | 0.0521 | 0.846 | -30.1 |  | 1 | 1 |  | -2.24 | -2.24 | -2.24 | [6 empty] | 1 |  | -30.1 | -30.1 | -30.1 | [6 empty] | bisulfate for TP but water for TN | 38
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | water blank 4 | 23 | 1 | 1 | TP | ug P/L | 0.00855 | 0.207 | -2.46 | TN | ug N/L | 0.0438 | 0.687 | -40.7 |  | 1 | 1 |  | -2.46 | -2.46 | -2.46 | [6 empty] | 1 |  | -40.7 | -40.7 | -40.7 | [6 empty] | bisulfate for TP but water for TN | 39
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | water blank 5 | 24 | 1 | 1 | TP | ug P/L | 0.00832 | 0.235 | -1.78 | TN | ug N/L | 0.0502 | 0.802 | -33 |  | 1 | 1 |  | -1.78 | -1.78 | -1.78 | [6 empty] | 1 |  | -33 | -33 | -33 | [6 empty] | bisulfate for TP but water for TN | 40
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank1 | 20 | 1 | 1 | TP | ug P/L | 0.0136 | 0.308 | -0.453 | TN | ug N/L | 0.0585 | 0.989 | -5.79 |  | 1 | 1 |  | -0.453 | -0.453 | -0.453 | [6 empty] | 1 |  | -5.79 | -5.79 | -5.79 | [6 empty] | bisulfate for TP but water for TN | 41
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank2 | 21 | 1 | 1 | TP | ug P/L | 0.0125 | 0.345 | 0.576 | TN | ug N/L | 0.0608 | 1.01 | -4.53 |  | 1 | 1 |  | 0.576 | 0.576 | 0.576 | [6 empty] | 1 |  | -4.53 | -4.53 | -4.53 | [6 empty] | bisulfate for TP but water for TN | 42
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank3 | 22 | 1 | 1 | TP | ug P/L | 0.0134 | 0.348 | 0.675 | TN | ug N/L | 0.0589 | 0.973 | -6.93 |  | 1 | 1 |  | 0.675 | 0.675 | 0.675 | [6 empty] | 1 |  | -6.93 | -6.93 | -6.93 | [6 empty] | bisulfate for TP but water for TN | 43
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank4 | 23 | 1 | 1 | TP | ug P/L | 0.0126 | 0.337 | 0.353 | TN | ug N/L | 0.0623 | 0.972 | -7 |  | 1 | 1 |  | 0.353 | 0.353 | 0.353 | [6 empty] | 1 |  | -7 | -7 | -7 | [6 empty] | bisulfate for TP but water for TN | 44
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank5 | 24 | 1 | 1 | TP | ug P/L | 0.0115 | 0.301 | -0.66 | TN | ug N/L | 0.0674 | 1.09 | 1.59 |  | 1 | 1 |  | -0.66 | -0.66 | -0.66 | [6 empty] | 1 |  | 1.59 | 1.59 | 1.59 | [6 empty] | bisulfate for TP but water for TN | 45
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | waterblank1 | 20 | 1 | 1 | TP | ug P/L | 0.0218 | 0.366 | 0.0149 | TN | ug N/L | 0.063 | 0.948 | -6.52 |  | 1 | 1 |  | 0.0149 | 0.0149 | 0.0149 | [6 empty] | 1 |  | -6.52 | -6.52 | -6.52 | [6 empty] | bisulfate for TP but water for TN | 46
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | waterblank2 | 21 | 1 | 1 | TP | ug P/L | 0.0213 | 0.366 | 0.0252 | TN | ug N/L | 0.07 | 0.967 | -5.22 |  | 1 | 1 |  | 0.0252 | 0.0252 | 0.0252 | [6 empty] | 1 |  | -5.22 | -5.22 | -5.22 | [6 empty] | bisulfate for TP but water for TN | 47
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | waterblank3 | 22 | 1 | 1 | TP | ug P/L | 0.022 | 0.363 | -0.0623 | TN | ug N/L | 0.0629 | 0.961 | -5.61 |  | 1 | 1 |  | -0.0623 | -0.0623 | -0.0623 | [6 empty] | 1 |  | -5.61 | -5.61 | -5.61 | [6 empty] | bisulfate for TP but water for TN | 48
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | waterblank4 | 23 | 1 | 1 | TP | ug P/L | 0.0218 | 0.364 | -0.047 | TN | ug N/L | 0.071 | 0.975 | -4.68 |  | 1 | 1 |  | -0.047 | -0.047 | -0.047 | [6 empty] | 1 |  | -4.68 | -4.68 | -4.68 | [6 empty] | bisulfate for TP but water for TN | 49
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | waterblank5 | 24 | 1 | 1 | TP | ug P/L | 0.0209 | 0.365 | -7.23e-05 | TN | ug N/L | 0.0644 | 0.938 | -7.2 |  | 1 | 1 |  | -7.23e-05 | -7.23e-05 | -7.23e-05 | [6 empty] | 1 |  | -7.2 | -7.2 | -7.2 | [6 empty] | bisulfate for TP but water for TN | 50
[20 empty] | TP | [10 empty] | TN
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Carrier | Observation #
[19 empty] | Mean |  |  | 0.50831 | [7 empty] | Mean |  |  | -7.4089 | [6 empty] | Min | 1 | Mean
[19 empty] | Std |  |  | 2.4912 | [7 empty] | Std |  |  | 14.7971 | [6 empty] | Max | 50
[19 empty] | CV |  |  | 490.0911 | [7 empty] | CV |  |  | -199.7202
[19 empty] | S/N |  |  | 0.20404 | [7 empty] | S/N |  |  | -0.5007
[18 empty] | Warning Limits | Xbar + 2SD |  |  | 5.4907 | [6 empty] | Warning Limits | Xbar + 2SD |  |  | 22.1853
[19 empty] | Xbar - 2SD |  |  | -4.4741 | [7 empty] | Xbar - 2SD |  |  | -37.0031
[18 empty] | Control Limits | Xbar + 3SD |  |  | 7.9819 | [6 empty] | Control Limits | Xbar + 3SD |  |  | 36.9824
[19 empty] | Xbar - 3SD |  |  | -6.9652 | [7 empty] | Xbar - 3SD |  |  | -51.8002
[18 empty] | Count |  |  |  | 49 | [6 empty] | Count |  |  |  | 50
[15 empty] | MDLb if N >100 |  |  | 99th percentile.inc |  |  |  | 7.2742 |  |  |  | MDLb if N >100 |  |  | 99th percentile.inc |  |  |  | 15.0873
[15 empty] | MDLb if N <100 |  |  | Maximum |  |  |  | 7.6698 |  |  |  | MDLb if N <100 |  |  | Maximum |  |  |  | 22.2139
[15 empty] | MDLb if all are >0 |  |  | MDL |  |  |  | 5.9952 | [6 empty] | MDL |  |  |  | 35.5854
[18 empty] | LOQ |  |  |  | 24.9119 | [6 empty] | LOQ |  |  |  | 147.971

## rolling spike PR
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 030 | 30 | 1 | 1 | Total Phosphorus | ppb | 0.0949 | 2.08 | 37.1 | Total Nitrogen | ppb | 2.39 | 36.8 | 1090 |  | 1.5 | 1 |  | 47.9029 | 48.3829 | 47.9029 |  |  |  |  | 116.6233 | PASS | 1 |  | 920.2523 | 936.1023 | 920.2523 |  |  |  |  | 90.904 | PASS |  | 1
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 042 | 42 | 1 | 1 | Total Phosphorus | ppb | 0.0698 | 1.6 | 28.2 | Total Nitrogen | ppb | 1.45 | 22.8 | 671 |  | 1.5 | 1 |  | 32.7393 | 33.2193 | 32.7393 |  |  |  |  | 124.4217 | FAIL | 1 |  | 386.469 | 402.319 | 386.469 |  |  |  |  | 84.575 | PASS |  | 2
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 054 | 54 | 1 | 1 | Total Phosphorus | ppb | 0.0798 | 1.24 | 21.4 | Total Nitrogen | ppb | 1.65 | 26 | 766 |  | 1.5 | 1 |  | 20.6354 | 21.1154 | 20.6354 |  |  |  |  | -6.1798 | FAIL | 1 |  | 509.6928 | 525.5428 | 509.6928 |  |  |  |  | 6.6329 | FAIL |  | 3
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 066 | 66 | 1 | 1 | Total Phosphorus | ppb | 0.117 | 2.5 | 45 | Total Nitrogen | ppb | 2.07 | 32.5 | 962 |  | 1.5 | 1 |  | 60.257 | 60.737 | 60.257 |  |  |  |  | 131.9564 | FAIL | 1 |  | 757.7723 | 773.6223 | 757.7723 |  |  |  |  | 94.5322 | PASS |  | 4
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 090 | 90 | 1 | 1 | Total Phosphorus | ppb | 0.112 | 2.27 | 40.7 | Total Nitrogen | ppb | 2.08 | 32.7 | 967 |  | 1.5 | 1 |  | 53.5973 | 54.0773 | 53.5973 |  |  |  |  | 103.3194 | PASS | 1 |  | 765.3584 | 781.2084 | 765.3584 |  |  |  |  | 85.3832 | PASS |  | 5
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 102 | 102 | 1 | 1 | Total Phosphorus | ppb | 0.0902 | 2.02 | 36 | Total Nitrogen | ppb | 1.65 | 22.4 | 660 |  | 1.5 | 1 |  | 46.0684 | 46.5484 | 46.0684 |  |  |  |  | 137.16 | FAIL | 3 | chk std failed | 371.0153 | 386.8653 | 371.0153 | [7 empty] | 6
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 114 | 114 | 1 | 1 | Total Phosphorus | ppb | 0.101 | 2.2 | 39.4 | Total Nitrogen | ppb | 1.5 | 17.8 | 523 |  | 1.5 | 1 |  | 51.5197 | 51.9997 | 51.5197 |  |  |  |  | 64.692 | FAIL | 3 | chk std failed | 192.4891 | 208.3391 | 192.4891 | [7 empty] | 7
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 126 | 126 | 1 | 1 | Total Phosphorus | ppb | 0.116 | 2.44 | 43.8 | Total Nitrogen | ppb | 1.7 | 18.2 | 535 |  | 1.5 | 1 |  | 58.5444 | 59.0244 | 58.5444 |  |  |  |  | 111.6986 | PASS | 3 | chk std failed | 208.0722 | 223.9222 | 208.0722 | [7 empty] | 8
2020-03-26 | OM_3-26-2020_01-54-43PM.OMN | Sample 138 | 138 | 1 | 1 | Total Phosphorus | ppb | 0.115 | 2.43 | 43.7 | Total Nitrogen | ppb | 1.5 | 14.9 | 436 |  | 1.5 | 1 |  | 58.2572 | 58.7372 | 58.2572 |  |  |  |  | 106.5882 | PASS | 3 | chk std failed | 79.1745 | 95.0245 | 79.1745 | [7 empty] | 9
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | Sample 102 | 102 | 1 | 1 | Total Phosphorus | ppb | 0.0832 | 1.74 | 36.5 | Total Nitrogen | ppb | 1.48 | 25.6 | 668 |  | 1.5 | 1 |  | 37.2771 | 37.7571 | 37.2771 |  |  |  |  | 136.3946 | FAIL | 1 |  | 494.3292 | 510.1792 | 494.3292 |  |  |  |  | 93.005 | PASS |  | 10
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | Sample 114 | 114 | 1 | 1 | Total Phosphorus | ppb | 0.0925 | 1.91 | 39.4 | Total Nitrogen | ppb | 1.5 | 25.9 | 674 |  | 1.5 | 1 |  | 42.6599 | 43.1399 | 42.6599 |  |  |  |  | 70.5362 | FAIL | 1 |  | 505.8529 | 521.7029 | 505.8529 |  |  |  |  | 69.822 | FAIL |  | 11
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | Sample 126 | 126 | 1 | 1 | Total Phosphorus | ppb | 0.11 | 2.2 | 44.5 | Total Nitrogen | ppb | 1.84 | 31.9 | 829 |  | 1.5 | 1 |  | 51.5197 | 51.9997 | 51.5197 |  |  |  |  | 82.8974 | PASS | 1 |  | 734.9972 | 750.8472 | 734.9972 |  |  |  |  | 80.5951 | PASS |  | 12
2020-03-27 | OM_3-27-2020_12-13-42PM.OMN | Sample 138 | 138 | 1 | 1 | Total Phosphorus | ppb | 0.109 | 2.19 | 44.3 | Total Nitrogen | ppb | 1.86 | 32.2 | 836 |  | 1.5 | 1 |  | 51.2209 | 51.7009 | 51.2209 |  |  |  |  | 113.7639 | PASS | 1 |  | 746.3879 | 762.2379 | 746.3879 |  |  |  |  | 83.7327 | PASS |  | 13
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | F 04oct19 50 t r1 SPK | 30 | 1 | 1 | Total Phosphorus | ppb | 0.125 | 2.6 | 43.8 | Total Nitrogen | ppb | 2.29 | 40.1 | 1080 |  | 1.5 | 1 |  | 58.6996 | 56.6696 | 58.6996 |  |  |  |  | 117.4128 | PASS | 1 |  | 1180.8045 | 1174.5645 | 1180.8045 |  |  |  |  | 105.9369 | PASS |  | 14
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | B 22aug19 100 t r2 SPK | 42 | 1 | 1 | Total Phosphorus | ppb | 0.222 | 4.28 | 73.9 | Total Nitrogen | ppb | 1.29 | 22.6 | 613 |  | 1.5 | 1 |  | 103.0449 | 101.0149 | 103.0449 |  |  |  |  | 108.8578 | PASS | 1 |  | 452.7102 | 446.4702 | 452.7102 |  |  |  |  | 94.4114 | PASS |  | 15
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | F 19aug19 t 3.8 SPK | 54 | 1 | 1 | Total Phosphorus | ppb | 0.154 | 2.64 | 44.6 | Total Nitrogen | ppb | 1.36 | 24 | 650 |  | 1.5 | 1 |  | 59.7718 | 57.7418 | 59.7718 |  |  |  |  | 115.2531 | PASS | 1 |  | 514.3004 | 508.0604 | 514.3004 |  |  |  |  | 99.6344 | PASS |  | 16
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | F 19aug19 t 9 SPK | 66 | 1 | 1 | Total Phosphorus | ppb | 0.124 | 2.52 | 42.6 | Total Nitrogen | ppb | 1.81 | 31.9 | 863 |  | 1.5 | 1 |  | 56.5527 | 54.5227 | 56.5527 |  |  |  |  | 100.9628 | PASS | 1 |  | 850.9497 | 844.7097 | 850.9497 |  |  |  |  | 94.5946 | PASS |  | 17
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | B 04oct19 50 t r1 SPK | 78 | 1 | 1 | Total Phosphorus | ppb | 0.0971 | 2.02 | 33.8 | Total Nitrogen | ppb | 1.32 | 23.1 | 625 |  | 1.5 | 1 |  | 43.0622 | 41.0322 | 43.0622 |  |  |  |  | 92.7858 | PASS | 1 |  | 474.7734 | 468.5334 | 474.7734 |  |  |  |  | 96.1574 | PASS |  | 18
2020-03-31 | OM_3-31-2020_11-13-11AM.OMN | F 23oct19 t isco r2 SPK | 90 | 1 | 1 | Total Phosphorus | ppb | 0.3 | 5.69 | 100 | Total Nitrogen | ppb | 1.71 | 29.9 | 808 |  | 1.5 | 1 |  | 139.174 | 137.144 | 139.174 |  |  |  |  | 97.4787 | PASS | 1 |  | 767.4719 | 761.2319 | 767.4719 |  |  |  |  | 89.7495 | PASS |  | 19
2020-10-22 | OM_10-22-2020_10-45-05AM.OMN | Sample 044 | 44 | 1 | 1 | TP | ug P/L | 0.1 | 2.04 | 55.3 | TN | ug N/L | 0.736 | 12.4 | 341 |  | 1 | 1 |  | 55.3 | 55.3 | 55.3 |  |  |  |  | 111.93 | PASS | 1 |  | 341 | 341 | 341 |  |  |  |  | 171.65 | FAIL |  | 20
2020-10-22 | OM_10-22-2020_10-45-05AM.OMN | Sample 056 | 56 | 1 | 1 | TP | ug P/L | 0.0721 | 1.55 | 40.1 | TN | ug N/L | 0.502 | 8.56 | 72.4 |  | 1 | 1 |  | 40.1 | 40.1 | 40.1 |  |  |  |  | 114.81 | PASS | 1 |  | 72.4 | 72.4 | 72.4 |  |  |  |  | 105.684 | PASS |  | 21
2020-10-22 | OM_10-22-2020_11-52-24AM.OMN | Sample 092 | 92 | 1 | 1 | TP | ug P/L | 0.219 | 4.21 | 115 | TN | ug N/L | 1.02 | 17.4 | 743 |  | 1 | 1 |  | 115 | 115 | 115 |  |  |  |  | 83.5 | PASS | 1 |  | 743 | 743 | 743 |  |  |  |  | 146.23 | FAIL |  | 22
2020-10-22 | OM_10-22-2020_11-52-24AM.OMN | Sample 104 | 104 | 1 | 1 | TP | ug P/L | 0.108 | 2.18 | 57.9 | TN | ug N/L | 1.05 | 18 | 786 |  | 1 | 1 |  | 57.9 | 57.9 | 57.9 |  |  |  |  | 110.99 | PASS | 1 |  | 786 | 786 | 786 |  |  |  |  | 187.46 | FAIL |  | 23
2020-10-22 | OM_10-22-2020_11-52-24AM.OMN | Sample 116 | 116 | 1 | 1 | TP | ug P/L | 0.124 | 2.55 | 68.6 | TN | ug N/L | 0.748 | 12.8 | 394 |  | 1 | 1 |  | 68.6 | 68.6 | 68.6 |  |  |  |  | 119.26 | PASS | 1 |  | 394 | 394 | 394 |  |  |  |  | 95.54 | PASS |  | 24
2020-10-22 | OM_10-22-2020_11-52-24AM.OMN | Sample 140 | 140 | 1 | 1 | TP | ug P/L | 0.0542 | 1.2 | 28.1 | TN | ug N/L | 0.563 | 9.51 | 141 |  | 1 | 1 |  | 28.1 | 28.1 | 28.1 |  |  |  |  | -17.19 | FAIL | 1 |  | 141 | 141 | 141 |  |  |  |  | -39.39 | FAIL |  | 25
2020-10-22 | OM_10-22-2020_11-52-24AM.OMN | Sample 152 | 152 | 1 | 1 | TP | ug P/L | 0.057 | 1.32 | 31.9 | TN | ug N/L | 0.869 | 14.9 | 553 |  | 1 | 1 |  | 31.9 | 31.9 | 31.9 |  |  |  |  | 6.79 | FAIL | 1 |  | 553 | 553 | 553 |  |  |  |  | 13.93 | FAIL |  | 26
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 046 | 46 | 1 | 1 | TP | ug P/L | 0.124 | 2.55 | 63.5 | TN | ug N/L | 0.642 | 10.9 | 654 |  | 1 | 1 |  | 63.5 | 63.5 | 63.5 |  |  |  |  | 78.35 | FAIL | 1 |  | 654 | 654 | 654 |  |  |  |  | 86.14 | PASS |  | 27
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 058 | 58 | 1 | 1 | TP | ug P/L | 0.0889 | 1.95 | 45.9 | TN | ug N/L | 0.473 | 8.02 | 433 |  | 1 | 1 |  | 45.9 | 45.9 | 45.9 |  |  |  |  | 98.99 | PASS | 1 |  | 433 | 433 | 433 |  |  |  |  | 90.33 | PASS |  | 28
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 070 | 70 | 1 | 1 | TP | ug P/L | 0.119 | 2.46 | 60.9 | TN | ug N/L | 0.801 | 13.6 | 856 |  | 1 | 1 |  | 60.9 | 60.9 | 60.9 |  |  |  |  | 100.49 | PASS | 1 |  | 856 | 856 | 856 |  |  |  |  | 78.96 | FAIL |  | 29
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 082 | 82 | 1 | 1 | TP | ug P/L | 0.174 | 3.44 | 89.1 | TN | ug N/L | 1.61 | 27.2 | 1950 |  | 1 | 1 |  | 89.1 | 89.1 | 89.1 |  |  |  |  | 100.51 | PASS | 1 |  | 1950 | 1950 | 1950 |  |  |  |  | 91.5 | PASS |  | 30
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 094 | 94 | 1 | 1 | TP | ug P/L | 0.103 | 2.17 | 52.4 | TN | ug N/L | 0.906 | 15.3 | 994 |  | 1 | 1 |  | 52.4 | 52.4 | 52.4 |  |  |  |  | 101.24 | PASS | 1 |  | 994 | 994 | 994 |  |  |  |  | 82.34 | PASS |  | 31
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 106 | 106 | 1 | 1 | TP | ug P/L | 0.134 | 2.7 | 67.7 | TN | ug N/L | 0.669 | 11.4 | 691 |  | 1 | 1 |  | 67.7 | 67.7 | 67.7 |  |  |  |  | 95.97 | PASS | 1 |  | 691 | 691 | 691 |  |  |  |  | 89.31 | PASS |  | 32
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 118 | 118 | 1 | 1 | TP | ug P/L | 0.0994 | 2.01 | 47.5 | TN | ug N/L | 1.26 | 21.5 | 1480 |  | 1 | 1 |  | 47.5 | 47.5 | 47.5 |  |  |  |  | 96.35 | PASS | 1 |  | 1480 | 1480 | 1480 |  |  |  |  | 86.8 | PASS |  | 33
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 130 | 130 | 1 | 1 | TP | ug P/L | 0.101 | 2.1 | 50.3 | TN | ug N/L | 0.618 | 10.5 | 621 |  | 1 | 1 |  | 50.3 | 50.3 | 50.3 |  |  |  |  | 96.23 | PASS | 1 |  | 621 | 621 | 621 |  |  |  |  | 94.61 | PASS |  | 34
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 142 | 142 | 1 | 1 | TP | ug P/L | 0.112 | 2.36 | 57.7 | TN | ug N/L | 0.701 | 12 | 734 |  | 1 | 1 |  | 57.7 | 57.7 | 57.7 |  |  |  |  | 100.57 | PASS | 1 |  | 734 | 734 | 734 |  |  |  |  | 91.74 | PASS |  | 35
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | Sample 154 | 154 | 1 | 1 | TP | ug P/L | 0.103 | 2.11 | 50.6 | TN | ug N/L | 0.348 | 5.99 | 282 |  | 1 | 1 |  | 50.6 | 50.6 | 50.6 |  |  |  |  | 111.06 | PASS | 1 |  | 282 | 282 | 282 |  |  |  |  | 97.14 | PASS |  | 36
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 46 | 1 | 1 | TP | ug P/L | 0.153 | 3.11 | 69.8 | TN | ug N/L | 0.29 | 5.25 | 372 |  | 1 | 2 | reworked calibration | 81.1145 | 81.1145 | 81.1145 |  |  |  |  | 96.0675 | PASS | 1 |  | 397.1273 | 397.1273 | 397.1273 |  |  |  |  | 94.882 | PASS |  | 37
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 58 | 1 | 1 | TP | ug P/L | 0.0973 | 2.05 | 39.5 | TN | ug N/L | 0.355 | 6.39 | 472 |  | 1 | 2 | reworked calibration | 48.4233 | 48.4233 | 48.4233 |  |  |  |  | 98.9386 | PASS | 1 |  | 508.191 | 508.191 | 508.191 |  |  |  |  | 104.264 | PASS |  | 38
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 70 | 1 | 1 | TP | ug P/L | 0.0924 | 1.95 | 36.8 | TN | ug N/L | 0.345 | 6.23 | 458 |  | 1 | 2 | reworked calibration | 45.3822 | 45.3822 | 45.3822 |  |  |  |  | 98.4033 | PASS | 1 |  | 492.4996 | 492.4996 | 492.4996 |  |  |  |  | 118.9796 | PASS |  | 39
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 82 | 1 | 1 | TP | ug P/L | 0.118 | 2.4 | 49.3 | TN | ug N/L | 0.564 | 10.2 | 801 |  | 1 | 2 | reworked calibration | 59.1255 | 59.1255 | 59.1255 |  |  |  |  | 100.818 | PASS | 1 |  | 891.8231 | 891.8231 | 891.8231 |  |  |  |  | 106.4747 | PASS |  | 40
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 94 | 1 | 1 | TP | ug P/L | 0.117 | 2.42 | 49.9 | TN | ug N/L | 0.525 | 9.4 | 734 |  | 1 | 2 | reworked calibration | 59.7398 | 59.7398 | 59.7398 |  |  |  |  | 102.1313 | PASS | 1 |  | 809.6814 | 809.6814 | 809.6814 |  |  |  |  | 107.4098 | PASS |  | 41
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 106 | 1 | 1 | TP | ug P/L | 0.0941 | 2.02 | 38.6 | TN | ug N/L | 0.374 | 6.83 | 511 |  | 1 | 2 | reworked calibration | 47.5102 | 47.5102 | 47.5102 |  |  |  |  | 110.7022 | PASS | 1 |  | 551.5165 | 551.5165 | 551.5165 |  |  |  |  | 104.1497 | PASS |  | 42
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 118 | 1 | 1 | TP | ug P/L | 0.0801 | 1.73 | 30.6 | TN | ug N/L | 0.362 | 6.67 | 497 |  | 1 | 2 | reworked calibration | 38.7178 | 38.7178 | 38.7178 |  |  |  |  | 103.743 | PASS | 1 |  | 535.7322 | 535.7322 | 535.7322 |  |  |  |  | 101.4917 | PASS |  | 43
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 130 | 1 | 1 | TP | ug P/L | 0.0842 | 1.79 | 32.4 | TN | ug N/L | 0.267 | 4.92 | 343 |  | 1 | 2 | reworked calibration | 40.5318 | 40.5318 | 40.5318 |  |  |  |  | 96.3611 | PASS | 1 |  | 365.2974 | 365.2974 | 365.2974 |  |  |  |  | 78.8375 | FAIL |  | 44
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 142 | 1 | 1 | TP | ug P/L | 0.108 | 2.25 | 45.2 | TN | ug N/L | 0.462 | 8.49 | 656 |  | 1 | 2 | reworked calibration | 54.5277 | 54.5277 | 54.5277 |  |  |  |  | 117.989 | PASS | 1 |  | 717.2721 | 717.2721 | 717.2721 |  |  |  |  | 98.6487 | PASS |  | 45
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | SPK | 154 | 1 | 1 | TP | ug P/L | 0.109 | 2.25 | 45.1 | TN | ug N/L | 0.482 | 8.87 | 688 |  | 1 | 2 | reworked calibration | 54.5277 | 54.5277 | 54.5277 |  |  |  |  | 102.4123 | PASS | 1 |  | 755.7277 | 755.7277 | 755.7277 |  |  |  |  | 90.0317 | PASS |  | 46
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 46 | 1 | 1 | TP | ug P/L | 0.0959 | 1.95 | 52.1 | TN | ug N/L | 0.645 | 9.21 | 552 |  | 1 | 1 |  | 52.1 | 52.1 | 52.1 |  |  |  |  | 96.01 | PASS | 1 |  | 552 | 552 | 552 |  |  |  |  | 92.72 | PASS |  | 47
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 58 | 1 | 1 | TP | ug P/L | 0.105 | 2.13 | 57.7 | TN | ug N/L | 0.682 | 9.74 | 587 |  | 1 | 1 |  | 57.7 | 57.7 | 57.7 |  |  |  |  | 113.37 | PASS | 1 |  | 587 | 587 | 587 |  |  |  |  | 85.07 | PASS |  | 48
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 70 | 1 | 1 | TP | ug P/L | 0.0757 | 1.61 | 41.6 | TN | ug N/L | 0.603 | 8.63 | 512 |  | 1 | 1 |  | 41.6 | 41.6 | 41.6 |  |  |  |  | 92.16 | PASS | 1 |  | 512 | 512 | 512 |  |  |  |  | 88.72 | PASS |  | 49
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 82 | 1 | 1 | TP | ug P/L | 0.0611 | 1.3 | 31.8 | TN | ug N/L | 0.424 | 6.27 | 351 |  | 1 | 1 |  | 31.8 | 31.8 | 31.8 |  |  |  |  | 97.3 | PASS | 1 |  | 351 | 351 | 351 |  |  |  |  | 108.19 | PASS |  | 50
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 94 | 1 | 1 | TP | ug P/L | 0.178 | 3.38 | 95.1 | TN | ug N/L | 0.849 | 12 | 738 |  | 1 | 1 |  | 95.1 | 95.1 | 95.1 |  |  |  |  | 91.11 | PASS | 1 |  | 738 | 738 | 738 |  |  |  |  | 81.38 | PASS |  | 51
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 106 | 1 | 1 | TP | ug P/L | 0.0709 | 1.56 | 39.8 | TN | ug N/L | 0.525 | 7.51 | 437 |  | 1 | 1 |  | 39.8 | 39.8 | 39.8 |  |  |  |  | 105.58 | PASS | 1 |  | 437 | 437 | 437 |  |  |  |  | 87.97 | PASS |  | 52
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 118 | 1 | 1 | TP | ug P/L | 0.0988 | 2.01 | 53.8 | TN | ug N/L | 1.1 | 15.5 | 963 |  | 1 | 1 |  | 53.8 | 53.8 | 53.8 |  |  |  |  | 69.78 | FAIL | 1 |  | 963 | 963 | 963 |  |  |  |  | 90.03 | PASS |  | 53
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 130 | 1 | 1 | TP | ug P/L | 0.0784 | 1.71 | 44.7 | TN | ug N/L | 0.499 | 7.17 | 413 |  | 1 | 1 |  | 44.7 | 44.7 | 44.7 |  |  |  |  | 120.07 | FAIL | 1 |  | 413 | 413 | 413 |  |  |  |  | 83.73 | PASS |  | 54
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 142 | 1 | 1 | TP | ug P/L | 0.0629 | 1.37 | 34 | TN | ug N/L | 0.572 | 8.2 | 483 |  | 1 | 1 |  | 34 | 34 | 34 |  |  |  |  | 103.32 | PASS | 1 |  | 483 | 483 | 483 |  |  |  |  | 106.83 | PASS |  | 55
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 154 | 1 | 1 | TP | ug P/L | 0.0766 | 1.59 | 40.8 | TN | ug N/L | 0.612 | 8.7 | 517 |  | 1 | 1 |  | 40.8 | 40.8 | 40.8 |  |  |  |  | 108.88 | PASS | 1 |  | 517 | 517 | 517 |  |  |  |  | 112.77 | PASS |  | 56
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK | 166 | 1 | 1 | TP | ug P/L | 0.0773 | 1.49 | 37.8 | TN | ug N/L | 0.582 | 8.3 | 491 |  | 1 | 1 |  | 37.8 | 37.8 | 37.8 |  |  |  |  | 111.78 | PASS | 1 |  | 491 | 491 | 491 |  |  |  |  | 95.71 | PASS |  | 57
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | SPK F 17aug20 9.0 | 178 | 1 | 1 | TP | ug P/L | 0.119 | 2.29 | 62.5 | TN | ug N/L | 0.833 | 11.8 | 724 |  | 1 | 1 |  | 62.5 | 62.5 | 62.5 |  |  |  |  | 99.85 | PASS | 1 |  | 724 | 724 | 724 |  |  |  |  | 86.44 | PASS |  | 58
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 46 | 1 | 1 | TP | ug P/L | 0.082 | 1.62 | 35.8 | TN | ug N/L | 0.491 | 7.24 | 396 |  | 1 | 1 |  | 35.8 | 35.8 | 35.8 |  |  |  |  | 103.18 | PASS | 1 |  | 396 | 396 | 396 |  |  |  |  | 97.56 | PASS |  | 59
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 58 | 1 | 1 | TP | ug P/L | 0.104 | 2.04 | 47.3 | TN | ug N/L | 0.519 | 7.63 | 422 |  | 1 | 1 |  | 47.3 | 47.3 | 47.3 |  |  |  |  | 98.33 | PASS | 1 |  | 422 | 422 | 422 |  |  |  |  | 90.22 | PASS |  | 60
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 82 | 1 | 1 | TP | ug P/L | 0.11 | 2.16 | 50.7 | TN | ug N/L | 0.64 | 9.34 | 532 |  | 1 | 1 |  | 50.7 | 50.7 | 50.7 |  |  |  |  | 109.87 | PASS | 1 |  | 532 | 532 | 532 |  |  |  |  | 120.92 | FAIL |  | 61
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 94 | 1 | 1 | TP | ug P/L | 0.123 | 2.34 | 55.7 | TN | ug N/L | 0.709 | 10.3 | 594 |  | 1 | 1 |  | 55.7 | 55.7 | 55.7 |  |  |  |  | 101.17 | PASS | 1 |  | 594 | 594 | 594 |  |  |  |  | 103.54 | PASS |  | 62
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 106 | 1 | 1 | TP | ug P/L | 0.107 | 2.05 | 47.7 | TN | ug N/L | 0.55 | 8.05 | 448 |  | 1 | 1 |  | 47.7 | 47.7 | 47.7 |  |  |  |  | 98.77 | PASS | 1 |  | 448 | 448 | 448 |  |  |  |  | 88.88 | PASS |  | 63
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 118 | 1 | 1 | TP | ug P/L | 0.123 | 2.33 | 55.3 | TN | ug N/L | 0.528 | 7.66 | 424 |  | 1 | 1 |  | 55.3 | 55.3 | 55.3 |  |  |  |  | 104.73 | PASS | 1 |  | 424 | 424 | 424 |  |  |  |  | 89.84 | PASS |  | 64
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 130 | 1 | 1 | TP | ug P/L | 0.133 | 2.53 | 61 | TN | ug N/L | 0.563 | 8.19 | 458 |  | 1 | 1 |  | 61 | 61 | 61 |  |  |  |  | 105.3 | PASS | 1 |  | 458 | 458 | 458 |  |  |  |  | 69.38 | FAIL |  | 65
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 142 | 1 | 1 | TP | ug P/L | 0.0964 | 1.86 | 42.5 | TN | ug N/L | 0.628 | 9.15 | 520 |  | 1 | 1 |  | 42.5 | 42.5 | 42.5 |  |  |  |  | 98.65 | PASS | 1 |  | 520 | 520 | 520 |  |  |  |  | 90.4 | PASS |  | 66
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | SPK | 154 | 1 | 1 | TP | ug P/L | 0.0777 | 1.53 | 33.3 | TN | ug N/L | 0.482 | 7.04 | 383 |  | 1 | 1 |  | 33.3 | 33.3 | 33.3 |  |  |  |  | 100.93 | PASS | 1 |  | 383 | 383 | 383 |  |  |  |  | 97.83 | PASS |  | 67
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 046 | 46 | 1 | 1 | TP | ug P/L | 0.115 | 2.08 | 51.9 | TN | ug N/L | 1.03 | 13.9 | 853 |  | 1 | 1 |  | 51.9 | 51.9 | 51.9 |  |  |  |  | 103.59 | PASS | 1 |  | 853 | 853 | 853 |  |  |  |  | 76.93 | FAIL |  | 68
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 058 | 58 | 1 | 1 | TP | ug P/L | 0.105 | 1.89 | 46.4 | TN | ug N/L | 0.63 | 8.52 | 484 |  | 1 | 1 |  | 46.4 | 46.4 | 46.4 |  |  |  |  | 97.84 | PASS | 1 |  | 484 | 484 | 484 |  |  |  |  | 89.24 | PASS |  | 69
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 070 | 70 | 1 | 1 | TP | ug P/L | 0.121 | 2.1 | 52.5 | TN | ug N/L | 0.824 | 11.1 | 658 |  | 1 | 1 |  | 52.5 | 52.5 | 52.5 |  |  |  |  | 104.05 | PASS | 1 |  | 658 | 658 | 658 |  |  |  |  | 104.58 | PASS |  | 70
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 082 | 82 | 1 | 1 | TP | ug P/L | 0.195 | 3.15 | 82.2 | TN | ug N/L | 0.995 | 13.4 | 818 |  | 1 | 1 |  | 82.2 | 82.2 | 82.2 |  |  |  |  | 110.22 | PASS | 1 |  | 818 | 818 | 818 |  |  |  |  | 25.78 | FAIL |  | 71
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 094 | 94 | 1 | 1 | TP | ug P/L | 0.118 | 2.06 | 51.3 | TN | ug N/L | 0.631 | 8.54 | 486 |  | 1 | 1 |  | 51.3 | 51.3 | 51.3 |  |  |  |  | 101.53 | PASS | 1 |  | 486 | 486 | 486 |  |  |  |  | 110.46 | PASS |  | 72
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 106 | 106 | 1 | 1 | TP | ug P/L | 0.0987 | 1.75 | 42.4 | TN | ug N/L | 0.664 | 9.01 | 517 |  | 1 | 1 |  | 42.4 | 42.4 | 42.4 |  |  |  |  | 104.64 | PASS | 1 |  | 517 | 517 | 517 |  |  |  |  | 77.97 | FAIL |  | 73
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 118 | 118 | 1 | 1 | TP | ug P/L | 0.116 | 2.06 | 51.4 | TN | ug N/L | 0.844 | 11.4 | 678 |  | 1 | 1 |  | 51.4 | 51.4 | 51.4 |  |  |  |  | 107.54 | PASS | 1 |  | 678 | 678 | 678 |  |  |  |  | 89.18 | PASS |  | 74
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 130 | 130 | 1 | 1 | TP | ug P/L | 0.0743 | 1.35 | 30.9 | TN | ug N/L | 0.466 | 6.35 | 341 |  | 1 | 1 |  | 30.9 | 30.9 | 30.9 |  |  |  |  | 98.33 | PASS | 1 |  | 341 | 341 | 341 |  |  |  |  | 87.41 | PASS |  | 75
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 142 | 142 | 1 | 1 | TP | ug P/L | 0.0944 | 1.74 | 42.1 | TN | ug N/L | 1.23 | 16.6 | 1040 |  | 1 | 1 |  | 42.1 | 42.1 | 42.1 |  |  |  |  | 108.21 | PASS | 1 |  | 1040 | 1040 | 1040 |  |  |  |  | 75.2 | FAIL |  | 76
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 154 | 154 | 1 | 1 | TP | ug P/L | 0.13 | 2.26 | 57 | TN | ug N/L | 1.02 | 13.8 | 844 |  | 1 | 1 |  | 57 | 57 | 57 |  |  |  |  | 91.7 | PASS | 1 |  | 844 | 844 | 844 |  |  |  |  | 78.84 | FAIL |  | 77
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample46 | 46 | 1 | 1 | TP | ug P/L | 0.0996 | 1.75 | 35.1 | TN | ug N/L | 0.513 | 6.96 | 379 |  | 1 | 1 |  | 35.1 | 35.1 | 35.1 |  |  |  |  | 81.11 | PASS | 1 |  | 379 | 379 | 379 |  |  |  |  | 43.79 | FAIL |  | 78
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample58 | 58 | 1 | 1 | TP | ug P/L | 0.105 | 1.93 | 39.4 | TN | ug N/L | 0.513 | 6.74 | 364 |  | 1 | 1 |  | 39.4 | 39.4 | 39.4 |  |  |  |  | 88.74 | PASS | 1 |  | 364 | 364 | 364 |  |  |  |  | 91.24 | PASS |  | 79
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample70 | 70 | 1 | 1 | TP | ug P/L | 0.106 | 1.89 | 38.4 | TN | ug N/L | 0.672 | 9.02 | 517 |  | 1 | 1 |  | 38.4 | 38.4 | 38.4 |  |  |  |  | 67.84 | FAIL | 1 |  | 517 | 517 | 517 |  |  |  |  | 77.97 | FAIL |  | 80
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample82 | 82 | 1 | 1 | TP | ug P/L | 0.123 | 2.16 | 45.3 | TN | ug N/L | 0.623 | 8.29 | 468 |  | 1 | 1 |  | 45.3 | 45.3 | 45.3 |  |  |  |  | 104.13 | PASS | 1 |  | 468 | 468 | 468 |  |  |  |  | 82.28 | PASS |  | 81
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample94 | 94 | 1 | 1 | TP | ug P/L | 0.201 | 3.48 | 79.3 | TN | ug N/L | 1.38 | 18.4 | 1150 |  | 1 | 1 |  | 79.3 | 79.3 | 79.3 |  |  |  |  | 62.33 | FAIL | 1 |  | 1150 | 1150 | 1150 |  |  |  |  | -16.5 | FAIL |  | 82
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample106 | 106 | 1 | 1 | TP | ug P/L | 0.132 | 2.33 | 49.6 | TN | ug N/L | 1.13 | 15.3 | 939 |  | 1 | 1 |  | 49.6 | 49.6 | 49.6 |  |  |  |  | 80.16 | PASS | 1 |  | 939 | 939 | 939 |  |  |  |  | 45.39 | FAIL |  | 83
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample118 | 118 | 1 | 1 | TP | ug P/L | 0.108 | 1.92 | 39.3 | TN | ug N/L | 0.745 | 10.1 | 588 |  | 1 | 1 |  | 39.3 | 39.3 | 39.3 |  |  |  |  | 67.93 | FAIL | 1 |  | 588 | 588 | 588 |  |  |  |  | 79.88 | FAIL |  | 84
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample130 | 130 | 1 | 1 | TP | ug P/L | 0.0823 | 1.48 | 28.3 | TN | ug N/L | 0.448 | 6.06 | 318 |  | 1 | 1 |  | 28.3 | 28.3 | 28.3 |  |  |  |  | 80.19 | PASS | 1 |  | 318 | 318 | 318 |  |  |  |  | 90.46 | PASS |  | 85
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample142 | 142 | 1 | 1 | TP | ug P/L | 0.0821 | 1.47 | 28 | TN | ug N/L | 0.486 | 6.56 | 352 |  | 1 | 1 |  | 28 | 28 | 28 |  |  |  |  | 88.88 | PASS | 1 |  | 352 | 352 | 352 |  |  |  |  | 105.84 | PASS |  | 86
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | Sample154 | 154 | 1 | 1 | TP | ug P/L | 0.0985 | 1.75 | 35 | TN | ug N/L | 0.639 | 8.51 | 482 |  | 1 | 1 |  | 35 | 35 | 35 |  |  |  |  | 87.9 | PASS | 1 |  | 482 | 482 | 482 |  |  |  |  | 95.22 | PASS |  | 87
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 046 | 46 | 1 | 1 | TP | ug P/L | 0.12 | 2.07 | 48.9 | TN | ug N/L | 1.91 | 26.2 | 1730 |  | 1 | 1 |  | 48.9 | 48.9 | 48.9 |  |  |  |  | 92.89 | PASS | 1 |  | 1730 | 1730 | 1730 |  |  |  |  | 61.3 | FAIL |  | 88
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 058 | 58 | 1 | 1 | TP | ug P/L | 0.12 | 2.06 | 48.7 | TN | ug N/L | 0.587 | 8.02 | 491 |  | 1 | 1 |  | 48.7 | 48.7 | 48.7 |  |  |  |  | 130.87 | FAIL | 1 |  | 491 | 491 | 491 |  |  |  |  | 91.71 | PASS |  | 89
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 070 | 70 | 1 | 1 | TP | ug P/L | 0.0988 | 1.71 | 38.9 | TN | ug N/L | 0.538 | 7.37 | 445 |  | 1 | 1 |  | 38.9 | 38.9 | 38.9 |  |  |  |  | 129.09 | FAIL | 1 |  | 445 | 445 | 445 |  |  |  |  | 102.45 | PASS |  | 90
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 082 | 82 | 1 | 1 | TP | ug P/L | 0.16 | 2.76 | 67.9 | TN | ug N/L | 1.03 | 14.1 | 913 |  | 1 | 1 |  | 67.9 | 67.9 | 67.9 |  |  |  |  | 83.59 | PASS | 1 |  | 913 | 913 | 913 |  |  |  |  | 177.93 | FAIL |  | 91
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 094 | 94 | 1 | 1 | TP | ug P/L | 0.119 | 2.04 | 48.1 | TN | ug N/L | 0.565 | 7.83 | 477 |  | 1 | 1 |  | 48.1 | 48.1 | 48.1 |  |  |  |  | 132.41 | FAIL | 1 |  | 477 | 477 | 477 |  |  |  |  | 104.37 | PASS |  | 92
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 106 | 106 | 1 | 1 | TP | ug P/L | 0.108 | 1.86 | 43.2 | TN | ug N/L | 0.5 | 6.86 | 409 |  | 1 | 1 |  | 43.2 | 43.2 | 43.2 |  |  |  |  | 135.52 | FAIL | 1 |  | 409 | 409 | 409 |  |  |  |  | 97.29 | PASS |  | 93
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 118 | 118 | 1 | 1 | TP | ug P/L | 0.16 | 2.74 | 67.6 | TN | ug N/L | 0.503 | 6.95 | 415 |  | 1 | 1 |  | 67.6 | 67.6 | 67.6 |  |  |  |  | 123.16 | FAIL | 1 |  | 415 | 415 | 415 |  |  |  |  | 125.35 | FAIL |  | 94
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 130 | 130 | 1 | 1 | TP | ug P/L | 0.15 | 2.58 | 63.1 | TN | ug N/L | 2.02 | 27.9 | 1840 |  | 1 | 1 |  | 63.1 | 63.1 | 63.1 |  |  |  |  | 153.91 | FAIL | 1 |  | 1840 | 1840 | 1840 |  |  |  |  | 66.4 | FAIL |  | 95
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 142 | 142 | 1 | 1 | TP | ug P/L | 0.142 | 2.46 | 59.6 | TN | ug N/L | 0.768 | 10.6 | 667 |  | 1 | 1 |  | 59.6 | 59.6 | 59.6 |  |  |  |  | 137.16 | FAIL | 1 |  | 667 | 667 | 667 |  |  |  |  | 109.07 | PASS |  | 96
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 154 | 154 | 1 | 1 | TP | ug P/L | 0.153 | 2.59 | 63.4 | TN | ug N/L | 1.09 | 14.9 | 969 |  | 1 | 1 |  | 63.4 | 63.4 | 63.4 |  |  |  |  | 133.14 | FAIL | 1 |  | 969 | 969 | 969 |  |  |  |  | 90.49 | PASS |  | 97
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 046 | 46 | 1 | 1 | TP | ug P/L | 0.134 | 2.26 | 54.1 | TN | ug N/L | 0.748 | 9.05 | 548 |  | 1 | 1 |  | 54.1 | 54.1 | 54.1 |  |  |  |  | 116.21 | PASS | 1 |  | 548 | 548 | 548 |  |  |  |  | 98.68 | PASS |  | 98
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 058 | 58 | 1 | 1 | TP | ug P/L | 0.124 | 2.03 | 47.2 | TN | ug N/L | 0.817 | 9.88 | 605 |  | 1 | 1 |  | 47.2 | 47.2 | 47.2 |  |  |  |  | 107.52 | PASS | 1 |  | 605 | 605 | 605 |  |  |  |  | 113.25 | PASS |  | 99
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 070 | 70 | 1 | 1 | TP | ug P/L | 0.128 | 2 | 46.5 | TN | ug N/L | 0.715 | 8.73 | 526 |  | 1 | 1 |  | 46.5 | 46.5 | 46.5 |  |  |  |  | 113.05 | PASS | 1 |  | 526 | 526 | 526 |  |  |  |  | 96.46 | PASS |  | 100
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 082 | 82 | 1 | 1 | TP | ug P/L | 0.101 | 1.68 | 37.3 | TN | ug N/L | 0.644 | 7.97 | 474 |  | 1 | 1 |  | 37.3 | 37.3 | 37.3 |  |  |  |  | 108.53 | PASS | 1 |  | 474 | 474 | 474 |  |  |  |  | 97.94 | PASS |  | 101
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 094 | 94 | 1 | 1 | TP | ug P/L | 0.161 | 2.66 | 65.5 | TN | ug N/L | 1.76 | 21.6 | 1420 |  | 1 | 1 |  | 65.5 | 65.5 | 65.5 |  |  |  |  | 115.75 | PASS | 1 |  | 1420 | 1420 | 1420 |  |  |  |  | 106.2 | PASS |  | 102
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 106 | 106 | 1 | 1 | TP | ug P/L | 0.119 | 1.97 | 45.6 | TN | ug N/L | 0.958 | 11.7 | 734 |  | 1 | 1 |  | 45.6 | 45.6 | 45.6 |  |  |  |  | 113.76 | PASS | 1 |  | 734 | 734 | 734 |  |  |  |  | 149.34 | FAIL |  | 103
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 118 | 118 | 1 | 1 | TP | ug P/L | 0.144 | 2.4 | 58.2 | TN | ug N/L | 0.775 | 9.68 | 591 |  | 1 | 1 |  | 58.2 | 58.2 | 58.2 |  |  |  |  | 115.02 | PASS | 1 |  | 591 | 591 | 591 |  |  |  |  | 106.71 | PASS |  | 104
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 130 | 130 | 1 | 1 | TP | ug P/L | 0.124 | 2.03 | 47.4 | TN | ug N/L | 0.707 | 8.89 | 537 |  | 1 | 1 |  | 47.4 | 47.4 | 47.4 |  |  |  |  | 105.94 | PASS | 1 |  | 537 | 537 | 537 |  |  |  |  | 104.17 | PASS |  | 105
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 142 | 142 | 1 | 1 | TP | ug P/L | 0.165 | 2.73 | 67.6 | TN | ug N/L | 1.31 | 16.2 | 1050 |  | 1 | 1 |  | 67.6 | 67.6 | 67.6 |  |  |  |  | 105.16 | PASS | 1 |  | 1050 | 1050 | 1050 |  |  |  |  | 104.5 | PASS |  | 106
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 154 | 154 | 1 | 1 | TP | ug P/L | 0.18 | 3 | 75.6 | TN | ug N/L | 0.452 | 5.78 | 323 |  | 1 | 1 |  | 75.6 | 75.6 | 75.6 |  |  |  |  | 132.36 | FAIL | 1 |  | 323 | 323 | 323 |  |  |  |  | 69.63 | FAIL |  | 107
[27 empty] | PO4-P- PR of spikes | [10 empty] | NO3-N- PR of spikes
[25 empty] | Mean |  | 101.0849 | [10 empty] | 91.2363 |  | Min | 1
[25 empty] | Std |  | 24.5831 | [10 empty] | 30.3034 |  | Max | 107
[25 empty] | CV |  | 24.3193 | [10 empty] | 33.2142 |  | Mean
[25 empty] | S/N |  | 4.112 | [10 empty] | 3.0108 |  | Whole series
[24 empty] | Warning Limits | Xbar + 2SD |  | 150.2511 | [10 empty] | 151.8432
[25 empty] | Xbar - 2SD |  | 51.9186 | [10 empty] | 30.6294
[24 empty] | Control Limits | Xbar + 3SD |  | 174.8343 | [10 empty] | 182.1466
[25 empty] | Xbar - 3SD |  | 27.3354 | [10 empty] | 0.32598
Count
99th percentile.inc
Maximum
Known
MDL
LOQ
ratio mean/mdl

## rolling dup RPD
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 043 | 43 | 1 | 1 | TP | ug P/L | 0.0646 | 1.46 | 28.1 | TN | ug N/L | 0.565 | 9.58 | 29.8 |  | 1 | 1 |  | 28.1 | 28.1 | 28.1 |  |  | 1.7953 | PASS |  |  | 2 | using modeled decline in efficiency | 381.7128 | 381.7128 | 346.237 |  |  | 8.5056 | PASS |  |  |  | 1
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 055 | 55 | 1 | 1 | TP | ug P/L | 0.0666 | 1.46 | 28.1 | TN | ug N/L | 0.466 | 7.85 | 14.9 |  | 1 | 1 |  | 28.1 | 28.1 | 28.1 |  |  | 16.6395 | PASS |  |  | 2 | using modeled decline in efficiency | 207.3709 | 207.3709 | 198.9491 |  |  | 19.969 | PASS |  |  |  | 2
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 067 | 67 | 1 | 1 | TP | ug P/L | 0.0722 | 1.61 | 32.6 | TN | ug N/L | 0.856 | 14.5 | 72.5 |  | 1 | 1 |  | 32.6 | 32.6 | 32.6 |  |  | 4.2042 | PASS |  |  | 2 | using modeled decline in efficiency | 855.8403 | 855.8403 | 839.2601 |  |  | 2.2851 | PASS |  |  |  | 3
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 079 | 79 | 1 | 1 | TP | ug P/L | 0.0495 | 1.17 | 20 | TN | ug N/L | 0.527 | 8.93 | 24.2 |  | 1 | 1 |  | 20 | 20 | 20 |  |  | 5.1282 | PASS |  |  | 2 | using modeled decline in efficiency | 316.674 | 316.674 | 333.3864 |  |  | 13.7071 | PASS |  |  |  | 4
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 091 | 91 | 1 | 1 | TP | ug P/L | 0.0628 | 1.38 | 25.9 | TN | ug N/L | 0.624 | 10.6 | 38.7 |  | 1 | 1 |  | 25.9 | 25.9 | 25.9 |  |  | 4.1588 | PASS |  |  | 2 | using modeled decline in efficiency | 482.6446 | 482.6446 | 511.4898 |  |  | 7.8651 | PASS |  |  |  | 5
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 103 | 103 | 1 | 1 | TP | ug P/L | 0.0488 | 1.12 | 18.7 | TN | ug N/L | 0.503 | 8.52 | 20.7 |  | 1 | 1 |  | 18.7 | 18.7 | 18.7 |  |  | 0.53619 | PASS |  |  | 2 | using modeled decline in efficiency | 275.3614 | 275.3614 | 320.3854 |  |  | 3.9719 | PASS |  |  |  | 6
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 115 | 115 | 1 | 1 | TP | ug P/L | 0.0566 | 1.29 | 23.5 | TN | ug N/L | 0.465 | 8 | 16.2 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 |  |  | 3.0238 | PASS |  |  | 2 | using modeled decline in efficiency | 222.6444 | 222.6444 | 281.6251 |  |  | 1.6705 | PASS |  |  |  | 7
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 127 | 127 | 1 | 1 | TP | ug P/L | 0.0733 | 1.59 | 31.9 | TN | ug N/L | 0.82 | 14.1 | 69 |  | 1 | 1 |  | 31.9 | 31.9 | 31.9 |  |  | 0 | PASS |  |  | 2 | using modeled decline in efficiency | 818.4919 | 818.4919 | 927.8125 |  |  | 1.1337 | PASS |  |  |  | 8
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | Sample 139 | 139 | 1 | 1 | TP | ug P/L | 0.0471 | 1.1 | 18 | TN | ug N/L | 0.405 | 6.93 | 6.99 |  | 1 | 1 |  | 18 | 18 | 18 |  |  | 1.6529 | PASS |  |  | 2 | using modeled decline in efficiency | 113.0412 | 113.0412 | 198.8718 |  |  | 9.7193 | PASS |  |  |  | 9
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 45 | 1 | 1 | TP | ug P/L | 0.0632 | 1.48 | 32.1 | TN | ug N/L | 0.858 | 14.7 | 943 |  | 1 | 1 |  | 32.1 | 32.1 | 32.1 |  |  | 3.1646 | PASS |  |  | 1 |  | 943 | 943 | 943 |  |  | 5.7673 | PASS |  |  |  | 10
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 57 | 1 | 1 | TP | ug P/L | 0.0738 | 1.65 | 36.9 | TN | ug N/L | 0.881 | 15 | 967 |  | 1 | 1 |  | 36.9 | 36.9 | 36.9 |  |  | 5.0198 | PASS |  |  | 1 |  | 967 | 967 | 967 |  |  | 5.6353 | PASS |  |  |  | 11
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 69 | 1 | 1 | TP | ug P/L | 0.0712 | 1.6 | 35.5 | TN | ug N/L | 1.22 | 20.6 | 1410 |  | 1 | 1 |  | 35.5 | 35.5 | 35.5 |  |  | 1.9915 | PASS |  |  | 1 |  | 1410 | 1410 | 1410 |  |  | 12.0301 | PASS |  |  |  | 12
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 81 | 1 | 1 | TP | ug P/L | 0.0674 | 1.49 | 32.4 | TN | ug N/L | 1.24 | 20.9 | 1440 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 |  |  | 3.0395 | PASS |  |  | 1 |  | 1440 | 1440 | 1440 |  |  | 11.7647 | PASS |  |  |  | 13
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 93 | 1 | 1 | TP | ug P/L | 0.065 | 1.5 | 32.6 | TN | ug N/L | 0.597 | 10.2 | 597 |  | 1 | 1 |  | 32.6 | 32.6 | 32.6 |  |  | 7.6433 | PASS |  |  | 1 |  | 597 | 597 | 597 |  |  | 2.8887 | PASS |  |  |  | 14
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 105 | 1 | 1 | TP | ug P/L | 0.0967 | 2.08 | 49.5 | TN | ug N/L | 0.408 | 7 | 357 |  | 1 | 1 |  | 49.5 | 49.5 | 49.5 |  |  | 11.9914 | PASS |  |  | 1 |  | 357 | 357 | 357 |  |  | 10.3093 | PASS |  |  |  | 15
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 117 | 1 | 1 | TP | ug P/L | 0.0505 | 1.12 | 21.2 | TN | ug N/L | 0.275 | 4.71 | 188 |  | 1 | 1 |  | 21.2 | 21.2 | 21.2 |  |  | 2.7907 | PASS |  |  | 1 |  | 188 | 188 | 188 |  |  | 11.0553 | PASS |  |  |  | 16
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 129 | 1 | 1 | TP | ug P/L | 0.0779 | 1.66 | 37.4 | TN | ug N/L | 0.766 | 13 | 812 |  | 1 | 1 |  | 37.4 | 37.4 | 37.4 |  |  | 0.80537 | PASS |  |  | 1 |  | 812 | 812 | 812 |  |  | 0.74166 | PASS |  |  |  | 17
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 141 | 1 | 1 | TP | ug P/L | 0.0453 | 1.1 | 20.7 | TN | ug N/L | 0.336 | 5.79 | 267 |  | 1 | 1 |  | 20.7 | 20.7 | 20.7 |  |  | 17.2185 | PASS |  |  | 1 |  | 267 | 267 | 267 |  |  | 1.8553 | PASS |  |  |  | 18
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | DUP | 153 | 1 | 1 | TP | ug P/L | 0.0722 | 1.67 | 37.6 | TN | ug N/L | 0.514 | 8.77 | 489 |  | 1 | 1 |  | 37.6 | 37.6 | 37.6 |  |  | 7.7348 | PASS |  |  | 1 |  | 489 | 489 | 489 |  |  | 5.7597 | PASS |  |  |  | 19
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 45 | 1 | 1 | TP | ug P/L | 0.0485 | 1.09 | 13.3 | TN | ug N/L | 0.167 | 3.05 | 178 |  | 1 | 2 | reworked calibration | 19.5346 | 19.5346 | 19.5346 |  |  | 0 | PASS |  |  | 1 |  | 187.6427 | 187.6427 | 187.6427 |  |  | 2.0185 | PASS |  |  |  | 20
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 57 | 1 | 1 | TP | ug P/L | 0.0487 | 1.12 | 13.9 | TN | ug N/L | 0.176 | 3.22 | 192 |  | 1 | 2 | reworked calibration | 20.427 | 20.427 | 20.427 |  |  | 0 | PASS |  |  | 1 |  | 203.6024 | 203.6024 | 203.6024 |  |  | 1.3943 | PASS |  |  |  | 21
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 69 | 1 | 1 | TP | ug P/L | 0.0334 | 0.713 | 3.3 | TN | ug N/L | 0.154 | 2.86 | 160 |  | 1 | 2 | reworked calibration | 8.3765 | 8.3765 | 8.3765 |  |  | 4.3103 | PASS |  |  | 1 |  | 169.8505 | 169.8505 | 169.8505 |  |  | 17.3138 | PASS |  |  |  | 22
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 81 | 1 | 1 | TP | ug P/L | 0.0648 | 1.41 | 21.9 | TN | ug N/L | 0.25 | 4.51 | 307 |  | 1 | 2 | reworked calibration | 29.0883 | 29.0883 | 29.0883 |  |  | 8.5891 | PASS |  |  | 1 |  | 325.9513 | 325.9513 | 325.9513 |  |  | 15.1065 | PASS |  |  |  | 23
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 93 | 1 | 1 | TP | ug P/L | 0.0341 | 0.763 | 4.63 | TN | ug N/L | 0.158 | 2.86 | 160 |  | 1 | 2 | reworked calibration | 9.8503 | 9.8503 | 9.8503 |  |  | 3.6583 | PASS |  |  | 1 |  | 169.8505 | 169.8505 | 169.8505 |  |  | 1.6652 | PASS |  |  |  | 24
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 105 | 1 | 1 | TP | ug P/L | 0.0733 | 1.53 | 25.1 | TN | ug N/L | 1.21 | 22.2 | 1800 |  | 1 | 2 | reworked calibration | 32.6905 | 32.6905 | 32.6905 |  |  | 18.4182 | PASS |  |  | 1 |  | 2225.3094 | 2225.3094 | 2225.3094 |  |  | 0.53609 | PASS |  |  |  | 25
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 117 | 1 | 1 | TP | ug P/L | 0.0651 | 1.46 | 23.3 | TN | ug N/L | 0.22 | 4.08 | 269 |  | 1 | 2 | reworked calibration | 30.5879 | 30.5879 | 30.5879 |  |  | 2.9862 | PASS |  |  | 1 |  | 284.9243 | 284.9243 | 284.9243 |  |  | 3.393 | PASS |  |  |  | 26
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 129 | 1 | 1 | TP | ug P/L | 0.0687 | 1.52 | 24.9 | TN | ug N/L | 0.459 | 8.44 | 651 |  | 1 | 2 | reworked calibration | 32.3899 | 32.3899 | 32.3899 |  |  | 10.5626 | PASS |  |  | 1 |  | 712.2263 | 712.2263 | 712.2263 |  |  | 15.4123 | PASS |  |  |  | 27
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 141 | 1 | 1 | TP | ug P/L | 0.0364 | 0.804 | 5.69 | TN | ug N/L | 0.188 | 3.46 | 214 |  | 1 | 2 | reworked calibration | 11.0602 | 11.0602 | 11.0602 |  |  | 4.9222 | PASS |  |  | 1 |  | 226.1987 | 226.1987 | 226.1987 |  |  | 45.8326 | FAIL |  |  |  | 28
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | DUP | 153 | 1 | 1 | TP | ug P/L | 0.0784 | 1.77 | 31.7 | TN | ug N/L | 0.293 | 5.47 | 391 |  | 1 | 2 | reworked calibration | 39.9269 | 39.9269 | 39.9269 |  |  | 37.5961 | FAIL |  |  | 1 |  | 418.4271 | 418.4271 | 418.4271 |  |  | 48.3522 | FAIL |  |  |  | 29
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 45 | 1 | 1 | TP | ug P/L | 0.0538 | 1.2 | 28.7 | TN | ug N/L | 0.804 | 11.4 | 700 |  | 1 | 1 |  | 28.7 | 28.7 | 28.7 |  |  | 2.4096 | PASS |  |  | 1 |  | 700 | 700 | 700 |  |  | 4.0816 | PASS |  |  |  | 30
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 57 | 1 | 1 | TP | ug P/L | 0.0342 | 0.745 | 14.2 | TN | ug N/L | 0.239 | 3.56 | 162 |  | 1 | 1 |  | 14.2 | 14.2 | 14.2 |  |  | 6.1433 | PASS |  |  | 1 |  | 162 | 162 | 162 |  |  | 8.8496 | PASS |  |  |  | 31
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 69 | 1 | 1 | TP | ug P/L | 0.0518 | 1.12 | 26.1 | TN | ug N/L | 0.464 | 6.67 | 379 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 |  |  | 1.5209 | PASS |  |  | 1 |  | 379 | 379 | 379 |  |  | 0.7947 | PASS |  |  |  | 32
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 81 | 1 | 1 | TP | ug P/L | 0.073 | 1.6 | 41.1 | TN | ug N/L | 0.908 | 12.9 | 796 |  | 1 | 1 |  | 41.1 | 41.1 | 41.1 |  |  | 6.015 | PASS |  |  | 1 |  | 796 | 796 | 796 |  |  | 2.6735 | PASS |  |  |  | 33
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 93 | 1 | 1 | TP | ug P/L | 0.0705 | 1.5 | 38 | TN | ug N/L | 1.05 | 14.8 | 918 |  | 1 | 1 |  | 38 | 38 | 38 |  |  | 3.871 | PASS |  |  | 1 |  | 918 | 918 | 918 |  |  | 11.0345 | PASS |  |  |  | 34
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 105 | 1 | 1 | TP | ug P/L | 0.103 | 2.13 | 57.5 | TN | ug N/L | 0.933 | 13.3 | 820 |  | 1 | 1 |  | 57.5 | 57.5 | 57.5 |  |  | 0.5231 | PASS |  |  | 1 |  | 820 | 820 | 820 |  |  | 15.3644 | PASS |  |  |  | 35
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 117 | 1 | 1 | TP | ug P/L | 0.0869 | 1.83 | 48.3 | TN | ug N/L | 0.36 | 5.24 | 280 |  | 1 | 1 |  | 48.3 | 48.3 | 48.3 |  |  | 5.0454 | PASS |  |  | 1 |  | 280 | 280 | 280 |  |  | 0.35778 | PASS |  |  |  | 36
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 129 | 1 | 1 | TP | ug P/L | 0.0277 | 0.588 | 9.21 | TN | ug N/L | 0.283 | 4.13 | 203 |  | 1 | 1 |  | 9.21 | 9.21 | 9.21 |  |  | 0.32626 | PASS |  |  | 1 |  | 203 | 203 | 203 |  |  | 6.6158 | PASS |  |  |  | 37
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 141 | 1 | 1 | TP | ug P/L | 0.0399 | 0.944 | 20.6 | TN | ug N/L | 0.378 | 5.5 | 298 |  | 1 | 1 |  | 20.6 | 20.6 | 20.6 |  |  | 61.5873 | FAIL |  |  | 1 |  | 298 | 298 | 298 |  |  | 4.4597 | PASS |  |  |  | 38
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 153 | 1 | 1 | TP | ug P/L | 0.0457 | 1.04 | 23.8 | TN | ug N/L | 0.368 | 5.31 | 285 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 |  |  | 7.4074 | PASS |  |  | 1 |  | 285 | 285 | 285 |  |  | 15.235 | PASS |  |  |  | 39
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP | 165 | 1 | 1 | TP | ug P/L | 0.104 | 1.97 | 52.7 | TN | ug N/L | 0.388 | 5.59 | 305 |  | 1 | 1 |  | 52.7 | 52.7 | 52.7 |  |  | 2.8874 | PASS |  |  | 1 |  | 305 | 305 | 305 |  |  | 25.0923 | FAIL |  |  |  | 40
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | DUP F 17aug20 0.1 | 177 | 1 | 1 | TP | ug P/L | 0.0642 | 1.32 | 32.5 | TN | ug N/L | 0.752 | 10.7 | 649 |  | 1 | 1 |  | 32.5 | 32.5 | 32.5 |  |  | 3.4429 | PASS |  |  | 1 |  | 649 | 649 | 649 |  |  | 8.2718 | PASS |  |  |  | 41
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | DUP | 45 | 1 | 1 | TP | ug P/L | 0.0542 | 1.12 | 21.8 | TN | ug N/L | 0.345 | 5.15 | 262 |  | 1 | 1 |  | 21.8 | 21.8 | 21.8 |  |  | 0.45977 | PASS |  |  | 1 |  | 262 | 262 | 262 |  |  | 39.8169 | FAIL |  |  |  | 42
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | DUP | 57 | 1 | 1 | TP | ug P/L | 0.0448 | 0.957 | 17.4 | TN | ug N/L | 0.248 | 3.7 | 170 |  | 1 | 1 |  | 17.4 | 17.4 | 17.4 |  |  | 3.3898 | PASS |  |  | 1 |  | 170 | 170 | 170 |  |  | 2.3256 | PASS |  |  |  | 43
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | DUP | 69 | 1 | 1 | TP | ug P/L | 0.0424 | 0.912 | 16.1 | TN | ug N/L | 0.256 | 3.81 | 177 |  | 1 | 1 |  | 16.1 | 16.1 | 16.1 |  |  | 1.25 | PASS |  |  | 1 |  | 177 | 177 | 177 |  |  | 0.56338 | PASS |  |  |  | 44
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 81 | 1 | 1 | TP | ug P/L | 0.0484 | 1.05 | 19.9 | TN | ug N/L | 0.28 | 4.17 | 200 |  | 1 | 1 |  | 19.9 | 19.9 | 19.9 |  |  | 1.0101 | PASS |  |  | 1 |  | 200 | 200 | 200 |  |  | 3.9216 | PASS |  |  |  | 45
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 93 | 1 | 1 | TP | ug P/L | 0.0397 | 0.868 | 14.9 | TN | ug N/L | 0.248 | 3.72 | 172 |  | 1 | 1 |  | 14.9 | 14.9 | 14.9 |  |  | 5.8632 | PASS |  |  | 1 |  | 172 | 172 | 172 |  |  | 3.9886 | PASS |  |  |  | 46
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 105 | 1 | 1 | TP | ug P/L | 0.0508 | 1.08 | 20.8 | TN | ug N/L | 0.273 | 4.08 | 194 |  | 1 | 1 |  | 20.8 | 20.8 | 20.8 |  |  | 2.3753 | PASS |  |  | 1 |  | 194 | 194 | 194 |  |  | 3.1414 | PASS |  |  |  | 47
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 117 | 1 | 1 | TP | ug P/L | 0.154 | 2.91 | 71.5 | TN | ug N/L | 0.631 | 9.12 | 517 |  | 1 | 1 |  | 71.5 | 71.5 | 71.5 |  |  | 2.4062 | PASS |  |  | 1 |  | 517 | 517 | 517 |  |  | 15.6413 | PASS |  |  |  | 48
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 129 | 1 | 1 | TP | ug P/L | 0.152 | 2.88 | 70.6 | TN | ug N/L | 0.531 | 7.74 | 428 |  | 1 | 1 |  | 70.6 | 70.6 | 70.6 |  |  | 10.9772 | PASS |  |  | 1 |  | 428 | 428 | 428 |  |  | 2.5375 | PASS |  |  |  | 49
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 141 | 1 | 1 | TP | ug P/L | 0.0638 | 1.34 | 27.9 | TN | ug N/L | 0.324 | 4.8 | 240 |  | 1 | 1 |  | 27.9 | 27.9 | 27.9 |  |  | 13.3779 | PASS |  |  | 1 |  | 240 | 240 | 240 |  |  | 17.4905 | PASS |  |  |  | 50
44357 | OM_6-10-2021_10-17-10AM.OMN | DUP | 153 | 1 | 1 | TP | ug P/L | 0.0316 | 0.73 | 11 | TN | ug N/L | 0.315 | 4.67 | 232 |  | 1 | 1 |  | 11 | 11 | 11 |  |  | 11.6402 | PASS |  |  | 1 |  | 232 | 232 | 232 |  |  | 0.8658 | PASS |  |  |  | 51
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 045 | 45 | 1 | 1 | TP | ug P/L | 0.0778 | 1.53 | 36.3 | TN | ug N/L | 1.35 | 18.1 | 1150 |  | 1 | 1 |  | 36.3 | 36.3 | 36.3 |  |  | 17.0404 | PASS |  |  | 1 |  | 1150 | 1150 | 1150 |  |  | 10.0457 | PASS |  |  |  | 52
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 057 | 57 | 1 | 1 | TP | ug P/L | 0.0314 | 0.654 | 11.3 | TN | ug N/L | 0.308 | 4.28 | 206 |  | 1 | 1 |  | 11.3 | 11.3 | 11.3 |  |  | 4.329 | PASS |  |  | 1 |  | 206 | 206 | 206 |  |  | 2.3981 | PASS |  |  |  | 53
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 069 | 69 | 1 | 1 | TP | ug P/L | 0.0802 | 1.46 | 34.3 | TN | ug N/L | 0.633 | 8.63 | 492 |  | 1 | 1 |  | 34.3 | 34.3 | 34.3 |  |  | 1.173 | PASS |  |  | 1 |  | 492 | 492 | 492 |  |  | 1.8462 | PASS |  |  |  | 54
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 081 | 81 | 1 | 1 | TP | ug P/L | 0.0764 | 1.4 | 32.5 | TN | ug N/L | 0.675 | 9.14 | 526 |  | 1 | 1 |  | 32.5 | 32.5 | 32.5 |  |  | 9.0032 | PASS |  |  | 1 |  | 526 | 526 | 526 |  |  | 11.245 | PASS |  |  |  | 55
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 105 | 105 | 1 | 1 | TP | ug P/L | 0.101 | 1.79 | 43.7 | TN | ug N/L | 0.792 | 10.8 | 639 |  | 1 | 1 |  | 43.7 | 43.7 | 43.7 |  |  | 3.7296 | PASS |  |  | 1 |  | 639 | 639 | 639 |  |  | 18.275 | PASS |  |  |  | 56
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 117 | 117 | 1 | 1 | TP | ug P/L | 0.0687 | 1.27 | 28.9 | TN | ug N/L | 0.515 | 7.07 | 388 |  | 1 | 1 |  | 28.9 | 28.9 | 28.9 |  |  | 2.3932 | PASS |  |  | 1 |  | 388 | 388 | 388 |  |  | 7.9208 | PASS |  |  |  | 57
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 129 | 129 | 1 | 1 | TP | ug P/L | 0.0473 | 0.944 | 19.5 | TN | ug N/L | 0.301 | 4.19 | 200 |  | 1 | 1 |  | 19.5 | 19.5 | 19.5 |  |  | 5.8047 | PASS |  |  | 1 |  | 200 | 200 | 200 |  |  | 9.9738 | PASS |  |  |  | 58
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 141 | 141 | 1 | 1 | TP | ug P/L | 0.0295 | 0.607 | 9.98 | TN | ug N/L | 0.228 | 3.2 | 136 |  | 1 | 1 |  | 9.98 | 9.98 | 9.98 |  |  | 0.90589 | PASS |  |  | 1 |  | 136 | 136 | 136 |  |  | 1.4599 | PASS |  |  |  | 59
44615 | OM_2-23-2022_11-17-46AM edit 1.omn | Vial 153 | 153 | 1 | 1 | TP | ug P/L | 0.0587 | 1.11 | 24.3 | TN | ug N/L | 0.362 | 5 | 253 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 |  |  | 3.3473 | PASS |  |  | 1 |  | 253 | 253 | 253 |  |  | 0.79365 | PASS |  |  |  | 60
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample45 | 45 | 1 | 1 | TP | ug P/L | 0.056 | 1.07 | 18.3 | TN | ug N/L | 0.477 | 6.48 | 347 |  | 1 | 1 |  | 18.3 | 18.3 | 18.3 |  |  | 15.6171 | PASS |  |  | 1 |  | 347 | 347 | 347 |  |  | 3.3994 | PASS |  |  |  | 61
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample57 | 57 | 1 | 1 | TP | ug P/L | 0.0864 | 1.61 | 31.5 | TN | ug N/L | 1.21 | 16 | 984 |  | 1 | 1 |  | 31.5 | 31.5 | 31.5 |  |  | 3.2258 | PASS |  |  | 1 |  | 984 | 984 | 984 |  |  | 1.4127 | PASS |  |  |  | 62
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample69 | 69 | 1 | 1 | TP | ug P/L | 0.112 | 2.03 | 42 | TN | ug N/L | 0.248 | 3.45 | 143 |  | 1 | 1 |  | 42 | 42 | 42 |  |  | 15.1088 | PASS |  |  | 1 |  | 143 | 143 | 143 |  |  | 40.5383 | FAIL |  |  |  | 63
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample81 | 81 | 1 | 1 | TP | ug P/L | 0.0882 | 1.57 | 30.6 | TN | ug N/L | 0.793 | 10.5 | 619 |  | 1 | 1 |  | 30.6 | 30.6 | 30.6 |  |  | 11.0345 | PASS |  |  | 1 |  | 619 | 619 | 619 |  |  | 1.2841 | PASS |  |  |  | 64
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample93 | 93 | 1 | 1 | TP | ug P/L | 0.0883 | 1.58 | 30.7 | TN | ug N/L | 1.41 | 18.7 | 1170 |  | 1 | 1 |  | 30.7 | 30.7 | 30.7 |  |  | 3.8339 | PASS |  |  | 1 |  | 1170 | 1170 | 1170 |  |  | 13.6986 | PASS |  |  |  | 65
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample105 | 105 | 1 | 1 | TP | ug P/L | 0.0459 | 0.9 | 14.1 | TN | ug N/L | 0.252 | 3.48 | 146 |  | 1 | 1 |  | 14.1 | 14.1 | 14.1 |  |  | 10.101 | PASS |  |  | 1 |  | 146 | 146 | 146 |  |  | 0.68259 | PASS |  |  |  | 66
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample117 | 117 | 1 | 1 | TP | ug P/L | 0.0803 | 1.5 | 28.9 | TN | ug N/L | 0.329 | 4.5 | 214 |  | 1 | 1 |  | 28.9 | 28.9 | 28.9 |  |  | 3.4014 | PASS |  |  | 1 |  | 214 | 214 | 214 |  |  | 3.2184 | PASS |  |  |  | 67
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample129 | 129 | 1 | 1 | TP | ug P/L | 0.127 | 2.26 | 47.7 | TN | ug N/L | 0.592 | 7.97 | 447 |  | 1 | 1 |  | 47.7 | 47.7 | 47.7 |  |  | 1.0428 | PASS |  |  | 1 |  | 447 | 447 | 447 |  |  | 6.2839 | PASS |  |  |  | 68
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample141 | 141 | 1 | 1 | TP | ug P/L | 0.101 | 1.81 | 36.6 | TN | ug N/L | 1.01 | 13.7 | 829 |  | 1 | 1 |  | 36.6 | 36.6 | 36.6 |  |  | 2.1622 | PASS |  |  | 1 |  | 829 | 829 | 829 |  |  | 11.4796 | PASS |  |  |  | 69
44643 | OM_3-23-2022_02-00-59PM.OMN | Sample153 | 153 | 1 | 1 | TP | ug P/L | 0.0525 | 0.975 | 16 | TN | ug N/L | 0.662 | 9.02 | 517 |  | 1 | 1 |  | 16 | 16 | 16 |  |  | 25.6131 | FAIL |  |  | 1 |  | 517 | 517 | 517 |  |  | 4.3478 | PASS |  |  |  | 70
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 045 | 45 | 1 | 1 | TP | ug P/L | 0.0553 | 1.05 | 20.3 | TN | ug N/L | 0.551 | 7.61 | 462 |  | 1 | 1 |  | 20.3 | 20.3 | 20.3 |  |  | 4.534 | PASS |  |  | 1 |  | 462 | 462 | 462 |  |  | 3.5242 | PASS |  |  |  | 71
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 057 | 57 | 1 | 1 | TP | ug P/L | 0.0639 | 1.17 | 23.9 | TN | ug N/L | 2.16 | 29.6 | 1960 |  | 1 | 1 |  | 23.9 | 23.9 | 23.9 |  |  | 17.9048 | PASS |  |  | 2 | oor | 1960 | 1960 | 1960 |  |  | 0.51151 | PASS |  |  |  | 72
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 069 | 69 | 1 | 1 | TP | ug P/L | 0.0561 | 1.01 | 19.4 | TN | ug N/L | 0.408 | 5.64 | 324 |  | 1 | 1 |  | 19.4 | 19.4 | 19.4 |  |  | 1.5584 | PASS |  |  | 1 |  | 324 | 324 | 324 |  |  | 7.7151 | PASS |  |  |  | 73
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 081 | 81 | 1 | 1 | TP | ug P/L | 0.0225 | 0.466 | 3.99 | TN | ug N/L | 0.247 | 3.48 | 171 |  | 1 | 1 |  | 3.99 | 3.99 | 3.99 |  |  | 2.5381 | PASS |  |  | 1 |  | 171 | 171 | 171 |  |  | 4.7904 | PASS |  |  |  | 74
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 093 | 93 | 1 | 1 | TP | ug P/L | 0.0567 | 1.04 | 20.3 | TN | ug N/L | 0.411 | 5.74 | 330 |  | 1 | 1 |  | 20.3 | 20.3 | 20.3 |  |  | 2.4331 | PASS |  |  | 1 |  | 330 | 330 | 330 |  |  | 0.30349 | PASS |  |  |  | 75
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 105 | 105 | 1 | 1 | TP | ug P/L | 0.066 | 1.22 | 25.2 | TN | ug N/L | 1.55 | 21.3 | 1400 |  | 1 | 1 |  | 25.2 | 25.2 | 25.2 |  |  | 4.6512 | PASS |  |  | 1 |  | 1400 | 1400 | 1400 |  |  | 4.1958 | PASS |  |  |  | 76
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 117 | 117 | 1 | 1 | TP | ug P/L | 0.199 | 3.45 | 86.9 | TN | ug N/L | 0.433 | 6.04 | 352 |  | 1 | 1 |  | 86.9 | 86.9 | 86.9 |  |  | 42.6437 | FAIL |  |  | 1 |  | 352 | 352 | 352 |  |  | 59.3407 | FAIL |  |  |  | 77
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 129 | 129 | 1 | 1 | TP | ug P/L | 0.0668 | 1.2 | 24.6 | TN | ug N/L | 0.103 | 1.47 | 28.8 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 |  |  | 9.3023 | PASS |  |  | 1 |  | 28.8 | 28.8 | 28.8 |  |  | 75.4595 | FAIL |  |  |  | 78
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 141 | 141 | 1 | 1 | TP | ug P/L | 0.0547 | 0.998 | 19 | TN | ug N/L | 0.32 | 4.53 | 245 |  | 1 | 1 |  | 19 | 19 | 19 |  |  | 3.6176 | PASS |  |  | 1 |  | 245 | 245 | 245 |  |  | 12.1212 | PASS |  |  |  | 79
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Vial 153 | 153 | 1 | 1 | TP | ug P/L | 0.0509 | 0.979 | 18.4 | TN | ug N/L | 0.519 | 7.17 | 431 |  | 1 | 1 |  | 18.4 | 18.4 | 18.4 |  |  | 1.6438 | PASS |  |  | 1 |  | 431 | 431 | 431 |  |  | 0.69849 | PASS |  |  |  | 80
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 045 | 45 | 1 | 1 | TP | ug P/L | 0.0775 | 1.32 | 26.9 | TN | ug N/L | 0.479 | 5.92 | 333 |  | 1 | 1 |  | 26.9 | 26.9 | 26.9 |  |  | 0 | PASS |  |  | 1 |  | 333 | 333 | 333 |  |  | 1.7857 | PASS |  |  |  | 81
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 057 | 57 | 1 | 1 | TP | ug P/L | 0.0728 | 1.25 | 24.9 | TN | ug N/L | 0.448 | 5.55 | 308 |  | 1 | 1 |  | 24.9 | 24.9 | 24.9 |  |  | 9.2437 | PASS |  |  | 1 |  | 308 | 308 | 308 |  |  | 2.2989 | PASS |  |  |  | 82
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 069 | 69 | 1 | 1 | TP | ug P/L | 0.0646 | 1.11 | 21 | TN | ug N/L | 0.485 | 6.29 | 358 |  | 1 | 1 |  | 21 | 21 | 21 |  |  | 5.8824 | PASS |  |  | 1 |  | 358 | 358 | 358 |  |  | 3.6984 | PASS |  |  |  | 83
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 081 | 81 | 1 | 1 | TP | ug P/L | 0.0898 | 1.54 | 33.2 | TN | ug N/L | 1.25 | 15.3 | 977 |  | 1 | 1 |  | 33.2 | 33.2 | 33.2 |  |  | 6.5319 | PASS |  |  | 1 |  | 977 | 977 | 977 |  |  | 1.1323 | PASS |  |  |  | 84
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 093 | 93 | 1 | 1 | TP | ug P/L | 0.0882 | 1.47 | 31.2 | TN | ug N/L | 0.455 | 5.72 | 319 |  | 1 | 1 |  | 31.2 | 31.2 | 31.2 |  |  | 4.0816 | PASS |  |  | 1 |  | 319 | 319 | 319 |  |  | 1.1323 | PASS |  |  |  | 85
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 105 | 105 | 1 | 1 | TP | ug P/L | 0.063 | 1.06 | 19.4 | TN | ug N/L | 0.448 | 5.73 | 320 |  | 1 | 1 |  | 19.4 | 19.4 | 19.4 |  |  | 11.6505 | PASS |  |  | 1 |  | 320 | 320 | 320 |  |  | 3.0769 | PASS |  |  |  | 86
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 117 | 117 | 1 | 1 | TP | ug P/L | 0.0853 | 1.42 | 29.9 | TN | ug N/L | 0.457 | 5.76 | 322 |  | 1 | 1 |  | 29.9 | 29.9 | 29.9 |  |  | 1.6584 | PASS |  |  | 1 |  | 322 | 322 | 322 |  |  | 0.31008 | PASS |  |  |  | 87
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 129 | 129 | 1 | 1 | TP | ug P/L | 0.0695 | 1.17 | 22.6 | TN | ug N/L | 0.427 | 5.45 | 301 |  | 1 | 1 |  | 22.6 | 22.6 | 22.6 |  |  | 5.1724 | PASS |  |  | 1 |  | 301 | 301 | 301 |  |  | 11.9718 | PASS |  |  |  | 88
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 141 | 141 | 1 | 1 | TP | ug P/L | 0.075 | 1.29 | 26.1 | TN | ug N/L | 0.482 | 6.11 | 346 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 |  |  | 1.5444 | PASS |  |  | 1 |  | 346 | 346 | 346 |  |  | 0.57971 | PASS |  |  |  | 89
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | Vial 153 | 153 | 1 | 1 | TP | ug P/L | 0.0873 | 1.48 | 31.6 | TN | ug N/L | 3.14 | 38.9 | 2650 |  | 1 | 1 |  | 31.6 | 31.6 | 31.6 |  |  | 6.7278 | PASS |  |  | 3 | oor | 2650 | 2650 | 2650 |  |  | 5.8252 | PASS |  |  |  | 90
[25 empty] | PO4-P- RPD of duplicates | [10 empty] | NO3-N- RPD of duplicates |  |  |  | OBSERVATION #
[19 empty] | Mean | [5 empty] | 6.9522 | [10 empty] | 9.2395 |  |  |  | Min | 1
[19 empty] | Std | [5 empty] | 9.2712 | [10 empty] | 13.0336 |  |  |  | Max | 90
[19 empty] | CV | [5 empty] | 133.3565 | [10 empty] | 141.0634
[19 empty] | S/N | [5 empty] | 0.74987 | [10 empty] | 0.7089
[18 empty] | Warning Limits | Xbar + 2SD | [5 empty] | 25.4946 | [10 empty] | 35.3067
[19 empty] | Xbar - 2SD | [5 empty] | -11.5902 | [10 empty] | -16.8277
[18 empty] | Control Limits | Xbar + 3SD | [5 empty] | 34.7658 | [10 empty] | 48.3403
[19 empty] | Xbar - 3SD | [5 empty] | -20.8614 | [10 empty] | -29.8612
[18 empty] | Count | [6 empty] | 90 | [10 empty] | 90
[18 empty] | 99th percentile.inc | [6 empty] | 44.7275 | [10 empty] | 61.1137
[18 empty] | Maximum | [6 empty] | 61.5873 | [10 empty] | 75.4595

## rolling CCC
check standard
continuing calibration check
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 15 | 1 | 1 | TP | ug P/L | 0.0571 | 1.3 | 23.8 | TN | ug N/L | 0.493 | 8.3 | 18.8 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | -4.8 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 253.1018 | 253.1018 | 214.3084 | -14.2767 | PASS | [5 empty] | 1
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 16 | 1 | 1 | TP | ug P/L | 0.0732 | 1.32 | 24.3 | TN | ug N/L | 0.499 | 8.41 | 19.7 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 264.2396 | 264.2396 | 233.6407 | -6.5437 | PASS | [5 empty] | 2
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 17 | 1 | 1 | TP | ug P/L | 0.0569 | 1.31 | 24 | TN | ug N/L | 0.497 | 8.46 | 20.1 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 269.297 | 269.297 | 247.9649 | -0.81405 | PASS | [5 empty] | 3
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 15 | 1 | 1 | TP | ug P/L | 0.057 | 1.29 | 23.5 | TN | ug N/L | 0.485 | 8.35 | 19.2 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 | -6 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 258.1665 | 258.1665 | 247.781 | -0.8876 | PASS | [5 empty] | 4
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 16 | 1 | 1 | TP | ug P/L | 0.0577 | 1.35 | 25 | TN | ug N/L | 0.502 | 8.46 | 20.2 |  | 1 | 1 |  | 25 | 25 | 25 | 0 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 269.297 | 269.297 | 268.3253 | 7.3301 | PASS | [5 empty] | 5
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0561 | 1.3 | 23.6 | TN | ug N/L | 0.48 | 8.13 | 17.3 |  | 1 | 1 |  | 23.6 | 23.6 | 23.6 | [6 empty] | 2 | using modeled decline in efficiency | 235.8573 | 235.8573 | 249.9837 | [7 empty] | 6
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 17 | 1 | 1 | TP | ug P/L | 0.0564 | 1.32 | 24.2 | TN | ug N/L | 0.487 | 8.34 | 19.1 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 257.1538 | 257.1538 | 284.5594 | 13.8238 | PASS | [5 empty] | 7
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 15 | 1 | 1 | TP | ug P/L | 0.0588 | 1.38 | 25.8 | TN | ug N/L | 0.461 | 7.84 | 14.8 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | 3.2 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 206.3517 | 206.3517 | 246.2969 | -1.4812 | PASS | [5 empty] | 8
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 16 | 1 | 1 | TP | ug P/L | 0.059 | 1.34 | 24.7 | TN | ug N/L | 0.466 | 7.97 | 15.9 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | -1.2 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 219.5921 | 219.5921 | 270.1384 | 8.0554 | PASS | [5 empty] | 9
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 17 | 1 | 1 | TP | ug P/L | 0.056 | 1.32 | 24.2 | TN | ug N/L | 0.461 | 7.83 | 14.8 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 205.3322 | 205.3322 | 267.1999 | 6.88 | PASS | [5 empty] | 10
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 15 | 1 | 1 | TP | ug P/L | 0.0563 | 1.23 | 21.6 | TN | ug N/L | 0.438 | 7.5 | 11.9 |  | 1 | 1 |  | 21.6 | 21.6 | 21.6 | -13.6 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 171.6169 | 171.6169 | 244.5824 | -2.167 | PASS | [5 empty] | 11
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 16 | 1 | 1 | TP | ug P/L | 0.0568 | 1.32 | 24.1 | TN | ug N/L | 0.431 | 7.37 | 10.7 |  | 1 | 1 |  | 24.1 | 24.1 | 24.1 | -3.6 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 158.2954 | 158.2954 | 242.2342 | -3.1063 | PASS | [5 empty] | 12
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 17 | 1 | 1 | TP | ug P/L | 0.0568 | 1.32 | 24.2 | TN | ug N/L | 0.381 | 6.54 | 3.63 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 2 | using modeled decline in efficiency | 72.7149 | 72.7149 | 165.5676 | -33.7729 | FAIL | [5 empty] | 13
2020-10-15 | OM_10-15-2020_10-22-59AMedit1.omn | digested chk 25 | 8 | 1 | 1 | TP | ug P/L | 0.0489 | 1.15 | 19.3 | TN | ug N/L | 0.451 | 7.79 | 14.3 |  | 1 | 1 |  | 19.3 | 19.3 | 19.3 | [6 empty] | 2 | using modeled decline in efficiency | 201.2532 | 201.2532 | 309.588 | [7 empty] | 14
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.056 | 1.25 | 25.1 | TN | ug N/L | 0.289 | 4.97 | 206 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | 0.4 | PASS |  |  |  |  | 1 |  | 206 | 206 | 206 | -17.6 | PASS | [5 empty] | 15
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.067 | 1.24 | 24.8 | TN | ug N/L | 0.335 | 5.74 | 263 |  | 1 | 1 |  | 24.8 | 24.8 | 24.8 | -0.8 | PASS |  |  |  |  | 1 |  | 263 | 263 | 263 | 5.2 | PASS | [5 empty] | 16
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0542 | 1.24 | 24.9 | TN | ug N/L | 0.31 | 5.3 | 231 |  | 1 | 1 |  | 24.9 | 24.9 | 24.9 | -0.4 | PASS |  |  |  |  | 1 |  | 231 | 231 | 231 | -7.6 | PASS | [5 empty] | 17
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0519 | 1.23 | 24.5 | TN | ug N/L | 0.316 | 5.42 | 239 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | -2 | PASS |  |  |  |  | 1 |  | 239 | 239 | 239 | -4.4 | PASS | [5 empty] | 18
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0539 | 1.25 | 25.3 | TN | ug N/L | 0.32 | 5.47 | 244 |  | 1 | 1 |  | 25.3 | 25.3 | 25.3 | 1.2 | PASS |  |  |  |  | 1 |  | 244 | 244 | 244 | -2.4 | PASS | [5 empty] | 19
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0547 | 1.23 | 24.7 | TN | ug N/L | 0.307 | 5.27 | 229 |  | 1 | 1 |  | 24.7 | 24.7 | 24.7 | -1.2 | PASS |  |  |  |  | 1 |  | 229 | 229 | 229 | -8.4 | PASS | [5 empty] | 20
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0556 | 1.31 | 27 | TN | ug N/L | 0.315 | 5.44 | 241 |  | 1 | 1 |  | 27 | 27 | 27 | 8 | PASS |  |  |  |  | 1 |  | 241 | 241 | 241 | -3.6 | PASS | [5 empty] | 21
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0515 | 1.22 | 24.2 | TN | ug N/L | 0.353 | 5.97 | 280 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 280 | 280 | 280 | 12 | PASS | [5 empty] | 22
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.057 | 1.31 | 27.1 | TN | ug N/L | 0.301 | 5.16 | 220 |  | 1 | 1 |  | 27.1 | 27.1 | 27.1 | 8.4 | PASS |  |  |  |  | 1 |  | 220 | 220 | 220 | -12 | PASS | [5 empty] | 23
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0524 | 1.23 | 24.6 | TN | ug N/L | 0.312 | 5.38 | 237 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 237 | 237 | 237 | -5.2 | PASS | [5 empty] | 24
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0562 | 1.28 | 26.1 | TN | ug N/L | 0.303 | 5.26 | 228 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | 4.4 | PASS |  |  |  |  | 1 |  | 228 | 228 | 228 | -8.8 | PASS | [5 empty] | 25
2020-10-29 | OM_10-29-2020_11-50-32AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.052 | 1.19 | 23.3 | TN | ug N/L | 0.306 | 5.22 | 225 |  | 1 | 1 |  | 23.3 | 23.3 | 23.3 | -6.8 | PASS |  |  |  |  | 1 |  | 225 | 225 | 225 | -10 | PASS | [5 empty] | 26
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0551 | 1.28 | 18.3 | TN | ug N/L | 0.196 | 3.57 | 224 |  | 1 | 2 | reworked calibration | 25.1979 | 25.1979 | 25.1979 | 0.79179 | PASS |  |  |  |  | 1 |  | 236.5807 | 236.5807 | 236.5807 | -5.3677 | PASS | [5 empty] | 27
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0552 | 1.26 | 17.8 | TN | ug N/L | 0.191 | 3.49 | 216 |  | 1 | 2 | reworked calibration | 24.6005 | 24.6005 | 24.6005 | -1.5978 | PASS |  |  |  |  | 1 |  | 229.0286 | 229.0286 | 229.0286 | -8.3886 | PASS | [5 empty] | 28
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.052 | 1.19 | 16 | TN | ug N/L | 0.188 | 3.4 | 209 |  | 1 | 2 | reworked calibration | 22.512 | 22.512 | 22.512 | -9.9521 | PASS |  |  |  |  | 1 |  | 220.5425 | 220.5425 | 220.5425 | -11.783 | PASS | [5 empty] | 29
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0521 | 1.22 | 16.6 | TN | ug N/L | 0.177 | 3.25 | 195 |  | 1 | 2 | reworked calibration | 23.4066 | 23.4066 | 23.4066 | -6.3735 | PASS |  |  |  |  | 1 |  | 206.4228 | 206.4228 | 206.4228 | -17.4309 | PASS | [5 empty] | 30
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0518 | 1.18 | 15.6 | TN | ug N/L | 0.19 | 3.45 | 213 |  | 1 | 2 | reworked calibration | 22.2139 | 22.2139 | 22.2139 | -11.1444 | PASS |  |  |  |  | 1 |  | 225.2557 | 225.2557 | 225.2557 | -9.8977 | PASS | [5 empty] | 31
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0515 | 1.18 | 15.7 | TN | ug N/L | 0.188 | 3.41 | 210 |  | 1 | 2 | reworked calibration | 22.2139 | 22.2139 | 22.2139 | -11.1444 | PASS |  |  |  |  | 1 |  | 221.4849 | 221.4849 | 221.4849 | -11.4061 | PASS | [5 empty] | 32
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0495 | 1.14 | 14.5 | TN | ug N/L | 0.188 | 3.4 | 209 |  | 1 | 2 | reworked calibration | 21.0224 | 21.0224 | 21.0224 | -15.9106 | PASS |  |  |  |  | 1 |  | 220.5425 | 220.5425 | 220.5425 | -11.783 | PASS | [5 empty] | 33
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0506 | 1.19 | 15.9 | TN | ug N/L | 0.176 | 3.27 | 197 |  | 1 | 2 | reworked calibration | 22.512 | 22.512 | 22.512 | -9.9521 | PASS |  |  |  |  | 1 |  | 208.3037 | 208.3037 | 208.3037 | -16.6785 | PASS | [5 empty] | 34
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0498 | 1.17 | 15.3 | TN | ug N/L | 0.175 | 3.11 | 183 |  | 1 | 2 | reworked calibration | 21.9159 | 21.9159 | 21.9159 | -12.3364 | PASS |  |  |  |  | 1 |  | 193.2712 | 193.2712 | 193.2712 | -22.6915 | FAIL | [5 empty] | 35
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0513 | 1.19 | 16 | TN | ug N/L | 0.18 | 3.32 | 202 |  | 1 | 2 | reworked calibration | 22.512 | 22.512 | 22.512 | -9.9521 | PASS |  |  |  |  | 1 |  | 213.0083 | 213.0083 | 213.0083 | -14.7967 | PASS | [5 empty] | 36
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0495 | 1.08 | 13 | TN | ug N/L | 0.173 | 3.2 | 191 |  | 1 | 2 | reworked calibration | 19.2373 | 19.2373 | 19.2373 | -23.0509 | FAIL |  |  |  |  | 1 |  | 201.7228 | 201.7228 | 201.7228 | -19.3109 | PASS | [5 empty] | 37
2020-12-01 | OM_12-1-2020_11-34-15AM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0511 | 1.17 | 15.3 | TN | ug N/L | 0.172 | 3.19 | 189 |  | 1 | 2 | reworked calibration | 21.9159 | 21.9159 | 21.9159 | -12.3364 | PASS |  |  |  |  | 1 |  | 200.7832 | 200.7832 | 200.7832 | -19.6867 | PASS | [5 empty] | 38
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0531 | 1.12 | 26 | TN | ug N/L | 0.337 | 4.92 | 258 |  | 1 | 1 |  | 26 | 26 | 26 | 4 | PASS |  |  |  |  | 1 |  | 258 | 258 | 258 | 3.2 | PASS | [5 empty] | 39
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0466 | 1.06 | 24.3 | TN | ug N/L | 0.269 | 3.94 | 189 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 1 |  | 189 | 189 | 189 | -24.4 | FAIL | [5 empty] | 40
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0495 | 1.12 | 26.1 | TN | ug N/L | 0.281 | 4.13 | 202 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | [6 empty] | 1 |  | 202 | 202 | 202 | [7 empty] | 41
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0511 | 1.16 | 27.5 | TN | ug N/L | 0.323 | 4.71 | 243 |  | 1 | 1 |  | 27.5 | 27.5 | 27.5 | 10 | PASS |  |  |  |  | 1 |  | 243 | 243 | 243 | -2.8 | PASS | [5 empty] | 42
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0506 | 1.11 | 26 | TN | ug N/L | 0.367 | 5.33 | 286 |  | 1 | 1 |  | 26 | 26 | 26 | 4 | PASS |  |  |  |  | 1 |  | 286 | 286 | 286 | 14.4 | PASS | [5 empty] | 43
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0503 | 1.13 | 26.3 | TN | ug N/L | 0.335 | 4.85 | 253 |  | 1 | 1 |  | 26.3 | 26.3 | 26.3 | [6 empty] | 1 |  | 253 | 253 | 253 | [7 empty] | 44
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0489 | 1.11 | 25.8 | TN | ug N/L | 0.33 | 4.8 | 250 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | 3.2 | PASS |  |  |  |  | 1 |  | 250 | 250 | 250 | 0 | PASS | [5 empty] | 45
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0492 | 1.12 | 26.2 | TN | ug N/L | 0.309 | 4.52 | 230 |  | 1 | 1 |  | 26.2 | 26.2 | 26.2 | 4.8 | PASS |  |  |  |  | 1 |  | 230 | 230 | 230 | -8 | PASS | [5 empty] | 46
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0477 | 0.998 | 22.3 | TN | ug N/L | 0.32 | 4.66 | 240 |  | 1 | 1 |  | 22.3 | 22.3 | 22.3 | [6 empty] | 1 |  | 240 | 240 | 240 | [7 empty] | 47
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0475 | 1.07 | 24.5 | TN | ug N/L | 0.31 | 4.49 | 228 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | -2 | PASS |  |  |  |  | 1 |  | 228 | 228 | 228 | -8.8 | PASS | [5 empty] | 48
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0583 | 1.17 | 27.7 | TN | ug N/L | 0.353 | 5.09 | 270 |  | 1 | 1 |  | 27.7 | 27.7 | 27.7 | 10.8 | PASS |  |  |  |  | 1 |  | 270 | 270 | 270 | 8 | PASS | [5 empty] | 49
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.049 | 1.12 | 26.1 | TN | ug N/L | 0.356 | 5.11 | 271 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | 4.4 | PASS |  |  |  |  | 1 |  | 271 | 271 | 271 | 8.4 | PASS | [5 empty] | 50
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0514 | 1.15 | 26.9 | TN | ug N/L | 0.351 | 5.1 | 271 |  | 1 | 1 |  | 26.9 | 26.9 | 26.9 | 7.6 | PASS |  |  |  |  | 1 |  | 271 | 271 | 271 | 8.4 | PASS | [5 empty] | 51
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 18 | 1 | 1 | TP | ug P/L | 0.0585 | 1.2 | 24.2 | TN | ug N/L | 0.264 | 4.01 | 190 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | [6 empty] | 1 |  | 190 | 190 | 190 | [7 empty] | 52
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 19 | 1 | 1 | TP | ug P/L | 0.0602 | 1.22 | 24.8 | TN | ug N/L | 0.317 | 4.73 | 235 |  | 1 | 1 |  | 24.8 | 24.8 | 24.8 | -0.8 | PASS |  |  |  |  | 1 |  | 235 | 235 | 235 | -6 | PASS | [5 empty] | 53
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 17 | 1 | 1 | TP | ug P/L | 0.0596 | 1.22 | 24.6 | TN | ug N/L | 0.318 | 4.7 | 234 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 234 | 234 | 234 | -6.4 | PASS | [5 empty] | 54
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 18 | 1 | 1 | TP | ug P/L | 0.0594 | 1.24 | 25.2 | TN | ug N/L | 0.313 | 4.63 | 230 |  | 1 | 1 |  | 25.2 | 25.2 | 25.2 | 0.8 | PASS |  |  |  |  | 1 |  | 230 | 230 | 230 | -8 | PASS | [5 empty] | 55
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 19 | 1 | 1 | TP | ug P/L | 0.0596 | 1.2 | 24.2 | TN | ug N/L | 0.301 | 4.49 | 221 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 221 | 221 | 221 | -11.6 | PASS | [5 empty] | 56
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 7 | 1 | 1 | TP | ug P/L | 0.0597 | 1.23 | 24.9 | TN | ug N/L | 0.321 | 4.77 | 238 |  | 1 | 1 |  | 24.9 | 24.9 | 24.9 | -0.4 | PASS |  |  |  |  | 1 |  | 238 | 238 | 238 | -4.8 | PASS | [5 empty] | 57
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 17 | 1 | 1 | TP | ug P/L | 0.0612 | 1.22 | 24.6 | TN | ug N/L | 0.314 | 4.69 | 233 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 233 | 233 | 233 | -6.8 | PASS | [5 empty] | 58
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 18 | 1 | 1 | TP | ug P/L | 0.0602 | 1.23 | 24.9 | TN | ug N/L | 0.301 | 4.44 | 217 |  | 1 | 1 |  | 24.9 | 24.9 | 24.9 | -0.4 | PASS |  |  |  |  | 1 |  | 217 | 217 | 217 | -13.2 | PASS | [5 empty] | 59
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 19 | 1 | 1 | TP | ug P/L | 0.0605 | 1.23 | 25.1 | TN | ug N/L | 0.308 | 4.55 | 224 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | 0.4 | PASS |  |  |  |  | 1 |  | 224 | 224 | 224 | -10.4 | PASS | [5 empty] | 60
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 17 | 1 | 1 | TP | ug P/L | 0.0604 | 1.21 | 24.3 | TN | ug N/L | 0.319 | 4.74 | 236 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 1 |  | 236 | 236 | 236 | -5.6 | PASS | [5 empty] | 61
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | chk std 25 | 7 | 1 | 1 | TP | ug P/L | 0.0606 | 1.22 | 24.6 | TN | ug N/L | 0.313 | 4.65 | 230 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 230 | 230 | 230 | -8 | PASS | [5 empty] | 62
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.058 | 1.1 | 24 | TN | ug N/L | 0.333 | 4.59 | 226 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 1 |  | 226 | 226 | 226 | -9.6 | PASS | [5 empty] | 63
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0591 | 1.12 | 24.4 | TN | ug N/L | 0.329 | 4.53 | 222 |  | 1 | 1 |  | 24.4 | 24.4 | 24.4 | -2.4 | PASS |  |  |  |  | 1 |  | 222 | 222 | 222 | -11.2 | PASS | [5 empty] | 64
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0579 | 1.11 | 24.3 | TN | ug N/L | 0.323 | 4.49 | 220 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | [6 empty] | 1 |  | 220 | 220 | 220 | [7 empty] | 65
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0591 | 1.11 | 24.3 | TN | ug N/L | 0.328 | 4.52 | 221 |  | 1 | 1 |  | 24.3 | 24.3 | 24.3 | -2.8 | PASS |  |  |  |  | 1 |  | 221 | 221 | 221 | -11.6 | PASS | [5 empty] | 66
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.058 | 1.09 | 23.8 | TN | ug N/L | 0.32 | 4.39 | 213 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | -4.8 | PASS |  |  |  |  | 1 |  | 213 | 213 | 213 | -14.8 | PASS | [5 empty] | 67
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0588 | 1.11 | 24.2 | TN | ug N/L | 0.311 | 4.31 | 208 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 208 | 208 | 208 | -16.8 | PASS | [5 empty] | 68
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0582 | 1.07 | 23.2 | TN | ug N/L | 0.313 | 4.31 | 208 |  | 1 | 1 |  | 23.2 | 23.2 | 23.2 | -7.2 | PASS |  |  |  |  | 1 |  | 208 | 208 | 208 | -16.8 | PASS | [5 empty] | 69
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0597 | 1.12 | 24.6 | TN | ug N/L | 0.319 | 4.43 | 216 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 216 | 216 | 216 | -13.6 | PASS | [5 empty] | 70
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0592 | 1.1 | 24 | TN | ug N/L | 0.301 | 4.19 | 200 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 1 |  | 200 | 200 | 200 | -20 | PASS | [5 empty] | 71
2022-02-23 | OM_2-23-2022_11-17-46AM edit 1.omn | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0591 | 1.14 | 25.1 | TN | ug N/L | 0.314 | 4.32 | 209 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | [6 empty] | 1 |  | 209 | 209 | 209 | [7 empty] | 72
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0662 | 1.26 | 22.9 | TN | ug N/L | 0.378 | 5.09 | 253 |  | 1 | 1 |  | 22.9 | 22.9 | 22.9 | [6 empty] | 1 |  | 253 | 253 | 253 | [7 empty] | 73
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0646 | 1.2 | 21.5 | TN | ug N/L | 0.304 | 4.13 | 190 |  | 1 | 1 |  | 21.5 | 21.5 | 21.5 | -14 | PASS |  |  |  |  | 1 |  | 190 | 190 | 190 | -24 | FAIL | [5 empty] | 74
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0649 | 1.22 | 21.8 | TN | ug N/L | 0.355 | 4.8 | 234 |  | 1 | 1 |  | 21.8 | 21.8 | 21.8 | -12.8 | PASS |  |  |  |  | 1 |  | 234 | 234 | 234 | -6.4 | PASS | [5 empty] | 75
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0633 | 1.19 | 21.2 | TN | ug N/L | 0.338 | 4.56 | 218 |  | 1 | 1 |  | 21.2 | 21.2 | 21.2 | -15.2 | PASS |  |  |  |  | 1 |  | 218 | 218 | 218 | -12.8 | PASS | [5 empty] | 76
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0663 | 1.26 | 22.8 | TN | ug N/L | 0.339 | 4.61 | 221 |  | 1 | 1 |  | 22.8 | 22.8 | 22.8 | -8.8 | PASS |  |  |  |  | 1 |  | 221 | 221 | 221 | -11.6 | PASS | [5 empty] | 77
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0641 | 1.2 | 21.4 | TN | ug N/L | 0.343 | 4.65 | 224 |  | 1 | 1 |  | 21.4 | 21.4 | 21.4 | -14.4 | PASS |  |  |  |  | 1 |  | 224 | 224 | 224 | -10.4 | PASS | [5 empty] | 78
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0643 | 1.19 | 21.1 | TN | ug N/L | 0.323 | 4.38 | 206 |  | 1 | 1 |  | 21.1 | 21.1 | 21.1 | -15.6 | PASS |  |  |  |  | 1 |  | 206 | 206 | 206 | -17.6 | PASS | [5 empty] | 79
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.066 | 1.19 | 21.3 | TN | ug N/L | 0.265 | 3.64 | 157 |  | 1 | 1 |  | 21.3 | 21.3 | 21.3 | -14.8 | PASS |  |  |  |  | 1 |  | 157 | 157 | 157 | -37.2 | FAIL | [5 empty] | 80
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.068 | 1.28 | 23.5 | TN | ug N/L | 0.331 | 4.58 | 219 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 | -6 | PASS |  |  |  |  | 1 |  | 219 | 219 | 219 | -12.4 | PASS | [5 empty] | 81
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0698 | 1.3 | 23.9 | TN | ug N/L | 0.327 | 4.46 | 211 |  | 1 | 1 |  | 23.9 | 23.9 | 23.9 | -4.4 | PASS |  |  |  |  | 1 |  | 211 | 211 | 211 | -15.6 | PASS | [5 empty] | 82
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0726 | 1.33 | 24.6 | TN | ug N/L | 0.361 | 4.9 | 240 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 240 | 240 | 240 | -4 | PASS | [5 empty] | 83
2022-03-23 | OM_3-23-2022_02-00-59PM.OMN | chk 25 | 7 | 1 | 1 | TP | ug P/L | 0.0738 | 1.36 | 25.4 | TN | ug N/L | 0.357 | 4.86 | 238 |  | 1 | 1 |  | 25.4 | 25.4 | 25.4 | 1.6 | PASS |  |  |  |  | 1 |  | 238 | 238 | 238 | -4.8 | PASS | [5 empty] | 84
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0702 | 1.25 | 26.1 | TN | ug N/L | 0.327 | 4.56 | 247 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | 4.4 | PASS |  |  |  |  | 1 |  | 247 | 247 | 247 | -1.2 | PASS | [5 empty] | 85
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0687 | 1.19 | 24.2 | TN | ug N/L | 0.32 | 4.47 | 241 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 241 | 241 | 241 | -3.6 | PASS | [5 empty] | 86
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0692 | 1.22 | 25.3 | TN | ug N/L | 0.321 | 4.46 | 240 |  | 1 | 1 |  | 25.3 | 25.3 | 25.3 | 1.2 | PASS |  |  |  |  | 1 |  | 240 | 240 | 240 | -4 | PASS | [5 empty] | 87
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0717 | 1.28 | 27 | TN | ug N/L | 0.313 | 4.36 | 233 |  | 1 | 1 |  | 27 | 27 | 27 | 8 | PASS |  |  |  |  | 1 |  | 233 | 233 | 233 | -6.8 | PASS | [5 empty] | 88
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk250 | 7 | 1 | 1 | TP | ug P/L | 0.0695 | 1.25 | 26 | TN | ug N/L | 0.303 | 4.24 | 225 |  | 1 | 1 |  | 26 | 26 | 26 | 4 | PASS |  |  |  |  | 1 |  | 225 | 225 | 225 | -10 | PASS | [5 empty] | 89
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0688 | 1.23 | 25.5 | TN | ug N/L | 0.301 | 4.23 | 224 |  | 1 | 1 |  | 25.5 | 25.5 | 25.5 | 2 | PASS |  |  |  |  | 1 |  | 224 | 224 | 224 | -10.4 | PASS | [5 empty] | 90
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0672 | 1.19 | 24.2 | TN | ug N/L | 0.301 | 4.2 | 222 |  | 1 | 1 |  | 24.2 | 24.2 | 24.2 | -3.2 | PASS |  |  |  |  | 1 |  | 222 | 222 | 222 | -11.2 | PASS | [5 empty] | 91
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.069 | 1.2 | 24.6 | TN | ug N/L | 0.289 | 4.06 | 212 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 212 | 212 | 212 | -15.2 | PASS | [5 empty] | 92
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0695 | 1.27 | 26.5 | TN | ug N/L | 0.3 | 4.24 | 224 |  | 1 | 1 |  | 26.5 | 26.5 | 26.5 | 6 | PASS |  |  |  |  | 1 |  | 224 | 224 | 224 | -10.4 | PASS | [5 empty] | 93
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0679 | 1.2 | 24.5 | TN | ug N/L | 0.3 | 4.2 | 222 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | -2 | PASS |  |  |  |  | 1 |  | 222 | 222 | 222 | -11.2 | PASS | [5 empty] | 94
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | chk250 | 7 | 1 | 1 | TP | ug P/L | 0.0689 | 1.24 | 25.8 | TN | ug N/L | 0.341 | 4.77 | 262 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | 3.2 | PASS |  |  |  |  | 1 |  | 262 | 262 | 262 | 4.8 | PASS | [5 empty] | 95
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0742 | 1.28 | 25.8 | TN | ug N/L | 0.39 | 4.8 | 257 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | 3.2 | PASS |  |  |  |  | 1 |  | 257 | 257 | 257 | 2.8 | PASS | [5 empty] | 96
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0738 | 1.24 | 24.6 | TN | ug N/L | 0.387 | 4.78 | 255 |  | 1 | 1 |  | 24.6 | 24.6 | 24.6 | -1.6 | PASS |  |  |  |  | 1 |  | 255 | 255 | 255 | 2 | PASS | [5 empty] | 97
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0765 | 1.29 | 26 | TN | ug N/L | 0.374 | 4.61 | 243 |  | 1 | 1 |  | 26 | 26 | 26 | 4 | PASS |  |  |  |  | 1 |  | 243 | 243 | 243 | -2.8 | PASS | [5 empty] | 98
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0774 | 1.35 | 27.7 | TN | ug N/L | 0.371 | 4.65 | 246 |  | 1 | 1 |  | 27.7 | 27.7 | 27.7 | 10.8 | PASS |  |  |  |  | 1 |  | 246 | 246 | 246 | -1.6 | PASS | [5 empty] | 99
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0764 | 1.33 | 27.4 | TN | ug N/L | 0.405 | 5.08 | 275 |  | 1 | 1 |  | 27.4 | 27.4 | 27.4 | 9.6 | PASS |  |  |  |  | 1 |  | 275 | 275 | 275 | 10 | PASS | [5 empty] | 100
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 250 | 7 | 1 | 1 | TP | ug P/L | 0.0732 | 1.22 | 24 | TN | ug N/L | 0.348 | 4.43 | 231 |  | 1 | 1 |  | 24 | 24 | 24 | -4 | PASS |  |  |  |  | 1 |  | 231 | 231 | 231 | -7.6 | PASS | [5 empty] | 101
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0751 | 1.26 | 25.3 | TN | ug N/L | 0.356 | 4.51 | 237 |  | 1 | 1 |  | 25.3 | 25.3 | 25.3 | 1.2 | PASS |  |  |  |  | 1 |  | 237 | 237 | 237 | -5.2 | PASS | [5 empty] | 102
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.075 | 1.28 | 25.7 | TN | ug N/L | 0.37 | 4.76 | 254 |  | 1 | 1 |  | 25.7 | 25.7 | 25.7 | 2.8 | PASS |  |  |  |  | 1 |  | 254 | 254 | 254 | 1.6 | PASS | [5 empty] | 103
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0734 | 1.25 | 24.8 | TN | ug N/L | 0.361 | 4.65 | 246 |  | 1 | 1 |  | 24.8 | 24.8 | 24.8 | -0.8 | PASS |  |  |  |  | 1 |  | 246 | 246 | 246 | -1.6 | PASS | [5 empty] | 104
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk25/250 | 7 | 1 | 1 | TP | ug P/L | 0.0751 | 1.29 | 26.1 | TN | ug N/L | 0.361 | 4.65 | 246 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | 4.4 | PASS |  |  |  |  | 1 |  | 246 | 246 | 246 | -1.6 | PASS | [5 empty] | 105
2022-05-04 | OM_5-4-2022_10-40-30AM.OMN | chk 250 | 7 | 1 | 1 | TP | ug P/L | 0.0762 | 1.28 | 25.9 | TN | ug N/L | 0.29 | 3.78 | 187 |  | 1 | 1 |  | 25.9 | 25.9 | 25.9 | 3.6 | PASS |  |  |  |  | 1 |  | 187 | 187 | 187 | -25.2 | FAIL | [5 empty] | 106
[20 empty] | TP | [10 empty] | TN
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | [6 empty] | OBSERVATION #
[19 empty] | Mean |  |  | 24.4015 | [7 empty] | Mean |  |  | 231.8313 | [6 empty] | Min | 1
[19 empty] | Std |  |  | 1.6942 | [7 empty] | Std |  |  | 24.7569 | [6 empty] | Max | 106
[19 empty] | CV |  |  | 6.9429 | [7 empty] | CV |  |  | 10.6789 | [6 empty] | Mean
[19 empty] | S/N |  |  | 14.4033 | limits as percent of known | [6 empty] | S/N |  |  | 9.3643 | limits as percent of known | [5 empty] | Whole series
[18 empty] | Warning Limits | Xbar + 2SD |  |  | 27.7898 | 111.1593 | [5 empty] | Warning Limits | Xbar + 2SD |  |  | 281.3452 | 112.5381
[19 empty] | Xbar - 2SD |  |  | 21.0132 | 84.0528 | [6 empty] | Xbar - 2SD |  |  | 182.3174 | 72.927
[18 empty] | Control Limits | Xbar + 3SD |  |  | 29.484 | 117.936 | [5 empty] | Control Limits | Xbar + 3SD |  |  | 306.1021 | 122.4408
[19 empty] | Xbar - 3SD |  |  | 19.319 | 77.2761 | [6 empty] | Xbar - 3SD |  |  | 157.5605 | 63.0242
[18 empty] | Count |  |  |  | 106 | [6 empty] | Count |  |  |  | 106
[18 empty] | 99th percentile.inc |  |  |  | 27.69 | [6 empty] | 99th percentile.inc |  |  |  | 285.928
[18 empty] | Maximum |  |  |  | 27.7 | [6 empty] | Maximum |  |  |  | 309.588
[18 empty] | Known |  |  |  | 40 | [6 empty] | Known |  |  |  | 250
[18 empty] | MDL |  |  |  | 4.0023 | [6 empty] | MDL |  |  |  | 58.4855
[18 empty] | LOQ |  |  |  | 16.9416 | [6 empty] | LOQ |  |  |  | 247.5694
[18 empty] | ratio mean/mdl |  |  |  | 6.0969 | [6 empty] | ratio mean/mdl |  |  |  | 3.9639
[21 empty] | Known Conc | 25 | [9 empty] | Known Conc | 250
[21 empty] | Mean | 24.4015 | [9 empty] | Mean | 231.8313
[21 empty] | Std Dev | 1.6942 | [9 empty] | Std Dev | 24.7569
[21 empty] | CV or %RSD | 6.9429 | [9 empty] | CV or %RSD | 10.6789
[21 empty] | T value | 2.3624 | [9 empty] | T value | 2.3624
[21 empty] | MDL | 4.0023 | [9 empty] | MDL | 58.4855
[21 empty] | LOQ | 16.9416 | [9 empty] | LOQ | 247.5694
[21 empty] | Error as % | -2.394 | [9 empty] | Error as % | -7.2675
[21 empty] | Known/MDL | 6.2465 | [9 empty] | Known/MDL | 4.2746
[21 empty] | Mean PR | 97.606 | [9 empty] | Mean PR | 92.7325
[21 empty] | S/N | 14.4033 | [9 empty] | S/N | 9.3643

## efficiency
Recoveries for N in particular seem low last time through.
True for both organic and inorganic.
Actions:
Remake stock solutions
Look at alternate organic N substrate
Pipet more carefully
Up the temperature? Shouldn't make a difference
Pay particular attention to tight caps and volume after digestion
Would loose caps and low volume lead to lower N? Not sure. Has a gas phase so it's possible. But wouldn't P go up?
Hagedorn and Schleppi 2000 say extra head space will increase N concentrations (atmospheric N gets captured and N goes up)
Matched inorganic and organic digestions
Inorganic
100 mls of 40 P and 400 N
using primary stock
|  | stock ppb | mls | target ppb | mls | uL
|  | 1000000 | 100 | 400 | 0.04 | 40
|  | need 2ndary stock |  | 1000 | ppb | for P but not for N
|  | stock ppb | mls | target ppb | mls | uL
|  | 1000 | 100 | 40 | 4 | 4000
|  | 1000 | 100 | 400 | 40 | 40000
Organic - unifomly use glycerophosphate stock for P
|  | stock ppb | mls | target ppb | mls | uL
|  | 400000 | 100 | 40 | 0.01 | 10
| a 10x dilution | 40000 | 100 | 40 | 0.1 | 100
USGS digest check
| Glycine hydrochloride (as per USGS; 7.96 g/L = 1000 mgN/L) |  |  | C2H5NO2-HCl | FW = 111.5 | 6000-43-7
| Glycerophosphate |  |  | C3H7O6PNa2-5H2O | FW = 306.1 | 13408-09-8
| Glucose |  |  | C6H12O6 | FW = 180.2 | 50-99-7
|  | stock ppb | mls | target ppb | mls | uL
| N | 1000000 | 100 | 400 | 0.04 | 40
| C | 1714000
| P 10x dilution | 40000 | 100 | 40 | 0.1
| C 10x dilution | 46500 | 100 | 46.5 | 0.1
C from Glucose
| Urea (0.214g/L = 100 mgN/L) |  |  | CH4N2O | FW = 60
|  | Then dilute into 100 ml | [5 empty] | N |  | C
|  | stock ppb | mls | target ppb | mls | uL | 28 | 0.099867 | 12 | 0.0428 | g
| N | 100000 | 100 | 400 | 0.4 | 400 | 60 | 0.214 | 60 | 0.214 | g
| C | 42800 | 100 | 171.2 | 0.4
| P from Glycerophosphate |  |  | 40
| C from Glycerophosphate |  |  | 46.5
| Sum of C |  |  | 217.7
| L- Glutamic Acid (1.051g/L = 100mgN/L) |  |  | C5H9NO4 | FW = 147
|  | Then dilute into 100 ml | [5 empty] | N |  | C
|  | stock ppb | mls | target ppb | mls | uL | 14 | 0.1001 | 60 | 0.42898 | g
| N | 100000 | 100 | 400 | 0.4 | 400 | 147 | 1.051 | 147 | 1.051 | g
| C | 428980 | 100 | 1715.92 | 0.4 | 400
| P from Glycerophosphate |  |  | 40
| C from Glycerophosphate |  |  | 46.5
| Sum of C |  |  | 1762.42
Using CN standard solution for an alternate digestion check?
1 g Urea and 5 g Glucose per liter
| 7.75 | %N | 0.0775 | mg/L | 77.5
| 36.62 | %C | 0.3662 | mg/L | 366.2
urea | 28 | just N
| 60.6 | formula weight
How about going the other way?
Use the TN TP standard to do CN?
|  | mg N | mg C
| MDL is | 0.08 | 0.1
|  | mg N/L | mgC/L | mgP/L | whole compound
| Glycine | 1000 | 1714 | 0 | 7.96 | g per L in stock solution
| Glucose | 0 | 1250 | 0 | 3.128 | g per L in stock solution
| Glycerophosphate | 0 | 0 | 400 | 3.952 | g per L in stock solution
| mls | 0.08
| if I use 0.08 | 0.08 | 0.13712
bump it up to 0.1
So... just use Glycine stock solution
pipet 100 ul into each crucible and dry.
|  | 0.1 | 0.1714 | 0.796
|  | mg N | mg C | mg whole compound
glutamate
8g/L | 14 | just N |  | 0.1 | g N
| 147 | formula weight |  | 1.05 | g glutamate
TN TP |  | mgP/L
| Glycerophosphate | 400
maybe dilute 20x to
20

## more digest checks
NOTES
The recovery of nitrogen is close to 100% when compared to paired inorganic samples, but less than that when read directly off calibration curve. WHY?
Phosphate recovery in spiked blanks and spiked samples is over 100% this time through. Higher than usual. Higher temp?
Did I make urea solution incorrectly? TOC and persulfate agree that recovery is low.
Solutions made:
Regular USGS digest check as per SOP
Urea instead of glycine for organic nitrogen. Glycerophosphate for TP. No glucose.
urea solution @ 0.214g/L then used 0.4 mls into 100 mls for 400 ppb
L glutamic acid instead of glycine for organic nitrogen. Glycerophosphate for TP. No glucose.
L-glutamic acid @ 1.051g/L then used 0.4 mls into 100 mls for 400 ppb
DIGEST CHECKS
DATA FROM TOC ANALYZER run april 6 2022
|  | Name | NPOC Mean of 2 reps mgC/L | TNb Mean of 2 reps mgN/L | Mean | PR of TN compared to known | PR of TN compared to paired inorganic |  | Known C | Known P | Known N
|  | inorganic recovery check | -0.084024 | 0.45117 | 0.45553 | 113.8827 | 100 |  | 0 | 40 | 400
|  | inorganic recovery check | 0.078799 | 0.45989
|  | organic digest check glycine | 2.6903 | 0.40121 | 0.40178 | 100.4453 | 88.2007 |  | 2.73 | 40 | 400
|  | organic digest check glycine | 2.6874 | 0.40236
|  | organic digest check L glutamic | 1.8006 | 0.43079 | 0.44 | 109.9989 | 96.5897 |  | 1.762 | 40 | 400
|  | organic digest check L glutamic | 1.8747 | 0.4492
|  | organic digest check urea | 0.14519 | 0.21047 | 0.21553 | 53.8821 | 47.3137 |  | 0.218 | 40 | 400
|  | organic digest check urea | 0.1258 | 0.22059
SAME THING FOR PERSULFATE DIGESTION run april 13 2022
|  | inorganic recovery check |  |  | 346 | 86.5 | 100
inorganic recovery check
|  | organic digest check glycine |  |  | 362 | 90.5 | 104.6243
organic digest check glycine
|  | organic digest check L glutamic |  |  | 342.6667 | 85.6667 | 99.0366
organic digest check L glutamic
|  | organic digest check urea |  |  | 145 | 36.25 | 41.9075
organic digest check urea
DIGESTION CHECK - USGS = GLYCINE
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check1 | 25 | 1 | 1 | TP | ug P/L | 0.0941 | 1.67 | 37.7 | TN | ug N/L | 0.453 | 6.33 | 372 | 1 | 1 | 37.7 | 37.7 | 37.7 | 1 | 372 | 372 | 372
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check2 | 26 | 1 | 1 | TP | ug P/L | 0.0948 | 1.66 | 37.4 | TN | ug N/L | 0.436 | 6.12 | 357 | 1 | 1 | 37.4 | 37.4 | 37.4 | 1 | 357 | 357 | 357
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digestcheck3 | 27 | 1 | 1 | TP | ug P/L | 0.0908 | 1.58 | 35.2 | TN | ug N/L | 0.425 | 5.93 | 344 | 1 | 1 | 35.2 | 35.2 | 35.2 | 1 | 344 | 344 | 344
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digestcheck4 | 28 | 1 | 1 | TP | ug P/L | 0.0929 | 1.65 | 37.2 | TN | ug N/L | 0.448 | 6.21 | 363 | 1 | 1 | 37.2 | 37.2 | 37.2 | 1 | 363 | 363 | 363
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check5 | 29 | 1 | 1 | TP | ug P/L | 0.0941 | 1.65 | 37.3 | TN | ug N/L | 0.459 | 6.36 | 374 | 1 | 1 | 37.3 | 37.3 | 37.3 | 1 | 374 | 374 | 374
[9 empty] | Known Conc | 40 |  |  |  | Known Conc | 400 |  |  |  | Known Conc | 40 |  |  | Known Conc | 400
[9 empty] | Mean | 36.96 |  |  |  | Mean | 362 |  |  |  | Mean | 36.96 |  |  | Mean | 362
[9 empty] | Std Dev | 1.0015 |  |  |  | Std Dev | 12.1861 |  |  |  | Std Dev | 1.0015 |  |  | Std Dev | 12.1861
[9 empty] | CV or %RSD | 2.7097 |  |  |  | CV or %RSD | 3.3663 |  |  |  | CV or %RSD | 2.7097 |  |  | CV or %RSD | 3.3663
[9 empty] | T value | 3.7469 |  |  |  | T value | 3.7469 |  |  |  | T value | 3.7469 |  |  | T value | 3.7469
[9 empty] | MDL | 3.7526 |  |  |  | MDL | 45.6605 |  |  |  | MDL | 3.7526 |  |  | MDL | 45.6605
[9 empty] | LOQ | 10.015 |  |  |  | LOQ | 121.8606 |  |  |  | LOQ | 10.015 |  |  | LOQ | 121.8606
[9 empty] | Error as % | -7.6 |  |  |  | Error as % | -9.5 |  |  |  | Error as % | -7.6 |  |  | Error as % | -9.5
[9 empty] | Known/MDL |  |  |  |  | Known/MDL |  |  |  |  | Known/MDL |  |  |  | Known/MDL
[9 empty] | Mean PR | 92.4 |  |  |  | Mean PR | 90.5 |  |  |  | Mean PR |  |  |  | Mean PR
[9 empty] | S/N | 36.9047 |  |  |  | S/N | 29.7061 |  |  |  | S/N | 36.9047 |  |  | S/N | 29.7061
DIGESTION CHECK - PAIRED INORGANIC
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | inorganic 40 400 | 158 | 1 | 1 | TP | ug P/L | 0.0999 | 1.75 | 39.9 | TN | ug N/L | 0.421 | 5.87 | 340 | 1 | 1 | 39.9 | 39.9 | 39.9 | 1 | 340 | 340 | 340
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | inorganic 40 400 | 159 | 1 | 1 | TP | ug P/L | 0.101 | 1.74 | 39.8 | TN | ug N/L | 0.436 | 6.06 | 353 | 1 | 1 | 39.8 | 39.8 | 39.8 | 1 | 353 | 353 | 353
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | inorganic 40 400 | 160 | 1 | 1 | TP | ug P/L | 0.101 | 1.76 | 40.4 | TN | ug N/L | 0.428 | 5.95 | 345 | 1 | 1 | 40.4 | 40.4 | 40.4 | 1 | 345 | 345 | 345
[9 empty] | Known Conc | 40 |  |  |  | Known Conc | 400 |  |  |  | Known Conc | 40 |  |  | Known Conc | 400
[9 empty] | Mean | 40.0333 |  |  |  | Mean | 346 |  |  |  | Mean | 40.0333 |  |  | Mean | 346
[9 empty] | Std Dev | 0.32146 |  |  |  | Std Dev | 6.5574 |  |  |  | Std Dev | 0.32146 |  |  | Std Dev | 6.5574
[9 empty] | CV or %RSD | 0.80297 |  |  |  | CV or %RSD | 1.8952 |  |  |  | CV or %RSD | 0.80297 |  |  | CV or %RSD | 1.8952
[9 empty] | T value | 3.7469 |  |  |  | T value | 3.7469 |  |  |  | T value | 3.7469 |  |  | T value | 3.7469
[9 empty] | MDL | 1.2045 |  |  |  | MDL | 24.5704 |  |  |  | MDL | 1.2045 |  |  | MDL | 24.5704
[9 empty] | LOQ | 3.2146 |  |  |  | LOQ | 65.5744 |  |  |  | LOQ | 3.2146 |  |  | LOQ | 65.5744
[9 empty] | Error as % | 0.083333 |  |  |  | Error as % | -13.5 |  |  |  | Error as % | 0.083333 |  |  | Error as % | -13.5
[9 empty] | Known/MDL |  |  |  |  | Known/MDL |  |  |  |  | Known/MDL |  |  |  | Known/MDL
[9 empty] | Mean PR | 100.0833 |  |  |  | Mean PR | 86.5 |  |  |  | Mean PR |  |  |  | Mean PR
[9 empty] | S/N | 124.5379 |  |  |  | S/N | 52.7645 |  |  |  | S/N | 124.5379 |  |  | S/N | 52.7645
DIGESTION CHECK - UREA
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | urea 40 400 | 161 | 1 | 1 | TP | ug P/L | 0.0993 | 1.74 | 39.7 | TN | ug N/L | 0.228 | 3.22 | 152 | 1 | 1 | 39.7 | 39.7 | 39.7 | 1 | 152 | 152 | 152
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | urea 40 400 | 162 | 1 | 1 | TP | ug P/L | 0.1 | 1.75 | 40 | TN | ug N/L | 0.214 | 3 | 137 | 1 | 1 | 40 | 40 | 40 | 1 | 137 | 137 | 137
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | urea 40 400 | 163 | 1 | 1 | TP | ug P/L | 0.0996 | 1.72 | 39.3 | TN | ug N/L | 0.223 | 3.13 | 146 | 1 | 1 | 39.3 | 39.3 | 39.3 | 1 | 146 | 146 | 146
[9 empty] | Known Conc | 40 |  |  |  | Known Conc | 400 |  |  |  | Known Conc | 40 |  |  | Known Conc | 400
[9 empty] | Mean | 39.6667 |  |  |  | Mean | 145 |  |  |  | Mean | 39.6667 |  |  | Mean | 145
[9 empty] | Std Dev | 0.35119 |  |  |  | Std Dev | 7.5498 |  |  |  | Std Dev | 0.35119 |  |  | Std Dev | 7.5498
[9 empty] | CV or %RSD | 0.88535 |  |  |  | CV or %RSD | 5.2068 |  |  |  | CV or %RSD | 0.88535 |  |  | CV or %RSD | 5.2068
[9 empty] | T value | 3.7469 |  |  |  | T value | 3.7469 |  |  |  | T value | 3.7469 |  |  | T value | 3.7469
[9 empty] | MDL | 1.3159 |  |  |  | MDL | 28.2888 |  |  |  | MDL | 1.3159 |  |  | MDL | 28.2888
[9 empty] | LOQ | 3.5119 |  |  |  | LOQ | 75.4983 |  |  |  | LOQ | 3.5119 |  |  | LOQ | 75.4983
[9 empty] | Error as % | -0.83333 |  |  |  | Error as % | -63.75 |  |  |  | Error as % | -0.83333 |  |  | Error as % | -63.75
[9 empty] | Known/MDL |  |  |  |  | Known/MDL |  |  |  |  | Known/MDL |  |  |  | Known/MDL
[9 empty] | Mean PR | 99.1667 |  |  |  | Mean PR | 36.25 |  |  |  | Mean PR |  |  |  | Mean PR
[9 empty] | S/N | 112.9498 |  |  |  | S/N | 19.2057 |  |  |  | S/N | 112.9498 |  |  | S/N | 19.2057
DIGESTION CHECK - L glutamic acid
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | lgulatmic acid 40 400 | 164 | 1 | 1 | TP | ug P/L | 0.0962 | 1.69 | 38.3 | TN | ug N/L | 0.414 | 5.74 | 330 | 1 | 1 | 38.3 | 38.3 | 38.3 | 1 | 330 | 330 | 330
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | lgulatmic acid 40 400 | 165 | 1 | 1 | TP | ug P/L | 0.1 | 1.76 | 40.3 | TN | ug N/L | 0.425 | 5.91 | 342 | 1 | 1 | 40.3 | 40.3 | 40.3 | 1 | 342 | 342 | 342
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | lgulatmic acid 40 400 | 166 | 1 | 1 | TP | ug P/L | 0.0994 | 1.76 | 40.4 | TN | ug N/L | 0.441 | 6.1 | 356 | 1 | 1 | 40.4 | 40.4 | 40.4 | 1 | 356 | 356 | 356
[9 empty] | Known Conc | 40 |  |  |  | Known Conc | 400 |  |  |  | Known Conc | 40 |  |  | Known Conc | 400
[9 empty] | Mean | 39.6667 |  |  |  | Mean | 342.6667 |  |  |  | Mean | 39.6667 |  |  | Mean | 342.6667
[9 empty] | Std Dev | 1.1846 |  |  |  | Std Dev | 13.0128 |  |  |  | Std Dev | 1.1846 |  |  | Std Dev | 13.0128
[9 empty] | CV or %RSD | 2.9864 |  |  |  | CV or %RSD | 3.7975 |  |  |  | CV or %RSD | 2.9864 |  |  | CV or %RSD | 3.7975
[9 empty] | T value | 3.7469 |  |  |  | T value | 3.7469 |  |  |  | T value | 3.7469 |  |  | T value | 3.7469
[9 empty] | MDL | 4.4387 |  |  |  | MDL | 48.7583 |  |  |  | MDL | 4.4387 |  |  | MDL | 48.7583
[9 empty] | LOQ | 11.8462 |  |  |  | LOQ | 130.1281 |  |  |  | LOQ | 11.8462 |  |  | LOQ | 130.1281
[9 empty] | Error as % | -0.83333 |  |  |  | Error as % | -14.3333 |  |  |  | Error as % | -0.83333 |  |  | Error as % | -14.3333
[9 empty] | Known/MDL |  |  |  |  | Known/MDL |  |  |  |  | Known/MDL |  |  |  | Known/MDL
[9 empty] | Mean PR | 99.1667 |  |  |  | Mean PR | 85.6667 |  |  |  | Mean PR |  |  |  | Mean PR
[9 empty] | S/N | 33.4846 |  |  |  | S/N | 26.333 |  |  |  | S/N | 33.4846 |  |  | S/N | 26.333
RAW DATA
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank1 | 20 | 1 | 1 | TP | ug P/L | 0.0136 | 0.308 | -0.453 | TN | ug N/L | 0.0585 | 0.989 | -5.79 | 1 | 1 | -0.453 | -0.453 | -0.453 | 1 | -5.79 | -5.79 | -5.79
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank2 | 21 | 1 | 1 | TP | ug P/L | 0.0125 | 0.345 | 0.576 | TN | ug N/L | 0.0608 | 1.01 | -4.53 | 1 | 1 | 0.576 | 0.576 | 0.576 | 1 | -4.53 | -4.53 | -4.53
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank3 | 22 | 1 | 1 | TP | ug P/L | 0.0134 | 0.348 | 0.675 | TN | ug N/L | 0.0589 | 0.973 | -6.93 | 1 | 1 | 0.675 | 0.675 | 0.675 | 1 | -6.93 | -6.93 | -6.93
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank4 | 23 | 1 | 1 | TP | ug P/L | 0.0126 | 0.337 | 0.353 | TN | ug N/L | 0.0623 | 0.972 | -7 | 1 | 1 | 0.353 | 0.353 | 0.353 | 1 | -7 | -7 | -7
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | waterblank5 | 24 | 1 | 1 | TP | ug P/L | 0.0115 | 0.301 | -0.66 | TN | ug N/L | 0.0674 | 1.09 | 1.59 | 1 | 1 | -0.66 | -0.66 | -0.66 | 1 | 1.59 | 1.59 | 1.59
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check1 | 25 | 1 | 1 | TP | ug P/L | 0.0941 | 1.67 | 37.7 | TN | ug N/L | 0.453 | 6.33 | 372 | 1 | 1 | 37.7 | 37.7 | 37.7 | 1 | 372 | 372 | 372
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check2 | 26 | 1 | 1 | TP | ug P/L | 0.0948 | 1.66 | 37.4 | TN | ug N/L | 0.436 | 6.12 | 357 | 1 | 1 | 37.4 | 37.4 | 37.4 | 1 | 357 | 357 | 357
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digestcheck3 | 27 | 1 | 1 | TP | ug P/L | 0.0908 | 1.58 | 35.2 | TN | ug N/L | 0.425 | 5.93 | 344 | 1 | 1 | 35.2 | 35.2 | 35.2 | 1 | 344 | 344 | 344
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digestcheck4 | 28 | 1 | 1 | TP | ug P/L | 0.0929 | 1.65 | 37.2 | TN | ug N/L | 0.448 | 6.21 | 363 | 1 | 1 | 37.2 | 37.2 | 37.2 | 1 | 363 | 363 | 363
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | digest check5 | 29 | 1 | 1 | TP | ug P/L | 0.0941 | 1.65 | 37.3 | TN | ug N/L | 0.459 | 6.36 | 374 | 1 | 1 | 37.3 | 37.3 | 37.3 | 1 | 374 | 374 | 374
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank1 | 30 | 1 | 1 | TP | ug P/L | 0.0821 | 1.39 | 30 | TN | ug N/L | 0.329 | 4.65 | 254 | 1 | 1 | 30 | 30 | 30 | 1 | 254 | 254 | 254
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank2 | 31 | 1 | 1 | TP | ug P/L | 0.0841 | 1.48 | 32.4 | TN | ug N/L | 0.335 | 4.69 | 256 | 1 | 1 | 32.4 | 32.4 | 32.4 | 1 | 256 | 256 | 256
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank3 | 32 | 1 | 1 | TP | ug P/L | 0.0827 | 1.49 | 32.7 | TN | ug N/L | 0.328 | 4.6 | 250 | 1 | 1 | 32.7 | 32.7 | 32.7 | 1 | 250 | 250 | 250
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank4 | 33 | 1 | 1 | TP | ug P/L | 0.0846 | 1.45 | 31.6 | TN | ug N/L | 0.328 | 4.6 | 250 | 1 | 1 | 31.6 | 31.6 | 31.6 | 1 | 250 | 250 | 250
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank 5 | 34 | 1 | 1 | TP | ug P/L | 0.0825 | 1.42 | 30.8 | TN | ug N/L | 0.331 | 4.65 | 254 | 1 | 1 | 30.8 | 30.8 | 30.8 | 1 | 254 | 254 | 254
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spiked blank BRN | 155 | 1 | 1 | TP | ug P/L | 0.0775 | 1.35 | 28.9 | TN | ug N/L | 0.267 | 3.76 | 191 | 1 | 1 | 28.9 | 28.9 | 28.9 | 1 | 191 | 191 | 191
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Spike blank BRN | 156 | 1 | 1 | TP | ug P/L | 0.0817 | 1.42 | 30.8 | TN | ug N/L | 0.278 | 3.9 | 201 | 1 | 1 | 30.8 | 30.8 | 30.8 | 1 | 201 | 201 | 201
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spiked blank BRN | 157 | 1 | 1 | TP | ug P/L | 0.0773 | 1.33 | 28.2 | TN | ug N/L | 0.27 | 3.83 | 195 | 1 | 1 | 28.2 | 28.2 | 28.2 | 1 | 195 | 195 | 195
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | inorganic 40 400 | 158 | 1 | 1 | TP | ug P/L | 0.0999 | 1.75 | 39.9 | TN | ug N/L | 0.421 | 5.87 | 340 | 1 | 1 | 39.9 | 39.9 | 39.9 | 1 | 340 | 340 | 340
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | inorganic 40 400 | 159 | 1 | 1 | TP | ug P/L | 0.101 | 1.74 | 39.8 | TN | ug N/L | 0.436 | 6.06 | 353 | 1 | 1 | 39.8 | 39.8 | 39.8 | 1 | 353 | 353 | 353
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | inorganic 40 400 | 160 | 1 | 1 | TP | ug P/L | 0.101 | 1.76 | 40.4 | TN | ug N/L | 0.428 | 5.95 | 345 | 1 | 1 | 40.4 | 40.4 | 40.4 | 1 | 345 | 345 | 345
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | urea 40 400 | 161 | 1 | 1 | TP | ug P/L | 0.0993 | 1.74 | 39.7 | TN | ug N/L | 0.228 | 3.22 | 152 | 1 | 1 | 39.7 | 39.7 | 39.7 | 1 | 152 | 152 | 152
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | urea 40 400 | 162 | 1 | 1 | TP | ug P/L | 0.1 | 1.75 | 40 | TN | ug N/L | 0.214 | 3 | 137 | 1 | 1 | 40 | 40 | 40 | 1 | 137 | 137 | 137
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | urea 40 400 | 163 | 1 | 1 | TP | ug P/L | 0.0996 | 1.72 | 39.3 | TN | ug N/L | 0.223 | 3.13 | 146 | 1 | 1 | 39.3 | 39.3 | 39.3 | 1 | 146 | 146 | 146
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | lgulatmic acid 40 400 | 164 | 1 | 1 | TP | ug P/L | 0.0962 | 1.69 | 38.3 | TN | ug N/L | 0.414 | 5.74 | 330 | 1 | 1 | 38.3 | 38.3 | 38.3 | 1 | 330 | 330 | 330
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | lgulatmic acid 40 400 | 165 | 1 | 1 | TP | ug P/L | 0.1 | 1.76 | 40.3 | TN | ug N/L | 0.425 | 5.91 | 342 | 1 | 1 | 40.3 | 40.3 | 40.3 | 1 | 342 | 342 | 342
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | lgulatmic acid 40 400 | 166 | 1 | 1 | TP | ug P/L | 0.0994 | 1.76 | 40.4 | TN | ug N/L | 0.441 | 6.1 | 356 | 1 | 1 | 40.4 | 40.4 | 40.4 | 1 | 356 | 356 | 356

## r data
runday | obs | TP | TN | runday | TP |  | runday | TP | TN
1 | 1 | 24 | 250 | 1 | 24 | 2020-10-15 | 1 | 24 | 250
1 | 2 | 23.8 | 251.6436 | 1 | 23.8 | 2020-10-15 | 1 | 23.8 | 251.6436
1 | 3 | 22.9 | 227.2878 | 1 | 22.9 | 2020-10-15 | 1 | 22.9 | 227.2878
1 | 4 | 23.1 | 205.5903 | 1 | 23.1 | 2020-10-15 | 1 | 23.1 | 205.5903
1 | 5 | 24.1 | 254.9893 | 1 | 24.1 | 2020-10-15 | 1 | 24.1 | 254.9893
2 | 6 | 22.4 | 230 | 2 | 22.4 | 2020-10-29 | 2 | 22.4 | 230
2 | 7 | 24.1 | 262 | 2 | 24.1 | 2020-10-29 | 2 | 24.1 | 262
2 | 8 | 19.2 | 249 | 2 | 19.2 | 2020-10-29 | 2 | 19.2 | 249
2 | 9 | 24.1 | 212 | 2 | 24.1 | 2020-10-29 | 2 | 24.1 | 212
2 | 10 | 27.7 | 246 | 2 | 27.7 | 2020-10-29 | 2 | 27.7 | 246
3 | 11 | 25.4968 | 181.0821 | 3 | 25.4968 | 2020-12-01 | 3 | 25.4968 | 181.0821
3 | 12 | 26.6927 | 239.415 | 3 | 26.6927 | 2020-12-01 | 3 | 26.6927 | 239.415
3 | 13 | 26.9919 | 281.1202 | 3 | 26.9919 | 2020-12-01 | 3 | 26.9919 | 281.1202
3 | 14 | 26.9919 | 267.8227 | 3 | 26.9919 | 2020-12-01 | 3 | 26.9919 | 267.8227
3 | 15 | 25.4968 | 223.37 | 3 | 25.4968 | 2020-12-01 | 3 | 25.4968 | 223.37
4 | 16 | 23.5 | 237 | 4 | 23.5 | 2021-01-14 | 4 | 23.5 | 237
4 | 17 | 25.1 | 241 | 4 | 25.1 | 2021-01-14 | 4 | 25.1 | 241
4 | 18 | 32.7 | 240 | 4 | 32.7 | 2021-01-14 | 4 | 32.7 | 240
4 | 19 | 26.2 | 306 | 4 | 26.2 | 2021-01-14 | 4 | 26.2 | 306
4 | 20 | 23.8 | 272 | 4 | 23.8 | 2021-01-14 | 4 | 23.8 | 272
5 | 21 | 23.3 | 178 | 5 | 23.3 | 2021-06-10 | 5 | 23.3 | 178
5 | 22 | 24.5 | 183 | 5 | 24.5 | 2021-06-10 | 5 | 24.5 | 183
5 | 23 | 23.9 | 172 | 5 | 23.9 | 2021-06-10 | 5 | 23.9 | 172
5 | 24 | 25 | 185 | 5 | 25 | 2021-06-10 | 5 | 25 | 185
5 | 25 | 23.6 | 195 | 5 | 23.6 | 2021-06-10 | 5 | 23.6 | 195
6 | 26 | 24 | 210 | 6 | 24 | 2021-06-10 | 6 | 24 | 210
6 | 27 | 25.4 | 224 | 6 | 25.4 | 2021-06-10 | 6 | 25.4 | 224
6 | 28 | 25 | 218 | 6 | 25 | 2021-06-10 | 6 | 25 | 218
6 | 29 | 26.4 | 223 | 6 | 26.4 | 2021-06-10 | 6 | 26.4 | 223
6 | 30 | 25.1 | 216 | 6 | 25.1 | 2021-06-10 | 6 | 25.1 | 216
7 | 31 | 23.7 | 174 | 7 | 23.7 | 2022-02-23 | 7 | 23.7 | 174
7 | 32 | 22.6 | 230 | 7 | 22.6 | 2022-02-23 | 7 | 22.6 | 230
7 | 33 | 22.5 | 230 | 7 | 22.5 | 2022-02-23 | 7 | 22.5 | 230
7 | 34 | 25 | 229 | 7 | 25 | 2022-02-23 | 7 | 25 | 229
7 | 35 | 25.8 | 257 | 7 | 25.8 | 2022-02-23 | 7 | 25.8 | 257
8 | 36 | 21.4 | 155 | 8 | 21.4 | 2022-03-23 | 8 | 21.4 | 155
8 | 37 | 17.4 | 148 | 8 | 17.4 | 2022-03-23 | 8 | 17.4 | 148
8 | 38 | 17.2 | 210 | 8 | 17.2 | 2022-03-23 | 8 | 17.2 | 210
8 | 39 | 18 | 216 | 8 | 18 | 2022-03-23 | 8 | 18 | 216
8 | 40 | 18.4 | 194 | 8 | 18.4 | 2022-03-23 | 8 | 18.4 | 194
8 | 41 | 30 | 254 | 8 | 30 | 2022-04-13 | 9 | 30 | 254
8 | 42 | 32.4 | 256 | 8 | 32.4 | 2022-04-13 | 9 | 32.4 | 256
8 | 43 | 32.7 | 250 | 8 | 32.7 | 2022-04-13 | 9 | 32.7 | 250
9 | 44 | 31.6 | 250 | 9 | 31.6 | 2022-04-13 | 9 | 31.6 | 250
9 | 45 | 30.8 | 254 | 9 | 30.8 | 2022-04-13 | 9 | 30.8 | 254
9 | 46 | 28.9 | 191 | 9 | 28.9 | 2022-04-13 | 9 | 28.9 | 191
9 | 47 | 30.8 | 201 | 9 | 30.8 | 2022-04-13 | 9 | 30.8 | 201
9 | 48 | 28.2 | 195 | 9 | 28.2 | 2022-04-13 | 9 | 28.2 | 195
[6 empty] | 2022-05-04 | 10 | 26.5 | 245
[6 empty] | 2022-05-04 | 10 | 26.2 | 234
[6 empty] | 2022-05-04 | 10 | 27.7 | 251
[6 empty] | 2022-05-04 | 10 | 26.4 | 237
[6 empty] | 2022-05-04 | 10 | 25.4 | 253

## old style MDL calc
2020 Field Season Batch 4
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Observation # | PR P | PR N
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 30 | 1 | 1 | TP | ug P/L | 0.0448 | 1.04 | 23.5 | TN | ug N/L | 0.314 | 4.62 | 237 |  | 1 | 1 |  | 23.5 | 23.5 | 23.5 | [6 empty] | 1 |  | 237 | 237 | 237 | [6 empty] | 16 | #VALUE! | #DIV/0!
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 31 | 1 | 1 | TP | ug P/L | 0.0468 | 1.09 | 25.1 | TN | ug N/L | 0.32 | 4.67 | 241 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | [6 empty] | 1 |  | 241 | 241 | 241 | [6 empty] | 17 | #VALUE! | #DIV/0!
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 32 | 1 | 1 | TP | ug P/L | 0.0543 | 1.33 | 32.7 | TN | ug N/L | 0.32 | 4.67 | 240 |  | 1 | 1 |  | 32.7 | 32.7 | 32.7 | [6 empty] | 1 |  | 240 | 240 | 240 | [6 empty] | 18 | #VALUE! | #DIV/0!
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 33 | 1 | 1 | TP | ug P/L | 0.0498 | 1.12 | 26.2 | TN | ug N/L | 0.388 | 5.61 | 306 |  | 1 | 1 |  | 26.2 | 26.2 | 26.2 | [6 empty] | 1 |  | 306 | 306 | 306 | [6 empty] | 19 | #VALUE! | #DIV/0!
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 34 | 1 | 1 | TP | ug P/L | 0.0446 | 1.05 | 23.8 | TN | ug N/L | 0.35 | 5.12 | 272 |  | 1 | 1 |  | 23.8 | 23.8 | 23.8 | [6 empty] | 1 |  | 272 | 272 | 272 | [6 empty] | 20 | #VALUE! | #DIV/0!
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 179 | 1 | 1 | TP | ug P/L | 0.0558 | 1.23 | 29.6 | TN | ug N/L | 0.382 | 5.54 | 301 |  | 1 | 1 |  | 29.6 | 29.6 | 29.6 | [6 empty] | 1 |  | 301 | 301 | 301
2021-01-14 | OM_1-14-2021_12-13-03PM.OMN | spiked blank | 180 | 1 | 1 | TP | ug P/L | 0.0538 | 1.14 | 26.9 | TN | ug N/L | 0.298 | 4.35 | 218 |  | 1 | 1 |  | 26.9 | 26.9 | 26.9 | [6 empty] | 1 |  | 218 | 218 | 218
[21 empty] | Known Conc | 24.3902 | [9 empty] | Known Conc | 243.9024
[21 empty] | Mean | 26.8286 | [9 empty] | Mean | 259.2857
[21 empty] | Std Dev | 3.3115 | [9 empty] | Std Dev | 34.1356
[21 empty] | CV or %RSD | 12.343 | [9 empty] | CV or %RSD | 13.1652
[21 empty] | T value | 3.1427 | [9 empty] | T value | 3.1427
[21 empty] | MDL | 10.4068 | [9 empty] | MDL | 107.2768
[21 empty] | LOQ | 33.1145 | [9 empty] | LOQ | 341.3558
[21 empty] | Error as % | 9.9971 | [9 empty] | Error as % | 6.3071
[21 empty] | Known/MDL | 2.3437 | [9 empty] | Known/MDL | 2.2736
[21 empty] | Mean PR | 109.9971 | [9 empty] | Mean PR | 106.3071
[21 empty] | S/N | 8.1018 | [9 empty] | S/N | 7.5958
1) Is the concentration used to calculate the MDL between 1 and 5 x the resulting MDL?
2) Is the signal to noise ratio reasonable?
A CV (%) or %RSD between 10 and 40% (<40% more important), but APHA suggests <=20%
A S/N between 2.5 and 10 (>2.5 more important)
If there is too little noise, it might indicate that the concentration used for the determination is too high. Try again with a lower concentration.
3) Are the recoveries (Mean PR) reasonable?
APHA says between 50 and 150%
2021 Field Season Batch 1
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Observation # | PR P | PR N
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 30 | 1 | 1 | TP | ug P/L | 0.0592 | 1.17 | 23.3 | TN | ug N/L | 0.252 | 3.83 | 178 |  | 1 | 1 |  | 23.3 | 23.3 | 23.3 | [6 empty] | 1 |  | 178 | 178 | 178 | [6 empty] | 21 | #VALUE! | #DIV/0!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 31 | 1 | 1 | TP | ug P/L | 0.0608 | 1.21 | 24.5 | TN | ug N/L | 0.26 | 3.9 | 183 |  | 1 | 1 |  | 24.5 | 24.5 | 24.5 | [6 empty] | 1 |  | 183 | 183 | 183 | [6 empty] | 22 | #VALUE! | #DIV/0!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 32 | 1 | 1 | TP | ug P/L | 0.0593 | 1.19 | 23.9 | TN | ug N/L | 0.247 | 3.72 | 172 |  | 1 | 1 |  | 23.9 | 23.9 | 23.9 | [6 empty] | 1 |  | 172 | 172 | 172 | [6 empty] | 23 | #VALUE! | #DIV/0!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 33 | 1 | 1 | TP | ug P/L | 0.0597 | 1.23 | 25 | TN | ug N/L | 0.259 | 3.93 | 185 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 185 | 185 | 185 | [6 empty] | 24 | #VALUE! | #DIV/0!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 34 | 1 | 1 | TP | ug P/L | 0.0599 | 1.18 | 23.6 | TN | ug N/L | 0.274 | 4.1 | 195 |  | 1 | 1 |  | 23.6 | 23.6 | 23.6 | [6 empty] | 1 |  | 195 | 195 | 195 | [6 empty] | 25 | #VALUE! | #DIV/0!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 156 | 1 | 1 | TP | ug P/L | 0.0609 | 1.2 | 24 | TN | ug N/L | 0.293 | 4.33 | 210 |  | 1 | 1 |  | 24 | 24 | 24 | [6 empty] | 1 |  | 210 | 210 | 210 | [6 empty] | 26 | #VALUE! | #DIV/0!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 157 | 1 | 1 | TP | ug P/L | 0.0625 | 1.25 | 25.4 | TN | ug N/L | 0.307 | 4.54 | 224 |  | 1 | 1 |  | 25.4 | 25.4 | 25.4 | [6 empty] | 1 |  | 224 | 224 | 224 | [6 empty] | 27 | #VALUE! | #DIV/0!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 158 | 1 | 1 | TP | ug P/L | 0.0616 | 1.23 | 25 | TN | ug N/L | 0.301 | 4.45 | 218 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 218 | 218 | 218 | [6 empty] | 28 | #VALUE! | #DIV/0!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 159 | 1 | 1 | TP | ug P/L | 0.0621 | 1.28 | 26.4 | TN | ug N/L | 0.307 | 4.53 | 223 |  | 1 | 1 |  | 26.4 | 26.4 | 26.4 | [6 empty] | 1 |  | 223 | 223 | 223 | [6 empty] | 29 | #VALUE! | #DIV/0!
2021-06-10 | OM_6-10-2021_10-17-10AM.OMN | spiked blank | 160 | 1 | 1 | TP | ug P/L | 0.0619 | 1.24 | 25.1 | TN | ug N/L | 0.299 | 4.42 | 216 |  | 1 | 1 |  | 25.1 | 25.1 | 25.1 | [6 empty] | 1 |  | 216 | 216 | 216 | [6 empty] | 30 | #VALUE! | #DIV/0!
[21 empty] | Known Conc | 24.3902 | [9 empty] | Known Conc | 243.9024
[21 empty] | Mean | 24.62 | [9 empty] | Mean | 200.4
[21 empty] | Std Dev | 0.94257 | [9 empty] | Std Dev | 19.9733
[21 empty] | CV or %RSD | 3.8285 | [9 empty] | CV or %RSD | 9.9667
[21 empty] | T value | 3.1427 | [9 empty] | T value | 3.1427
[21 empty] | MDL | 2.9622 | [9 empty] | MDL | 62.7695
[21 empty] | LOQ | 9.4257 | [9 empty] | LOQ | 199.7332
[21 empty] | Error as % | -1.9124 | [9 empty] | Error as % | -7.2222
[21 empty] | Known/MDL | 8.4734 | [9 empty] | Known/MDL | 3.4412
[21 empty] | Mean PR | 98.0876 | [9 empty] | Mean PR | 92.7778
[21 empty] | S/N | 26.12 | [9 empty] | S/N | 10.0334
1) Is the concentration used to calculate the MDL between 1 and 5 x the resulting MDL?
2) Is the signal to noise ratio reasonable?
A CV (%) or %RSD between 10 and 40% (<40% more important), but APHA suggests <=20%
A S/N between 2.5 and 10 (>2.5 more important)
If there is too little noise, it might indicate that the concentration used for the determination is too high. Try again with a lower concentration.
3) Are the recoveries (Mean PR) reasonable?
APHA says between 50 and 150%
2021 Field Season Batch 4
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- whatever we're using | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- whatever we're using | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | Observation # | PR P | PR N
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank1 | 30 | 1 | 1 | TP | ug P/L | 0.0821 | 1.39 | 30 | TN | ug N/L | 0.329 | 4.65 | 254 |  | 1 | 1 |  | 30 | 30 | 30 | [6 empty] | 1 |  | 254 | 254 | 254 | [6 empty] | 41 | #VALUE! | #DIV/0!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank2 | 31 | 1 | 1 | TP | ug P/L | 0.0841 | 1.48 | 32.4 | TN | ug N/L | 0.335 | 4.69 | 256 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 | [6 empty] | 1 |  | 256 | 256 | 256 | [6 empty] | 42 | #VALUE! | #DIV/0!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank3 | 32 | 1 | 1 | TP | ug P/L | 0.0827 | 1.49 | 32.7 | TN | ug N/L | 0.328 | 4.6 | 250 |  | 1 | 1 |  | 32.7 | 32.7 | 32.7 | [6 empty] | 1 |  | 250 | 250 | 250 | [6 empty] | 43 | #VALUE! | #DIV/0!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank4 | 33 | 1 | 1 | TP | ug P/L | 0.0846 | 1.45 | 31.6 | TN | ug N/L | 0.328 | 4.6 | 250 |  | 1 | 1 |  | 31.6 | 31.6 | 31.6 | [6 empty] | 1 |  | 250 | 250 | 250 | [6 empty] | 44 | #VALUE! | #DIV/0!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spike blank 5 | 34 | 1 | 1 | TP | ug P/L | 0.0825 | 1.42 | 30.8 | TN | ug N/L | 0.331 | 4.65 | 254 |  | 1 | 1 |  | 30.8 | 30.8 | 30.8 | [6 empty] | 1 |  | 254 | 254 | 254 | [6 empty] | 45 | #VALUE! | #DIV/0!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spiked blank BRN | 155 | 1 | 1 | TP | ug P/L | 0.0775 | 1.35 | 28.9 | TN | ug N/L | 0.267 | 3.76 | 191 |  | 1 | 1 |  | 28.9 | 28.9 | 28.9 | [6 empty] | 1 |  | 191 | 191 | 191 | [6 empty] | 46 | #VALUE! | #DIV/0!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | Spike blank BRN | 156 | 1 | 1 | TP | ug P/L | 0.0817 | 1.42 | 30.8 | TN | ug N/L | 0.278 | 3.9 | 201 |  | 1 | 1 |  | 30.8 | 30.8 | 30.8 | [6 empty] | 1 |  | 201 | 201 | 201 | [6 empty] | 47 | #VALUE! | #DIV/0!
2022-04-13 | OM_4-13-2022_11-02-03AM.OMN | spiked blank BRN | 157 | 1 | 1 | TP | ug P/L | 0.0773 | 1.33 | 28.2 | TN | ug N/L | 0.27 | 3.83 | 195 |  | 1 | 1 |  | 28.2 | 28.2 | 28.2 | [6 empty] | 1 |  | 195 | 195 | 195 | [6 empty] | 48 | #VALUE! | #DIV/0!
[21 empty] | Known Conc | 24.3902 | [9 empty] | Known Conc | 243.9024
[21 empty] | Mean | 30.675 | [9 empty] | Mean | 231.375
[21 empty] | Std Dev | 1.5899 | [9 empty] | Std Dev | 29.7607
[21 empty] | CV or %RSD | 5.1831 | [9 empty] | CV or %RSD | 12.8625
[21 empty] | T value | 3.1427 | [9 empty] | T value | 3.1427
[21 empty] | MDL | 4.9966 | [9 empty] | MDL | 93.5279
[21 empty] | LOQ | 15.8992 | [9 empty] | LOQ | 297.6065
[21 empty] | Error as % | 25.7675 | [9 empty] | Error as % | -5.1363
[21 empty] | Known/MDL | 4.8814 | [9 empty] | Known/MDL | 2.6078
[21 empty] | Mean PR | 125.7675 | [9 empty] | Mean PR | 94.8637
[21 empty] | S/N | 19.2934 | [9 empty] | S/N | 7.7745
1) Is the concentration used to calculate the MDL between 1 and 5 x the resulting MDL?
2) Is the signal to noise ratio reasonable?
A CV (%) or %RSD between 10 and 40% (<40% more important), but APHA suggests <=20%
A S/N between 2.5 and 10 (>2.5 more important)
If there is too little noise, it might indicate that the concentration used for the determination is too high. Try again with a lower concentration.
3) Are the recoveries (Mean PR) reasonable?
APHA says between 50 and 150%