## over days
First, calibration curves
More variance in TNb over day than for TIC or TC
That is consistent with statements in manual
Second, samples run independently on different days.
For all forms of C, max RSD is 10%
For TNb, 4 higher than 10% but all less than 20%
2022-08-15
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.5 | 0.19539 | 0.27974 | 0.08435 | 0.020139
TIC | TC | TNb |  | 0.0006 | 1193.3 | 0.0012 | 2372 | 6e-05 | 1117.4 | 0.2 | 3.2166 | 6.2254 | 3.0089 | 0.3393
483 | 1165 | 471.5 |  | 0.0018 | 3583.9 | 0.0036 | 7269.5 | 0.00018 | 3091.7 | 0.6 | 2.894 | 5.8784 | 2.9844 | 0.28328
|  |  |  | 0.002997 | 6049.322 | 0.005994 | 12058.11 | 0.0002997 | 5570.981 | 0.333 | 8.5996 | 17.2921 | 8.6925 | 0.89548
|  |  |  | 0.004203 | 8761.378 | 0.008406 | 17595.89 | 0.0004203 | 7358.39 | 0.467 | 8.7874 | 17.8556 | 9.0682 | 0.83649
|  |  |  | 0.0054 | 11935.9 | 0.0108 | 23289 | 0.00054 | 10696.2 | 0.6 | 9.2587 | 18.3187 | 9.06 | 0.93878
|  |  | Slope |  | 4.5724e-07 |  | 4.6594e-07 |  | 5.1719e-08
|  |  | Intercept |  | 9.7694e-05 |  | 0.00013987 |  | 1.0069e-05
|  |  | RSQ |  | 0.99693 |  | 0.99859 |  | 0.9923
2022-08-16
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.5 | 0.14512 | 0.30984 | 0.16472 | 0.015056
TIC | TC | TNb |  | 0.0006 | 1143.7 | 0.0012 | 2273.7 | 6e-05 | 1165.5 | 0.2 | 2.9569 | 5.9723 | 3.0155 | 0.30234
509.5 | 1129.5 | 495 |  | 0.0018 | 3681.6 | 0.0036 | 7399.1 | 0.00018 | 3589 | 0.6 | 2.9044 | 5.8964 | 2.992 | 0.28425
|  |  |  | 0.002997 | 6254.173 | 0.005994 | 12273.253 | 0.0002997 | 6253.83 | 0.333 | 8.7376 | 17.3162 | 8.5786 | 0.87567
|  |  |  | 0.004203 | 9267.127 | 0.008406 | 18064.547 | 0.0004203 | 9347.547 | 0.467 | 9.1571 | 18.0173 | 8.8602 | 0.92532
|  |  |  | 0.0054 | 11760.6 | 0.0108 | 23572.6 | 0.00054 | 11672.5 | 0.6 | 9.0124 | 18.2207 | 9.2082 | 0.89622
|  |  | Slope |  | 4.5362e-07 |  | 4.572e-07 |  | 4.5423e-08
|  |  | Intercept |  | 7.2562e-05 |  | 0.00015492 |  | 7.5278e-06
|  |  | RSQ |  | 0.99901 |  | 0.99889 |  | 0.99837
2022-08-18
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.5 | 0.15098 | 0.31167 | 0.16069 | 0.044059
TIC | TC | TNb |  | 0.0006 | 1118.6 | 0.0012 | 2244.1 | 6e-05 | 979.2 | 0.2 | 3.014 | 5.9409 | 2.9269 | 0.32188
698.5 | 1146 | 552 |  | 0.0018 | 3461.3 | 0.0036 | 7269.8 | 0.00018 | 3352.6 | 0.6 | 2.8452 | 5.8335 | 2.9883 | 0.27836
|  |  |  | 0.002997 | 6087.299 | 0.005994 | 12300.764 | 0.0002997 | 6146.368 | 0.333 | 8.8439 | 17.4609 | 8.617 | 0.86436
|  |  |  | 0.004203 | 8858.601 | 0.008406 | 18016.136 | 0.0004203 | 8481.636 | 0.467 | 9.1036 | 18.0807 | 8.9771 | 0.83259
|  |  |  | 0.0054 | 11333.8 | 0.0108 | 23350.8 | 0.00054 | 12669.6 | 0.6 | 9.0303 | 18.1629 | 9.1326 | 0.94988
|  |  | Slope |  | 4.714e-07 |  | 4.6002e-07 |  | 4.3245e-08
|  |  | Intercept |  | 7.5489e-05 |  | 0.00015583 |  | 2.2029e-05
|  |  | RSQ |  | 0.9991 |  | 0.99911 |  | 0.98767
2022-09-06
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.5 | 0.0072705 | -0.051863 | -0.059133 | 0.0073546
TIC | TC | TNb |  | 0.0006 | 1302 | 0.0012 | 2392 | 6e-05 | 973.7 | 0.2 | 3.2521 | 6.0721 | 2.82 | 0.25938
332.5 | 545 | 314.5 |  | 0.0018 | 3539.5 | 0.0036 | 7369 | 0.00018 | 3549.1 | 0.6 | 2.9365 | 6.3253 | 3.3888 | 0.29894
|  |  |  | 0.002997 | 5885.055 | 0.005994 | 11257.03 | 0.0002997 | 6340.043 | 0.333 | 8.79 | 17.4513 | 8.6613 | 0.95351
|  |  |  | 0.004203 | 8545.445 | 0.008406 | 16031.47 | 0.0004203 | 8151.47 | 0.467 | 9.0977 | 17.7452 | 8.6475 | 0.87192
|  |  |  | 0.0054 | 10880 | 0.0108 | 21105.5 | 0.00054 | 10842.1 | 0.6 | 9.0139 | 18.1968 | 9.1829 | 0.90063
|  |  | Slope |  | 4.9676e-07 |  | 5.1854e-07 |  | 4.9501e-08
|  |  | Intercept |  | 3.6352e-06 |  | -2.5931e-05 |  | 3.6773e-06
|  |  | RSQ |  | 0.9995 |  | 0.99885 |  | 0.9974
Accidentally ran many of the same samples twice... so... check precision
Name | TIC Mean of 2 reps mg/L | TC Mean of 2 reps mg/L | TOC Mean of 2 reps mg/L | TNb Mean of 2 reps mg/L
MDL | 0.79876 | 1.0554 | 0.70274 | 0.087475
LOQ | 3.2304 | 4.2682 | 2.842 | 0.35377 |  | RSD | RSD | RSD | RSD
B50 08aug22 0.1 | 3 | 7.16 | 4.16 | 0.31
B50 08aug22 0.1 | 3.09 | 7.21 | 4.12 | 0.32 |  | 2.9557 | 0.69589 | -0.96618 | 3.1746
B50 08aug22 3.0 | 3.49 | 7.68 | 4.19 | 0.36
B50 08aug22 3.0 | 3.64 | 8.1 | 4.46 | 0.4 |  | 4.2076 | 5.3232 | 6.2428 | 10.5263
B50 08aug22 6.0 | 5.08 | 8.18 | 3.09 | 1.16
B50 08aug22 6.0 | 5.27 | 8.49 | 3.22 | 1.3 |  | 3.6715 | 3.7193 | 4.1204 | 11.3821
B50 18jul22 0.1 | 3.12 | 8.69 | 5.57 | 0.7
B50 18jul22 0.1 | 3.35 | 8.79 | 5.44 | 0.76 |  | 7.1097 | 1.1442 | -2.3615 | 8.2192
B50 18jul22 3.0 | 3.49 | 7.35 | 3.86 | 0.41
B50 18jul22 3.0 | 3.86 | 7.76 | 3.9 | 0.43 |  | 10.068 | 5.4269 | 1.0309 | 4.7619
B50 25jul22 0.1 | 3.3 | 8.54 | 5.24 | 0.58
B50 25jul22 0.1 | 3.29 | 8.6 | 5.31 | 0.61 |  | -0.30349 | 0.70012 | 1.327 | 5.042
B50 25jul22 3.0 | 3.76 | 7.65 | 3.89 | 0.35
B50 25jul222 3.0 | 3.93 | 7.82 | 3.89 | 0.4 |  | 4.4213 | 2.1978 | 0 | 13.3333
C 14jul22 0.1 | 8.75 | 11.93 | 3.18 | 0.19
C 14jul22 0.1 | 8.58 | 11.77 | 3.19 | 0.2 |  | -1.9619 | -1.3502 | 0.31397 | 5.1282
C 14JUL22 9 | 9.74 | 12.84 | 3.1 | 0.23
C 14jul22 9.0 | 9.78 | 12.51 | 2.73 | 0.19
C50 28jul22 15.0 | 10.39 | 13.27 | 2.88 | 0.26
C50 28jul22 15.0 | 10.6 | 13.27 | 2.68 | 0.27 |  | 2.001 | 0 | -7.1942 | 3.7736
C50 28jul22 20.0 | 12.72 | 16.25 | 3.52 | 0.53
C50 28jul22 20.0 | 13.03 | 16.6 | 3.57 | 0.58 |  | 2.4078 | 2.1309 | 1.4104 | 9.009
C50 28jul22 6.0 | 8.66 | 11.74 | 3.08 | 0.18
C50 28jul22 6.0 | 8.57 | 11.67 | 3.1 | 0.21 |  | -1.0447 | -0.59804 | 0.64725 | 15.3846
F 01aug22 8.0 | 6.47 | 9.05 | 2.58 | 0.81
F 01aug22 8.0 | 6.52 | 9.16 | 2.65 | 0.89 |  | 0.76982 | 1.2081 | 2.6769 | 9.4118
F 11jul22 3.8 | 4.76 | 8.34 | 3.58 | 0.39
F 11jul22 3.8 | 4.65 | 8.54 | 3.89 | 0.43 |  | -2.3379 | 2.3697 | 8.2999 | 9.7561
F50 08aug22 3.0 | 3.75 | 10.84 | 7.09 | 0.62
F50 18jul22 0.1 | 3.74 | 10.79 | 7.05 | 0.65
F50 18jul22 5.0 | 5.56 | 8.49 | 2.93 | 0.43
F50 18jul22 5.0 | 5.71 | 8.75 | 3.04 | 0.5 |  | 2.6619 | 3.0162 | 3.6851 | 15.0538
F50 18jul22 6.2 | 6.35 | 8.8 | 2.45 | 0.58
F50 18jul22 6.2 | 6.22 | 8.8 | 2.58 | 0.64 |  | -2.0684 | 0 | 5.169 | 9.8361
F50 18jul22 9.0 | 5.9 | 8.26 | 2.36 | 0.53
F50 18jul22 9.0 | 5.99 | 8.2 | 2.2 | 0.59 |  | 1.5139 | -0.72904 | -7.0175 | 10.7143
F50 25jul22 0.1 | 4.2 | 10.39 | 6.19 | 0.47
F50 25jul22 0.1 | 4.04 | 10.11 | 6.08 | 0.51 |  | -3.8835 | -2.7317 | -1.793 | 8.1633
F50 25jul22 3.8 | 5.31 | 8.75 | 3.43 | 0.42
F50 25jul22 3.8 | 5.15 | 8.75 | 3.59 | 0.45 |  | -3.0593 | 0 | 4.5584 | 6.8966
F50 25jul22 8.0 | 6.48 | 8.67 | 2.19 | 0.64
F50 25jul22 8.0 | 6.31 | 8.61 | 2.3 | 0.77 |  | -2.6583 | -0.69444 | 4.8998 | 18.4397
F50 25jul22 9.0 | 6.1 | 10.85 | 4.76 | 1.26
F50 25jul22 9.0 | 6.11 | 10.93 | 4.81 | 1.28 |  | 0.1638 | 0.73462 | 1.0449 | 1.5748
SMB 28jul22 0.1 | 2.68 | 4.9 | 2.22 | 0.19
SMB 28jul22 0.1 | 2.72 | 4.87 | 2.15 | 0.2 |  | 1.4815 | -0.61412 | -3.2037 | 5.1282

## over run
2nd day starts at Run Order 121
only using 2nd and 3rd reps
3rd day starts at Run Order 251
[6 empty] | Mixed Check | [11 empty] | Water Blank | [10 empty] | Spiked Tap
[6 empty] | TIC signal is increasing over run time | [11 empty] | Messy (but signal is low) | [10 empty] | stable on 15aug... maybe slight drop on 16aug?
[6 empty] | TOC signal is stable | [11 empty] | Trending down over day | [10 empty] | signal is high but fairly similar to a typical sample
THIS IS WHAT WE HAVE SEEN REPEATEDLY
[6 empty] | FOR UNCOVERED SAMPLES | [22 empty] | rolling chart...
DON'T PANIC IF MIXED CHECK IS >20%
DO NOT USE RUN IN
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] |  | TIC Absolute value Percent error for check standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC Absolute value Percent error for check standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes |  | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4207 | 8679 |  | 3707 | 3.036 | 6.359 | 3.324 |  | 0.226 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 13:25 |  | 1 |  | 3.3688 | 6.973 | 3.6042 | 0.33632 | 4472 | 12.2938 |  |  |  | 16.2162 |  |  |  | 20.1387 |  |  |  |  | 12.1067 |  |  |  | 3.3357 | 6.964 | 3.6284 | 0.33205
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4120 | 8656 |  | 3608 | 2.979 | 6.343 | 3.364 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 13:36 |  | 1 |  | 3.3025 | 6.9551 | 3.6526 | 0.32779 | 4536 | 10.0838 |  |  |  | 15.9186 |  |  |  | 21.7533 |  |  |  |  | 9.2621
68 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 6115 | 10445 |  | 3775 | 4.255 | 7.606 | 3.351 |  | 0.232 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 2:09 |  | 1 |  | 4.8228 | 8.3444 | 3.5216 | 0.34218 | 4330 | 60.7609 | 1.4744 |  |  | 39.0733 | 0.07448 |  |  | 17.3858 | 2.2353 |  |  |  | 14.0605 | 0.30183 |  |  | 4.8586 | 8.3413 | 3.4826 | 0.3427
69 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 6209 | 10437 |  | 3787 | 4.315 | 7.601 | 3.285 |  | 0.233 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 2:21 |  | 1 |  | 4.8945 | 8.3382 | 3.4437 | 0.34322 | 4228
113 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 6843 | 11312 |  | 4203 | 4.721 | 8.218 | 3.497 |  | 0.27 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 12:15 |  | 1 |  | 5.3776 | 9.0177 | 3.6401 | 0.37907 | 4469 | 79.2536 | 0.3549 |  |  | 50.2948 | 0.97693 |  |  | 21.336 | 2.9124 |  |  |  | 26.3582 | 0.59297 |  |  | 5.3681 | 9.062 | 3.6939 | 0.37795
114 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 6818 | 11426 |  | 4177 | 4.705 | 8.299 | 3.594 |  | 0.267 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 12:26 |  | 1 |  | 5.3586 | 9.1062 | 3.7477 | 0.37683 | 4608
131 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4341 | 8814 |  | 4222 | 3.121 | 6.454 | 3.334 |  | 0.271 | 1 | 0 | 0 |  | 0 |  |  | 44789 | 0.68429 |  | 1 |  | 3.4029 | 6.9745 | 3.5716 | 0.33217 | 4473 | 13.4304 |  |  |  | 16.2421 |  |  |  | 19.0537 |  |  |  |  | 10.7249 |  |  |  | 3.3666 | 6.9292 | 3.5626 | 0.32922 |  | 251
132 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4245 | 8695 |  | 4144 | 3.06 | 6.37 | 3.311 |  | 0.264 | 1 | 0 | 0 |  | 0 |  |  | 44789 | 0.69212 |  | 1 |  | 3.3303 | 6.8838 | 3.5535 | 0.32627 | 4450 | 11.0111 |  |  |  | 14.7308 |  |  |  | 18.4504 |  |  |  |  | 8.7565 | [8 empty] | 252
188 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 5871 | 10103 |  | 4138 | 4.099 | 7.365 | 3.265 |  | 0.264 | 1 | 0 | 0 |  | 0 |  |  | 44790 | 0.22205 |  | 1 |  | 4.5597 | 7.9567 | 3.3971 | 0.32582 | 4232 | 51.9886 | 1.3353 |  |  | 32.6125 | 0.086155 |  |  | 13.2364 | 1.9629 |  |  |  | 8.6051 | 0.11611 |  |  | 4.5294 | 7.9602 | 3.4308 | 0.326 |  | 308
189 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 5791 | 10112 |  | 4143 | 4.048 | 7.371 | 3.323 |  | 0.264 | 1 | 0 | 0 |  | 0 |  |  | 44790 | 0.22992 |  | 1 |  | 4.4992 | 7.9636 | 3.4644 | 0.32619 | 4321 | [22 empty] | 309
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5914 | 12413 |  | 6633 | 8.253 | 17.991 | 9.737 |  | 0.963 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 11:32 |  | 1 |  | 9.3393 | 19.7454 | 10.4061 | 1.1771 | [18 empty] | 9.5885 | 19.7462 | 10.1577 | 1.1534
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6241 | 12414 |  | 6358 | 8.672 | 17.992 | 9.321 |  | 0.915 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 11:42 |  | 1 |  | 9.8377 | 19.7469 | 9.9092 | 1.1297
110 | 4 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6031 | 12509 |  | 6826 | 8.403 | 18.126 | 9.724 |  | 0.997 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 11:34 |  | 1 |  | 9.5176 | 19.8945 | 10.3769 | 1.2103 |  |  | 3.6165 |  |  |  | 0.41291 |  |  |  | 2.6187 |  |  |  |  | 0.65307 |  |  | 9.6929 | 19.9356 | 10.2427 | 1.2143
111 | 4 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6261 | 12562 |  | 6872 | 8.697 | 18.202 | 9.505 |  | 1.004 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 11:45 |  | 1 |  | 9.8682 | 19.9768 | 10.1086 | 1.2183
122 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7263 | 12756 |  | 7223 | 9.979 | 18.476 | 8.497 |  | 1.066 | 1 | 0 | 0 |  | 0 |  |  | 44789 | 0.60477 |  | 1 |  | 11.2241 | 19.9567 | 8.7326 | 1.1187 | [18 empty] | 11.1546 | 19.9796 | 8.825 | 1.1286 |  | 242
123 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7171 | 12786 |  | 7353 | 9.861 | 18.517 | 8.657 |  | 1.088 | 1 | 0 | 0 |  | 0 |  |  | 44789 | 0.61236 |  | 1 |  | 11.085 | 20.0024 | 8.9174 | 1.1384 | [23 empty] | 243
230 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5242 | 11250 |  | 7023 | 7.394 | 16.348 | 8.954 |  | 1.031 | 1 | 0 | 0 |  | 0 |  |  | 44790 | 0.61657 |  | 1 |  | 8.1682 | 17.6615 | 9.4933 | 1.0885 |  |  | 4.9292 |  |  |  | 0.53643 |  |  |  | 5.4903 |  |  |  |  | 0.27783 |  |  | 8.3746 | 17.6143 | 9.2397 | 1.09
231 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5515 | 11188 |  | 7043 | 7.743 | 16.262 | 8.519 |  | 1.034 | 1 | 0 | 0 |  | 0 |  |  | 44790 | 0.62417 |  | 1 |  | 8.581 | 17.5671 | 8.986 | 1.0915
245 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5405 | 11339 |  | 6944 | 7.602 | 16.475 | 8.872 |  | 1.017 | 1 | 0 | 0 |  | 0 |  |  | 44790 | 0.75938 |  | 1 |  | 8.4147 | 17.7972 | 9.3825 | 1.0765 |  |  | 1.9221 |  |  |  | 0.35901 |  |  |  | 1.0639 |  |  |  |  | 0.7008 |  |  | 8.4963 | 17.8292 | 9.3328 | 1.0803
246 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5513 | 11381 |  | 6994 | 7.741 | 16.533 | 8.793 |  | 1.026 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 18:24 |  | 1 |  | 8.578 | 17.8612 | 9.2832 | 1.0841
273 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5886 | 11611 |  | 6536 | 8.217 | 16.859 | 8.642 |  | 0.946 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 17:04 |  | 1 |  | 9.5004 | 18.3239 | 8.8235 | 1.0156 | [18 empty] | 9.4666 | 18.2848 | 8.8182 | 1.0163
274 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5843 | 11560 |  | 6545 | 8.162 | 16.786 | 8.624 |  | 0.948 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 17:15 |  | 1 |  | 9.4328 | 18.2457 | 8.8128 | 1.0169
318 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5150 | 10481 |  | 6215 | 7.276 | 15.263 | 7.987 |  | 0.89 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 3:04 |  | 1 |  | 8.3439 | 16.5911 | 8.2472 | 0.96933 | [18 empty] | 8.5812 | 16.6295 | 8.0483 | 0.97272
319 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5452 | 10531 |  | 6262 | 7.662 | 15.333 | 7.671 |  | 0.898 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 3:15 |  | 1 |  | 8.8185 | 16.6678 | 7.8493 | 0.97611
360 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5243 | 9802 |  | 5862 | 7.396 | 14.304 | 6.908 |  | 0.828 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 12:29 |  | 1 |  | 8.49 | 15.5499 | 7.0599 | 0.91844 | [18 empty] | 8.6236 | 15.553 | 6.9294 | 0.92854
361 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5413 | 9806 |  | 6002 | 7.612 | 14.31 | 6.698 |  | 0.853 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 12:39 |  | 1 |  | 8.7572 | 15.5561 | 6.7989 | 0.93863
22 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5337 | 11804 |  | 6516 | 7.516 | 17.131 | 9.615 |  | 0.942 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:35 |  | 1 |  | 8.8494 | 20.3163 | 11.4669 | 1.0874
23 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5366 | 12082 |  | 6436 | 7.552 | 17.524 | 9.972 |  | 0.929 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:46 |  | 1 |  | 8.8974 | 20.7968 | 11.8994 | 1.0742 |  | 100 |  | 0.92621 |  |  | 100 |  | 2.704 |  |  | 100 |  | 5.506 |  |  |  | 100 | 1.4544 |  |  |  | 8.9388 | 20.5194 | 11.5806 | 1.0665
24 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5416 | 11761 |  | 6342 | 7.617 | 17.07 | 9.453 |  | 0.912 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:56 |  | 1 |  | 8.9802 | 20.242 | 11.2618 | 1.0587
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2855 | 10731 |  | 6516 | 4.342 | 15.616 | 11.274 |  | 0.942 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:37 |  | 1 |  | 4.7396 | 18.4617 | 13.7221 | 1.0874
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4729 | 10802 |  | 6479 | 6.738 | 15.716 | 8.978 |  | 0.936 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:47 |  | 1 |  | 7.8426 | 18.5844 | 10.7417 | 1.0813 |  | 90.3821 |  | 5.8814 |  |  | 90.0726 |  | 1.1976 |  |  | 88.3581 |  | 6.701 |  |  |  | 101.4165 | 0.015258 |  |  |  | 8.0803 | 18.4738 | 10.3935 | 1.0814
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5016 | 10674 |  | 6480 | 7.106 | 15.536 | 8.43 |  | 0.936 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:58 |  | 1 |  | 8.3179 | 18.3632 | 10.0453 | 1.0815
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1013 | 1242 |  | 577 | 1.192 | 1.331 | 0.139 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 12:10 |  | 1 |  | 1.1217 | 1.4371 | 0.31539 | 0.079822 | [18 empty] | 0.98412 | 1.4795 | 0.49542 | 0.070358
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 712 | 1333 |  | 394 | 0.961 | 1.408 | 0.447 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 12:20 |  | 1 |  | 0.84649 | 1.5219 | 0.67545 | 0.060893
23 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 518 | 1113 |  | 491 | 0.812 | 1.221 | 0.409 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:09 |  | 1 |  | 0.66909 | 1.3169 | 0.64784 | 0.070927 |  |  | 10.0479 |  |  |  | 7.0981 |  |  |  | 22.0964 |  |  |  |  | 5.8542 |  |  | 0.63708 | 1.3654 | 0.72831 | 0.06891
24 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 448 | 1217 |  | 452 | 0.758 | 1.309 | 0.551 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:19 |  | 1 |  | 0.60507 | 1.4138 | 0.80877 | 0.066893
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 210 | 629 |  | 382 | 0.576 | 0.811 | 0.235 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 1:31 |  | 1 |  | 0.38743 | 0.86589 | 0.47847 | 0.059652 |  |  | 2.6306 |  |  |  | 1.0705 |  |  |  | 3.9697 |  |  |  |  | 5.1584 |  |  | 0.3824 | 0.87055 | 0.48816 | 0.058152
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 199 | 639 |  | 353 | 0.568 | 0.82 | 0.252 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 1:40 |  | 1 |  | 0.37737 | 0.87521 | 0.49784 | 0.056652
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 171 | 588 |  | 474 | 0.546 | 0.776 | 0.231 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 10:56 |  | 1 |  | 0.35176 | 0.82769 | 0.47592 | 0.069168 |  |  | 2.3127 |  |  |  | 4.6195 |  |  |  | 6.2903 |  |  |  |  | 4.2448 |  |  | 0.35588 | 0.84726 | 0.49138 | 0.070668
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 180 | 630 |  | 503 | 0.553 | 0.812 | 0.259 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 11:05 |  | 1 |  | 0.35999 | 0.86683 | 0.50683 | 0.072168
125 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 630 | 1078 |  | 515 | 0.898 | 1.192 | 0.294 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44789 | 0.6315 |  | 1 |  | 0.71669 | 1.2956 | 0.57888 | 0.061841 | [18 empty] | 0.67995 | 1.3235 | 0.64351 | 0.06275 |  | 245
126 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 549 | 1139 |  | 535 | 0.836 | 1.243 | 0.408 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44789 | 0.63847 |  | 1 |  | 0.6432 | 1.3514 | 0.70815 | 0.063658 | [23 empty] | 246
143 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 543 | 1140 |  | 502 | 0.831 | 1.244 | 0.413 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44789 | 0.80155 |  | 1 |  | 0.63776 | 1.3523 | 0.71451 | 0.06066 |  |  | 10.0081 |  |  |  | 1.4302 |  |  |  | 5.6553 |  |  |  |  | 2.1189 |  |  | 0.60737 | 1.3427 | 0.7353 | 0.060025 |  | 263
144 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 476 | 1119 |  | 488 | 0.78 | 1.227 | 0.446 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44789 | 0.8085 |  | 1 |  | 0.57698 | 1.3331 | 0.75609 | 0.059389 | [23 empty] | 264
185 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 408 | 552 |  | 508 | 0.728 | 0.746 | 0.018 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44790 | 0.19497 |  | 1 |  | 0.51528 | 0.81459 | 0.29931 | 0.061206 |  |  | 15.5609 |  |  |  | 1.6697 |  |  |  | 25.6607 |  |  |  |  | 0.59196 |  |  | 0.47809 | 0.82145 | 0.34337 | 0.061387 |  | 305
186 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 326 | 567 |  | 512 | 0.665 | 0.759 | 0.094 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44790 | 0.20159 |  | 1 |  | 0.44089 | 0.82831 | 0.38742 | 0.061569 | [23 empty] | 306
227 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 274 | 560 |  | 487 | 0.625 | 0.753 | 0.128 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44790 | 0.59019 |  | 1 |  | 0.39371 | 0.82191 | 0.4282 | 0.059298 |  |  | 5.2014 |  |  |  | 1.5455 |  |  |  | 7.369 |  |  |  |  | 1.8555 |  |  | 0.38373 | 0.82831 | 0.44458 | 0.058753
228 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 252 | 574 |  | 475 | 0.608 | 0.765 | 0.156 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44790 | 0.59683 |  | 1 |  | 0.37375 | 0.83471 | 0.46096 | 0.058208
255 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 723 | 1142 |  | 564 | 0.969 | 1.246 | 0.276 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44791 | 0.5405 |  | 1 |  | 0.83262 | 1.3624 | 0.52974 | 0.09284 | [18 empty] | 0.80952 | 1.366 | 0.55652 | 0.091802
256 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 674 | 1150 |  | 540 | 0.932 | 1.252 | 0.32 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44791 | 0.54738 |  | 1 |  | 0.78642 | 1.3697 | 0.5833 | 0.090764
315 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 278 | 534 |  | 496 | 0.628 | 0.731 | 0.103 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44792 | 0.10128 |  | 1 |  | 0.41307 | 0.80297 | 0.3899 | 0.086958 | [18 empty] | 0.40977 | 0.79883 | 0.38906 | 0.085228
316 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 271 | 525 |  | 456 | 0.623 | 0.724 | 0.101 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44792 | 0.10785 |  | 1 |  | 0.40648 | 0.79469 | 0.38822 | 0.083499
357 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 295 | 515 |  | 789 | 0.641 | 0.714 | 0.073 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44792 | 0.49346 |  | 1 |  | 0.4291 | 0.78549 | 0.35639 | 0.1123 | [18 empty] | 0.43099 | 0.80067 | 0.36968 | 0.11195
358 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 299 | 548 |  | 781 | 0.644 | 0.743 | 0.099 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 44792 | 0.50003 |  | 1 |  | 0.43287 | 0.81585 | 0.38298 | 0.11161

## template
RUN NOTES:
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area |  | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | -476 | 0 | -1897 | 0 | -352 |  | 0.5 | -0.10446 | -0.53662 | -0.43215 | 0.008251
TIC | TC | TNb |  | 0.0006 | 1111.6 | 0.0012 | 1633.2 | 6e-05 | 832.9 |  | 0.2 | 3.7325 | 8.0684 | 4.3359 | 0.32596
808.5 | 2442 | 666.5 |  | 0.0018 | 2968.3 | 0.0036 | 5092.6 | 0.00018 | 3126.7 |  | 0.6 | 2.801 | 5.7632 | 2.9622 | 0.30568
|  |  |  | 0.002997 | 5568.039 | 0.005994 | 9993.628 | 0.0002997 | 6105.611 |  | 0.333 | 8.9747 | 18.2304 | 9.2557 | 1.0118
|  |  |  | 0.004203 | 8100.861 | 0.008406 | 14259.672 | 0.0004203 | 6379.672 |  | 0.467 | 9.1282 | 17.8694 | 8.7412 | 0.75174
|  |  |  | 0.0054 | 10308.8 | 0.0108 | 18829.1 | 0.00054 | 10419.7 |  | 0.6 | 8.9561 | 17.9683 | 9.0122 | 0.93212
|  |  | Slope |  | 5.0311e-07 |  | 5.3311e-07 |  | 5.1538e-08
|  |  | Intercept |  | 0.00018725 |  | 0.000743 |  | 2.2267e-05
|  |  | RSQ |  | 0.99805 |  | 0.99688 |  | 0.96984
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 0 | 4301 |  | 7612 | 0 | 6.538 | 6.538 |  | 1.134 | 1 | 0 | 0 |  | 0 |  | TICnp | 2022-09-06 | 15:37 |  | 1 |  | 0.62416 | 10.1197 | 9.4956 | 1.3819
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 34 | 3333 |  | 7031 | 0.735 | 5.17 | 4.435 |  | 1.032 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 15:45 |  | 1 |  | 0.68118 | 8.3995 | 7.7184 | 1.2821 |  |  | 1.4663 |  |  |  |  | 14.2458 |  |  |  |  | 15.2951 |  |  |  |  | 2.2767 |  |  | 0.68621 | 9.0437 | 8.3575 | 1.2677
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 40 | 4058 |  | 6863 | 0.743 | 6.194 | 5.45 |  | 1.003 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 15:54 |  | 1 |  | 0.69124 | 9.6879 | 8.9966 | 1.2532
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3519 | 2119 |  | 1165 | 3.114 | 2.074 | 0 |  | 0.006 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 16:16 |  | 1 |  | 3.9154 | 3.7453 | -0.17004 | 0.16462
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 837 | 3048 |  | 858 | 1.057 | 2.861 | 1.804 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 16:27 |  | 1 |  | 1.2167 | 4.7358 | 3.5192 | 0.13297 |  |  | 4.8277 |  |  |  |  | 31.5978 |  |  |  |  | 42.5581 |  |  |  |  | 34.8642 |  |  | 1.188 | 4.0897 | 2.9017 | 0.11323
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 780 | 1836 |  | 475 | 1.013 | 1.834 | 0.82 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 16:38 |  | 1 |  | 1.1593 | 3.4436 | 2.2842 | 0.093494
7 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 787 | 2738 |  | 1439 | 2.546 | 6.496 | 3.95 |  | 0.086 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:00 |  | 1 |  | 2.916 | 11.0133 | 8.0973 | 0.48215
8 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1407 | 2688 |  | 1211 | 3.736 | 6.388 | 2.652 |  | 0.027 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:09 |  | 1 |  | 4.4756 | 10.88 | 6.4044 | 0.4234 |  | 51.5345 | 3.0987 |  |  |  | 77.8685 | 3.8964 |  |  |  | 104.2025 | 9.0873 |  |  |  | 31.5543 | 14.5604 |  |  | 4.546 | 10.6721 | 6.1261 | 0.39466
9 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1463 | 2532 |  | 988 | 3.843 | 6.058 | 2.215 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:19 |  | 1 |  | 4.6165 | 10.4642 | 5.8477 | 0.36593
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3792 | 8006 |  | 4043 | 2.77 | 5.885 | 3.114 |  | 0.256 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:42 |  | 1 |  | 3.4917 | 8.3518 | 4.8601 | 0.38439
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3900 | 8043 |  | 3960 | 2.839 | 5.911 | 3.072 |  | 0.248 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:52 |  | 1 |  | 3.5823 | 8.3847 | 4.8024 | 0.37726 |  | 20.4854 | 1.7863 |  |  |  | 39.4487 | 0.42478 |  |  |  | 58.4119 | 2.1065 |  |  |  | 24.7943 | 1.5372 |  |  | 3.6146 | 8.3669 | 4.7524 | 0.37438
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3977 | 8003 |  | 3893 | 2.888 | 5.882 | 2.994 |  | 0.243 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 18:03 |  | 1 |  | 3.6468 | 8.3491 | 4.7023 | 0.37151
13 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 4493 | 11671 |  | 6650 | 5.798 | 15.264 | 9.466 |  | 0.87 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 18:25 |  | 1 |  | 7.3505 | 20.9157 | 13.5653 | 1.0961
14 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6051 | 11610 |  | 6672 | 7.594 | 15.186 | 7.593 |  | 0.874 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 18:35 |  | 1 |  | 9.7044 | 20.8181 | 11.1137 | 1.0995 |  | 8.7579 | 1.7133 |  |  |  | 15.7449 | 0.15368 |  |  |  | 22.732 | 1.2284 |  |  |  | 20.0579 | 3.5093 |  |  | 9.7882 | 20.8341 | 11.0459 | 1.0805
15 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6162 | 11630 |  | 6427 | 7.721 | 15.212 | 7.492 |  | 0.835 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 18:46 |  | 1 |  | 9.8721 | 20.8501 | 10.978 | 1.0616
16 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8734 | 16958 |  | 8850 | 7.618 | 15.68 | 8.062 |  | 0.867 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 19:08 |  | 1 |  | 9.8103 | 20.9497 | 11.1394 | 1.0244
17 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8908 | 16464 |  | 8737 | 7.761 | 15.232 | 7.471 |  | 0.854 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 19:20 |  | 1 |  | 9.9977 | 20.3857 | 10.388 | 1.0119 |  | 10.4632 | 1.127 |  |  |  | 13.7393 | 0.85312 |  |  |  | 17.0153 | 2.7223 |  |  |  | 11.4941 | 1.6827 |  |  | 9.9417 | 20.4731 | 10.5314 | 1.0034
18 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8804 | 16617 |  | 8584 | 7.676 | 15.37 | 7.695 |  | 0.837 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 19:31 |  | 1 |  | 9.8857 | 20.5604 | 10.6747 | 0.995
19 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11213 | 21958 |  | 11361 | 7.515 | 15.734 | 8.22 |  | 0.894 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 19:53 |  | 1 |  | 9.7143 | 20.7484 | 11.0341 | 1.013
20 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11284 | 21878 |  | 11273 | 7.56 | 15.678 | 8.118 |  | 0.886 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:04 |  | 1 |  | 9.7739 | 20.6773 | 10.9035 | 1.0054 |  | 8.5518 | 0.085828 |  |  |  | 14.2892 | 1.0236 |  |  |  | 20.0265 | 1.8718 |  |  |  | 11.2027 | 0.91833 |  |  | 9.7697 | 20.572 | 10.8024 | 1.0008
21 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11274 | 21641 |  | 11166 | 7.553 | 15.51 | 7.957 |  | 0.877 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:16 |  | 1 |  | 9.7655 | 20.4668 | 10.7013 | 0.99623
22 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5337 | 11804 |  | 6516 | 7.516 | 17.131 | 9.615 |  | 0.942 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:35 |  | 1 |  | 9.5744 | 23.4528 | 13.8784 | 1.1936
23 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5366 | 12082 |  | 6436 | 7.552 | 17.524 | 9.972 |  | 0.929 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:46 |  | 1 |  | 9.6231 | 23.9468 | 14.3238 | 1.1799 | 100 |  | 0.86758 |  |  | 100 |  | 2.4108 |  |  | 100 |  | 4.6746 |  |  | 100 |  | 1.3781 |  |  | 9.665 | 23.6616 | 13.9966 | 1.1718
24 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5416 | 11761 |  | 6342 | 7.617 | 17.07 | 9.453 |  | 0.912 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:56 |  | 1 |  | 9.7069 | 23.3764 | 13.6695 | 1.1637
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5043 | 6907 |  | 3345 | 4.284 | 6.13 | 1.847 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:16 |  | 1 |  | 5.4488 | 8.8504 | 3.4016 | 0.38932
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3545 | 7179 |  | 3324 | 3.135 | 6.36 | 3.225 |  | 0.232 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:26 |  | 1 |  | 3.9415 | 9.1404 | 5.1989 | 0.38716 |  |  | 1.5954 |  |  |  |  | 3.2842 |  |  |  |  | 4.5836 |  |  |  |  | 1.2053 |  |  | 3.9103 | 8.9927 | 5.0824 | 0.38484
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3483 | 6902 |  | 3279 | 3.087 | 6.126 | 3.039 |  | 0.227 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:37 |  | 1 |  | 3.8791 | 8.8451 | 4.9659 | 0.38252
28 | 9 | F 17mar22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3859 | 7164 |  | 1995 | 3.376 | 6.348 | 2.972 |  | 0.093 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:56 |  | 1 |  | 4.2575 | 9.1244 | 4.8669 | 0.25017
29 | 9 | F 17mar22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4123 | 7239 |  | 1934 | 3.578 | 6.411 | 2.834 |  | 0.086 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 22:06 |  | 1 |  | 4.5231 | 9.2044 | 4.6813 | 0.24388 |  |  | 0.84894 |  |  |  |  | 1.6703 |  |  |  |  | 2.4704 |  |  |  |  | 3.5268 |  |  | 4.504 | 9.1282 | 4.6241 | 0.23966
30 | 9 | F 17mar22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4085 | 7096 |  | 1852 | 3.549 | 6.29 | 2.741 |  | 0.078 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 22:17 |  | 1 |  | 4.4849 | 9.0519 | 4.567 | 0.23543
31 | 10 | B 24feb22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3323 | 5839 |  | 2511 | 2.964 | 5.225 | 2.261 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 22:35 |  | 1 |  | 3.7181 | 7.7117 | 3.9935 | 0.30336
32 | 10 | B 24feb22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3065 | 5850 |  | 2443 | 2.766 | 5.234 | 2.468 |  | 0.139 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 22:45 |  | 1 |  | 3.4585 | 7.7234 | 4.2649 | 0.29635 |  |  | 1.1282 |  |  |  |  | 1.3574 |  |  |  |  | 1.5429 |  |  |  |  | 2.0313 |  |  | 3.4782 | 7.7762 | 4.298 | 0.29939
33 | 10 | B 24feb22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3104 | 5949 |  | 2502 | 2.796 | 5.318 | 2.522 |  | 0.146 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 22:56 |  | 1 |  | 3.4978 | 7.829 | 4.3312 | 0.30243
34 | 11 | C 28feb22 CCT 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9848 | 14266 |  | 2088 | 7.97 | 12.364 | 4.394 |  | 0.102 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 23:15 |  | 1 |  | 10.2837 | 16.6967 | 6.413 | 0.25976
35 | 11 | C 28feb22 CCT 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12457 | 14406 |  | 1905 | 9.972 | 12.483 | 2.511 |  | 0.083 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 23:26 |  | 1 |  | 12.9089 | 16.846 | 3.9371 | 0.24089 |  |  | 2.0521 |  |  |  |  | 1.4728 |  |  |  |  | 13.9657 |  |  |  |  | 1.7263 |  |  | 13.0427 | 16.7228 | 3.6801 | 0.23883
36 | 11 | C 28feb22 CCT 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12723 | 14175 |  | 1865 | 10.176 | 12.287 | 2.112 |  | 0.079 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 23:38 |  | 1 |  | 13.1766 | 16.5997 | 3.4231 | 0.23677
37 | 12 | B50 13jun22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6009 | 6941 |  | 2535 | 5.025 | 6.159 | 1.134 |  | 0.149 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 23:57 |  | 1 |  | 6.4208 | 8.8867 | 2.4658 | 0.30583
38 | 12 | B50 13jun22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3384 | 6942 |  | 2459 | 3.011 | 6.16 | 3.149 |  | 0.141 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 0:07 |  | 1 |  | 3.7795 | 8.8877 | 5.1082 | 0.298 |  |  | 0.31998 |  |  |  |  | 0.42076 |  |  |  |  | 0.49539 |  |  |  |  | 0.48543 |  |  | 3.7735 | 8.8691 | 5.0956 | 0.29727
39 | 12 | B50 13jun22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3372 | 6907 |  | 2445 | 3.002 | 6.13 | 3.128 |  | 0.14 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 0:17 |  | 1 |  | 3.7675 | 8.8504 | 5.0829 | 0.29655
40 | 13 | C 20apr22 21.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7690 | 11465 |  | 2041 | 6.314 | 9.991 | 3.677 |  | 0.097 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 0:37 |  | 1 |  | 8.1123 | 13.7102 | 5.5979 | 0.25491
41 | 13 | C 20apr22 21.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9335 | 11612 |  | 1937 | 7.577 | 10.116 | 2.539 |  | 0.087 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 0:47 |  | 1 |  | 9.7675 | 13.867 | 4.0995 | 0.24419 |  |  | 1.8979 |  |  |  |  | 1.0357 |  |  |  |  | 8.3882 |  |  |  |  | 0.59271 |  |  | 9.8611 | 13.7955 | 3.9344 | 0.24347
42 | 13 | C 20apr22 21.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9521 | 11478 |  | 1923 | 7.719 | 10.003 | 2.284 |  | 0.085 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 0:59 |  | 1 |  | 9.9547 | 13.7241 | 3.7694 | 0.24275
43 | 14 | B50 23may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4927 | 5910 |  | 1738 | 4.195 | 5.285 | 1.091 |  | 0.066 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 1:18 |  | 1 |  | 5.3321 | 7.7874 | 2.4553 | 0.22368
44 | 14 | B50 23may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3171 | 5970 |  | 1662 | 2.847 | 5.336 | 2.489 |  | 0.058 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 1:28 |  | 1 |  | 3.5652 | 7.8513 | 4.2861 | 0.21585 |  |  | 1.3926 |  |  |  |  | 0.55834 |  |  |  |  | 0.13032 |  |  |  |  | 0 |  |  | 3.5406 | 7.8295 | 4.2889 | 0.21585
45 | 14 | B50 23may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3122 | 5929 |  | 1662 | 2.81 | 5.302 | 2.492 |  | 0.058 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 1:38 |  | 1 |  | 3.5159 | 7.8076 | 4.2917 | 0.21585
46 | 15 | F50 23may22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4263 | 12295 |  | 2543 | 3.685 | 10.694 | 7.009 |  | 0.15 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 1:58 |  | 1 |  | 4.664 | 14.5952 | 9.9312 | 0.30665
47 | 15 | F50 23may22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4680 | 12393 |  | 2515 | 4.006 | 10.778 | 6.772 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 2:08 |  | 1 |  | 5.0836 | 14.6997 | 9.6161 | 0.30377 |  |  | 0.33592 |  |  |  |  | 0.55278 |  |  |  |  | 1.0258 |  |  |  |  | 0.74931 |  |  | 5.0921 | 14.6592 | 9.567 | 0.30263
48 | 15 | F50 23may22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4697 | 12317 |  | 2493 | 4.019 | 10.713 | 6.694 |  | 0.145 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 2:20 |  | 1 |  | 5.1007 | 14.6187 | 9.518 | 0.3015
49 | 16 | F 17mar22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4306 | 7812 |  | 2138 | 3.718 | 6.897 | 3.178 |  | 0.108 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 2:38 |  | 1 |  | 4.7073 | 9.8153 | 5.1081 | 0.26491
50 | 16 | F 17mar22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4098 | 7831 |  | 2072 | 3.559 | 6.912 | 3.353 |  | 0.101 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 2:49 |  | 1 |  | 4.498 | 9.8356 | 5.3376 | 0.25811 |  |  | 0.13413 |  |  |  |  | 0.63073 |  |  |  |  | 1.2798 |  |  |  |  | 0.36007 |  |  | 4.501 | 9.8047 | 5.3037 | 0.25764
51 | 16 | F 17mar22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4104 | 7773 |  | 2063 | 3.563 | 6.864 | 3.301 |  | 0.1 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 3:00 |  | 1 |  | 4.504 | 9.7737 | 5.2697 | 0.25718
52 | 17 | F 24feb22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4142 | 7402 |  | 2209 | 3.592 | 6.55 | 2.957 |  | 0.115 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 3:18 |  | 1 |  | 4.5422 | 9.3782 | 4.8359 | 0.27223
53 | 17 | F 24feb22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4094 | 7421 |  | 2247 | 3.556 | 6.565 | 3.009 |  | 0.119 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 3:29 |  | 1 |  | 4.4939 | 9.3984 | 4.9045 | 0.27614 |  |  | 0.936 |  |  |  |  | 0.21578 |  |  |  |  | 1.2829 |  |  |  |  | 0.37396 |  |  | 4.5151 | 9.3883 | 4.8732 | 0.27563
54 | 17 | F 24feb22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4136 | 7402 |  | 2237 | 3.588 | 6.55 | 2.962 |  | 0.118 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 3:40 |  | 1 |  | 4.5362 | 9.3782 | 4.842 | 0.27511
55 | 18 | F 31jan22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4392 | 7856 |  | 2330 | 3.785 | 6.934 | 3.15 |  | 0.128 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 3:59 |  | 1 |  | 4.7938 | 9.8622 | 5.0685 | 0.2847
56 | 18 | F 31jan22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4521 | 7867 |  | 2362 | 3.883 | 6.943 | 3.06 |  | 0.131 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 4:09 |  | 1 |  | 4.9236 | 9.874 | 4.9504 | 0.288 |  |  | 0.12254 |  |  |  |  | 0.1295 |  |  |  |  | 0.13641 |  |  |  |  | 1.1519 |  |  | 4.9266 | 9.8804 | 4.9538 | 0.28635
57 | 18 | F 31jan22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4527 | 7879 |  | 2330 | 3.888 | 6.954 | 3.066 |  | 0.128 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 4:20 |  | 1 |  | 4.9296 | 9.8868 | 4.9571 | 0.2847
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10012 | 16883 |  | 4902 | 8.096 | 14.581 | 6.485 |  | 0.397 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 4:40 |  | 1 |  | 10.4487 | 19.487 | 9.0383 | 0.54981
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12270 | 17005 |  | 4945 | 9.829 | 14.685 | 4.857 |  | 0.401 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 4:50 |  | 1 |  | 12.7208 | 19.6171 | 6.8964 | 0.55424 |  |  | 0.29224 |  | 99.7644 |  |  | 1.1974 |  | 98.0054 |  |  | 4.0043 |  | 96.2463 |  |  | 0.1858 |  | 105.427 | 12.7394 | 19.5004 | 6.761 | 0.55476
60 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12307 | 16786 |  | 4955 | 9.856 | 14.5 | 4.643 |  | 0.402 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 5:02 |  | 1 |  | 12.758 | 19.3836 | 6.6256 | 0.55527
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7004 | 8324 |  | 2514 | 5.788 | 7.331 | 1.543 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 5:21 |  | 1 |  | 7.422 | 10.3612 | 2.9392 | 0.30367
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4965 | 9537 |  | 2599 | 4.224 | 8.358 | 4.134 |  | 0.156 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 5:32 |  | 1 |  | 5.3704 | 11.6546 | 6.2842 | 0.31243 |  |  | 0.62022 | 8.3098 |  |  |  | 12.2322 | 10.5686 |  |  |  | 23.2761 | 12.7628 |  |  |  | 6.5768 | 5.4791 |  | 5.3538 | 10.9828 | 5.6291 | 0.30248
63 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4932 | 8277 |  | 2406 | 4.199 | 7.291 | 3.092 |  | 0.136 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 5:43 |  | 1 |  | 5.3371 | 10.3111 | 4.974 | 0.29253
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1623 | 587 |  | 422 | 1.66 | 0.776 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:01 |  | 1 |  | 2.0076 | 2.1119 | 0.1043 | 0.088031
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 326 | 530 |  | 321 | 0.665 | 0.727 | 0.062 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:10 |  | 1 |  | 0.70252 | 2.0511 | 1.3486 | 0.077621 |  |  | 1.8448 |  |  |  |  | 1.5474 |  |  |  |  | 1.3921 |  |  |  |  | 1.7413 |  |  | 0.70906 | 2.0671 | 1.358 | 0.076951
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 339 | 560 |  | 308 | 0.675 | 0.753 | 0.078 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:19 |  | 1 |  | 0.7156 | 2.0831 | 1.3675 | 0.076281
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2855 | 10731 |  | 6516 | 4.342 | 15.616 | 11.274 |  | 0.942 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:37 |  | 1 |  | 5.4121 | 21.5461 | 16.134 | 1.1936
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4729 | 10802 |  | 6479 | 6.738 | 15.716 | 8.978 |  | 0.936 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:47 |  | 1 |  | 8.5548 | 21.6722 | 13.1174 | 1.1873 | 90.3821 |  | 5.4722 |  |  | 90.0726 |  | 1.0551 |  |  | 88.3581 |  | 5.5533 |  |  | 101.4165 |  | 0.014469 |  |  | 8.7955 | 21.5585 | 12.763 | 1.1874
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5016 | 10674 |  | 6480 | 7.106 | 15.536 | 8.43 |  | 0.936 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:58 |  | 1 |  | 9.0361 | 21.4448 | 12.4086 | 1.1874
70 | 21 | B50 31may22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5144 | 7702 |  | 8635 | 4.362 | 6.804 | 2.442 |  | 0.787 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 7:17 |  | 1 |  | 5.5505 | 9.698 | 4.1476 | 0.93459
71 | 21 | B50 31may22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4108 | 7712 |  | 8695 | 3.567 | 6.812 | 3.245 |  | 0.793 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 7:27 |  | 1 |  | 4.508 | 9.7087 | 5.2007 | 0.94078 |  |  | 0.8446 |  |  |  |  | 0.021962 |  |  |  |  | 0.69663 |  |  |  |  | 0.70368 |  |  | 4.5271 | 9.7098 | 5.1826 | 0.93748
72 | 21 | B50 31may22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4146 | 7714 |  | 8631 | 3.596 | 6.814 | 3.218 |  | 0.787 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 7:38 |  | 1 |  | 4.5463 | 9.7108 | 5.1646 | 0.93418
73 | 22 | C50 09jun22 1.5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7114 | 11220 |  | 2418 | 5.873 | 9.784 | 3.911 |  | 0.137 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 7:57 |  | 1 |  | 7.5327 | 13.449 | 5.9163 | 0.29377
74 | 22 | C50 09jun22 1.5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8310 | 11241 |  | 2313 | 6.79 | 9.802 | 3.012 |  | 0.126 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 8:07 |  | 1 |  | 8.7361 | 13.4714 | 4.7353 | 0.28295 |  |  | 0.78016 |  |  |  |  | 0.34091 |  |  |  |  | 2.4426 |  |  |  |  | 0.10935 |  |  | 8.7704 | 13.4485 | 4.6781 | 0.28279
75 | 22 | C50 09jun22 1.5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8378 | 11198 |  | 2310 | 6.842 | 9.766 | 2.924 |  | 0.126 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 8:18 |  | 1 |  | 8.8046 | 13.4256 | 4.621 | 0.28264
76 | 23 | B 24feb22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4715 | 5977 |  | 2513 | 4.032 | 5.342 | 1.31 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 8:38 |  | 1 |  | 5.1188 | 7.8588 | 2.74 | 0.30356
77 | 23 | B 24feb22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3322 | 6074 |  | 2562 | 2.963 | 5.424 | 2.46 |  | 0.152 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 8:48 |  | 1 |  | 3.7171 | 7.9622 | 4.2451 | 0.30861 |  |  | 2.134 |  |  |  |  | 0.46759 |  |  |  |  | 2.6912 |  |  |  |  | 1.6501 |  |  | 3.6779 | 7.9809 | 4.303 | 0.30609
78 | 23 | B 24feb22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3244 | 6109 |  | 2513 | 2.904 | 5.454 | 2.55 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 8:58 |  | 1 |  | 3.6387 | 7.9996 | 4.3609 | 0.30356
79 | 24 | B50 31may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3240 | 5606 |  | 1663 | 2.901 | 5.027 | 2.127 |  | 0.058 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 9:17 |  | 1 |  | 3.6346 | 7.4632 | 3.8286 | 0.21595
80 | 24 | B50 31may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3246 | 5634 |  | 1655 | 2.905 | 5.051 | 2.146 |  | 0.057 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 9:27 |  | 1 |  | 3.6407 | 7.4931 | 3.8524 | 0.21512 |  |  | 1.0448 |  |  |  |  | 0.042679 |  |  |  |  | 0.91365 |  |  |  |  | 1.9351 |  |  | 3.6598 | 7.4947 | 3.8349 | 0.21306
81 | 24 | B50 31may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3284 | 5637 |  | 1615 | 2.934 | 5.055 | 2.121 |  | 0.053 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 9:38 |  | 1 |  | 3.6789 | 7.4963 | 3.8174 | 0.211
82 | 25 | F50 13jun22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3810 | 9895 |  | 2832 | 3.338 | 8.661 | 5.323 |  | 0.18 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 9:57 |  | 1 |  | 4.2082 | 12.0363 | 7.8281 | 0.33644
83 | 25 | F50 13jun22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4032 | 10009 |  | 2891 | 3.508 | 8.758 | 5.25 |  | 0.186 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 10:08 |  | 1 |  | 4.4316 | 12.1578 | 7.7263 | 0.34252 |  |  | 0.18181 |  |  |  |  | 0.33381 |  |  |  |  | 0.4211 |  |  |  |  | 0.72485 |  |  | 4.4275 | 12.1376 | 7.71 | 0.34129
84 | 25 | F50 13jun22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4024 | 9971 |  | 2867 | 3.502 | 8.726 | 5.224 |  | 0.184 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 10:19 |  | 1 |  | 4.4235 | 12.1173 | 7.6938 | 0.34005
85 | 26 | B50 09may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2450 | 3674 |  | 1300 | 2.295 | 3.391 | 1.096 |  | 0.02 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 10:37 |  | 1 |  | 2.8397 | 5.4033 | 2.5636 | 0.17853
86 | 26 | B50 09may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1779 | 3682 |  | 1282 | 1.78 | 3.397 | 1.618 |  | 0.018 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 10:47 |  | 1 |  | 2.1646 | 5.4118 | 3.2473 | 0.17668 |  |  | 1.0749 |  |  |  |  | 0.98026 |  |  |  |  | 2.327 |  |  |  |  | 3.3839 |  |  | 2.153 | 5.4385 | 3.2855 | 0.17972
87 | 26 | B50 09may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1756 | 3732 |  | 1341 | 1.762 | 3.44 | 1.678 |  | 0.024 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 10:57 |  | 1 |  | 2.1414 | 5.4651 | 3.3237 | 0.18276
88 | 27 | B50 31may22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2476 | 6157 |  | 2203 | 2.315 | 5.495 | 3.18 |  | 0.114 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 11:16 |  | 1 |  | 2.8659 | 8.0507 | 5.1848 | 0.27161
89 | 27 | B50 31may22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2697 | 6173 |  | 2193 | 2.484 | 5.509 | 3.025 |  | 0.113 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 11:26 |  | 1 |  | 3.0883 | 8.0678 | 4.9795 | 0.27058 |  |  | 0.97271 |  |  |  |  | 0.64968 |  |  |  |  | 1.6692 |  |  |  |  | 0.64971 |  |  | 3.1034 | 8.0417 | 4.9383 | 0.2697
90 | 27 | B50 31may22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2727 | 6124 |  | 2176 | 2.507 | 5.467 | 2.96 |  | 0.112 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 11:36 |  | 1 |  | 3.1184 | 8.0155 | 4.8971 | 0.26883
91 | 28 | F200 07jun22 wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4522 | 7998 |  | 1766 | 3.884 | 7.054 | 3.171 |  | 0.069 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 11:55 |  | 1 |  | 4.9246 | 10.0136 | 5.089 | 0.22656
92 | 28 | F200 07jun22 wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5284 | 7968 |  | 1724 | 4.468 | 7.029 | 2.561 |  | 0.064 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 12:05 |  | 1 |  | 5.6913 | 9.9817 | 4.2903 | 0.22224 |  |  | 0.56416 |  |  |  |  | 0.29864 |  |  |  |  | 0.054665 |  |  |  |  | 1.2903 |  |  | 5.7074 | 9.9966 | 4.2892 | 0.22368
93 | 28 | F200 07jun22 wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5316 | 7996 |  | 1752 | 4.493 | 7.053 | 2.559 |  | 0.067 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 12:16 |  | 1 |  | 5.7235 | 10.0115 | 4.288 | 0.22512
94 | 29 | ? 13jun22 weir 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3856 | 5533 |  | 2260 | 3.373 | 4.966 | 1.593 |  | 0.12 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 12:35 |  | 1 |  | 4.2545 | 7.3854 | 3.1309 | 0.27748
95 | 29 | ? 13jun22 weir 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3347 | 5517 |  | 2220 | 2.983 | 4.953 | 1.97 |  | 0.116 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 12:45 |  | 1 |  | 3.7423 | 7.3684 | 3.6261 | 0.27336 |  |  | 1.4898 |  |  |  |  | 0.15905 |  |  |  |  | 1.8327 |  |  |  |  | 1.0875 |  |  | 3.7146 | 7.3742 | 3.6596 | 0.27486
96 | 29 | ? 13jun22 weir 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3292 | 5528 |  | 2249 | 2.94 | 4.961 | 2.021 |  | 0.119 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 12:55 |  | 1 |  | 3.687 | 7.3801 | 3.6931 | 0.27635
97 | 30 | F 31jan22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4162 | 7711 |  | 2593 | 3.608 | 6.811 | 3.204 |  | 0.155 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 13:14 |  | 1 |  | 4.5624 | 9.7076 | 5.1453 | 0.31181
98 | 30 | F 31jan22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4551 | 7701 |  | 2598 | 3.906 | 6.802 | 2.896 |  | 0.156 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 13:25 |  | 1 |  | 4.9538 | 9.697 | 4.7432 | 0.31232 |  |  | 1.4919 |  |  |  |  | 0.53733 |  |  |  |  | 0.46946 |  |  |  |  | 1.2463 |  |  | 4.991 | 9.7231 | 4.7321 | 0.31428
99 | 30 | F 31jan22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4625 | 7750 |  | 2636 | 3.963 | 6.844 | 2.881 |  | 0.16 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 13:36 |  | 1 |  | 5.0282 | 9.7492 | 4.721 | 0.31624
100 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7109 | 16768 |  | 5900 | 5.869 | 14.485 | 8.616 |  | 0.501 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 13:56 |  | 1 |  | 7.5277 | 19.3644 | 11.8367 | 0.65268
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8101 | 16869 |  | 6008 | 6.63 | 14.57 | 7.94 |  | 0.512 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:06 |  | 1 |  | 8.5258 | 19.4721 | 10.9463 | 0.66381 |  |  | 0.55316 |  | 139.9635 |  |  | 0.50503 |  | 124.3383 |  |  | 1.337 |  | 108.7132 |  |  | 0.077609 |  | 109.5859 | 8.5495 | 19.4231 | 10.8736 | 0.66407
102 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8148 | 16777 |  | 6013 | 6.666 | 14.492 | 7.825 |  | 0.513 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:18 |  | 1 |  | 8.5731 | 19.374 | 10.8009 | 0.66433
103 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5734 | 8331 |  | 2900 | 4.814 | 7.337 | 2.523 |  | 0.187 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:37 |  | 1 |  | 6.1441 | 10.3687 | 4.2246 | 0.34345
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4821 | 8318 |  | 2841 | 4.113 | 7.325 | 3.212 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:47 |  | 1 |  | 5.2255 | 10.3548 | 5.1294 | 0.33737 |  |  | 0.096234 | 33.8458 |  |  |  | 0.27763 | 33.7591 |  |  |  | 0.46208 | 33.6711 |  |  |  | 0.73596 | 20.0587 |  | 5.228 | 10.3692 | 5.1413 | 0.33613
105 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4826 | 8345 |  | 2817 | 4.117 | 7.348 | 3.231 |  | 0.179 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:59 |  | 1 |  | 5.2305 | 10.3836 | 5.1531 | 0.3349
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1598 | 502 |  | 500 | 1.641 | 0.704 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:17 |  | 1 |  | 1.9824 | 2.0212 | 0.038822 | 0.096071
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 323 | 554 |  | 472 | 0.663 | 0.748 | 0.086 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:26 |  | 1 |  | 0.6995 | 2.0767 | 1.3772 | 0.093185 |  |  | 0.85937 |  |  |  |  | 1.6296 |  |  |  |  | 2.0185 |  |  |  |  | 2.1881 |  |  | 0.70252 | 2.0938 | 1.3912 | 0.094216
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 329 | 586 |  | 492 | 0.667 | 0.775 | 0.108 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:35 |  | 1 |  | 0.70554 | 2.1108 | 1.4053 | 0.095247
109 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2859 | 10984 |  | 6743 | 4.347 | 15.974 | 11.627 |  | 0.982 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:53 |  | 1 |  | 5.4188 | 21.9956 | 16.5769 | 1.2326
110 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4840 | 11094 |  | 6993 | 6.88 | 16.129 | 9.249 |  | 1.026 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 16:03 |  | 1 |  | 8.741 | 22.1911 | 13.4502 | 1.2756 | 91.6806 |  | 3.8572 |  |  | 92.8574 |  | 0.38512 |  |  | 85.7445 |  | 3.2419 |  |  | 108.4207 |  | 1.7937 |  |  | 8.9129 | 22.1485 | 13.2356 | 1.2642
111 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5045 | 11046 |  | 6861 | 7.142 | 16.061 | 8.918 |  | 1.003 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 16:14 |  | 1 |  | 9.0848 | 22.1058 | 13.0211 | 1.2529
112 | 6 | [15 empty] | 1 | [9 empty] | 1 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## 06sep22
RUN NOTES:
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area |  | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0.0072705 | -0.051863 | -0.059133 | 0.0073546
TIC | TC | TNb |  | 0.0006 | 1302 | 0.0012 | 2392 | 6e-05 | 973.7 |  | 0.2 | 3.2521 | 6.0721 | 2.82 | 0.25938
332.5 | 545 | 314.5 |  | 0.0018 | 3539.5 | 0.0036 | 7369 | 0.00018 | 3549.1 |  | 0.6 | 2.9365 | 6.3253 | 3.3888 | 0.29894
|  |  |  | 0.002997 | 5885.055 | 0.005994 | 11257.03 | 0.0002997 | 6340.043 |  | 0.333 | 8.79 | 17.4513 | 8.6613 | 0.95351
|  |  |  | 0.004203 | 8545.445 | 0.008406 | 16031.47 | 0.0004203 | 8151.47 |  | 0.467 | 9.0977 | 17.7452 | 8.6475 | 0.87192
|  |  |  | 0.0054 | 10880 | 0.0108 | 21105.5 | 0.00054 | 10842.1 |  | 0.6 | 9.0139 | 18.1968 | 9.1829 | 0.90063
|  |  | Slope |  | 4.9676e-07 |  | 5.1854e-07 |  | 4.9501e-08
|  |  | Intercept |  | 3.6352e-06 |  | -2.5931e-05 |  | 3.6773e-06
|  |  | RSQ |  | 0.9995 |  | 0.99885 |  | 0.9974
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 0 | 4301 |  | 7612 | 0 | 6.538 | 6.538 |  | 1.134 | 1 | 0 | 0 |  | 0 |  | TICnp | 2022-09-06 | 15:37 |  | 1 |  | 0.012117 | 7.3477 | 7.3356 | 1.2683
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 34 | 3333 |  | 7031 | 0.735 | 5.17 | 4.435 |  | 1.032 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 15:45 |  | 1 |  | 0.068417 | 5.6745 | 5.6061 | 1.1724 |  |  | 13.5385 |  |  |  |  | 19.8876 |  |  |  |  | 19.9624 |  |  |  |  | 2.3927 |  |  | 0.073384 | 6.3011 | 6.2277 | 1.1585
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 40 | 4058 |  | 6863 | 0.743 | 6.194 | 5.45 |  | 1.003 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 15:54 |  | 1 |  | 0.078352 | 6.9277 | 6.8493 | 1.1447
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3519 | 2119 |  | 1165 | 3.114 | 2.074 | 0 |  | 0.006 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 16:16 |  | 1 |  | 3.5034 | 2.1457 | -1.3577 | 0.12269
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 837 | 3048 |  | 858 | 1.057 | 2.861 | 1.804 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 16:27 |  | 1 |  | 0.83884 | 3.1091 | 2.2703 | 0.092299 |  |  | 6.9869 |  |  |  |  | 50.6691 |  |  |  |  | 71.8681 |  |  |  |  | 51.7018 |  |  | 0.81053 | 2.4807 | 1.6702 | 0.07334
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 780 | 1836 |  | 475 | 1.013 | 1.834 | 0.82 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 16:38 |  | 1 |  | 0.78221 | 1.8522 | 1.07 | 0.054381
7 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 787 | 2738 |  | 1439 | 2.546 | 6.496 | 3.95 |  | 0.086 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:00 |  | 1 |  | 1.9729 | 6.9691 | 4.9962 | 0.37455
8 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1407 | 2688 |  | 1211 | 3.736 | 6.388 | 2.652 |  | 0.027 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:09 |  | 1 |  | 3.5129 | 6.8395 | 3.3266 | 0.31812 |  | 19.4134 | 3.8826 |  |  |  | 10.6211 | 6.0938 |  |  |  | 1.8288 | 17.793 |  |  |  | 3.1597 | 18.9983 |  |  | 3.5824 | 6.6373 | 3.0549 | 0.29052
9 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1463 | 2532 |  | 988 | 3.843 | 6.058 | 2.215 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:19 |  | 1 |  | 3.6519 | 6.435 | 2.7831 | 0.26292
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3792 | 8006 |  | 4043 | 2.77 | 5.885 | 3.114 |  | 0.256 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:42 |  | 1 |  | 3.1456 | 6.8758 | 3.7303 | 0.33969
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3900 | 8043 |  | 3960 | 2.839 | 5.911 | 3.072 |  | 0.248 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:52 |  | 1 |  | 3.235 | 6.9078 | 3.6728 | 0.33284 |  | 8.8949 | 1.9514 |  |  |  | 14.8417 | 0.50169 |  |  |  | 20.7884 | 2.7133 |  |  |  | 10.0249 | 1.6747 |  |  | 3.2668 | 6.8905 | 3.6237 | 0.33007
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3977 | 8003 |  | 3893 | 2.888 | 5.882 | 2.994 |  | 0.243 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 18:03 |  | 1 |  | 3.2987 | 6.8732 | 3.5745 | 0.32731
13 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 4493 | 11671 |  | 6650 | 5.798 | 15.264 | 9.466 |  | 0.87 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 18:25 |  | 1 |  | 6.7134 | 18.0959 | 11.3825 | 0.99959
14 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6051 | 11610 |  | 6672 | 7.594 | 15.186 | 7.593 |  | 0.874 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 18:35 |  | 1 |  | 9.0376 | 18.0009 | 8.9633 | 1.0029 |  | 1.3372 | 1.8156 |  |  |  | 0.091464 | 0.17286 |  |  |  | 1.1543 | 1.5112 |  |  |  | 9.4051 | 3.6988 |  |  | 9.1203 | 18.0165 | 8.8961 | 0.98465
15 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6162 | 11630 |  | 6427 | 7.721 | 15.212 | 7.492 |  | 0.835 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 18:46 |  | 1 |  | 9.2031 | 18.032 | 8.8289 | 0.96644
16 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8734 | 16958 |  | 8850 | 7.618 | 15.68 | 8.062 |  | 0.867 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 19:08 |  | 1 |  | 9.2983 | 18.774 | 9.4757 | 0.94596
17 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8908 | 16464 |  | 8737 | 7.761 | 15.232 | 7.471 |  | 0.854 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 19:20 |  | 1 |  | 9.4834 | 18.2254 | 8.7421 | 0.93399 |  | 4.7563 | 1.1734 |  |  |  | 1.7244 | 0.92781 |  |  |  | 1.3076 | 3.1581 |  |  |  | 2.8753 | 1.7516 |  |  | 9.4281 | 18.3104 | 8.8823 | 0.92588
18 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8804 | 16617 |  | 8584 | 7.676 | 15.37 | 7.695 |  | 0.837 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 19:31 |  | 1 |  | 9.3728 | 18.3953 | 9.0226 | 0.91777
19 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11213 | 21958 |  | 11361 | 7.515 | 15.734 | 8.22 |  | 0.894 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 19:53 |  | 1 |  | 9.2896 | 18.9336 | 9.6439 | 0.94344
20 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11284 | 21878 |  | 11273 | 7.56 | 15.678 | 8.118 |  | 0.886 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:04 |  | 1 |  | 9.3484 | 18.8644 | 9.516 | 0.93618 |  | 3.8249 | 0.088603 |  |  |  | 4.2333 | 1.0917 |  |  |  | 4.6417 | 2.0869 |  |  |  | 3.5294 | 0.94742 |  |  | 9.3442 | 18.762 | 9.4178 | 0.93177
21 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11274 | 21641 |  | 11166 | 7.553 | 15.51 | 7.957 |  | 0.877 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:16 |  | 1 |  | 9.3401 | 18.6596 | 9.3195 | 0.92735
22 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5337 | 11804 |  | 6516 | 7.516 | 17.131 | 9.615 |  | 0.942 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:35 |  | 1 |  | 8.8494 | 20.3163 | 11.4669 | 1.0874
23 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5366 | 12082 |  | 6436 | 7.552 | 17.524 | 9.972 |  | 0.929 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:46 |  | 1 |  | 8.8974 | 20.7968 | 11.8994 | 1.0742 | 100 |  | 0.92621 |  |  | 100 |  | 2.704 |  |  | 100 |  | 5.506 |  |  | 100 |  | 1.4544 |  |  | 8.9388 | 20.5194 | 11.5806 | 1.0665
24 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5416 | 11761 |  | 6342 | 7.617 | 17.07 | 9.453 |  | 0.912 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:56 |  | 1 |  | 8.9802 | 20.242 | 11.2618 | 1.0587
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5043 | 6907 |  | 3345 | 4.284 | 6.13 | 1.847 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:16 |  | 1 |  | 5.0176 | 7.1112 | 2.0937 | 0.33852
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3545 | 7179 |  | 3324 | 3.135 | 6.36 | 3.225 |  | 0.232 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:26 |  | 1 |  | 3.5293 | 7.3933 | 3.864 | 0.33644 |  |  | 1.7607 |  |  |  |  | 3.9625 |  |  |  |  | 6.016 |  |  |  |  | 1.333 |  |  | 3.4985 | 7.2497 | 3.7512 | 0.33421
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3483 | 6902 |  | 3279 | 3.087 | 6.126 | 3.039 |  | 0.227 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:37 |  | 1 |  | 3.4677 | 7.106 | 3.6384 | 0.33199
28 | 9 | F 17mar22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3859 | 7164 |  | 1995 | 3.376 | 6.348 | 2.972 |  | 0.093 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:56 |  | 1 |  | 3.8412 | 7.3778 | 3.5365 | 0.20487
29 | 9 | F 17mar22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4123 | 7239 |  | 1934 | 3.578 | 6.411 | 2.834 |  | 0.086 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 22:06 |  | 1 |  | 4.1035 | 7.4555 | 3.352 | 0.19883 |  |  | 0.92428 |  |  |  |  | 2.0091 |  |  |  |  | 3.3533 |  |  |  |  | 4.1682 |  |  | 4.0846 | 7.3814 | 3.2967 | 0.19477
30 | 9 | F 17mar22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4085 | 7096 |  | 1852 | 3.549 | 6.29 | 2.741 |  | 0.078 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 22:17 |  | 1 |  | 4.0658 | 7.3072 | 3.2415 | 0.19071
31 | 10 | B 24feb22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3323 | 5839 |  | 2511 | 2.964 | 5.225 | 2.261 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 22:35 |  | 1 |  | 3.3087 | 6.0036 | 2.6949 | 0.25595
32 | 10 | B 24feb22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3065 | 5850 |  | 2443 | 2.766 | 5.234 | 2.468 |  | 0.139 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 22:45 |  | 1 |  | 3.0524 | 6.015 | 2.9626 | 0.24922 |  |  | 1.2614 |  |  |  |  | 1.6925 |  |  |  |  | 2.1346 |  |  |  |  | 2.3166 |  |  | 3.0718 | 6.0664 | 2.9946 | 0.25214
33 | 10 | B 24feb22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3104 | 5949 |  | 2502 | 2.796 | 5.318 | 2.522 |  | 0.146 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 22:56 |  | 1 |  | 3.0911 | 6.1177 | 3.0266 | 0.25506
34 | 11 | C 28feb22 CCT 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9848 | 14266 |  | 2088 | 7.97 | 12.364 | 4.394 |  | 0.102 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 23:15 |  | 1 |  | 9.7914 | 14.7431 | 4.9517 | 0.21407
35 | 11 | C 28feb22 CCT 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12457 | 14406 |  | 1905 | 9.972 | 12.483 | 2.511 |  | 0.083 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 23:26 |  | 1 |  | 12.3835 | 14.8883 | 2.5048 | 0.19596 |  |  | 2.1116 |  |  |  |  | 1.6221 |  |  |  |  | 22.3641 |  |  |  |  | 2.0416 |  |  | 12.5156 | 14.7685 | 2.2529 | 0.19398
36 | 11 | C 28feb22 CCT 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12723 | 14175 |  | 1865 | 10.176 | 12.287 | 2.112 |  | 0.079 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 23:38 |  | 1 |  | 12.6477 | 14.6487 | 2.001 | 0.192
37 | 12 | B50 13jun22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6009 | 6941 |  | 2535 | 5.025 | 6.159 | 1.134 |  | 0.149 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 23:57 |  | 1 |  | 5.9773 | 7.1465 | 1.1692 | 0.25833
38 | 12 | B50 13jun22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3384 | 6942 |  | 2459 | 3.011 | 6.16 | 3.149 |  | 0.141 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 0:07 |  | 1 |  | 3.3693 | 7.1475 | 3.7782 | 0.2508 |  |  | 0.35447 |  |  |  |  | 0.50913 |  |  |  |  | 0.64725 |  |  |  |  | 0.55417 |  |  | 3.3634 | 7.1294 | 3.766 | 0.25011
39 | 12 | B50 13jun22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3372 | 6907 |  | 2445 | 3.002 | 6.13 | 3.128 |  | 0.14 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 0:17 |  | 1 |  | 3.3574 | 7.1112 | 3.7538 | 0.24942
40 | 13 | C 20apr22 21.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7690 | 11465 |  | 2041 | 6.314 | 9.991 | 3.677 |  | 0.097 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 0:37 |  | 1 |  | 7.6474 | 11.8382 | 4.1908 | 0.20942
41 | 13 | C 20apr22 21.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9335 | 11612 |  | 1937 | 7.577 | 10.116 | 2.539 |  | 0.087 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 0:47 |  | 1 |  | 9.2817 | 11.9907 | 2.709 | 0.19912 |  |  | 1.9713 |  |  |  |  | 1.1657 |  |  |  |  | 12.7111 |  |  |  |  | 0.6985 |  |  | 9.3741 | 11.9212 | 2.5471 | 0.19843
42 | 13 | C 20apr22 21.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9521 | 11478 |  | 1923 | 7.719 | 10.003 | 2.284 |  | 0.085 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 0:59 |  | 1 |  | 9.4665 | 11.8517 | 2.3852 | 0.19774
43 | 14 | B50 23may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4927 | 5910 |  | 1738 | 4.195 | 5.285 | 1.091 |  | 0.066 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 1:18 |  | 1 |  | 4.9023 | 6.0773 | 1.175 | 0.17942
44 | 14 | B50 23may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3171 | 5970 |  | 1662 | 2.847 | 5.336 | 2.489 |  | 0.058 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 1:28 |  | 1 |  | 3.1577 | 6.1395 | 2.9818 | 0.1719 |  |  | 1.5537 |  |  |  |  | 0.69498 |  |  |  |  | 0.20644 |  |  |  |  | 0 |  |  | 3.1334 | 6.1182 | 2.9849 | 0.1719
45 | 14 | B50 23may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3122 | 5929 |  | 1662 | 2.81 | 5.302 | 2.492 |  | 0.058 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 1:38 |  | 1 |  | 3.109 | 6.097 | 2.9879 | 0.1719
46 | 15 | F50 23may22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4263 | 12295 |  | 2543 | 3.685 | 10.694 | 7.009 |  | 0.15 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 1:58 |  | 1 |  | 4.2426 | 12.699 | 8.4564 | 0.25912
47 | 15 | F50 23may22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4680 | 12393 |  | 2515 | 4.006 | 10.778 | 6.772 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 2:08 |  | 1 |  | 4.6569 | 12.8006 | 8.1437 | 0.25635 |  |  | 0.36202 |  |  |  |  | 0.61764 |  |  |  |  | 1.1822 |  |  |  |  | 0.85328 |  |  | 4.6654 | 12.7612 | 8.0959 | 0.25526
48 | 15 | F50 23may22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4697 | 12317 |  | 2493 | 4.019 | 10.713 | 6.694 |  | 0.145 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 2:20 |  | 1 |  | 4.6738 | 12.7218 | 8.048 | 0.25417
49 | 16 | F 17mar22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4306 | 7812 |  | 2138 | 3.718 | 6.897 | 3.178 |  | 0.108 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 2:38 |  | 1 |  | 4.2853 | 8.0498 | 3.7644 | 0.21902
50 | 16 | F 17mar22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4098 | 7831 |  | 2072 | 3.559 | 6.912 | 3.353 |  | 0.101 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 2:49 |  | 1 |  | 4.0787 | 8.0695 | 3.9908 | 0.21249 |  |  | 0.14605 |  |  |  |  | 0.74819 |  |  |  |  | 1.6704 |  |  |  |  | 0.42021 |  |  | 4.0817 | 8.0394 | 3.9577 | 0.21204
51 | 16 | F 17mar22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4104 | 7773 |  | 2063 | 3.563 | 6.864 | 3.301 |  | 0.1 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 3:00 |  | 1 |  | 4.0846 | 8.0093 | 3.9247 | 0.2116
52 | 17 | F 24feb22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4142 | 7402 |  | 2209 | 3.592 | 6.55 | 2.957 |  | 0.115 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 3:18 |  | 1 |  | 4.1224 | 7.6246 | 3.5022 | 0.22605
53 | 17 | F 24feb22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4094 | 7421 |  | 2247 | 3.556 | 6.565 | 3.009 |  | 0.119 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 3:29 |  | 1 |  | 4.0747 | 7.6443 | 3.5696 | 0.22981 |  |  | 1.0188 |  |  |  |  | 0.2581 |  |  |  |  | 1.7359 |  |  |  |  | 0.43173 |  |  | 4.0956 | 7.6344 | 3.5389 | 0.22932
54 | 17 | F 24feb22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4136 | 7402 |  | 2237 | 3.588 | 6.55 | 2.962 |  | 0.118 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 3:40 |  | 1 |  | 4.1164 | 7.6246 | 3.5081 | 0.22882
55 | 18 | F 31jan22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4392 | 7856 |  | 2330 | 3.785 | 6.934 | 3.15 |  | 0.128 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 3:59 |  | 1 |  | 4.3708 | 8.0954 | 3.7246 | 0.23803
56 | 18 | F 31jan22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4521 | 7867 |  | 2362 | 3.883 | 6.943 | 3.06 |  | 0.131 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 4:09 |  | 1 |  | 4.4989 | 8.1068 | 3.6079 | 0.2412 |  |  | 0.13241 |  |  |  |  | 0.15339 |  |  |  |  | 0.17955 |  |  |  |  | 1.3222 |  |  | 4.5019 | 8.113 | 3.6111 | 0.23962
57 | 18 | F 31jan22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4527 | 7879 |  | 2330 | 3.888 | 6.954 | 3.066 |  | 0.128 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 4:20 |  | 1 |  | 4.5049 | 8.1193 | 3.6144 | 0.23803
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10012 | 16883 |  | 4902 | 8.096 | 14.581 | 6.485 |  | 0.397 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 4:40 |  | 1 |  | 9.9543 | 17.4571 | 7.5028 | 0.49267
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12270 | 17005 |  | 4945 | 9.829 | 14.685 | 4.857 |  | 0.401 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 4:50 |  | 1 |  | 12.1977 | 17.5836 | 5.386 | 0.49692 |  |  | 0.30092 |  | 98.3962 |  |  | 1.3001 |  | 95.1018 |  |  | 5.0224 |  | 91.8074 |  |  | 0.19903 |  | 101.1552 | 12.216 | 17.4701 | 5.254 | 0.49742
60 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12307 | 16786 |  | 4955 | 9.856 | 14.5 | 4.643 |  | 0.402 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 5:02 |  | 1 |  | 12.2344 | 17.3565 | 5.1221 | 0.49791
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7004 | 8324 |  | 2514 | 5.788 | 7.331 | 1.543 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 5:21 |  | 1 |  | 6.9658 | 8.5808 | 1.6149 | 0.25625
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4965 | 9537 |  | 2599 | 4.224 | 8.358 | 4.134 |  | 0.156 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 5:32 |  | 1 |  | 4.9401 | 9.8387 | 4.8987 | 0.26466 |  |  | 0.66588 | 8.949 |  |  |  | 14.226 | 12.3981 |  |  |  | 29.8925 | 16.5275 |  |  |  | 7.4899 | 6.2637 |  | 4.9237 | 9.1854 | 4.2617 | 0.25511
63 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4932 | 8277 |  | 2406 | 4.199 | 7.291 | 3.092 |  | 0.136 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 5:43 |  | 1 |  | 4.9073 | 8.532 | 3.6247 | 0.24556
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1623 | 587 |  | 422 | 1.66 | 0.776 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:01 |  | 1 |  | 1.6197 | 0.5569 | -1.0628 | 0.049134
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 326 | 530 |  | 321 | 0.665 | 0.727 | 0.062 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:10 |  | 1 |  | 0.33116 | 0.49779 | 0.16663 | 0.039135 |  |  | 3.8256 |  |  |  |  | 6.0607 |  |  |  |  | 10.3549 |  |  |  |  | 3.3437 |  |  | 0.33761 | 0.51334 | 0.17573 | 0.038491
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 339 | 560 |  | 308 | 0.675 | 0.753 | 0.078 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:19 |  | 1 |  | 0.34407 | 0.5289 | 0.18483 | 0.037847
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2855 | 10731 |  | 6516 | 4.342 | 15.616 | 11.274 |  | 0.942 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:37 |  | 1 |  | 4.7396 | 18.4617 | 13.7221 | 1.0874
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4729 | 10802 |  | 6479 | 6.738 | 15.716 | 8.978 |  | 0.936 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:47 |  | 1 |  | 7.8426 | 18.5844 | 10.7417 | 1.0813 | 90.3821 |  | 5.8814 |  |  | 90.0726 |  | 1.1976 |  |  | 88.3581 |  | 6.701 |  |  | 101.4165 |  | 0.015258 |  |  | 8.0803 | 18.4738 | 10.3935 | 1.0814
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5016 | 10674 |  | 6480 | 7.106 | 15.536 | 8.43 |  | 0.936 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:58 |  | 1 |  | 8.3179 | 18.3632 | 10.0453 | 1.0815
70 | 21 | B50 31may22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5144 | 7702 |  | 8635 | 4.362 | 6.804 | 2.442 |  | 0.787 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 7:17 |  | 1 |  | 5.1179 | 7.9357 | 2.8178 | 0.86224
71 | 21 | B50 31may22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4108 | 7712 |  | 8695 | 3.567 | 6.812 | 3.245 |  | 0.793 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 7:27 |  | 1 |  | 4.0886 | 7.9461 | 3.8575 | 0.86819 |  |  | 0.91914 |  |  |  |  | 0.026099 |  |  |  |  | 0.92924 |  |  |  |  | 0.73249 |  |  | 4.1075 | 7.9471 | 3.8396 | 0.86502
72 | 21 | B50 31may22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4146 | 7714 |  | 8631 | 3.596 | 6.814 | 3.218 |  | 0.787 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 7:38 |  | 1 |  | 4.1264 | 7.9481 | 3.8218 | 0.86185
73 | 22 | C50 09jun22 1.5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7114 | 11220 |  | 2418 | 5.873 | 9.784 | 3.911 |  | 0.137 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 7:57 |  | 1 |  | 7.0751 | 11.5841 | 4.509 | 0.24674
74 | 22 | C50 09jun22 1.5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8310 | 11241 |  | 2313 | 6.79 | 9.802 | 3.012 |  | 0.126 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 8:07 |  | 1 |  | 8.2634 | 11.6059 | 3.3426 | 0.23635 |  |  | 0.81424 |  |  |  |  | 0.38498 |  |  |  |  | 3.4126 |  |  |  |  | 0.12574 |  |  | 8.2971 | 11.5836 | 3.2865 | 0.2362
75 | 22 | C50 09jun22 1.5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8378 | 11198 |  | 2310 | 6.842 | 9.766 | 2.924 |  | 0.126 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 8:18 |  | 1 |  | 8.3309 | 11.5613 | 3.2304 | 0.23605
76 | 23 | B 24feb22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4715 | 5977 |  | 2513 | 4.032 | 5.342 | 1.31 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 8:38 |  | 1 |  | 4.6917 | 6.1467 | 1.4551 | 0.25615
77 | 23 | B 24feb22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3322 | 6074 |  | 2562 | 2.963 | 5.424 | 2.46 |  | 0.152 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 8:48 |  | 1 |  | 3.3077 | 6.2473 | 2.9396 | 0.261 |  |  | 2.3706 |  |  |  |  | 0.57933 |  |  |  |  | 3.7975 |  |  |  |  | 1.8761 |  |  | 3.269 | 6.2655 | 2.9965 | 0.25857
78 | 23 | B 24feb22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3244 | 6109 |  | 2513 | 2.904 | 5.454 | 2.55 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 8:58 |  | 1 |  | 3.2302 | 6.2836 | 3.0534 | 0.25615
79 | 24 | B50 31may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3240 | 5606 |  | 1663 | 2.901 | 5.027 | 2.127 |  | 0.058 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 9:17 |  | 1 |  | 3.2262 | 5.762 | 2.5357 | 0.172
80 | 24 | B50 31may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3246 | 5634 |  | 1655 | 2.905 | 5.051 | 2.146 |  | 0.057 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 9:27 |  | 1 |  | 3.2322 | 5.791 | 2.5588 | 0.1712 |  |  | 1.1613 |  |  |  |  | 0.053711 |  |  |  |  | 1.3631 |  |  |  |  | 2.3402 |  |  | 3.2511 | 5.7926 | 2.5415 | 0.16922
81 | 24 | B50 31may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3284 | 5637 |  | 1615 | 2.934 | 5.055 | 2.121 |  | 0.053 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 9:38 |  | 1 |  | 3.27 | 5.7941 | 2.5242 | 0.16724
82 | 25 | F50 13jun22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3810 | 9895 |  | 2832 | 3.338 | 8.661 | 5.323 |  | 0.18 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 9:57 |  | 1 |  | 3.7926 | 10.21 | 6.4175 | 0.28773
83 | 25 | F50 13jun22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4032 | 10009 |  | 2891 | 3.508 | 8.758 | 5.25 |  | 0.186 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 10:08 |  | 1 |  | 4.0131 | 10.3282 | 6.3151 | 0.29357 |  |  | 0.19825 |  |  |  |  | 0.38229 |  |  |  |  | 0.49943 |  |  |  |  | 0.81265 |  |  | 4.0091 | 10.3085 | 6.2994 | 0.29238
84 | 25 | F50 13jun22 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4024 | 9971 |  | 2867 | 3.502 | 8.726 | 5.224 |  | 0.184 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 10:19 |  | 1 |  | 4.0052 | 10.2888 | 6.2837 | 0.2912
85 | 26 | B50 09may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2450 | 3674 |  | 1300 | 2.295 | 3.391 | 1.096 |  | 0.02 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 10:37 |  | 1 |  | 2.4414 | 3.7584 | 1.317 | 0.13606
86 | 26 | B50 09may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1779 | 3682 |  | 1282 | 1.78 | 3.397 | 1.618 |  | 0.018 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 10:47 |  | 1 |  | 1.7747 | 3.7667 | 1.9919 | 0.13428 |  |  | 1.2959 |  |  |  |  | 1.3672 |  |  |  |  | 3.6813 |  |  |  |  | 4.2575 |  |  | 1.7633 | 3.7926 | 2.0293 | 0.1372
87 | 26 | B50 09may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1756 | 3732 |  | 1341 | 1.762 | 3.44 | 1.678 |  | 0.024 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 10:57 |  | 1 |  | 1.7519 | 3.8185 | 2.0666 | 0.14012
88 | 27 | B50 31may22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2476 | 6157 |  | 2203 | 2.315 | 5.495 | 3.18 |  | 0.114 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 11:16 |  | 1 |  | 2.4672 | 6.3334 | 3.8662 | 0.22546
89 | 27 | B50 31may22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2697 | 6173 |  | 2193 | 2.484 | 5.509 | 3.025 |  | 0.113 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 11:26 |  | 1 |  | 2.6868 | 6.35 | 3.6632 | 0.22447 |  |  | 1.1032 |  |  |  |  | 0.80348 |  |  |  |  | 2.2253 |  |  |  |  | 0.75262 |  |  | 2.7017 | 6.3246 | 3.6229 | 0.22363
90 | 27 | B50 31may22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2727 | 6124 |  | 2176 | 2.507 | 5.467 | 2.96 |  | 0.112 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 11:36 |  | 1 |  | 2.7166 | 6.2992 | 3.5826 | 0.22278
91 | 28 | F200 07jun22 wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4522 | 7998 |  | 1766 | 3.884 | 7.054 | 3.171 |  | 0.069 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 11:55 |  | 1 |  | 4.4999 | 8.2427 | 3.7427 | 0.18219
92 | 28 | F200 07jun22 wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5284 | 7968 |  | 1724 | 4.468 | 7.029 | 2.561 |  | 0.064 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 12:05 |  | 1 |  | 5.257 | 8.2116 | 2.9546 | 0.17804 |  |  | 0.60294 |  |  |  |  | 0.353 |  |  |  |  | 0.093263 |  |  |  |  | 1.545 |  |  | 5.2729 | 8.2261 | 2.9532 | 0.17942
93 | 28 | F200 07jun22 wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5316 | 7996 |  | 1752 | 4.493 | 7.053 | 2.559 |  | 0.067 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 12:16 |  | 1 |  | 5.2888 | 8.2406 | 2.9518 | 0.18081
94 | 29 | ? 13jun22 weir 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3856 | 5533 |  | 2260 | 3.373 | 4.966 | 1.593 |  | 0.12 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 12:35 |  | 1 |  | 3.8383 | 5.6863 | 1.848 | 0.2311
95 | 29 | ? 13jun22 weir 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3347 | 5517 |  | 2220 | 2.983 | 4.953 | 1.97 |  | 0.116 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 12:45 |  | 1 |  | 3.3326 | 5.6697 | 2.3371 | 0.22714 |  |  | 1.6532 |  |  |  |  | 0.20101 |  |  |  |  | 2.7868 |  |  |  |  | 1.2561 |  |  | 3.3052 | 5.6754 | 2.3702 | 0.22858
96 | 29 | ? 13jun22 weir 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3292 | 5528 |  | 2249 | 2.94 | 4.961 | 2.021 |  | 0.119 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 12:55 |  | 1 |  | 3.2779 | 5.6811 | 2.4032 | 0.23001
97 | 30 | F 31jan22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4162 | 7711 |  | 2593 | 3.608 | 6.811 | 3.204 |  | 0.155 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 13:14 |  | 1 |  | 4.1423 | 7.945 | 3.8028 | 0.26407
98 | 30 | F 31jan22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4551 | 7701 |  | 2598 | 3.906 | 6.802 | 2.896 |  | 0.156 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 13:25 |  | 1 |  | 4.5287 | 7.9347 | 3.4059 | 0.26456 |  |  | 1.6103 |  |  |  |  | 0.6384 |  |  |  |  | 0.66881 |  |  |  |  | 1.412 |  |  | 4.5655 | 7.9601 | 3.3946 | 0.26645
99 | 30 | F 31jan22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4625 | 7750 |  | 2636 | 3.963 | 6.844 | 2.881 |  | 0.16 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 13:36 |  | 1 |  | 4.6023 | 7.9855 | 3.3832 | 0.26833
100 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7109 | 16768 |  | 5900 | 5.869 | 14.485 | 8.616 |  | 0.501 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 13:56 |  | 1 |  | 7.0701 | 17.3378 | 10.2677 | 0.59147
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8101 | 16869 |  | 6008 | 6.63 | 14.57 | 7.94 |  | 0.512 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:06 |  | 1 |  | 8.0557 | 17.4426 | 9.3869 | 0.60216 |  |  | 0.57798 |  | 138.0878 |  |  | 0.5485 |  | 120.7149 |  |  | 1.5254 |  | 103.3421 |  |  | 0.082172 |  | 105.1498 | 8.0791 | 17.3949 | 9.3158 | 0.60241
102 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8148 | 16777 |  | 6013 | 6.666 | 14.492 | 7.825 |  | 0.513 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:18 |  | 1 |  | 8.1024 | 17.3472 | 9.2448 | 0.60266
103 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5734 | 8331 |  | 2900 | 4.814 | 7.337 | 2.523 |  | 0.187 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:37 |  | 1 |  | 5.7041 | 8.588 | 2.884 | 0.29446
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4821 | 8318 |  | 2841 | 4.113 | 7.325 | 3.212 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:47 |  | 1 |  | 4.797 | 8.5745 | 3.7775 | 0.28862 |  |  | 0.1035 | 4.9967 |  |  |  | 0.32603 | 7.5954 |  |  |  | 0.60789 | 10.9832 |  |  |  | 0.82665 | 7.5788 |  | 4.7995 | 8.5885 | 3.7891 | 0.28743
105 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4826 | 8345 |  | 2817 | 4.117 | 7.348 | 3.231 |  | 0.179 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:59 |  | 1 |  | 4.802 | 8.6025 | 3.8006 | 0.28625
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1598 | 502 |  | 500 | 1.641 | 0.704 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:17 |  | 1 |  | 1.5949 | 0.46875 | -1.1262 | 0.056856
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 323 | 554 |  | 472 | 0.663 | 0.748 | 0.086 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:26 |  | 1 |  | 0.32817 | 0.52268 | 0.1945 | 0.054084 |  |  | 1.8001 |  |  |  |  | 6.1539 |  |  |  |  | 13.0819 |  |  |  |  | 3.5953 |  |  | 0.33116 | 0.53927 | 0.20812 | 0.055074
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 329 | 586 |  | 492 | 0.667 | 0.775 | 0.108 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:35 |  | 1 |  | 0.33414 | 0.55586 | 0.22173 | 0.056064
109 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2859 | 10984 |  | 6743 | 4.347 | 15.974 | 11.627 |  | 0.982 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:53 |  | 1 |  | 4.7462 | 18.899 | 14.1528 | 1.1249
110 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4840 | 11094 |  | 6993 | 6.88 | 16.129 | 9.249 |  | 1.026 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 16:03 |  | 1 |  | 8.0264 | 19.0891 | 11.0627 | 1.1661 | 91.6806 |  | 4.1416 |  |  | 92.8574 |  | 0.43557 |  |  | 85.7445 |  | 3.8927 |  |  | 108.4207 |  | 1.8854 |  |  | 8.1962 | 19.0476 | 10.8514 | 1.1552
111 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5045 | 11046 |  | 6861 | 7.142 | 16.061 | 8.918 |  | 1.003 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 16:14 |  | 1 |  | 8.3659 | 19.0061 | 10.6402 | 1.1444
112 | 6 | [15 empty] | 1 | [9 empty] | 1 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## 07sep22
RUN NOTES:
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area |  | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | -251 | 0 | -740.5 | 0 | -241.5 |  | 0.5 | -0.029025 | -0.012827 | 0.016198 | 0.0056453
TIC | TC | TNb |  | 0.0006 | 1081.2 | 0.0012 | 1803.4 | 6e-05 | 1084.8 |  | 0.2 | 3.2119 | 6.53 | 3.3181 | 0.33133
512 | 1189 | 583 |  | 0.0018 | 3335.1 | 0.0036 | 6203.7 | 0.00018 | 3108.9 |  | 0.6 | 2.9229 | 5.9602 | 3.0373 | 0.27181
|  |  |  | 0.002997 | 5838.008 | 0.005994 | 10636.126 | 0.0002997 | 6033.722 |  | 0.333 | 8.9727 | 17.6061 | 8.6335 | 0.90989
|  |  |  | 0.004203 | 8424.792 | 0.008406 | 15480.474 | 0.0004203 | 8705.474 |  | 0.467 | 9.1293 | 17.9059 | 8.7766 | 0.92247
|  |  |  | 0.0054 | 10663.1 | 0.0108 | 20398.7 | 0.00054 | 10863.9 |  | 0.6 | 8.9451 | 18.1657 | 9.2205 | 0.89007
|  |  | Slope |  | 4.9309e-07 |  | 5.159e-07 |  | 4.7834e-08
|  |  | Intercept |  | 0.00010925 |  | 0.00037561 |  | 1.4375e-05
|  |  | RSQ |  | 0.99959 |  | 0.99953 |  | 0.99776
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5039 | 11427 |  | 6955 | 7.135 | 16.599 | 9.465 |  | 1.019 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 16:38 |  | 1 |  | 8.6464 | 20.9028 | 12.2565 | 1.1569
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5114 | 11310 |  | 6802 | 7.23 | 16.433 | 9.203 |  | 0.992 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 16:48 |  | 1 |  | 8.7697 | 20.7016 | 11.932 | 1.1325 |  |  | 1.2637 |  |  |  |  | 0.10805 |  |  |  |  | 0.73286 |  |  |  |  | 0.098605 |  |  | 8.7146 | 20.6905 | 11.9759 | 1.1319
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5047 | 11297 |  | 6795 | 7.145 | 16.415 | 9.27 |  | 0.991 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 16:59 |  | 1 |  | 8.6595 | 20.6793 | 12.0197 | 1.1314
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2687 | 1073 |  | 652 | 2.476 | 1.188 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 17:17 |  | 1 |  | 2.8684 | 1.8584 | -1.01 | 0.091125
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 534 | 1080 |  | 613 | 0.825 | 1.194 | 0.369 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 17:26 |  | 1 |  | 0.74512 | 1.8656 | 1.1205 | 0.087394 |  |  | 5.9981 |  |  |  |  | 11.3715 |  |  |  |  | 21.387 |  |  |  |  | 6.7911 |  |  | 0.72343 | 1.978 | 1.2546 | 0.084524
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 490 | 1298 |  | 553 | 0.791 | 1.378 | 0.587 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 17:36 |  | 1 |  | 0.70173 | 2.0905 | 1.3888 | 0.081654
7 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 625 | 2253 |  | 1324 | 2.237 | 5.469 | 3.232 |  | 0.056 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 17:53 |  | 1 |  | 2.0872 | 7.6897 | 5.6026 | 0.38854
8 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1285 | 2282 |  | 1349 | 3.502 | 5.531 | 2.029 |  | 0.063 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 18:02 |  | 1 |  | 3.7143 | 7.7645 | 4.0502 | 0.39451 |  | 23.8936 | 0.13266 |  |  |  | 29.28 | 0.19953 |  |  |  | 34.6663 | 0.50515 |  |  |  | 29.0333 | 3.8307 |  |  | 3.7168 | 7.7568 | 4.04 | 0.3871
9 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1287 | 2276 |  | 1287 | 3.506 | 5.518 | 2.012 |  | 0.046 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 18:12 |  | 1 |  | 3.7193 | 7.7491 | 4.0298 | 0.37969
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3881 | 7652 |  | 3905 | 2.827 | 5.635 | 2.808 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 18:31 |  | 1 |  | 3.3715 | 7.2055 | 3.834 | 0.33528
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3965 | 7657 |  | 3822 | 2.881 | 5.638 | 2.757 |  | 0.236 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 18:41 |  | 1 |  | 3.4406 | 7.2098 | 3.7693 | 0.32866 |  | 14.261 | 0.74321 |  |  |  | 19.7839 | 0.63408 |  |  |  | 25.3069 | 0.53457 |  |  |  | 9.195 | 0.65709 |  |  | 3.4278 | 7.187 | 3.7592 | 0.32758
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3934 | 7604 |  | 3795 | 2.861 | 5.601 | 2.74 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 18:52 |  | 1 |  | 3.4151 | 7.1643 | 3.7492 | 0.32651
13 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 4606 | 11362 |  | 6382 | 5.928 | 14.872 | 8.943 |  | 0.828 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 19:11 |  | 1 |  | 7.1484 | 18.7307 | 11.5823 | 0.95992
14 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6143 | 11516 |  | 6427 | 7.699 | 15.067 | 7.368 |  | 0.835 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 19:22 |  | 1 |  | 9.4243 | 18.9693 | 9.545 | 0.96638 |  | 5.3065 | 1.1249 |  |  |  | 4.6274 | 1.4478 |  |  |  | 3.9484 | 4.0542 |  |  |  | 7.2958 | 0.14875 |  |  | 9.4776 | 18.8329 | 9.3554 | 0.96566
15 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6215 | 11340 |  | 6417 | 7.782 | 14.843 | 7.061 |  | 0.834 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 19:33 |  | 1 |  | 9.5309 | 18.6966 | 9.1657 | 0.96494
16 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8805 | 16593 |  | 9727 | 7.676 | 15.349 | 7.673 |  | 0.965 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 19:52 |  | 1 |  | 9.5308 | 19.1349 | 9.6041 | 1.0271
17 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8885 | 16567 |  | 9779 | 7.742 | 15.325 | 7.583 |  | 0.971 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 20:03 |  | 1 |  | 9.6153 | 19.1062 | 9.4909 | 1.0324 |  | 7.0474 | 0.39454 |  |  |  | 6.2929 | 0.27715 |  |  |  | 5.5384 | 0.15808 |  |  |  | 15.1355 | 0.73148 |  |  | 9.6343 | 19.1327 | 9.4985 | 1.0362
18 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8921 | 16615 |  | 9853 | 7.771 | 15.369 | 7.597 |  | 0.979 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 20:15 |  | 1 |  | 9.6533 | 19.1592 | 9.506 | 1.04
19 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11259 | 21872 |  | 11610 | 7.544 | 15.674 | 8.13 |  | 0.915 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 20:36 |  | 1 |  | 9.4349 | 19.4325 | 9.9976 | 0.94955
20 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11274 | 21969 |  | 11621 | 7.554 | 15.742 | 8.189 |  | 0.916 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 20:47 |  | 1 |  | 9.4472 | 19.5159 | 10.0687 | 0.95043 |  | 5.0007 | 0.060874 |  |  |  | 7.736 | 1.2725 |  |  |  | 10.4712 | 2.5399 |  |  |  | 5.0935 | 0.96932 |  |  | 9.4501 | 19.3925 | 9.9424 | 0.94584
21 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11281 | 21682 |  | 11506 | 7.558 | 15.539 | 7.981 |  | 0.906 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 20:58 |  | 1 |  | 9.4529 | 19.2691 | 9.8161 | 0.94126
22 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5311 | 11789 |  | 6557 | 7.483 | 17.11 | 9.628 |  | 0.95 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 21:17 |  | 1 |  | 9.0935 | 21.5254 | 12.4319 | 1.0934
23 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5221 | 11718 |  | 6497 | 7.367 | 17.01 | 9.643 |  | 0.939 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 21:27 |  | 1 |  | 8.9455 | 21.4033 | 12.4577 | 1.0838 | 100 |  | 0.1103 |  |  | 100 |  | 0.24939 |  |  | 100 |  | 0.34938 |  |  | 100 |  | 0.85691 |  |  | 8.9406 | 21.3766 | 12.436 | 1.0792
24 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5215 | 11687 |  | 6439 | 7.359 | 16.966 | 9.607 |  | 0.929 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 21:38 |  | 1 |  | 8.9357 | 21.35 | 12.4143 | 1.0746
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5011 | 6659 |  | 3259 | 4.259 | 5.92 | 1.66 |  | 0.225 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 21:57 |  | 1 |  | 5.1602 | 7.622 | 2.4618 | 0.34053
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3683 | 6721 |  | 3188 | 3.24 | 5.973 | 2.732 |  | 0.217 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 22:07 |  | 1 |  | 3.8506 | 7.686 | 3.8354 | 0.33374 |  |  | 0.74549 |  |  |  |  | 0.12075 |  |  |  |  | 0.98292 |  |  |  |  | 1.0663 |  |  | 3.8363 | 7.6907 | 3.8544 | 0.33197
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3654 | 6730 |  | 3151 | 3.218 | 5.98 | 2.763 |  | 0.214 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 22:18 |  | 1 |  | 3.822 | 7.6953 | 3.8733 | 0.3302
28 | 9 | F 23may22 wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5025 | 8678 |  | 2014 | 4.27 | 7.63 | 3.36 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 22:37 |  | 1 |  | 5.174 | 9.7053 | 4.5312 | 0.22143
29 | 9 | F 23may22 wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5611 | 8776 |  | 2022 | 4.719 | 7.714 | 2.994 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 22:48 |  | 1 |  | 5.7519 | 9.8064 | 4.0545 | 0.22219 |  |  | 0.034284 |  |  |  |  | 1.2174 |  |  |  |  | 3.0202 |  |  |  |  | 5.2131 |  |  | 5.7529 | 9.7471 | 3.9941 | 0.21655
30 | 9 | F 23may22 wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5613 | 8661 |  | 1904 | 4.721 | 7.616 | 2.895 |  | 0.083 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 22:59 |  | 1 |  | 5.7539 | 9.6877 | 3.9338 | 0.2109
31 | 10 | B 02may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3795 | 6047 |  | 1612 | 3.326 | 5.402 | 2.075 |  | 0.053 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 23:18 |  | 1 |  | 3.961 | 6.9906 | 3.0295 | 0.18297
32 | 10 | B 02may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3043 | 6047 |  | 1626 | 2.749 | 5.401 | 2.652 |  | 0.054 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 23:28 |  | 1 |  | 3.2194 | 6.9906 | 3.7711 | 0.18431 |  |  | 1.2638 |  |  |  |  | 0.22165 |  |  |  |  | 0.65958 |  |  |  |  | 2.9497 |  |  | 3.1992 | 6.9828 | 3.7836 | 0.18163
33 | 10 | B 02may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3002 | 6032 |  | 1570 | 2.718 | 5.388 | 2.671 |  | 0.048 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 23:38 |  | 1 |  | 3.179 | 6.9751 | 3.7961 | 0.17895
34 | 11 | B50 07jun22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3953 | 7400 |  | 9055 | 3.448 | 6.548 | 3.1 |  | 0.831 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 23:57 |  | 1 |  | 4.1168 | 8.3866 | 4.2698 | 0.89503
35 | 11 | B50 07jun22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4280 | 7327 |  | 9014 | 3.698 | 6.486 | 2.788 |  | 0.827 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 0:07 |  | 1 |  | 4.4393 | 8.3113 | 3.872 | 0.8911 |  |  | 0.1332 |  |  |  |  | 1.1357 |  |  |  |  | 2.2727 |  |  |  |  | 1.3013 |  |  | 4.4423 | 8.3588 | 3.9165 | 0.89694
36 | 11 | B50 07jun22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4286 | 7419 |  | 9136 | 3.703 | 6.564 | 2.861 |  | 0.839 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 0:18 |  | 1 |  | 4.4452 | 8.4062 | 3.961 | 0.90277
37 | 12 | B50 09may22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3276 | 6014 |  | 2088 | 2.928 | 5.374 | 2.446 |  | 0.102 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 0:36 |  | 1 |  | 3.4492 | 6.9565 | 3.5073 | 0.2285
38 | 12 | B50 09may22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2907 | 5974 |  | 2026 | 2.645 | 5.34 | 2.695 |  | 0.096 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 0:46 |  | 1 |  | 3.0853 | 6.9153 | 3.8299 | 0.22257 |  |  | 1.1897 |  |  |  |  | 0.23845 |  |  |  |  | 1.3743 |  |  |  |  | 0.94117 |  |  | 3.0671 | 6.9235 | 3.8564 | 0.22363
39 | 12 | B50 09may22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2870 | 5990 |  | 2048 | 2.617 | 5.353 | 2.737 |  | 0.098 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 0:57 |  | 1 |  | 3.0488 | 6.9318 | 3.8829 | 0.22468
40 | 13 | F 17may22 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3763 | 7725 |  | 1808 | 3.302 | 6.823 | 3.521 |  | 0.073 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 1:16 |  | 1 |  | 3.9295 | 8.722 | 4.7925 | 0.20172
41 | 13 | F 17may22 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4066 | 7691 |  | 1786 | 3.535 | 6.794 | 3.26 |  | 0.071 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 1:26 |  | 1 |  | 4.2283 | 8.6869 | 4.4586 | 0.19961 |  |  | 0.74357 |  |  |  |  | 0.28547 |  |  |  |  | 1.2712 |  |  |  |  | 2.3273 |  |  | 4.2441 | 8.6745 | 4.4304 | 0.19732
42 | 13 | F 17may22 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4098 | 7667 |  | 1738 | 3.559 | 6.774 | 3.215 |  | 0.066 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 1:37 |  | 1 |  | 4.2598 | 8.6621 | 4.4023 | 0.19502
43 | 14 | B50 13jun22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3070 | 8811 |  | 3787 | 2.77 | 7.744 | 4.974 |  | 0.28 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 1:56 |  | 1 |  | 3.2461 | 9.8425 | 6.5964 | 0.39105
44 | 14 | B50 13jun22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2763 | 8859 |  | 3898 | 2.534 | 7.784 | 5.249 |  | 0.292 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 2:06 |  | 1 |  | 2.9433 | 9.892 | 6.9487 | 0.40166 |  |  | 1.2814 |  |  |  |  | 0.38668 |  |  |  |  | 0.010108 |  |  |  |  | 1.9966 |  |  | 2.9246 | 9.8729 | 6.9484 | 0.39769
45 | 14 | B50 13jun22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2725 | 8822 |  | 3815 | 2.505 | 7.752 | 5.247 |  | 0.283 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 2:17 |  | 1 |  | 2.9058 | 9.8538 | 6.948 | 0.39372
46 | 15 | B 12apr22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2841 | 6592 |  | 2164 | 2.595 | 5.863 | 3.269 |  | 0.11 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 2:36 |  | 1 |  | 3.0202 | 7.5529 | 4.5327 | 0.23578
47 | 15 | B 12apr22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2895 | 6497 |  | 2077 | 2.636 | 5.782 | 3.146 |  | 0.101 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 2:46 |  | 1 |  | 3.0735 | 7.4549 | 4.3814 | 0.22745 |  |  | 0.35358 |  |  |  |  | 0.44192 |  |  |  |  | 0.99619 |  |  |  |  | 1.0044 |  |  | 3.0681 | 7.4714 | 4.4033 | 0.2286
48 | 15 | B 12apr22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2884 | 6529 |  | 2101 | 2.628 | 5.81 | 3.182 |  | 0.104 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 2:57 |  | 1 |  | 3.0626 | 7.4879 | 4.4253 | 0.22975
49 | 16 | B 25jan22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3087 | 5936 |  | 1912 | 2.783 | 5.308 | 2.524 |  | 0.084 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 3:15 |  | 1 |  | 3.2628 | 6.876 | 3.6132 | 0.21167
50 | 16 | B 25jan22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3199 | 5934 |  | 1899 | 2.869 | 5.306 | 2.436 |  | 0.083 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 3:25 |  | 1 |  | 3.3733 | 6.874 | 3.5007 | 0.21042 |  |  | 0.14628 |  |  |  |  | 0 |  |  |  |  | 0.14075 |  |  |  |  | 1.1751 |  |  | 3.3708 | 6.874 | 3.5032 | 0.21167
51 | 16 | B 25jan22 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3194 | 5934 |  | 1925 | 2.865 | 5.306 | 2.441 |  | 0.085 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 3:36 |  | 1 |  | 3.3683 | 6.874 | 3.5056 | 0.21291
52 | 17 | B50 13jun22 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3814 | 7442 |  | 10910 | 3.341 | 6.583 | 3.242 |  | 1.025 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 3:54 |  | 1 |  | 3.9798 | 8.4299 | 4.4502 | 1.0725
53 | 17 | B50 13jun22 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4126 | 7371 |  | 10968 | 3.58 | 6.523 | 2.943 |  | 1.031 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 4:05 |  | 1 |  | 4.2875 | 8.3567 | 4.0692 | 1.078 |  |  | 0.57339 |  |  |  |  | 0.47029 |  |  |  |  | 1.5818 |  |  |  |  | 0.25702 |  |  | 4.2998 | 8.3371 | 4.0373 | 1.0794
54 | 17 | B50 13jun22 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4151 | 7333 |  | 10997 | 3.6 | 6.491 | 2.892 |  | 1.034 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 4:15 |  | 1 |  | 4.3121 | 8.3175 | 4.0054 | 1.0808
55 | 18 | B 12apr22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3295 | 6221 |  | 2003 | 2.943 | 5.549 | 2.606 |  | 0.093 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 4:34 |  | 1 |  | 3.4679 | 7.1701 | 3.7022 | 0.22037
56 | 18 | B 12apr22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2954 | 6053 |  | 1963 | 2.681 | 5.407 | 2.725 |  | 0.089 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 4:44 |  | 1 |  | 3.1317 | 6.9968 | 3.8651 | 0.21655 |  |  | 0.65912 |  |  |  |  | 0.6028 |  |  |  |  | 0.55715 |  |  |  |  | 1.738 |  |  | 3.142 | 7.0179 | 3.8759 | 0.21468
57 | 18 | B 12apr22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2975 | 6094 |  | 1924 | 2.697 | 5.442 | 2.744 |  | 0.085 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 4:55 |  | 1 |  | 3.1524 | 7.0391 | 3.8867 | 0.21281
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5767 | 13296 |  | 4299 | 4.839 | 11.542 | 6.703 |  | 0.334 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 5:14 |  | 1 |  | 5.9058 | 14.4702 | 8.5644 | 0.44003
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6885 | 13292 |  | 4232 | 5.697 | 11.54 | 5.842 |  | 0.327 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 5:25 |  | 1 |  | 7.0083 | 14.466 | 7.4577 | 0.43362 |  |  | 0.25297 |  | 94.5425 |  |  | 0.60094 |  | 97.9669 |  |  | 1.4101 |  | 101.3912 |  |  | 2.0742 |  | 81.5961 | 7.0172 | 14.4227 | 7.4055 | 0.43816
60 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6903 | 13208 |  | 4327 | 5.711 | 11.468 | 5.757 |  | 0.337 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 5:36 |  | 1 |  | 7.0261 | 14.3794 | 7.3533 | 0.44271
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4132 | 6443 |  | 1969 | 3.585 | 5.737 | 2.152 |  | 0.09 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 5:55 |  | 1 |  | 4.2934 | 7.3992 | 3.1058 | 0.21712
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3039 | 6460 |  | 1953 | 2.746 | 5.751 | 3.005 |  | 0.088 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 6:05 |  | 1 |  | 3.2155 | 7.4167 | 4.2012 | 0.21559 |  |  | 0.06132 | 2.3419 |  |  |  | 1.1895 | 4.9329 |  |  |  | 2.1575 | 6.984 |  |  |  | 3.3386 | 1.233 |  | 3.2165 | 7.3729 | 4.1564 | 0.21205
63 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3041 | 6375 |  | 1879 | 2.748 | 5.679 | 2.931 |  | 0.081 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 6:16 |  | 1 |  | 3.2175 | 7.329 | 4.1115 | 0.20851
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1063 | 471 |  | 327 | 1.23 | 0.677 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 6:34 |  | 1 |  | 1.2668 | 1.2372 | -0.029598 | 0.060033
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 266 | 442 |  | 336 | 0.619 | 0.653 | 0.034 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 6:42 |  | 1 |  | 0.48083 | 1.2073 | 0.72646 | 0.060894 |  |  | 2.0722 |  |  |  |  | 1.1049 |  |  |  |  | 3.1534 |  |  |  |  | 1.7134 |  |  | 0.4759 | 1.214 | 0.7381 | 0.06142
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 256 | 455 |  | 347 | 0.612 | 0.664 | 0.053 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 6:52 |  | 1 |  | 0.47096 | 1.2207 | 0.74974 | 0.061946
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 1318 | 10881 |  | 6258 | 2.377 | 15.828 | 13.452 |  | 0.898 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 7:10 |  | 1 |  | 2.5305 | 19.9639 | 17.4334 | 1.0457
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4241 | 11265 |  | 6443 | 6.114 | 16.37 | 10.256 |  | 0.93 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 7:20 |  | 1 |  | 7.3348 | 20.6243 | 13.2895 | 1.0752 | 89.0667 |  | 16.6972 |  |  | 95.5522 |  | 1.3938 |  |  | 81.312 |  | 12.9961 |  |  | 98.4153 |  | 2.3252 |  |  | 8.0029 | 20.4815 | 12.4786 | 1.0629
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5054 | 11099 |  | 6288 | 7.154 | 16.135 | 8.981 |  | 0.903 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 7:31 |  | 1 |  | 8.671 | 20.3388 | 11.6677 | 1.0505
70 | 21 | F 29mar22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5361 | 8044 |  | 2547 | 4.528 | 7.094 | 2.565 |  | 0.15 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 7:49 |  | 1 |  | 5.5054 | 9.0511 | 3.5457 | 0.27242
71 | 21 | F 29mar22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4299 | 8012 |  | 2510 | 3.713 | 7.066 | 3.353 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 8:00 |  | 1 |  | 4.4581 | 9.0181 | 4.56 | 0.26888 |  |  | 0.26581 |  |  |  |  | 0.30845 |  |  |  |  | 0.86668 |  |  |  |  | 1.3252 |  |  | 4.4521 | 9.032 | 4.5799 | 0.26711
72 | 21 | F 29mar22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4287 | 8039 |  | 2473 | 3.704 | 7.089 | 3.385 |  | 0.143 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 8:11 |  | 1 |  | 4.4462 | 9.0459 | 4.5997 | 0.26534
73 | 22 | F 29mar22 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4207 | 7274 |  | 1965 | 3.642 | 6.441 | 2.799 |  | 0.09 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 8:30 |  | 1 |  | 4.3673 | 8.2566 | 3.8893 | 0.21674
74 | 22 | F 29mar22 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4221 | 7315 |  | 1973 | 3.653 | 6.476 | 2.823 |  | 0.09 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 8:40 |  | 1 |  | 4.3811 | 8.2989 | 3.9178 | 0.2175 |  |  | 0.72291 |  |  |  |  | 1.2007 |  |  |  |  | 1.7378 |  |  |  |  | 2.5841 |  |  | 4.3654 | 8.2494 | 3.884 | 0.21473
75 | 22 | F 29mar22 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4189 | 7219 |  | 1915 | 3.629 | 6.394 | 2.766 |  | 0.084 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 8:51 |  | 1 |  | 4.3496 | 8.1999 | 3.8503 | 0.21195
76 | 23 | F 31jan22 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4515 | 8067 |  | 2465 | 3.879 | 7.113 | 3.234 |  | 0.142 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 9:10 |  | 1 |  | 4.6711 | 9.0748 | 4.4038 | 0.26457
77 | 23 | F 31jan22 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4575 | 8021 |  | 2496 | 3.925 | 7.074 | 3.149 |  | 0.145 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 9:21 |  | 1 |  | 4.7302 | 9.0274 | 4.2971 | 0.26754 |  |  | 0.91313 |  |  |  |  | 0.30908 |  |  |  |  | 1.672 |  |  |  |  | 1.5495 |  |  | 4.7519 | 9.0134 | 4.2615 | 0.26548
78 | 23 | F 31jan22 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4619 | 7994 |  | 2453 | 3.958 | 7.051 | 3.092 |  | 0.141 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 9:32 |  | 1 |  | 4.7736 | 8.9995 | 4.2259 | 0.26342
79 | 24 | B50 17may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3731 | 5890 |  | 1695 | 3.278 | 5.269 | 1.991 |  | 0.061 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 9:51 |  | 1 |  | 3.8979 | 6.8286 | 2.9307 | 0.19091
80 | 24 | B50 17may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3441 | 5886 |  | 1680 | 3.055 | 5.265 | 2.21 |  | 0.06 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 10:01 |  | 1 |  | 3.6119 | 6.8245 | 3.2125 | 0.18947 |  |  | 1.1132 |  |  |  |  | 0.66304 |  |  |  |  | 0.15448 |  |  |  |  | 3.7165 |  |  | 3.6321 | 6.8472 | 3.215 | 0.19306
81 | 24 | B50 17may22 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3482 | 5930 |  | 1755 | 3.086 | 5.302 | 2.216 |  | 0.067 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 10:12 |  | 1 |  | 3.6524 | 6.8699 | 3.2175 | 0.19665
82 | 25 | F 31jan22 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4622 | 8107 |  | 4048 | 3.961 | 7.147 | 3.186 |  | 0.307 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 10:30 |  | 1 |  | 4.7766 | 9.1161 | 4.3395 | 0.41601
83 | 25 | F 31jan22 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5056 | 8106 |  | 4004 | 4.294 | 7.146 | 2.852 |  | 0.303 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 10:41 |  | 1 |  | 5.2046 | 9.1151 | 3.9105 | 0.41181 |  |  | 0.30363 |  |  |  |  | 0.93255 |  |  |  |  | 1.7758 |  |  |  |  | 0.99399 |  |  | 5.1967 | 9.0728 | 3.8761 | 0.41386
84 | 25 | F 31jan22 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5040 | 8024 |  | 4047 | 4.282 | 7.076 | 2.794 |  | 0.307 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 10:52 |  | 1 |  | 5.1888 | 9.0305 | 3.8416 | 0.41592
85 | 26 | F50 23may22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5131 | 8109 |  | 1767 | 4.352 | 7.149 | 2.797 |  | 0.069 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 11:11 |  | 1 |  | 5.2786 | 9.1182 | 3.8396 | 0.1978
86 | 26 | F50 23may22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5241 | 8076 |  | 1731 | 4.436 | 7.121 | 2.685 |  | 0.065 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 11:21 |  | 1 |  | 5.387 | 9.0841 | 3.6971 | 0.19435 |  |  | 0.073199 |  |  |  |  | 0.50102 |  |  |  |  | 1.3437 |  |  |  |  | 1.1884 |  |  | 5.389 | 9.0614 | 3.6724 | 0.1932
87 | 26 | F50 23may22 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5245 | 8032 |  | 1707 | 4.439 | 7.083 | 2.644 |  | 0.063 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 11:32 |  | 1 |  | 5.391 | 9.0387 | 3.6477 | 0.19205
88 | 27 | F 29mar22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4524 | 7419 |  | 1945 | 3.886 | 6.564 | 2.678 |  | 0.087 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 11:51 |  | 1 |  | 4.68 | 8.4062 | 3.7263 | 0.21482
89 | 27 | F 29mar22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4314 | 7400 |  | 1941 | 3.724 | 6.548 | 2.824 |  | 0.087 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 12:01 |  | 1 |  | 4.4729 | 8.3866 | 3.9138 | 0.21444 |  |  | 1.1618 |  |  |  |  | 0.62943 |  |  |  |  | 2.7164 |  |  |  |  | 1.7098 |  |  | 4.499 | 8.3603 | 3.8613 | 0.21262
90 | 27 | F 29mar22 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4367 | 7349 |  | 1903 | 3.765 | 6.505 | 2.74 |  | 0.083 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 12:12 |  | 1 |  | 4.5251 | 8.334 | 3.8089 | 0.21081
91 | 28 | B 22mar 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3396 | 5922 |  | 2597 | 3.02 | 5.296 | 2.276 |  | 0.156 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 12:31 |  | 1 |  | 3.5675 | 6.8616 | 3.294 | 0.2772
92 | 28 | B 22mar 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3091 | 5927 |  | 2636 | 2.786 | 5.3 | 2.514 |  | 0.16 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 12:41 |  | 1 |  | 3.2668 | 6.8668 | 3.6 | 0.28093 |  |  | 0.72715 |  |  |  |  | 1.5292 |  |  |  |  | 2.2627 |  |  |  |  | 0.51212 |  |  | 3.2549 | 6.8147 | 3.5597 | 0.28021
93 | 28 | B 22mar 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3067 | 5826 |  | 2621 | 2.768 | 5.214 | 2.447 |  | 0.158 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 12:52 |  | 1 |  | 3.2431 | 6.7625 | 3.5194 | 0.2795
94 | 29 | B50 17may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2920 | 5813 |  | 1782 | 2.655 | 5.203 | 2.548 |  | 0.07 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 13:10 |  | 1 |  | 3.0981 | 6.7491 | 3.651 | 0.19923
95 | 29 | B50 17may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2900 | 5745 |  | 1718 | 2.639 | 5.146 | 2.506 |  | 0.064 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 13:20 |  | 1 |  | 3.0784 | 6.679 | 3.6006 | 0.19311 |  |  | 1.3049 |  |  |  |  | 0.17008 |  |  |  |  | 1.4486 |  |  |  |  | 0.89574 |  |  | 3.0986 | 6.6733 | 3.5747 | 0.19225
96 | 29 | B50 17may22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2941 | 5734 |  | 1700 | 2.671 | 5.136 | 2.465 |  | 0.062 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 13:31 |  | 1 |  | 3.1188 | 6.6676 | 3.5488 | 0.19139
97 | 30 | F100 07jun22 weir | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3154 | 5492 |  | 1687 | 2.834 | 4.931 | 2.097 |  | 0.06 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 13:50 |  | 1 |  | 3.3289 | 6.4179 | 3.089 | 0.19014
98 | 30 | F100 07jun22 weir | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3265 | 5433 |  | 1679 | 2.92 | 4.882 | 1.962 |  | 0.06 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 14:00 |  | 1 |  | 3.4384 | 6.357 | 2.9187 | 0.18938 |  |  | 0.62901 |  |  |  |  | 0.01623 |  |  |  |  | 0.71051 |  |  |  |  | 5.0714 |  |  | 3.4492 | 6.3576 | 2.9084 | 0.1943
99 | 30 | F100 07jun22 weir | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3287 | 5434 |  | 1782 | 2.937 | 4.882 | 1.946 |  | 0.07 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 14:10 |  | 1 |  | 3.4601 | 6.3581 | 2.898 | 0.19923
100 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7049 | 14397 |  | 6562 | 5.823 | 12.475 | 6.653 |  | 0.57 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 14:30 |  | 1 |  | 7.17 | 15.6062 | 8.4361 | 0.65652
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8588 | 14345 |  | 6556 | 7.004 | 12.432 | 5.428 |  | 0.57 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 14:41 |  | 1 |  | 8.6878 | 15.5525 | 6.8648 | 0.65595 |  |  | 0.48692 |  | 119.6879 |  |  | 0.6189 |  | 109.522 |  |  | 2.0361 |  | 99.3561 |  |  | 0.058322 |  | 82.7284 | 8.709 | 15.5045 | 6.7956 | 0.65614
102 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8631 | 14252 |  | 6560 | 7.036 | 12.353 | 5.317 |  | 0.57 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 14:52 |  | 1 |  | 8.7302 | 15.4566 | 6.7264 | 0.65633
103 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4877 | 5841 |  | 1884 | 4.156 | 5.227 | 1.071 |  | 0.081 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 15:11 |  | 1 |  | 5.0281 | 6.778 | 1.7499 | 0.20899
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3491 | 5799 |  | 1932 | 3.093 | 5.192 | 2.099 |  | 0.086 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 15:22 |  | 1 |  | 3.6612 | 6.7347 | 3.0734 | 0.21358 |  |  | 0.8385 | 5.5458 |  |  |  | 0.44332 | 5.9828 |  |  |  | 1.9491 | 6.4985 |  |  |  | 0.584 | 9.1613 |  | 3.646 | 6.7496 | 3.1037 | 0.21296
105 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3460 | 5828 |  | 1919 | 3.069 | 5.216 | 2.147 |  | 0.085 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 15:32 |  | 1 |  | 3.6307 | 6.7646 | 3.1339 | 0.21234

## 08sep22
RUN NOTES:
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area |  | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | -183.5 | 0 | -498.5 | 0 | -266 |  | 0.5 | -0.027616 | 0.13254 | 0.16016 | -0.008749
TIC | TC | TNb |  | 0.0006 | 1130.4 | 0.0012 | 1904.9 | 6e-05 | 1028.4 |  | 0.2 | 3.1496 | 6.1908 | 3.0412 | 0.31275
501.5 | 1001.5 | 526.5 |  | 0.0018 | 3493.2 | 0.0036 | 6792.2 | 0.00018 | 3210.7 |  | 0.6 | 2.9793 | 6.0354 | 3.0561 | 0.2923
|  |  |  | 0.002997 | 5879.501 | 0.005994 | 11111.001 | 0.0002997 | 6014.351 |  | 0.333 | 8.879 | 17.1984 | 8.3193 | 0.96199
|  |  |  | 0.004203 | 8561.099 | 0.008406 | 16721.599 | 0.0004203 | 7701.599 |  | 0.467 | 9.1446 | 18.1216 | 8.977 | 0.87276
|  |  |  | 0.0054 | 10803.2 | 0.0108 | 21688.2 | 0.00054 | 10234.2 |  | 0.6 | 8.9484 | 18.1408 | 9.1924 | 0.89754
|  |  | Slope |  | 4.8994e-07 |  | 4.876e-07 |  | 5.1703e-08
|  |  | Intercept |  | 7.6097e-05 |  | 0.00030934 |  | 9.3786e-06
|  |  | RSQ |  | 0.99962 |  | 0.999 |  | 0.99709
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4080 | 11805 |  | 6509 | 5.909 | 17.133 | 11.225 |  | 0.941 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 16:04 |  | 1 |  | 6.9169 | 20.2182 | 13.3013 | 1.1531
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5490 | 11792 |  | 6540 | 7.711 | 17.114 | 9.403 |  | 0.947 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 16:15 |  | 1 |  | 9.2196 | 20.1971 | 10.9774 | 1.1584 |  |  | 2.2246 |  |  |  |  | 2.0569 |  |  |  |  | 5.7988 |  |  |  |  | 1.5896 |  |  | 9.3233 | 19.9914 | 10.6681 | 1.1493
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5617 | 11539 |  | 6434 | 7.874 | 16.757 | 8.884 |  | 0.928 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 16:25 |  | 1 |  | 9.427 | 19.7858 | 10.3588 | 1.1401
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2997 | 998 |  | 561 | 2.714 | 1.124 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 16:43 |  | 1 |  | 3.0889 | 1.5919 | -1.497 | 0.076768
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 519 | 996 |  | 556 | 0.813 | 1.122 | 0.309 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 16:53 |  | 1 |  | 0.66075 | 1.59 | 0.92922 | 0.076251 |  |  | 5.3287 |  |  |  |  | 0.67241 |  |  |  |  | 4.7306 |  |  |  |  | 8.3346 |  |  | 0.64361 | 1.5953 | 0.95174 | 0.073201
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 484 | 1007 |  | 497 | 0.787 | 1.132 | 0.345 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 17:02 |  | 1 |  | 0.62646 | 1.6007 | 0.97425 | 0.07015
7 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 704 | 2337 |  | 1204 | 2.387 | 5.646 | 3.259 |  | 0.025 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 17:19 |  | 1 |  | 2.1051 | 7.2443 | 5.1392 | 0.35815
8 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1320 | 2353 |  | 1288 | 3.568 | 5.68 | 2.112 |  | 0.047 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 17:29 |  | 1 |  | 3.6141 | 7.2833 | 3.6692 | 0.37986 |  | 21.3685 | 1.4802 |  |  |  | 19.4584 | 3.2314 |  |  |  | 17.5483 | 8.0961 |  |  |  | 22.3986 | 6.8995 |  |  | 3.6411 | 7.1675 | 3.5264 | 0.3672
9 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1342 | 2258 |  | 1190 | 3.611 | 5.478 | 1.867 |  | 0.021 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 17:38 |  | 1 |  | 3.668 | 7.0517 | 3.3837 | 0.35453
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3965 | 7983 |  | 3917 | 2.88 | 5.868 | 2.988 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 17:58 |  | 1 |  | 3.3645 | 7.0031 | 3.6385 | 0.35317
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4059 | 8066 |  | 3875 | 2.941 | 5.926 | 2.986 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 18:08 |  | 1 |  | 3.4413 | 7.0705 | 3.6292 | 0.34955 |  | 15.6896 | 1.694 |  |  |  | 16.867 | 1.6689 |  |  |  | 18.0443 | 4.9647 |  |  |  | 15.5828 | 1.6154 |  |  | 3.4707 | 7.012 | 3.5413 | 0.34675
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4131 | 7922 |  | 3810 | 2.987 | 5.825 | 2.838 |  | 0.235 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 18:20 |  | 1 |  | 3.5001 | 6.9535 | 3.4534 | 0.34395
13 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 4663 | 11919 |  | 6462 | 5.995 | 15.58 | 9.585 |  | 0.841 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 18:39 |  | 1 |  | 7.0892 | 18.3815 | 11.2923 | 1.0315
14 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6145 | 11837 |  | 6327 | 7.701 | 15.476 | 7.774 |  | 0.819 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 18:49 |  | 1 |  | 9.2697 | 18.2614 | 8.9918 | 1.0105 |  | 4.116 | 2.1511 |  |  |  | 0.97245 | 0.95066 |  |  |  | 2.1711 | 4.2518 |  |  |  | 12.9366 | 1.1609 |  |  | 9.3704 | 18.175 | 8.8046 | 1.0164
15 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6282 | 11719 |  | 6403 | 7.859 | 15.325 | 7.466 |  | 0.831 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 19:00 |  | 1 |  | 9.4712 | 18.0887 | 8.6174 | 1.0223
16 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8790 | 17796 |  | 8653 | 7.664 | 16.44 | 8.776 |  | 0.845 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 19:20 |  | 1 |  | 9.3848 | 19.2434 | 9.8586 | 0.97809
17 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 9042 | 17680 |  | 8634 | 7.871 | 16.335 | 8.463 |  | 0.843 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 19:31 |  | 1 |  | 9.6492 | 19.1223 | 9.4731 | 0.97599 |  | 7.0673 | 0.27219 |  |  |  | 6.1014 | 0.25148 |  |  |  | 5.1356 | 0.2304 |  |  |  | 8.48 | 0.06804 |  |  | 9.6361 | 19.0983 | 9.4622 | 0.97632
18 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 9017 | 17634 |  | 8640 | 7.851 | 16.293 | 8.443 |  | 0.843 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 19:42 |  | 1 |  | 9.6229 | 19.0742 | 9.4513 | 0.97665
19 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11329 | 22857 |  | 10772 | 7.588 | 16.369 | 8.781 |  | 0.842 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 20:03 |  | 1 |  | 9.3778 | 19.0907 | 9.7129 | 0.94388
20 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11424 | 22896 |  | 10897 | 7.649 | 16.396 | 8.747 |  | 0.853 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 20:15 |  | 1 |  | 9.4553 | 19.1224 | 9.667 | 0.95465 |  | 4.887 | 0.32871 |  |  |  | 6.2083 | 0.051011 |  |  |  | 7.5296 | 0.21986 |  |  |  | 5.7756 | 0.56122 |  |  | 9.4398 | 19.1175 | 9.6777 | 0.95198
21 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11386 | 22884 |  | 10835 | 7.625 | 16.388 | 8.763 |  | 0.848 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 20:27 |  | 1 |  | 9.4243 | 19.1126 | 9.6883 | 0.94931
22 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5436 | 11813 |  | 6235 | 7.643 | 17.144 | 9.501 |  | 0.894 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 20:46 |  | 1 |  | 9.1314 | 20.2312 | 11.0998 | 1.1058
23 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5318 | 11834 |  | 6169 | 7.492 | 17.174 | 9.682 |  | 0.882 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 20:56 |  | 1 |  | 8.9387 | 20.2653 | 11.3266 | 1.0945 | 100 |  | 0.091394 |  |  | 100 |  | 1.0967 |  |  | 100 |  | 1.8973 |  |  | 100 |  | 0.55267 |  |  | 8.9346 | 20.1548 | 11.2202 | 1.0914
24 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5313 | 11698 |  | 6134 | 7.485 | 16.982 | 9.497 |  | 0.876 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:08 |  | 1 |  | 8.9305 | 20.0443 | 11.1137 | 1.0884
25 | 2 | Spiked Blank 100ml + 300uL NOT COVERED | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5148 | 7473 |  | 3415 | 4.365 | 6.61 | 2.245 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:27 |  | 1 |  | 5.1966 | 7.9063 | 2.7097 | 0.37189
26 | 2 | Spiked Blank 100ml + 300uL NOT COVERED | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3783 | 7443 |  | 3416 | 3.317 | 6.584 | 3.266 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:37 |  | 1 |  | 3.8591 | 7.8771 | 4.018 | 0.372 |  |  | 1.8192 |  |  |  |  | 0.62093 |  |  |  |  | 0.51663 |  |  |  |  | 0.83047 |  |  | 3.8243 | 7.8527 | 4.0284 | 0.37355
27 | 2 | Spiked Blank 100ml + 300uL NOT COVERED | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3712 | 7393 |  | 3446 | 3.263 | 6.542 | 3.279 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:48 |  | 1 |  | 3.7895 | 7.8283 | 4.0388 | 0.3751
28 | 9 | F 29mar 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4111 | 7560 |  | 2001 | 3.569 | 6.684 | 3.115 |  | 0.093 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 22:07 |  | 1 |  | 4.1805 | 7.9912 | 3.8107 | 0.22567
29 | 9 | F 29mar 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4250 | 7491 |  | 1986 | 3.675 | 6.625 | 2.95 |  | 0.092 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 22:17 |  | 1 |  | 4.3167 | 7.9239 | 3.6072 | 0.22412 |  |  | 0.27203 |  |  |  |  | 0.049216 |  |  |  |  | 0.21808 |  |  |  |  | 1.9097 |  |  | 4.3226 | 7.9258 | 3.6033 | 0.222
30 | 9 | F 29mar 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4262 | 7495 |  | 1945 | 3.685 | 6.628 | 2.944 |  | 0.087 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 22:28 |  | 1 |  | 4.3285 | 7.9278 | 3.5993 | 0.21988
31 | 10 | F 31jan 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5071 | 8629 |  | 5017 | 4.305 | 7.589 | 3.284 |  | 0.409 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 22:47 |  | 1 |  | 5.1212 | 9.0337 | 3.9125 | 0.53755
32 | 10 | F 31jan 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5387 | 8586 |  | 5013 | 4.547 | 7.552 | 3.005 |  | 0.408 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 22:58 |  | 1 |  | 5.4308 | 8.9917 | 3.5609 | 0.53714 |  |  | 0.16252 |  |  |  |  | 1.497 |  |  |  |  | 3.5668 |  |  |  |  | 0.28919 |  |  | 5.4264 | 8.9249 | 3.4985 | 0.53636
33 | 10 | F 31jan 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5378 | 8449 |  | 4998 | 4.541 | 7.436 | 2.896 |  | 0.407 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 23:09 |  | 1 |  | 5.422 | 8.8581 | 3.4361 | 0.53558
34 | 11 | F 17may 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4654 | 8675 |  | 2649 | 3.985 | 7.628 | 3.643 |  | 0.161 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 23:28 |  | 1 |  | 4.7126 | 9.0785 | 4.366 | 0.29268
35 | 11 | F 17may 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4434 | 8680 |  | 2672 | 3.817 | 7.632 | 3.815 |  | 0.163 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 23:38 |  | 1 |  | 4.497 | 9.0834 | 4.5864 | 0.29506 |  |  | 0.8462 |  |  |  |  | 1.0468 |  |  |  |  | 2.9383 |  |  |  |  | 3.6034 |  |  | 4.5161 | 9.0361 | 4.52 | 0.28984
36 | 11 | F 17may 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4473 | 8583 |  | 2571 | 3.846 | 7.55 | 3.704 |  | 0.153 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 23:49 |  | 1 |  | 4.5352 | 8.9888 | 4.4536 | 0.28462
37 | 12 | C 20apr HPB | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2251 | 3601 |  | 1112 | 2.142 | 3.329 | 1.188 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 0:07 |  | 1 |  | 2.3579 | 4.1304 | 1.7725 | 0.13375
38 | 12 | C 20apr HPB | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1340 | 3533 |  | 1105 | 1.443 | 3.272 | 1.829 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 0:17 |  | 1 |  | 1.4652 | 4.0641 | 2.5988 | 0.13302 |  |  | 1.3946 |  |  |  |  | 0.5981 |  |  |  |  | 0.1462 |  |  |  |  | 0.46751 |  |  | 1.4755 | 4.0762 | 2.6007 | 0.13271
39 | 12 | C 20apr HPB | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1361 | 3558 |  | 1099 | 1.459 | 3.293 | 1.834 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 0:28 |  | 1 |  | 1.4858 | 4.0884 | 2.6026 | 0.1324
40 | 13 | F 17may wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3956 | 7527 |  | 1619 | 3.45 | 6.656 | 3.206 |  | 0.053 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 0:46 |  | 1 |  | 4.0286 | 7.959 | 3.9304 | 0.18617
41 | 13 | F 17may wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4971 | 7523 |  | 1623 | 4.228 | 6.652 | 2.424 |  | 0.054 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 0:57 |  | 1 |  | 5.0232 | 7.9551 | 2.9319 | 0.18659 |  |  | 1.9127 |  |  |  |  | 1.2831 |  |  |  |  | 7.005 |  |  |  |  | 1.2828 |  |  | 5.0717 | 7.9044 | 2.8327 | 0.1854
42 | 13 | F 17may wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5070 | 7419 |  | 1600 | 4.304 | 6.563 | 2.259 |  | 0.051 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 1:08 |  | 1 |  | 5.1202 | 7.8537 | 2.7335 | 0.18421
43 | 14 | F50 31may 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3978 | 7885 |  | 2776 | 3.467 | 6.959 | 3.492 |  | 0.174 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 1:27 |  | 1 |  | 4.0502 | 8.3081 | 4.2579 | 0.30581
44 | 14 | F50 31may 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3636 | 7843 |  | 2744 | 3.205 | 6.923 | 3.718 |  | 0.171 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 1:37 |  | 1 |  | 3.7151 | 8.2672 | 4.5521 | 0.30251 |  |  | 1.4662 |  |  |  |  | 1.0673 |  |  |  |  | 3.1834 |  |  |  |  | 4.7238 |  |  | 3.7425 | 8.2233 | 4.4808 | 0.29553
45 | 14 | F50 31may 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3692 | 7753 |  | 2609 | 3.247 | 6.847 | 3.599 |  | 0.157 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 1:48 |  | 1 |  | 3.7699 | 8.1794 | 4.4095 | 0.28855
46 | 15 | F 17may 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4577 | 7450 |  | 1802 | 3.926 | 6.59 | 2.664 |  | 0.072 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 2:07 |  | 1 |  | 4.6371 | 7.8839 | 3.2468 | 0.2051
47 | 15 | F 17may 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4945 | 7415 |  | 1723 | 4.209 | 6.56 | 2.351 |  | 0.064 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 2:18 |  | 1 |  | 4.9977 | 7.8498 | 2.8521 | 0.19693 |  |  | 0.098081 |  |  |  |  | 0.91103 |  |  |  |  | 2.3516 |  |  |  |  | 0.31556 |  |  | 4.9953 | 7.8142 | 2.8189 | 0.19662
48 | 15 | F 17may 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4940 | 7342 |  | 1717 | 4.205 | 6.498 | 2.294 |  | 0.064 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 2:29 |  | 1 |  | 4.9928 | 7.7786 | 2.7858 | 0.19631
49 | 16 | B50 31may 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3426 | 8212 |  | 3947 | 3.043 | 7.236 | 4.193 |  | 0.297 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 2:48 |  | 1 |  | 3.5093 | 8.627 | 5.1177 | 0.4269
50 | 16 | B50 31may 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2759 | 8068 |  | 3967 | 2.532 | 7.113 | 4.582 |  | 0.299 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 2:58 |  | 1 |  | 2.8557 | 8.4866 | 5.6309 | 0.42897 |  |  | 1.126 |  |  |  |  | 0.98337 |  |  |  |  | 0.91098 |  |  |  |  | 0.26552 |  |  | 2.8719 | 8.5285 | 5.6567 | 0.4284
51 | 16 | B50 31may 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2792 | 8154 |  | 3956 | 2.557 | 7.187 | 4.63 |  | 0.298 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 3:09 |  | 1 |  | 2.888 | 8.5705 | 5.6824 | 0.42783
52 | 17 | B50 12may 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2925 | 6001 |  | 1920 | 2.659 | 5.363 | 2.703 |  | 0.085 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 3:28 |  | 1 |  | 3.0184 | 6.4709 | 3.4525 | 0.2173
53 | 17 | B50 12may 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2867 | 5963 |  | 1892 | 2.615 | 5.331 | 2.716 |  | 0.082 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 3:38 |  | 1 |  | 2.9615 | 6.4338 | 3.4723 | 0.2144 |  |  | 0.49508 |  |  |  |  | 0.57433 |  |  |  |  | 0.64187 |  |  |  |  | 1.3414 |  |  | 2.9689 | 6.4523 | 3.4834 | 0.21585
54 | 17 | B50 12may 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2882 | 6001 |  | 1920 | 2.626 | 5.362 | 2.736 |  | 0.085 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 3:49 |  | 1 |  | 2.9762 | 6.4709 | 3.4946 | 0.2173
55 | 18 | F 24feb 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3781 | 7430 |  | 2273 | 3.316 | 6.573 | 3.257 |  | 0.122 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 4:07 |  | 1 |  | 3.8571 | 7.8644 | 4.0073 | 0.2538
56 | 18 | F 24feb 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4170 | 7370 |  | 2289 | 3.614 | 6.522 | 2.908 |  | 0.123 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 4:18 |  | 1 |  | 4.2383 | 7.8059 | 3.5676 | 0.25546 |  |  | 0.62619 |  |  |  |  | 1.0676 |  |  |  |  | 1.5945 |  |  |  |  | 2.4586 |  |  | 4.2251 | 7.7645 | 3.5394 | 0.25235
57 | 18 | F 24feb 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4143 | 7285 |  | 2229 | 3.593 | 6.45 | 2.857 |  | 0.117 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 4:28 |  | 1 |  | 4.2119 | 7.723 | 3.5111 | 0.24925
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7376 | 13876 |  | 4272 | 6.074 | 12.034 | 5.96 |  | 0.331 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 4:48 |  | 1 |  | 7.3798 | 14.1505 | 6.7707 | 0.46051
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8615 | 13793 |  | 4154 | 7.024 | 11.964 | 4.94 |  | 0.318 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 4:59 |  | 1 |  | 8.5939 | 14.0696 | 5.4757 | 0.44831 |  |  | 1.9868 |  | 122.8851 |  |  | 0.013863 |  | 104.8476 |  |  | 3.2367 |  | 86.8101 |  |  | 0.41605 |  | 88.6693 | 8.6801 | 14.0686 | 5.3885 | 0.44738
60 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8791 | 13791 |  | 4136 | 7.159 | 11.962 | 4.803 |  | 0.317 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 5:10 |  | 1 |  | 8.7664 | 14.0677 | 5.3013 | 0.44645
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5590 | 7819 |  | 2354 | 4.703 | 6.902 | 2.199 |  | 0.13 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 5:30 |  | 1 |  | 5.6298 | 8.2438 | 2.614 | 0.26218
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4495 | 7820 |  | 2279 | 3.864 | 6.903 | 3.04 |  | 0.122 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 5:40 |  | 1 |  | 4.5568 | 8.2447 | 3.688 | 0.25442 |  |  | 0.8205 | 7.1451 |  |  |  | 0.84334 | 5.5797 |  |  |  | 0.87157 | 3.6773 |  |  |  | 0.57064 | 0.53129 |  | 4.5382 | 8.2101 | 3.672 | 0.2537
63 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4457 | 7749 |  | 2265 | 3.834 | 6.843 | 3.009 |  | 0.121 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 5:51 |  | 1 |  | 4.5195 | 8.1755 | 3.656 | 0.25297
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1474 | 495 |  | 354 | 1.546 | 0.698 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:09 |  | 1 |  | 1.5965 | 1.1014 | -0.49514 | 0.055363
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 322 | 480 |  | 240 | 0.662 | 0.685 | 0.023 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:18 |  | 1 |  | 0.46772 | 1.0868 | 0.61906 | 0.043575 |  |  | 1.6902 |  |  |  |  | 4.0443 |  |  |  |  | 8.1651 |  |  |  |  | 9.2783 |  |  | 0.4638 | 1.1092 | 0.64541 | 0.045695
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 314 | 526 |  | 281 | 0.656 | 0.724 | 0.068 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:27 |  | 1 |  | 0.45988 | 1.1316 | 0.67176 | 0.047815
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2781 | 11134 |  | 5855 | 4.247 | 16.186 | 11.939 |  | 0.827 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:46 |  | 1 |  | 4.7954 | 19.1276 | 14.3322 | 1.0403
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4683 | 11098 |  | 5984 | 6.68 | 16.134 | 9.455 |  | 0.85 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:56 |  | 1 |  | 7.9017 | 19.0691 | 11.1674 | 1.0626 | 91.0733 |  | 6.3247 |  |  | 93.63 |  | 1.399 |  |  | 80.7706 |  | 7.247 |  |  | 96.4074 |  | 1.7507 |  |  | 8.1597 | 18.9366 | 10.7769 | 1.0534
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4999 | 10935 |  | 5877 | 7.083 | 15.905 | 8.821 |  | 0.831 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 7:07 |  | 1 |  | 8.4177 | 18.8041 | 10.3864 | 1.0441
70 | 21 | F50 13jun 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6507 | 9738 |  | 3076 | 5.407 | 8.528 | 3.121 |  | 0.206 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 7:26 |  | 1 |  | 6.5283 | 10.1152 | 3.5869 | 0.33684
71 | 21 | F50 13jun 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5918 | 9555 |  | 2992 | 4.955 | 8.373 | 3.418 |  | 0.197 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 7:37 |  | 1 |  | 5.9512 | 9.9367 | 3.9856 | 0.32815 |  |  | 0.78723 |  |  |  |  | 0.75855 |  |  |  |  | 3.1119 |  |  |  |  | 1.5241 |  |  | 5.9747 | 9.8992 | 3.9245 | 0.32567
72 | 21 | F50 13jun 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5966 | 9478 |  | 2944 | 4.992 | 8.308 | 3.317 |  | 0.192 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 7:48 |  | 1 |  | 5.9982 | 9.8616 | 3.8634 | 0.32319
73 | 22 | F 24feb 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4632 | 7385 |  | 2009 | 3.968 | 6.535 | 2.567 |  | 0.094 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 8:07 |  | 1 |  | 4.691 | 7.8205 | 3.1295 | 0.2265
74 | 22 | F 24feb 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4126 | 7394 |  | 2002 | 3.581 | 6.542 | 2.962 |  | 0.093 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 8:17 |  | 1 |  | 4.1952 | 7.8293 | 3.6341 | 0.22578 |  |  | 1.7136 |  |  |  |  | 1.8733 |  |  |  |  | 6.1788 |  |  |  |  | 1.8956 |  |  | 4.2315 | 7.7567 | 3.5252 | 0.22366
75 | 22 | F 24feb 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4200 | 7245 |  | 1961 | 3.637 | 6.417 | 2.78 |  | 0.089 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 8:28 |  | 1 |  | 4.2677 | 7.684 | 3.4163 | 0.22154
76 | 23 | B50 07jun 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3232 | 6851 |  | 2316 | 2.894 | 6.083 | 3.188 |  | 0.126 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 8:47 |  | 1 |  | 3.3192 | 7.2998 | 3.9806 | 0.25825
77 | 23 | B50 07jun 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2913 | 6855 |  | 2422 | 2.65 | 6.086 | 3.436 |  | 0.137 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 8:58 |  | 1 |  | 3.0066 | 7.3037 | 4.2971 | 0.26921 |  |  | 0.45732 |  |  |  |  | 1.1414 |  |  |  |  | 1.6228 |  |  |  |  | 3.9164 |  |  | 2.9997 | 7.2622 | 4.2625 | 0.26404
78 | 23 | B50 07jun 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2899 | 6770 |  | 2322 | 2.639 | 6.014 | 3.375 |  | 0.127 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 9:08 |  | 1 |  | 2.9929 | 7.2208 | 4.2279 | 0.25887
79 | 24 | C 28feb 21.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8211 | 12412 |  | 2021 | 6.715 | 10.794 | 4.079 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 9:28 |  | 1 |  | 8.198 | 12.7229 | 4.5248 | 0.22774
80 | 24 | C 28feb 21.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10267 | 12385 |  | 1985 | 8.292 | 10.771 | 2.479 |  | 0.092 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 9:39 |  | 1 |  | 10.2127 | 12.6965 | 2.4838 | 0.22402 |  |  | 0.33525 |  |  |  |  | 0.63955 |  |  |  |  | 4.7497 |  |  |  |  | 4.7761 |  |  | 10.2298 | 12.6561 | 2.4262 | 0.2295
81 | 24 | C 28feb 21.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10302 | 12302 |  | 2091 | 8.319 | 10.7 | 2.382 |  | 0.103 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 9:50 |  | 1 |  | 10.247 | 12.6156 | 2.3686 | 0.23498
82 | 25 | F 17mar 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6020 | 7416 |  | 1791 | 5.034 | 6.562 | 1.528 |  | 0.071 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 10:09 |  | 1 |  | 6.0511 | 7.8508 | 1.7997 | 0.20396
83 | 25 | F 17mar 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4412 | 7434 |  | 1776 | 3.8 | 6.576 | 2.777 |  | 0.07 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 10:20 |  | 1 |  | 4.4755 | 7.8683 | 3.3929 | 0.20241 |  |  | 0.19725 |  |  |  |  | 0.82136 |  |  |  |  | 1.6506 |  |  |  |  | 2.6924 |  |  | 4.471 | 7.8361 | 3.3651 | 0.19972
84 | 25 | F 17mar 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4403 | 7368 |  | 1724 | 3.793 | 6.52 | 2.728 |  | 0.064 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 10:31 |  | 1 |  | 4.4666 | 7.8039 | 3.3373 | 0.19703
85 | 26 | C 28feb 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8562 | 12502 |  | 2061 | 6.984 | 10.87 | 3.887 |  | 0.1 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 10:51 |  | 1 |  | 8.542 | 12.8106 | 4.2686 | 0.23188
86 | 26 | C 28feb 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10246 | 12385 |  | 2031 | 8.276 | 10.771 | 2.495 |  | 0.096 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 11:02 |  | 1 |  | 10.1921 | 12.6965 | 2.5044 | 0.22878 |  |  | 1.7441 |  |  |  |  | 0.046074 |  |  |  |  | 7.175 |  |  |  |  | 4.7222 |  |  | 10.2818 | 12.6994 | 2.4177 | 0.23431
87 | 26 | C 28feb 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10429 | 12391 |  | 2138 | 8.416 | 10.776 | 2.36 |  | 0.108 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 11:13 |  | 1 |  | 10.3714 | 12.7024 | 2.3309 | 0.23984
88 | 27 | F 17mar weir | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5676 | 5062 |  | 866 | 4.77 | 4.567 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 11:32 |  | 1 |  | 5.714 | 5.5551 | -0.15889 | 0.10831
89 | 27 | F 17mar weir | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3853 | 5020 |  | 781 | 3.371 | 4.531 | 1.161 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 11:42 |  | 1 |  | 3.9277 | 5.5142 | 1.5865 | 0.099518 |  |  | 0.85185 |  |  |  |  | 0.47865 |  |  |  |  | 0.43936 |  |  |  |  | 5.7524 |  |  | 3.911 | 5.501 | 1.59 | 0.10247
90 | 27 | F 17mar weir | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3819 | 4993 |  | 838 | 3.345 | 4.508 | 1.164 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 11:53 |  | 1 |  | 3.8944 | 5.4878 | 1.5935 | 0.10541
91 | 28 | F 12apr 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4036 | 7300 |  | 2216 | 3.512 | 6.463 | 2.951 |  | 0.116 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 12:12 |  | 1 |  | 4.107 | 7.7376 | 3.6306 | 0.24791
92 | 28 | F 12apr 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4164 | 7282 |  | 2243 | 3.609 | 6.447 | 2.838 |  | 0.119 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 12:23 |  | 1 |  | 4.2324 | 7.7201 | 3.4876 | 0.2507 |  |  | 0.50805 |  |  |  |  | 1.7328 |  |  |  |  | 4.5208 |  |  |  |  | 5.3793 |  |  | 4.2432 | 7.6538 | 3.4106 | 0.24413
93 | 28 | F 12apr 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4186 | 7146 |  | 2116 | 3.626 | 6.333 | 2.707 |  | 0.105 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 12:34 |  | 1 |  | 4.254 | 7.5875 | 3.3335 | 0.23757
94 | 29 | C 28feb SMB 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1599 | 1956 |  | 612 | 1.642 | 1.935 | 0.294 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 12:52 |  | 1 |  | 1.719 | 2.5262 | 0.80714 | 0.082042
95 | 29 | C 28feb SMB 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 577 | 1952 |  | 587 | 0.857 | 1.933 | 1.075 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 13:01 |  | 1 |  | 0.71759 | 2.5223 | 1.8047 | 0.079457 |  |  | 4.3247 |  |  |  |  | 0.1935 |  |  |  |  | 1.4031 |  |  |  |  | 4.9499 |  |  | 0.7024 | 2.5198 | 1.8174 | 0.081473
96 | 29 | C 28feb SMB 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 546 | 1947 |  | 626 | 0.834 | 1.928 | 1.094 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 13:11 |  | 1 |  | 0.68721 | 2.5174 | 1.8302 | 0.08349
97 | 30 | F50 07jun 9.,0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4144 | 10254 |  | 2237 | 3.594 | 8.966 | 5.372 |  | 0.118 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 13:30 |  | 1 |  | 4.2128 | 10.6184 | 6.4055 | 0.25008
98 | 30 | F50 07jun 9.,0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5515 | 10366 |  | 2247 | 4.646 | 9.06 | 4.415 |  | 0.119 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 13:41 |  | 1 |  | 5.5563 | 10.7276 | 5.1713 | 0.25111 |  |  | 0.66792 |  |  |  |  | 1.8626 |  |  |  |  | 4.654 |  |  |  |  | 0.454 |  |  | 5.5749 | 10.6286 | 5.0537 | 0.25054
99 | 30 | F50 07jun 9.,0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5553 | 10163 |  | 2236 | 4.675 | 8.888 | 4.213 |  | 0.118 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 13:52 |  | 1 |  | 5.5935 | 10.5296 | 4.9361 | 0.24998
100 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6610 | 12992 |  | 4494 | 5.486 | 11.285 | 5.799 |  | 0.354 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 14:12 |  | 1 |  | 6.6292 | 13.2885 | 6.6592 | 0.48347
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6954 | 12801 |  | 4472 | 5.75 | 11.123 | 5.373 |  | 0.352 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 14:23 |  | 1 |  | 6.9663 | 13.1022 | 6.1359 | 0.48119 |  |  | 0.30897 |  | 85.6284 |  |  | 0.10415 |  | 89.8481 |  |  | 0.12891 |  | 94.0677 |  |  | 1.0475 |  | 96.1202 | 6.9771 | 13.109 | 6.1319 | 0.48373
102 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6976 | 12815 |  | 4521 | 5.767 | 11.135 | 5.369 |  | 0.357 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 14:34 |  | 1 |  | 6.9879 | 13.1159 | 6.128 | 0.48626
103 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6150 | 10684 |  | 2370 | 5.133 | 9.33 | 4.197 |  | 0.132 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 14:53 |  | 1 |  | 6.1785 | 11.0377 | 4.8592 | 0.26383
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5891 | 10646 |  | 2380 | 4.934 | 9.298 | 4.363 |  | 0.133 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 15:04 |  | 1 |  | 5.9247 | 11.0007 | 5.0759 | 0.26487 |  |  | 0.54728 | 5.811 |  |  |  | 0.8815 | 3.0005 |  |  |  | 1.273 | 0.19579 |  |  |  | 1.4154 | 4.8527 |  | 5.9085 | 10.9524 | 5.0438 | 0.263
105 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5858 | 10547 |  | 2344 | 4.909 | 9.213 | 4.304 |  | 0.129 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 15:15 |  | 1 |  | 5.8924 | 10.9041 | 5.0117 | 0.26114
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2025 | 591 |  | 312 | 1.968 | 0.78 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 15:33 |  | 1 |  | 2.1365 | 1.195 | -0.94144 | 0.05102
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 439 | 606 |  | 307 | 0.751 | 0.792 | 0.04 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 15:42 |  | 1 |  | 0.58236 | 1.2096 | 0.62729 | 0.050503 |  |  | 12.8959 |  |  |  |  | 4.0287 |  |  |  |  | 17.4981 |  |  |  |  | 2.0687 |  |  | 0.54709 | 1.2345 | 0.68743 | 0.049986
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 367 | 657 |  | 297 | 0.697 | 0.835 | 0.138 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 15:52 |  | 1 |  | 0.51181 | 1.2594 | 0.74757 | 0.049469
109 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2842 | 10666 |  | 5450 | 4.326 | 15.525 | 11.199 |  | 0.757 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 16:10 |  | 1 |  | 4.895 | 18.3669 | 13.4719 | 0.97054
110 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4768 | 10742 |  | 5493 | 6.788 | 15.631 | 8.843 |  | 0.764 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 16:20 |  | 1 |  | 8.0405 | 18.4905 | 10.45 | 0.97795 | 91.5624 |  | 3.9424 |  |  | 91.0675 |  | 0.4758 |  |  | 80.9743 |  | 4.0132 |  |  | 89.8805 |  | 1.2609 |  |  | 8.2022 | 18.4466 | 10.2444 | 0.98416
111 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4966 | 10688 |  | 5565 | 7.041 | 15.555 | 8.514 |  | 0.777 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 16:31 |  | 1 |  | 8.3638 | 18.4027 | 10.0388 | 0.99036
112 | 6 | [15 empty] | 1 | [9 empty] | 1 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## 12sep22
RUN NOTES:
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area |  | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | -461 | 0 | -747.5 | 0 | -453 |  | 0.5 | -0.1036 | -0.0078899 | 0.095708 | -0.014171
TIC | TC | TNb |  | 0.0006 | 1074.9 | 0.0012 | 1966 | 6e-05 | 1142.8 |  | 0.2 | 3.5358 | 6.6971 | 3.1613 | 0.37661
739 | 1210 | 653 |  | 0.0018 | 3223.2 | 0.0036 | 6458 | 0.00018 | 2867.9 |  | 0.6 | 2.9479 | 5.9388 | 2.9909 | 0.27401
|  |  |  | 0.002997 | 5594.326 | 0.005994 | 10915.14 | 0.0002997 | 5987.102 |  | 0.333 | 8.8301 | 17.3269 | 8.4968 | 0.97744
|  |  |  | 0.004203 | 8205.274 | 0.008406 | 16345.36 | 0.0004203 | 7421.36 |  | 0.467 | 9.0591 | 18.1118 | 9.0527 | 0.85557
|  |  |  | 0.0054 | 10581.7 | 0.0108 | 21223.5 | 0.00054 | 10185.9 |  | 0.6 | 9.0082 | 18.122 | 9.1138 | 0.90386
|  |  | Slope |  | 4.9415e-07 |  | 4.9507e-07 |  | 5.1641e-08
|  |  | Intercept |  | 0.000176 |  | 0.00036612 |  | 1.6308e-05
|  |  | RSQ |  | 0.99913 |  | 0.9991 |  | 0.99256
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7969 | 10362 |  | 3979 | 10.881 | 15.095 | 4.215 |  | 0.5 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 12:14 |  | 1 |  | 13.7128 | 18.3201 | 4.6072 | 0.73929
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7124 | 10014 |  | 3852 | 9.801 | 14.604 | 4.803 |  | 0.478 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 12:24 |  | 1 |  | 12.321 | 17.7458 | 5.4248 | 0.71743 |  |  | 0.40026 |  |  |  |  | 0.074422 |  |  |  |  | 1.161 |  |  |  |  | 0 |  |  | 12.3457 | 17.7392 | 5.3935 | 0.71743
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7154 | 10006 |  | 3852 | 9.839 | 14.592 | 4.753 |  | 0.478 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 12:35 |  | 1 |  | 12.3704 | 17.7326 | 5.3622 | 0.71743
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3856 | 1301 |  | 691 | 3.373 | 1.381 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 12:53 |  | 1 |  | 4.1629 | 2.0204 | -2.1424 | 0.10398
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 753 | 1213 |  | 615 | 0.992 | 1.306 | 0.314 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 13:03 |  | 1 |  | 1.0962 | 1.9333 | 0.83709 | 0.096134 |  |  | 2.5567 |  |  |  |  | 0.30777 |  |  |  |  | 2.5628 |  |  |  |  | 7.8448 |  |  | 1.0824 | 1.9303 | 0.84795 | 0.10006
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 725 | 1207 |  | 691 | 0.971 | 1.301 | 0.33 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 13:12 |  | 1 |  | 1.0685 | 1.9273 | 0.85882 | 0.10398
7 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 698 | 2469 |  | 1412 | 2.377 | 5.926 | 3.549 |  | 0.079 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 13:30 |  | 1 |  | 2.6046 | 7.9422 | 5.3376 | 0.44612
8 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1327 | 2469 |  | 1424 | 3.581 | 5.925 | 2.344 |  | 0.082 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 13:39 |  | 1 |  | 4.1587 | 7.9422 | 3.7835 | 0.44922 |  | 42.2047 | 5.0386 |  |  |  | 31.5864 | 1.1914 |  |  |  | 20.968 | 8.5151 |  |  |  | 48.0191 | 2.3259 |  |  | 4.2661 | 7.8952 | 3.629 | 0.44406
9 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1414 | 2431 |  | 1384 | 3.75 | 5.845 | 2.095 |  | 0.072 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 13:49 |  | 1 |  | 4.3736 | 7.8481 | 3.4745 | 0.43889
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4058 | 7948 |  | 3661 | 2.94 | 5.843 | 2.903 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 14:08 |  | 1 |  | 3.6354 | 7.1682 | 3.5328 | 0.34227
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4108 | 7944 |  | 3611 | 2.972 | 5.841 | 2.869 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 14:19 |  | 1 |  | 3.6766 | 7.1649 | 3.4883 | 0.33797 |  | 22.6077 | 0.089562 |  |  |  | 18.9475 | 0.78617 |  |  |  | 15.2873 | 1.7175 |  |  |  | 13.8189 | 2.0417 |  |  | 3.6782 | 7.1368 | 3.4586 | 0.34146
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4112 | 7876 |  | 3692 | 2.975 | 5.793 | 2.818 |  | 0.225 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 14:30 |  | 1 |  | 3.6799 | 7.1088 | 3.4289 | 0.34494
13 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 4589 | 11788 |  | 6524 | 5.909 | 15.414 | 9.504 |  | 0.85 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 14:49 |  | 1 |  | 7.3382 | 18.6246 | 11.2863 | 1.0607
14 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6070 | 11785 |  | 6469 | 7.615 | 15.41 | 7.794 |  | 0.842 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 15:00 |  | 1 |  | 9.5359 | 18.6201 | 9.0842 | 1.0522 |  | 6.2268 | 0.51221 |  |  |  | 2.9165 | 1.0272 |  |  |  | 0.39379 | 2.669 |  |  |  | 16.0977 | 1.3951 |  |  | 9.5604 | 18.525 | 8.9646 | 1.0449
15 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6103 | 11657 |  | 6375 | 7.653 | 15.247 | 7.594 |  | 0.827 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 15:10 |  | 1 |  | 9.5849 | 18.4298 | 8.8449 | 1.0376
16 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8749 | 17672 |  | 8633 | 7.631 | 16.327 | 8.697 |  | 0.842 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 15:31 |  | 1 |  | 9.6344 | 19.5181 | 9.8837 | 0.98956
17 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8873 | 17501 |  | 8624 | 7.732 | 16.172 | 8.44 |  | 0.842 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 15:42 |  | 1 |  | 9.7656 | 19.3369 | 9.5712 | 0.98856 |  | 8.7716 | 0.4864 |  |  |  | 7.2768 | 0.27999 |  |  |  | 5.7819 | 1.068 |  |  |  | 8.9493 | 1.6352 |  |  | 9.7894 | 19.3098 | 9.5204 | 0.98054
18 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8918 | 17450 |  | 8479 | 7.769 | 16.126 | 8.357 |  | 0.825 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 15:54 |  | 1 |  | 9.8133 | 19.2828 | 9.4695 | 0.97253
19 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11341 | 22746 |  | 10906 | 7.596 | 16.29 | 8.694 |  | 0.854 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 16:15 |  | 1 |  | 9.6335 | 19.3782 | 9.7447 | 0.96584
20 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11484 | 22831 |  | 10970 | 7.688 | 16.351 | 8.663 |  | 0.859 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 16:27 |  | 1 |  | 9.7513 | 19.4484 | 9.6971 | 0.97134 |  | 8.2057 | 0.26216 |  |  |  | 7.3337 | 1.3282 |  |  |  | 6.4618 | 2.4117 |  |  |  | 7.9224 | 0.0088611 |  |  | 9.7385 | 19.3201 | 9.5816 | 0.9713
21 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11453 | 22520 |  | 10969 | 7.668 | 16.131 | 8.463 |  | 0.859 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 16:39 |  | 1 |  | 9.7257 | 19.1918 | 9.466 | 0.97126
22 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6125 | 9496 |  | 4258 | 8.523 | 13.873 | 5.35 |  | 0.549 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 16:58 |  | 1 |  | 10.6755 | 16.891 | 6.2155 | 0.78731
23 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6637 | 9495 |  | 4149 | 9.177 | 13.871 | 4.694 |  | 0.53 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 17:09 |  | 1 |  | 11.5188 | 16.8893 | 5.3705 | 0.76855 | 100 |  | 1.1516 |  |  | 100 |  | 0.38179 |  |  | 100 |  | 3.7518 |  |  | 100 |  | 0.26841 |  |  | 11.5855 | 16.8571 | 5.2716 | 0.76958
24 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6718 | 9456 |  | 4161 | 9.281 | 13.816 | 4.535 |  | 0.532 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 17:20 |  | 1 |  | 11.6522 | 16.825 | 5.1727 | 0.77062
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5885 | 7581 |  | 3621 | 4.93 | 6.701 | 1.771 |  | 0.263 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 17:39 |  | 1 |  | 6.1681 | 8.2385 | 2.0704 | 0.4066
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3969 | 7549 |  | 3514 | 3.46 | 6.674 | 3.215 |  | 0.252 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 17:50 |  | 1 |  | 4.2745 | 8.2068 | 3.9323 | 0.39555 |  |  | 0.90578 |  |  |  |  | 0.99424 |  |  |  |  | 1.0905 |  |  |  |  | 0.59876 |  |  | 4.2553 | 8.1662 | 3.9109 | 0.39673
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3930 | 7467 |  | 3537 | 3.43 | 6.605 | 3.175 |  | 0.254 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 18:01 |  | 1 |  | 4.236 | 8.1256 | 3.8896 | 0.39792
28 | 9 | B 26jan 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3546 | 6140 |  | 1976 | 3.135 | 5.48 | 2.345 |  | 0.091 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 18:20 |  | 1 |  | 3.8565 | 6.8117 | 2.9552 | 0.2367
29 | 9 | B 26jan 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3303 | 6178 |  | 1972 | 2.949 | 5.513 | 2.564 |  | 0.09 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 18:30 |  | 1 |  | 3.6163 | 6.8493 | 3.233 | 0.23629 |  |  | 0.98869 |  |  |  |  | 1.0025 |  |  |  |  | 1.0179 |  |  |  |  | 3.1073 |  |  | 3.5985 | 6.8151 | 3.2166 | 0.23267
30 | 9 | B 26jan 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3267 | 6109 |  | 1902 | 2.921 | 5.454 | 2.533 |  | 0.083 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 18:42 |  | 1 |  | 3.5807 | 6.781 | 3.2002 | 0.22906
31 | 10 | F50 31may 3.4 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3863 | 7410 |  | 1866 | 3.378 | 6.556 | 3.178 |  | 0.079 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 19:00 |  | 1 |  | 4.1698 | 8.0692 | 3.8994 | 0.22534
32 | 10 | F50 31may 3.4 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4050 | 7358 |  | 1778 | 3.522 | 6.512 | 2.99 |  | 0.07 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 19:11 |  | 1 |  | 4.3546 | 8.0177 | 3.6631 | 0.21625 |  |  | 0.33985 |  |  |  |  | 0.63181 |  |  |  |  | 1.7993 |  |  |  |  | 0.28697 |  |  | 4.362 | 7.9924 | 3.6304 | 0.21594
33 | 10 | F50 31may 3.4 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4065 | 7307 |  | 1772 | 3.534 | 6.469 | 2.935 |  | 0.069 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 19:22 |  | 1 |  | 4.3694 | 7.9672 | 3.5978 | 0.21563
34 | 11 | F 17mar 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4037 | 7383 |  | 1981 | 3.512 | 6.533 | 3.021 |  | 0.091 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 19:41 |  | 1 |  | 4.3417 | 8.0424 | 3.7007 | 0.23722
35 | 11 | F 17mar 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4070 | 7347 |  | 1987 | 3.537 | 6.503 | 2.966 |  | 0.092 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 19:51 |  | 1 |  | 4.3743 | 8.0068 | 3.6324 | 0.23784 |  |  | 0.22618 |  |  |  |  | 0.28483 |  |  |  |  | 0.3555 |  |  |  |  | 0.4352 |  |  | 4.3694 | 7.9954 | 3.626 | 0.23732
36 | 11 | F 17mar 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4060 | 7324 |  | 1977 | 3.53 | 6.483 | 2.954 |  | 0.091 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 20:02 |  | 1 |  | 4.3645 | 7.984 | 3.6195 | 0.2368
37 | 12 | F50 13jun 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5168 | 13209 |  | 3014 | 4.38 | 11.469 | 7.089 |  | 0.199 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 20:22 |  | 1 |  | 5.4595 | 13.811 | 8.3515 | 0.34391
38 | 12 | F50 13jun 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5675 | 12967 |  | 2973 | 4.768 | 11.264 | 6.496 |  | 0.195 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 20:32 |  | 1 |  | 5.9606 | 13.5713 | 7.6108 | 0.33967 |  |  | 1.3504 |  |  |  |  | 0.68816 |  |  |  |  | 2.3142 |  |  |  |  | 1.7483 |  |  | 6.0011 | 13.5248 | 7.5237 | 0.33673
39 | 12 | F50 13jun 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5757 | 12873 |  | 2916 | 4.832 | 11.184 | 6.353 |  | 0.189 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 20:44 |  | 1 |  | 6.0416 | 13.4783 | 7.4367 | 0.33378
40 | 13 | F 17may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4631 | 7947 |  | 2022 | 3.968 | 7.011 | 3.044 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 21:03 |  | 1 |  | 4.9288 | 8.6009 | 3.6721 | 0.24145
41 | 13 | F 17may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4216 | 7995 |  | 2017 | 3.649 | 7.052 | 3.402 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 21:13 |  | 1 |  | 4.5186 | 8.6484 | 4.1297 | 0.24093 |  |  | 1.2323 |  |  |  |  | 0.74695 |  |  |  |  | 0.21852 |  |  |  |  | 4.6047 |  |  | 4.491 | 8.6162 | 4.1252 | 0.23551
42 | 13 | F 17may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4160 | 7930 |  | 1912 | 3.607 | 6.997 | 3.39 |  | 0.084 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 21:24 |  | 1 |  | 4.4633 | 8.584 | 4.1207 | 0.23009
43 | 14 | C 28feb TCT 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9335 | 12420 |  | 1576 | 7.577 | 10.801 | 3.224 |  | 0.049 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 21:43 |  | 1 |  | 9.5777 | 13.0297 | 3.452 | 0.19539
44 | 14 | C 28feb TCT 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11251 | 12272 |  | 1580 | 9.047 | 10.675 | 1.628 |  | 0.049 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 21:54 |  | 1 |  | 11.4713 | 12.8832 | 1.4119 | 0.1958 |  |  | 1.3945 |  |  |  |  | 0.076825 |  |  |  |  | 11.3137 |  |  |  |  | 2.8353 |  |  | 11.5518 | 12.8881 | 1.3363 | 0.19306
45 | 14 | C 28feb TCT 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11414 | 12282 |  | 1527 | 9.171 | 10.684 | 1.512 |  | 0.044 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 22:06 |  | 1 |  | 11.6324 | 12.8931 | 1.2607 | 0.19033
46 | 15 | B50 05may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5634 | 6118 |  | 1572 | 4.737 | 5.462 | 0.725 |  | 0.048 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 22:25 |  | 1 |  | 5.92 | 6.7899 | 0.86986 | 0.19497
47 | 15 | B50 05may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3300 | 6177 |  | 1633 | 2.947 | 5.512 | 2.565 |  | 0.055 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 22:35 |  | 1 |  | 3.6134 | 6.8483 | 3.2349 | 0.20127 |  |  | 0.96189 |  |  |  |  | 0.63819 |  |  |  |  | 0.27785 |  |  |  |  | 5.6998 |  |  | 3.5961 | 6.8265 | 3.2305 | 0.1957
48 | 15 | B50 05may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3265 | 6133 |  | 1525 | 2.92 | 5.474 | 2.555 |  | 0.043 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 22:46 |  | 1 |  | 3.5788 | 6.8047 | 3.226 | 0.19012
49 | 16 | B50 13jun 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3908 | 6209 |  | 3700 | 3.413 | 5.539 | 2.126 |  | 0.271 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 23:05 |  | 1 |  | 4.2142 | 6.88 | 2.6658 | 0.41476
50 | 16 | B50 13jun 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4107 | 6125 |  | 3710 | 3.566 | 5.467 | 1.901 |  | 0.272 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 23:15 |  | 1 |  | 4.4109 | 6.7968 | 2.3859 | 0.41579 |  |  | 0.067194 |  |  |  |  | 0.89913 |  |  |  |  | 2.4191 |  |  |  |  | 0.52028 |  |  | 4.4124 | 6.8275 | 2.4151 | 0.41687
51 | 16 | B50 13jun 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4110 | 6187 |  | 3731 | 3.568 | 5.52 | 1.952 |  | 0.274 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 23:26 |  | 1 |  | 4.4139 | 6.8582 | 2.4443 | 0.41796
52 | 17 | B 24feb 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3457 | 5935 |  | 2111 | 3.067 | 5.307 | 2.24 |  | 0.105 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 23:45 |  | 1 |  | 3.7685 | 6.6087 | 2.8402 | 0.25064
53 | 17 | B 24feb 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3077 | 5767 |  | 2124 | 2.775 | 5.164 | 2.389 |  | 0.106 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-12 | 23:55 |  | 1 |  | 3.393 | 6.4424 | 3.0494 | 0.25199 |  |  | 0.87766 |  |  |  |  | 1.3617 |  |  |  |  | 1.903 |  |  |  |  | 0.16408 |  |  | 3.3781 | 6.3988 | 3.0206 | 0.25178
54 | 17 | B 24feb 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3047 | 5679 |  | 2120 | 2.752 | 5.09 | 2.337 |  | 0.106 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 0:06 |  | 1 |  | 3.3633 | 6.3552 | 2.9919 | 0.25157
55 | 18 | F 17mar 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3820 | 6966 |  | 1621 | 3.346 | 6.18 | 2.835 |  | 0.053 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 0:25 |  | 1 |  | 4.1273 | 7.6295 | 3.5023 | 0.20004
56 | 18 | F 17mar 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4050 | 6939 |  | 1585 | 3.522 | 6.157 | 2.635 |  | 0.05 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 0:35 |  | 1 |  | 4.3546 | 7.6028 | 3.2482 | 0.19632 |  |  | 0.27272 |  |  |  |  | 0.77134 |  |  |  |  | 1.4437 |  |  |  |  | 3.6668 |  |  | 4.3487 | 7.5736 | 3.2249 | 0.19998
57 | 18 | F 17mar 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4038 | 6880 |  | 1656 | 3.513 | 6.107 | 2.594 |  | 0.057 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 0:46 |  | 1 |  | 4.3427 | 7.5444 | 3.2017 | 0.20365
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6119 | 13300 |  | 4984 | 5.109 | 11.547 | 6.438 |  | 0.405 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 1:06 |  | 1 |  | 6.3994 | 13.9011 | 7.5017 | 0.54737
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7038 | 13285 |  | 4905 | 5.815 | 11.533 | 5.718 |  | 0.397 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 1:17 |  | 1 |  | 7.3076 | 13.8862 | 6.5786 | 0.53921 |  |  | 0.10813 |  | 96.2128 |  |  | 0.057059 |  | 89.8497 |  |  | 0.24088 |  | 83.4866 |  |  | 2.2911 |  | 104.9523 | 7.3115 | 13.8822 | 6.5707 | 0.54546
60 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7046 | 13277 |  | 5026 | 5.82 | 11.527 | 5.706 |  | 0.41 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 1:28 |  | 1 |  | 7.3155 | 13.8783 | 6.5628 | 0.55171
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5145 | 7343 |  | 1719 | 4.362 | 6.5 | 2.138 |  | 0.064 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 1:47 |  | 1 |  | 5.4368 | 8.0028 | 2.5661 | 0.21016
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4303 | 7251 |  | 1734 | 3.716 | 6.421 | 2.706 |  | 0.065 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 1:58 |  | 1 |  | 4.6046 | 7.9117 | 3.3071 | 0.21171 |  |  | 0.55649 | 5.9962 |  |  |  | 0.012516 | 4.3609 |  |  |  | 0.81019 | 2.1115 |  |  |  | 1.7718 | 4.8134 |  | 4.6175 | 7.9112 | 3.2938 | 0.20985
63 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4329 | 7250 |  | 1698 | 3.736 | 6.421 | 2.685 |  | 0.062 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 2:09 |  | 1 |  | 4.6303 | 7.9107 | 3.2804 | 0.20799
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1426 | 519 |  | 268 | 1.509 | 0.718 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 2:27 |  | 1 |  | 1.7613 | 1.2461 | -0.51519 | 0.060295
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 275 | 459 |  | 189 | 0.626 | 0.667 | 0.041 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 2:35 |  | 1 |  | 0.62378 | 1.1867 | 0.56293 | 0.052136 |  |  | 0.94611 |  |  |  |  | 0.58235 |  |  |  |  | 0.1777 |  |  |  |  | 4.2653 |  |  | 0.62675 | 1.1902 | 0.56343 | 0.053272
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 281 | 466 |  | 211 | 0.63 | 0.673 | 0.043 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 2:45 |  | 1 |  | 0.62971 | 1.1936 | 0.56393 | 0.054408
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3687 | 7557 |  | 3455 | 5.405 | 11.135 | 5.73 |  | 0.409 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 3:03 |  | 1 |  | 6.6597 | 13.6912 | 7.0314 | 0.64909
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6216 | 7601 |  | 3606 | 8.64 | 11.197 | 2.558 |  | 0.435 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 3:13 |  | 1 |  | 10.8254 | 13.7638 | 2.9384 | 0.67508 | 95.1254 |  | 4.0548 |  |  | 79.6422 |  | 1.3155 |  |  | 25.0435 |  | 23.9249 |  |  | 86.2816 |  | 1.0767 |  |  | 11.0494 | 13.6738 | 2.6245 | 0.67147
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6488 | 7492 |  | 3564 | 8.987 | 11.042 | 2.055 |  | 0.428 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 3:24 |  | 1 |  | 11.2734 | 13.5839 | 2.3105 | 0.66785
70 | 21 | B01 05may 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5367 | 6247 |  | 1951 | 4.532 | 5.571 | 1.039 |  | 0.088 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 3:43 |  | 1 |  | 5.6562 | 6.9176 | 1.2615 | 0.23412
71 | 21 | B01 05may 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3345 | 6231 |  | 1853 | 2.981 | 5.557 | 2.576 |  | 0.078 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 3:54 |  | 1 |  | 3.6578 | 6.9018 | 3.2439 | 0.224 |  |  | 1.3054 |  |  |  |  | 0.64767 |  |  |  |  | 0.088797 |  |  |  |  | 2.1437 |  |  | 3.6341 | 6.8795 | 3.2454 | 0.22162
72 | 21 | B01 05may 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3297 | 6186 |  | 1807 | 2.944 | 5.519 | 2.575 |  | 0.073 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 4:05 |  | 1 |  | 3.6104 | 6.8572 | 3.2468 | 0.21925
73 | 22 | F 31jan 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4091 | 8104 |  | 2454 | 3.553 | 7.144 | 3.591 |  | 0.141 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 4:23 |  | 1 |  | 4.3951 | 8.7563 | 4.3612 | 0.28607
74 | 22 | F 31jan 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4444 | 8089 |  | 2420 | 3.824 | 7.131 | 3.307 |  | 0.137 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 4:34 |  | 1 |  | 4.744 | 8.7415 | 3.9975 | 0.28256 |  |  | 0.33388 |  |  |  |  | 0.2495 |  |  |  |  | 0.14946 |  |  |  |  | 1.77 |  |  | 4.7361 | 8.7306 | 3.9945 | 0.28008
75 | 22 | F 31jan 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4428 | 8067 |  | 2372 | 3.812 | 7.113 | 3.301 |  | 0.132 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 4:45 |  | 1 |  | 4.7282 | 8.7197 | 3.9915 | 0.2776
76 | 23 | F50 07jun 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3976 | 8417 |  | 3195 | 3.465 | 7.409 | 3.944 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 5:04 |  | 1 |  | 4.2814 | 9.0662 | 4.7848 | 0.3626
77 | 23 | F50 07jun 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3802 | 8317 |  | 3068 | 3.332 | 7.325 | 3.993 |  | 0.205 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 5:14 |  | 1 |  | 4.1095 | 8.9672 | 4.8577 | 0.34948 |  |  | 0.28817 |  |  |  |  | 0.18753 |  |  |  |  | 0.10232 |  |  |  |  | 1.7866 |  |  | 4.1154 | 8.9756 | 4.8602 | 0.35263
78 | 23 | F50 07jun 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3814 | 8334 |  | 3129 | 3.341 | 7.339 | 3.998 |  | 0.211 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 5:25 |  | 1 |  | 4.1213 | 8.984 | 4.8627 | 0.35578
79 | 24 | B 26jan 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3394 | 5928 |  | 1791 | 3.018 | 5.301 | 2.283 |  | 0.071 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 5:44 |  | 1 |  | 3.7063 | 6.6018 | 2.8955 | 0.21759
80 | 24 | B 26jan 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3160 | 5948 |  | 1895 | 2.839 | 5.317 | 2.479 |  | 0.082 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 5:54 |  | 1 |  | 3.475 | 6.6216 | 3.1466 | 0.22833 |  |  | 0.085284 |  |  |  |  | 0.96161 |  |  |  |  | 2.1306 |  |  |  |  | 3.1239 |  |  | 3.4765 | 6.5899 | 3.1134 | 0.22482
81 | 24 | B 26jan 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3163 | 5884 |  | 1827 | 2.841 | 5.263 | 2.422 |  | 0.075 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 6:05 |  | 1 |  | 3.478 | 6.5582 | 3.0802 | 0.22131
82 | 25 | F50 23may 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3881 | 7780 |  | 1864 | 3.392 | 6.87 | 3.478 |  | 0.079 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 6:24 |  | 1 |  | 4.1876 | 8.4355 | 4.2479 | 0.22513
83 | 25 | F50 23may 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4161 | 7732 |  | 1844 | 3.607 | 6.829 | 3.222 |  | 0.077 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 6:34 |  | 1 |  | 4.4643 | 8.388 | 3.9237 | 0.22307 |  |  | 1.0012 |  |  |  |  | 0.73455 |  |  |  |  | 0.43204 |  |  |  |  | 0.50801 |  |  | 4.442 | 8.3573 | 3.9152 | 0.22363
84 | 25 | F50 23may 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4116 | 7670 |  | 1855 | 3.573 | 6.776 | 3.204 |  | 0.078 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 6:45 |  | 1 |  | 4.4198 | 8.3266 | 3.9068 | 0.2242
85 | 26 | B50 31may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3210 | 7834 |  | 4526 | 2.877 | 6.915 | 4.038 |  | 0.357 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 7:04 |  | 1 |  | 3.5244 | 8.489 | 4.9646 | 0.50007
86 | 26 | B50 31may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2880 | 7800 |  | 4571 | 2.625 | 6.887 | 4.262 |  | 0.362 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 7:14 |  | 1 |  | 3.1983 | 8.4553 | 5.257 | 0.50472 |  |  | 0.46459 |  |  |  |  | 0.16408 |  |  |  |  | 0.018306 |  |  |  |  | 0.14335 |  |  | 3.1909 | 8.4484 | 5.2575 | 0.50435
87 | 26 | B50 31may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2865 | 7786 |  | 4564 | 2.613 | 6.875 | 4.262 |  | 0.361 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 7:25 |  | 1 |  | 3.1835 | 8.4414 | 5.258 | 0.50399
88 | 27 | B50 07jun 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2798 | 7631 |  | 2850 | 2.561 | 6.744 | 4.182 |  | 0.182 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 7:44 |  | 1 |  | 3.1172 | 8.288 | 5.1707 | 0.32697
89 | 27 | B50 07jun 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2762 | 7595 |  | 2930 | 2.533 | 6.713 | 4.18 |  | 0.19 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 7:54 |  | 1 |  | 3.0817 | 8.2523 | 5.1707 | 0.33523 |  |  | 2.4017 |  |  |  |  | 0.13207 |  |  |  |  | 1.1966 |  |  |  |  | 4.3762 |  |  | 3.0451 | 8.2469 | 5.2018 | 0.32805
90 | 27 | B50 07jun 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2688 | 7584 |  | 2791 | 2.477 | 6.704 | 4.227 |  | 0.176 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 8:05 |  | 1 |  | 3.0085 | 8.2414 | 5.2329 | 0.32087
91 | 28 | F50 09may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3703 | 7592 |  | 1693 | 3.256 | 6.711 | 3.454 |  | 0.061 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 8:24 |  | 1 |  | 4.0116 | 8.2494 | 4.2377 | 0.20747
92 | 28 | F50 09may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4112 | 7530 |  | 1709 | 3.569 | 6.658 | 3.089 |  | 0.063 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 8:34 |  | 1 |  | 4.4159 | 8.188 | 3.7721 | 0.20912 |  |  | 0.56108 |  |  |  |  | 0.3513 |  |  |  |  | 0.10628 |  |  |  |  | 1.7939 |  |  | 4.4035 | 8.1736 | 3.7701 | 0.20726
93 | 28 | F50 09may 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4087 | 7501 |  | 1673 | 3.55 | 6.633 | 3.083 |  | 0.059 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 8:45 |  | 1 |  | 4.3911 | 8.1593 | 3.7681 | 0.20541
94 | 29 | F 31jan 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4242 | 7835 |  | 2070 | 3.669 | 6.917 | 3.247 |  | 0.1 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 9:05 |  | 1 |  | 4.5443 | 8.49 | 3.9456 | 0.24641
95 | 29 | F 31jan 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4368 | 7807 |  | 2110 | 3.766 | 6.892 | 3.126 |  | 0.105 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 9:15 |  | 1 |  | 4.6689 | 8.4622 | 3.7934 | 0.25054 |  |  | 1.8663 |  |  |  |  | 0.11708 |  |  |  |  | 2.6134 |  |  |  |  | 1.3697 |  |  | 4.7128 | 8.4573 | 3.7444 | 0.24884
96 | 29 | F 31jan 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4457 | 7797 |  | 2077 | 3.835 | 6.884 | 3.049 |  | 0.101 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 9:26 |  | 1 |  | 4.7568 | 8.4523 | 3.6955 | 0.24713
97 | 30 | F 17mar 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4196 | 7035 |  | 1527 | 3.634 | 6.239 | 2.605 |  | 0.044 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 9:45 |  | 1 |  | 4.4989 | 7.6978 | 3.199 | 0.19033
98 | 30 | F 17mar 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4048 | 7041 |  | 1564 | 3.52 | 6.244 | 2.723 |  | 0.048 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 9:56 |  | 1 |  | 4.3526 | 7.7038 | 3.3512 | 0.19415 |  |  | 1.7778 |  |  |  |  | 0.14148 |  |  |  |  | 2.6905 |  |  |  |  | 1.1234 |  |  | 4.3916 | 7.6983 | 3.3067 | 0.19306
99 | 30 | F 17mar 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4127 | 7030 |  | 1543 | 3.581 | 6.234 | 2.653 |  | 0.045 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 10:07 |  | 1 |  | 4.4307 | 7.6929 | 3.2622 | 0.19198
100 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6334 | 13676 |  | 4424 | 5.274 | 11.865 | 6.591 |  | 0.347 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 10:27 |  | 1 |  | 6.6118 | 14.2733 | 7.6615 | 0.48953
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7263 | 13615 |  | 4392 | 5.987 | 11.813 | 5.826 |  | 0.343 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 10:37 |  | 1 |  | 7.53 | 14.2129 | 6.683 | 0.48623 |  |  | 2.1038 |  | 107.8818 |  |  | 0.21619 |  | 99.4683 |  |  | 2.8963 |  | 91.0548 |  |  | 0.5538 |  | 88.5382 | 7.61 | 14.1976 | 6.5876 | 0.48489
102 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7425 | 13584 |  | 4366 | 6.111 | 11.787 | 5.675 |  | 0.341 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 10:49 |  | 1 |  | 7.6901 | 14.1823 | 6.4922 | 0.48354
103 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5247 | 7542 |  | 1806 | 4.44 | 6.668 | 2.228 |  | 0.073 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 11:08 |  | 1 |  | 5.5376 | 8.1998 | 2.6623 | 0.21914
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4331 | 7511 |  | 1855 | 3.738 | 6.642 | 2.904 |  | 0.078 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 11:18 |  | 1 |  | 4.6323 | 8.1692 | 3.5369 | 0.2242 |  |  | 0.27697 | 5.472 |  |  |  | 1.244 | 5.3147 |  |  |  | 3.2714 | 5.1053 |  |  |  | 1.2816 | 15.5645 |  | 4.6387 | 8.1187 | 3.4799 | 0.22565
105 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4344 | 7409 |  | 1883 | 3.748 | 6.555 | 2.807 |  | 0.081 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 11:29 |  | 1 |  | 4.6451 | 8.0682 | 3.423 | 0.22709
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1470 | 529 |  | 255 | 1.543 | 0.726 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 11:47 |  | 1 |  | 1.8048 | 1.256 | -0.54877 | 0.058952
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 280 | 490 |  | 238 | 0.63 | 0.694 | 0.064 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 11:56 |  | 1 |  | 0.62873 | 1.2174 | 0.58868 | 0.057197 |  |  | 8.181 |  |  |  |  | 0.40749 |  |  |  |  | 7.2783 |  |  |  |  | 4.8077 |  |  | 0.60402 | 1.2149 | 0.61091 | 0.055854
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 230 | 485 |  | 212 | 0.591 | 0.69 | 0.099 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 12:05 |  | 1 |  | 0.57931 | 1.2125 | 0.63314 | 0.054511
109 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3198 | 8496 |  | 3993 | 4.781 | 12.46 | 7.679 |  | 0.503 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 12:23 |  | 1 |  | 5.8543 | 15.2407 | 9.3865 | 0.7417
110 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5715 | 8715 |  | 4305 | 7.999 | 12.77 | 4.771 |  | 0.557 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 12:34 |  | 1 |  | 10.0001 | 15.6021 | 5.602 | 0.7954 | 87.6975 |  | 4.5395 |  |  | 90.9345 |  | 2.1056 |  |  | 53.7214 |  | 15.1635 |  |  | 101.071 |  | 4.673 |  |  | 10.2324 | 15.4396 | 5.2072 | 0.77724
111 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5997 | 8518 |  | 4094 | 8.359 | 12.491 | 4.132 |  | 0.52 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 12:44 |  | 1 |  | 10.4646 | 15.277 | 4.8124 | 0.75908
112 | 6 | [15 empty] | 1 | [9 empty] | 1 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## 13sep22
RUN NOTES:
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area |  | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0.044452 | 0.2624 | 0.21795 | -0.0055789
TIC | TC | TNb |  | 0.0006 | 1237.5 | 0.0012 | 2128.6 | 6e-05 | 1201.4 |  | 0.2 | 3.0899 | 5.9125 | 2.8227 | 0.30522
280 | 603.5 | 224 |  | 0.0018 | 3742.5 | 0.0036 | 6950.3 | 0.00018 | 3346.2 |  | 0.6 | 3.0398 | 5.9399 | 2.9 | 0.29167
|  |  |  | 0.002997 | 5916.02 | 0.005994 | 11412.069 | 0.0002997 | 5969.816 |  | 0.333 | 8.6194 | 17.3201 | 8.7007 | 0.94414
|  |  |  | 0.004203 | 8739.48 | 0.008406 | 16906.331 | 0.0004203 | 7892.831 |  | 0.467 | 9.0567 | 18.161 | 9.1043 | 0.89202
|  |  |  | 0.0054 | 11246 | 0.0108 | 21750.3 | 0.00054 | 10136.2 |  | 0.6 | 9.0603 | 18.1227 | 9.0624 | 0.89295
|  |  | Slope |  | 4.8141e-07 |  | 4.939e-07 |  | 5.3132e-08
|  |  | Intercept |  | 2.2226e-05 |  | 0.0001312 |  | -2.7895e-06
|  |  | RSQ |  | 0.99911 |  | 0.99907 |  | 0.99871
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6366 | 8820 |  | 4227 | 8.831 | 12.918 | 4.086 |  | 0.543 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 13:30 |  | 1 |  | 10.2896 | 14.9579 | 4.6683 | 0.73934
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6393 | 8800 |  | 4235 | 8.866 | 12.889 | 4.023 |  | 0.545 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 13:40 |  | 1 |  | 10.3329 | 14.925 | 4.592 | 0.74075 |  |  | 0.57296 |  |  |  |  | 1.3996 |  |  |  |  | 5.9841 |  |  |  |  | 0.023906 |  |  | 10.3626 | 14.8212 | 4.4586 | 0.74084
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6430 | 8674 |  | 4236 | 8.914 | 12.712 | 3.799 |  | 0.545 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 13:51 |  | 1 |  | 10.3923 | 14.7175 | 4.3252 | 0.74093
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3645 | 832 |  | 455 | 3.211 | 0.983 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 14:09 |  | 1 |  | 3.5539 | 1.0842 | -2.4697 | 0.042771
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 692 | 846 |  | 486 | 0.946 | 0.995 | 0.049 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 14:19 |  | 1 |  | 0.71072 | 1.0981 | 0.38735 | 0.046066 |  |  | 8.3255 |  |  |  |  | 0.90363 |  |  |  |  | 11.4232 |  |  |  |  | 17.8392 |  |  | 0.68232 | 1.0931 | 0.41081 | 0.042293
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 633 | 836 |  | 415 | 0.901 | 0.987 | 0.086 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 14:28 |  | 1 |  | 0.65392 | 1.0882 | 0.43428 | 0.038521
7 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 678 | 2343 |  | 1244 | 2.337 | 5.66 | 3.323 |  | 0.035 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 14:46 |  | 1 |  | 1.7431 | 6.442 | 4.6989 | 0.31654
8 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1341 | 2369 |  | 1284 | 3.609 | 5.713 | 2.104 |  | 0.046 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 14:55 |  | 1 |  | 3.339 | 6.5062 | 3.1672 | 0.32716 |  | 11.9813 | 1.2181 |  |  |  | 8.4779 | 0.075883 |  |  |  | 4.9745 | 1.1425 |  |  |  | 9.6738 | 1.1304 |  |  | 3.3594 | 6.5087 | 3.1492 | 0.32902
9 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1358 | 2371 |  | 1298 | 3.642 | 5.718 | 2.076 |  | 0.049 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 15:05 |  | 1 |  | 3.3799 | 6.5111 | 3.1312 | 0.33088
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4063 | 7521 |  | 3607 | 2.943 | 5.542 | 2.599 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 15:24 |  | 1 |  | 3.297 | 6.4097 | 3.1127 | 0.31476
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4056 | 7688 |  | 3670 | 2.939 | 5.66 | 2.721 |  | 0.223 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 15:35 |  | 1 |  | 3.2914 | 6.5471 | 3.2558 | 0.32034 |  | 10.3141 | 1.091 |  |  |  | 8.9336 | 0.34004 |  |  |  | 7.5531 | 1.8078 |  |  |  | 5.1576 | 3.0877 |  |  | 3.3094 | 6.536 | 3.2266 | 0.31547
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4101 | 7661 |  | 3560 | 2.968 | 5.641 | 2.673 |  | 0.214 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 15:47 |  | 1 |  | 3.3275 | 6.5249 | 3.1974 | 0.3106
13 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 4444 | 11691 |  | 6110 | 5.742 | 15.29 | 9.548 |  | 0.785 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 16:06 |  | 1 |  | 6.4913 | 17.7338 | 11.2425 | 0.96651
14 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6017 | 11775 |  | 6115 | 7.554 | 15.397 | 7.843 |  | 0.786 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 16:16 |  | 1 |  | 8.7654 | 17.8584 | 9.093 | 0.96731 |  | 1.2337 | 2.7811 |  |  |  | 0.46555 | 0.64571 |  |  |  | 0.30257 | 1.457 |  |  |  | 7.5498 | 0.13187 |  |  | 8.889 | 17.9162 | 9.0272 | 0.96795
15 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 6188 | 11853 |  | 6123 | 7.751 | 15.495 | 7.744 |  | 0.787 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 16:27 |  | 1 |  | 9.0126 | 17.974 | 8.9615 | 0.96859
16 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8825 | 17446 |  | 8430 | 7.693 | 16.123 | 8.43 |  | 0.82 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 16:48 |  | 1 |  | 9.1449 | 18.7317 | 9.5868 | 0.95314
17 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8947 | 17436 |  | 8420 | 7.793 | 16.113 | 8.32 |  | 0.819 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 16:59 |  | 1 |  | 9.2707 | 18.7212 | 9.4505 | 0.952 |  | 3.6258 | 1.1937 |  |  |  | 4.2062 | 0.38341 |  |  |  | 4.7866 | 0.41795 |  |  |  | 6.2392 | 0.86863 |  |  | 9.3263 | 18.7571 | 9.4308 | 0.95615
18 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 9055 | 17504 |  | 8493 | 7.882 | 16.176 | 8.294 |  | 0.827 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 17:11 |  | 1 |  | 9.382 | 18.7931 | 9.4111 | 0.96031
19 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11474 | 22284 |  | 10459 | 7.681 | 15.965 | 8.283 |  | 0.815 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 17:33 |  | 1 |  | 9.2432 | 18.562 | 9.3188 | 0.92153
20 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11512 | 22523 |  | 10393 | 7.706 | 16.133 | 8.428 |  | 0.809 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 17:46 |  | 1 |  | 9.2737 | 18.7587 | 9.485 | 0.91569 |  | 3.665 | 1.204 |  |  |  | 3.9933 | 0.42656 |  |  |  | 4.3217 | 2.0468 |  |  |  | 1.8614 | 0.23183 |  |  | 9.3298 | 18.7188 | 9.389 | 0.91675
21 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11652 | 22426 |  | 10417 | 7.795 | 16.065 | 8.27 |  | 0.811 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 17:59 |  | 1 |  | 9.386 | 18.6789 | 9.2929 | 0.91782
22 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6340 | 8828 |  | 4071 | 8.798 | 12.929 | 4.131 |  | 0.516 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 18:18 |  | 1 |  | 10.2479 | 14.9711 | 4.7232 | 0.71171
23 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6980 | 8808 |  | 4030 | 9.616 | 12.9 | 3.284 |  | 0.509 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 18:28 |  | 1 |  | 11.2749 | 14.9381 | 3.6633 | 0.70444 | 100 |  | 0.44024 |  |  | 100 |  | 0.44181 |  |  | 100 |  | 3.2062 |  |  | 100 |  | 1.0504 |  |  | 11.2998 | 14.9052 | 3.6055 | 0.70816
24 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7011 | 8768 |  | 4072 | 9.656 | 12.844 | 3.188 |  | 0.516 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 18:39 |  | 1 |  | 11.3246 | 14.8723 | 3.5477 | 0.71188
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6310 | 7457 |  | 3141 | 5.256 | 6.596 | 1.34 |  | 0.213 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 18:59 |  | 1 |  | 6.1198 | 7.6284 | 1.5085 | 0.3282
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4279 | 7246 |  | 3093 | 3.697 | 6.418 | 2.72 |  | 0.207 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 19:10 |  | 1 |  | 4.1644 | 7.4199 | 3.2556 | 0.3231 |  |  | 0.44026 |  |  |  |  | 0.50716 |  |  |  |  | 0.59281 |  |  |  |  | 0.55756 |  |  | 4.1552 | 7.4012 | 3.246 | 0.324
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4260 | 7208 |  | 3110 | 3.683 | 6.385 | 2.702 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 19:21 |  | 1 |  | 4.1461 | 7.3824 | 3.2364 | 0.3249
28 | 9 | B 02may 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3500 | 6894 |  | 2840 | 3.1 | 6.119 | 3.019 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 19:39 |  | 1 |  | 3.4143 | 7.0722 | 3.6579 | 0.29621
29 | 9 | B 02may 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3095 | 6849 |  | 2763 | 2.789 | 6.081 | 3.292 |  | 0.173 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 19:50 |  | 1 |  | 3.0244 | 7.0278 | 4.0034 | 0.28803 |  |  | 1.767 |  |  |  |  | 1.7725 |  |  |  |  | 4.5314 |  |  |  |  | 3.6056 |  |  | 3.0513 | 6.9661 | 3.9147 | 0.28293
30 | 9 | B 02may 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3151 | 6724 |  | 2667 | 2.832 | 5.975 | 3.143 |  | 0.163 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 20:01 |  | 1 |  | 3.0783 | 6.9043 | 3.826 | 0.27783
31 | 10 | B600 23may 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3637 | 6294 |  | 4705 | 3.205 | 5.611 | 2.406 |  | 0.376 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 20:20 |  | 1 |  | 3.5462 | 6.4796 | 2.9333 | 0.4944
32 | 10 | B600 23may 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3918 | 6197 |  | 4673 | 3.42 | 5.528 | 2.108 |  | 0.373 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 20:30 |  | 1 |  | 3.8168 | 6.3838 | 2.567 | 0.491 |  |  | 1.0289 |  |  |  |  | 0.93274 |  |  |  |  | 3.9221 |  |  |  |  | 1.2194 |  |  | 3.8365 | 6.3541 | 2.5176 | 0.48802
33 | 10 | B600 23may 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3959 | 6137 |  | 4617 | 3.452 | 5.478 | 2.026 |  | 0.367 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 20:41 |  | 1 |  | 3.8563 | 6.3245 | 2.4682 | 0.48504
34 | 11 | F 24feb 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4102 | 7220 |  | 2192 | 3.562 | 6.395 | 2.833 |  | 0.113 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 21:00 |  | 1 |  | 3.9939 | 7.3943 | 3.4003 | 0.22735
35 | 11 | F 24feb 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4183 | 7235 |  | 2135 | 3.624 | 6.408 | 2.784 |  | 0.107 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 21:10 |  | 1 |  | 4.0719 | 7.4091 | 3.3372 | 0.2213 |  |  | 1.1175 |  |  |  |  | 0.57493 |  |  |  |  | 0.083191 |  |  |  |  | 4.2668 |  |  | 4.0493 | 7.3878 | 3.3385 | 0.21667
36 | 11 | F 24feb 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4136 | 7192 |  | 2048 | 3.588 | 6.372 | 2.784 |  | 0.098 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 21:22 |  | 1 |  | 4.0267 | 7.3666 | 3.3399 | 0.21205
37 | 12 | F 24feb 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4135 | 7221 |  | 1902 | 3.587 | 6.396 | 2.809 |  | 0.083 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 21:41 |  | 1 |  | 4.0257 | 7.3953 | 3.3695 | 0.19654
38 | 12 | F 24feb 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4172 | 7272 |  | 1913 | 3.615 | 6.439 | 2.824 |  | 0.084 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 21:51 |  | 1 |  | 4.0613 | 7.4456 | 3.3843 | 0.19771 |  |  | 0.47302 |  |  |  |  | 0.53208 |  |  |  |  | 1.7517 |  |  |  |  | 0.48257 |  |  | 4.071 | 7.4259 | 3.3549 | 0.19818
39 | 12 | F 24feb 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4192 | 7232 |  | 1922 | 3.631 | 6.405 | 2.774 |  | 0.085 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 22:02 |  | 1 |  | 4.0806 | 7.4061 | 3.3255 | 0.19866
40 | 13 | F 24feb 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4136 | 7452 |  | 2027 | 3.588 | 6.592 | 3.004 |  | 0.096 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 22:21 |  | 1 |  | 4.0267 | 7.6234 | 3.5968 | 0.20982
41 | 13 | F 24feb 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4159 | 7295 |  | 2036 | 3.606 | 6.458 | 2.853 |  | 0.097 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 22:32 |  | 1 |  | 4.0488 | 7.4684 | 3.4195 | 0.21078 |  |  | 0.095076 |  |  |  |  | 1.4183 |  |  |  |  | 2.9626 |  |  |  |  | 0.50289 |  |  | 4.0507 | 7.5217 | 3.471 | 0.21131
42 | 13 | F 24feb 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4163 | 7403 |  | 2046 | 3.609 | 6.551 | 2.942 |  | 0.098 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 22:43 |  | 1 |  | 4.0527 | 7.575 | 3.5224 | 0.21184
43 | 14 | F 24feb 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4128 | 7224 |  | 1820 | 3.582 | 6.398 | 2.817 |  | 0.074 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 23:02 |  | 1 |  | 4.019 | 7.3982 | 3.3793 | 0.18782
44 | 14 | F 24feb 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4189 | 7222 |  | 1875 | 3.628 | 6.397 | 2.769 |  | 0.08 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 23:12 |  | 1 |  | 4.0777 | 7.3962 | 3.3185 | 0.19367 |  |  | 0.14177 |  |  |  |  | 0.68345 |  |  |  |  | 1.3531 |  |  |  |  | 5.121 |  |  | 4.0748 | 7.3711 | 3.2962 | 0.18883
45 | 14 | F 24feb 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4183 | 7171 |  | 1784 | 3.624 | 6.354 | 2.729 |  | 0.071 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 23:23 |  | 1 |  | 4.0719 | 7.3459 | 3.2739 | 0.184
46 | 15 | F50 31may 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5200 | 17705 |  | 2805 | 4.404 | 15.278 | 10.874 |  | 0.177 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 23:44 |  | 1 |  | 5.0511 | 17.7513 | 12.7002 | 0.29249
47 | 15 | F50 31may 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5694 | 17673 |  | 2784 | 4.783 | 15.251 | 10.467 |  | 0.175 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-13 | 23:55 |  | 1 |  | 5.5267 | 17.7197 | 12.1929 | 0.29026 |  |  | 1.3325 |  |  |  |  | 0.34503 |  |  |  |  | 0.1058 |  |  |  |  | 3.3122 |  |  | 5.5638 | 17.7503 | 12.1865 | 0.28553
48 | 15 | F50 31may 8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5771 | 17735 |  | 2695 | 4.843 | 15.304 | 10.461 |  | 0.166 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 0:07 |  | 1 |  | 5.6009 | 17.7809 | 12.18 | 0.2808
49 | 16 | F 25jan weir | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4601 | 5665 |  | 647 | 3.945 | 5.078 | 1.133 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 0:26 |  | 1 |  | 4.4744 | 5.8582 | 1.3839 | 0.063174
50 | 16 | F 25jan weir | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4021 | 5556 |  | 639 | 3.5 | 4.986 | 1.486 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 0:36 |  | 1 |  | 3.9159 | 5.7506 | 1.8346 | 0.062324 |  |  | 1.8272 |  |  |  |  | 0.51399 |  |  |  |  | 2.348 |  |  |  |  | 0.8489 |  |  | 3.9521 | 5.7654 | 1.8133 | 0.06259
51 | 16 | F 25jan weir | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4096 | 5586 |  | 644 | 3.557 | 5.011 | 1.454 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 0:47 |  | 1 |  | 3.9882 | 5.7802 | 1.7921 | 0.062855
52 | 17 | B 31may 1.8? 1.6? | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3148 | 6716 |  | 2217 | 2.83 | 5.968 | 3.138 |  | 0.116 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 1:06 |  | 1 |  | 3.0754 | 6.8964 | 3.821 | 0.23001
53 | 17 | B 31may 1.8? 1.6? | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2826 | 6692 |  | 2196 | 2.583 | 5.948 | 3.365 |  | 0.114 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 1:16 |  | 1 |  | 2.7654 | 6.8727 | 4.1073 | 0.22778 |  |  | 1.3671 |  |  |  |  | 0.10066 |  |  |  |  | 0.7431 |  |  |  |  | 1.0317 |  |  | 2.7466 | 6.8693 | 4.1227 | 0.22661
54 | 17 | B 31may 1.8? 1.6? | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2787 | 6685 |  | 2174 | 2.553 | 5.942 | 3.389 |  | 0.111 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 1:27 |  | 1 |  | 2.7278 | 6.8658 | 4.138 | 0.22544
55 | 18 | F15 13jun 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3455 | 8584 |  | 2880 | 3.065 | 7.551 | 4.485 |  | 0.185 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 1:46 |  | 1 |  | 3.371 | 8.7416 | 5.3706 | 0.30046
56 | 18 | F15 13jun 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3674 | 8626 |  | 2890 | 3.233 | 7.586 | 4.353 |  | 0.186 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 1:56 |  | 1 |  | 3.5818 | 8.7831 | 5.2013 | 0.30153 |  |  | 0.26844 |  |  |  |  | 0.63179 |  |  |  |  | 1.2565 |  |  |  |  | 0.99168 |  |  | 3.5867 | 8.7554 | 5.1688 | 0.30004
57 | 18 | F15 13jun 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3684 | 8570 |  | 2862 | 3.242 | 7.539 | 4.297 |  | 0.183 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 2:07 |  | 1 |  | 3.5915 | 8.7278 | 5.1363 | 0.29855
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6750 | 13545 |  | 4356 | 5.593 | 11.754 | 6.16 |  | 0.34 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 2:27 |  | 1 |  | 6.5435 | 13.6421 | 7.0986 | 0.45731
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7990 | 13481 |  | 4437 | 6.545 | 11.699 | 5.154 |  | 0.348 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 2:38 |  | 1 |  | 7.7374 | 13.5788 | 5.8415 | 0.46592 |  |  | 0.12451 |  | 125.0471 |  |  | 0.13086 |  | 103.1388 |  |  | 0.46811 |  | 81.2305 |  |  | 0.45511 |  | 86.6251 | 7.7326 | 13.5877 | 5.8552 | 0.46698
60 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7980 | 13499 |  | 4457 | 6.537 | 11.715 | 5.178 |  | 0.35 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 2:49 |  | 1 |  | 7.7277 | 13.5966 | 5.8689 | 0.46804
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4935 | 8665 |  | 3033 | 4.201 | 7.62 | 3.418 |  | 0.201 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 3:09 |  | 1 |  | 4.796 | 8.8216 | 4.0257 | 0.31672
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3709 | 8610 |  | 2959 | 3.26 | 7.573 | 4.313 |  | 0.193 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 3:19 |  | 1 |  | 3.6155 | 8.7673 | 5.1518 | 0.30886 |  |  | 0.64117 | 0.48203 |  |  |  | 0.25947 | 0.0056409 |  |  |  | 0.007539 | 0.32627 |  |  |  | 0.86386 | 2.4661 |  | 3.604 | 8.7559 | 5.1519 | 0.30753
63 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3685 | 8587 |  | 2934 | 3.242 | 7.553 | 4.311 |  | 0.191 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 3:30 |  | 1 |  | 3.5924 | 8.7446 | 5.1521 | 0.3062
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1281 | 646 |  | 280 | 1.398 | 0.826 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 3:48 |  | 1 |  | 1.2778 | 0.90051 | -0.37731 | 0.024175
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 266 | 618 |  | 214 | 0.619 | 0.802 | 0.183 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 3:57 |  | 1 |  | 0.30056 | 0.87286 | 0.57229 | 0.017162 |  |  | 8.5845 |  |  |  |  | 3.3366 |  |  |  |  | 10.2123 |  |  |  |  | 11.6618 |  |  | 0.31404 | 0.85853 | 0.54449 | 0.018224
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 294 | 589 |  | 234 | 0.641 | 0.777 | 0.137 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 4:07 |  | 1 |  | 0.32752 | 0.84421 | 0.51669 | 0.019287
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3421 | 7359 |  | 3391 | 5.066 | 10.854 | 5.788 |  | 0.398 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 4:25 |  | 1 |  | 5.5638 | 12.5526 | 6.9889 | 0.59127
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6147 | 7413 |  | 3505 | 8.551 | 10.931 | 2.38 |  | 0.418 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 4:35 |  | 1 |  | 9.9382 | 12.6415 | 2.7034 | 0.61146 | 89.1716 |  | 2.8962 |  |  | 83.8473 |  | 1.1658 |  |  | 63.0683 |  | 17.6556 |  |  | 86.1392 |  | 0.90195 |  |  | 10.0842 | 12.5683 | 2.4841 | 0.60872
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6329 | 7324 |  | 3474 | 8.783 | 10.806 | 2.023 |  | 0.412 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 4:46 |  | 1 |  | 10.2302 | 12.495 | 2.2648 | 0.60597
70 | 21 | B 24feb 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5406 | 6238 |  | 2476 | 4.563 | 5.563 | 1 |  | 0.143 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 5:05 |  | 1 |  | 5.2495 | 6.4243 | 1.1748 | 0.25753
71 | 21 | B 24feb 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3381 | 6294 |  | 2504 | 3.009 | 5.611 | 2.602 |  | 0.146 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 5:15 |  | 1 |  | 3.2997 | 6.4796 | 3.1798 | 0.26051 |  |  | 0.78473 |  |  |  |  | 0.87274 |  |  |  |  | 2.6221 |  |  |  |  | 1.0662 |  |  | 3.3127 | 6.4514 | 3.1387 | 0.25913
72 | 21 | B 24feb 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3408 | 6237 |  | 2478 | 3.029 | 5.563 | 2.533 |  | 0.143 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 5:26 |  | 1 |  | 3.3257 | 6.4233 | 3.0975 | 0.25774
73 | 22 | C 28feb 1.5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8188 | 12319 |  | 1823 | 6.697 | 10.715 | 4.019 |  | 0.075 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 5:46 |  | 1 |  | 7.928 | 12.431 | 4.503 | 0.18814
74 | 22 | C 28feb 1.5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9968 | 12309 |  | 1855 | 8.062 | 10.706 | 2.644 |  | 0.078 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 5:57 |  | 1 |  | 9.6418 | 12.4211 | 2.7793 | 0.19154 |  |  | 1.644 |  |  |  |  | 0.72631 |  |  |  |  | 9.4075 |  |  |  |  | 4.423 |  |  | 9.7217 | 12.3762 | 2.6545 | 0.1874
75 | 22 | C 28feb 1.5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10134 | 12218 |  | 1777 | 8.189 | 10.629 | 2.44 |  | 0.07 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 6:08 |  | 1 |  | 9.8017 | 12.3313 | 2.5296 | 0.18325
76 | 23 | C 28feb 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10262 | 12503 |  | 1723 | 8.288 | 10.871 | 2.584 |  | 0.064 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 6:28 |  | 1 |  | 9.9249 | 12.6128 | 2.6879 | 0.17751
77 | 23 | C 28feb 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10417 | 12489 |  | 1819 | 8.406 | 10.859 | 2.453 |  | 0.074 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 6:39 |  | 1 |  | 10.0741 | 12.5989 | 2.5248 | 0.18772 |  |  | 0.095528 |  |  |  |  | 0.3141 |  |  |  |  | 1.9654 |  |  |  |  | 5.1684 |  |  | 10.0789 | 12.5792 | 2.5002 | 0.18299
78 | 23 | C 28feb 9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10427 | 12449 |  | 1730 | 8.414 | 10.825 | 2.411 |  | 0.065 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 6:51 |  | 1 |  | 10.0838 | 12.5594 | 2.4757 | 0.17826
79 | 24 | B 25jan 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5530 | 6315 |  | 1835 | 4.658 | 5.628 | 0.971 |  | 0.076 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 7:10 |  | 1 |  | 5.3688 | 6.5003 | 1.1315 | 0.18942
80 | 24 | B 25jan 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3542 | 6301 |  | 1811 | 3.132 | 5.617 | 2.485 |  | 0.073 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 7:20 |  | 1 |  | 3.4548 | 6.4865 | 3.0317 | 0.18687 |  |  | 2.6259 |  |  |  |  | 0.060932 |  |  |  |  | 2.7839 |  |  |  |  | 0.28474 |  |  | 3.41 | 6.4845 | 3.0745 | 0.1866
81 | 24 | B 25jan 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3449 | 6297 |  | 1806 | 3.061 | 5.613 | 2.553 |  | 0.073 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 7:31 |  | 1 |  | 3.3652 | 6.4825 | 3.1173 | 0.18633
82 | 25 | C 28feb 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8071 | 12092 |  | 1765 | 6.607 | 10.522 | 3.915 |  | 0.069 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 7:51 |  | 1 |  | 7.8154 | 12.2068 | 4.3914 | 0.18198
83 | 25 | C 28feb 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10014 | 12155 |  | 1753 | 8.097 | 10.576 | 2.479 |  | 0.067 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 8:02 |  | 1 |  | 9.6861 | 12.269 | 2.5829 | 0.1807 |  |  | 1.3036 |  |  |  |  | 0.16089 |  |  |  |  | 4.2438 |  |  |  |  | 0.76741 |  |  | 9.7497 | 12.2789 | 2.5292 | 0.18001
84 | 25 | C 28feb 6.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10146 | 12175 |  | 1740 | 8.198 | 10.593 | 2.395 |  | 0.066 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 8:13 |  | 1 |  | 9.8132 | 12.2888 | 2.4756 | 0.17932
85 | 26 | F50 07jun 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6475 | 8200 |  | 2008 | 5.382 | 7.225 | 1.843 |  | 0.094 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 8:33 |  | 1 |  | 6.2787 | 8.3623 | 2.0836 | 0.2078
86 | 26 | F50 07jun 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5180 | 8166 |  | 2020 | 4.389 | 7.197 | 2.807 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 8:43 |  | 1 |  | 5.0319 | 8.3287 | 3.2969 | 0.20908 |  |  | 0.40263 |  |  |  |  | 0.23692 |  |  |  |  | 1.2052 |  |  |  |  | 0.20351 |  |  | 5.0217 | 8.3386 | 3.3169 | 0.20886
87 | 26 | F50 07jun 5.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5159 | 8186 |  | 2016 | 4.373 | 7.214 | 2.841 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 8:54 |  | 1 |  | 5.0116 | 8.3485 | 3.3368 | 0.20865
88 | 27 | F50 31may 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5016 | 8300 |  | 1866 | 4.263 | 7.31 | 3.047 |  | 0.079 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 9:14 |  | 1 |  | 4.874 | 8.4611 | 3.5871 | 0.19271
89 | 27 | F50 31may 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4990 | 8245 |  | 1843 | 4.243 | 7.263 | 3.02 |  | 0.077 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 9:24 |  | 1 |  | 4.8489 | 8.4068 | 3.5578 | 0.19027 |  |  | 1.0862 |  |  |  |  | 0.94444 |  |  |  |  | 3.7796 |  |  |  |  | 0.11176 |  |  | 4.8754 | 8.3672 | 3.4918 | 0.19016
90 | 27 | F50 31may 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5045 | 8165 |  | 1841 | 4.286 | 7.195 | 2.91 |  | 0.077 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 9:35 |  | 1 |  | 4.9019 | 8.3277 | 3.4259 | 0.19005
91 | 28 | C 28feb HPB 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2570 | 3492 |  | 912 | 2.387 | 3.237 | 0.85 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 9:54 |  | 1 |  | 2.5189 | 3.7118 | 1.1929 | 0.091334
92 | 28 | C 28feb HPB 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1603 | 3476 |  | 881 | 1.644 | 3.223 | 1.579 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 10:04 |  | 1 |  | 1.5879 | 3.696 | 2.1081 | 0.08804 |  |  | 0.7914 |  |  |  |  | 0.50909 |  |  |  |  | 0.29698 |  |  |  |  | 5.4022 |  |  | 1.5816 | 3.6866 | 2.105 | 0.090484
93 | 28 | C 28feb HPB 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1590 | 3457 |  | 927 | 1.635 | 3.207 | 1.573 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 10:15 |  | 1 |  | 1.5753 | 3.6772 | 2.1019 | 0.092928
94 | 29 | B 22mar 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2473 | 6011 |  | 1715 | 2.312 | 5.371 | 3.059 |  | 0.063 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 10:34 |  | 1 |  | 2.4255 | 6.2 | 3.7745 | 0.17666
95 | 29 | B 22mar 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2924 | 6010 |  | 1708 | 2.658 | 5.37 | 2.712 |  | 0.063 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 10:44 |  | 1 |  | 2.8597 | 6.199 | 3.3393 | 0.17592 |  |  | 2.6611 |  |  |  |  | 0.46317 |  |  |  |  | 1.3815 |  |  |  |  | 3.5015 |  |  | 2.8222 | 6.1847 | 3.3625 | 0.17906
96 | 29 | B 22mar 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2846 | 5981 |  | 1767 | 2.598 | 5.346 | 2.748 |  | 0.069 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 10:55 |  | 1 |  | 2.7846 | 6.1704 | 3.3858 | 0.18219
97 | 30 | F 17mar wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3378 | 5340 |  | 1231 | 3.006 | 4.802 | 1.796 |  | 0.013 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 11:14 |  | 1 |  | 3.2969 | 5.5372 | 2.2404 | 0.12523
98 | 30 | F 17mar wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3608 | 5347 |  | 1276 | 3.183 | 4.808 | 1.626 |  | 0.017 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 11:24 |  | 1 |  | 3.5183 | 5.5441 | 2.0258 | 0.13001 |  |  | 0.38239 |  |  |  |  | 0.1248 |  |  |  |  | 1.0118 |  |  |  |  | 0.97601 |  |  | 3.525 | 5.5407 | 2.0156 | 0.13065
99 | 30 | F 17mar wet | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3622 | 5340 |  | 1288 | 3.193 | 4.802 | 1.609 |  | 0.019 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 11:35 |  | 1 |  | 3.5318 | 5.5372 | 2.0054 | 0.13129
100 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10210 | 17205 |  | 4324 | 8.247 | 14.854 | 6.607 |  | 0.336 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 11:55 |  | 1 |  | 9.8748 | 17.2574 | 7.3826 | 0.45391
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12807 | 17236 |  | 4328 | 10.24 | 14.881 | 4.641 |  | 0.337 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 12:06 |  | 1 |  | 12.3753 | 17.288 | 4.9127 | 0.45433 |  |  | 0.89074 |  | 93.0948 |  |  | 0.062871 |  | 85.9868 |  |  | 2.506 |  | 78.8788 |  |  | 0.82198 |  | 92.1781 | 12.4306 | 17.2826 | 4.8519 | 0.45247
102 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12922 | 17225 |  | 4293 | 10.328 | 14.872 | 4.543 |  | 0.333 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 12:18 |  | 1 |  | 12.486 | 17.2771 | 4.7911 | 0.45061
103 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6760 | 5874 |  | 1340 | 5.601 | 5.255 | 0 |  | 0.024 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 12:38 |  | 1 |  | 6.5531 | 6.0647 | -0.48841 | 0.13682
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4190 | 5875 |  | 1336 | 3.63 | 5.256 | 1.626 |  | 0.024 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 12:48 |  | 1 |  | 4.0787 | 6.0657 | 1.987 | 0.13639 |  |  | 1.8584 | 13.6416 |  |  |  | 0.11406 | 8.9901 |  |  |  | 3.3736 | 0.2717 |  |  |  | 2.2853 | 3.1619 |  | 4.0411 | 6.0622 | 2.0211 | 0.13485
105 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4112 | 5868 |  | 1307 | 3.569 | 5.25 | 1.681 |  | 0.021 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 12:59 |  | 1 |  | 4.0036 | 6.0588 | 2.0552 | 0.13331
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1382 | 554 |  | 303 | 1.475 | 0.748 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 13:17 |  | 1 |  | 1.3751 | 0.80964 | -0.56543 | 0.026619
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 325 | 487 |  | 281 | 0.664 | 0.691 | 0.027 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 13:26 |  | 1 |  | 0.35737 | 0.74346 | 0.38609 | 0.024281 |  |  | 1.3562 |  |  |  |  | 0.26538 |  |  |  |  | 1.7433 |  |  |  |  | 6.3559 |  |  | 0.35496 | 0.74444 | 0.38948 | 0.025078
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 320 | 489 |  | 296 | 0.66 | 0.693 | 0.033 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 13:35 |  | 1 |  | 0.35255 | 0.74543 | 0.39288 | 0.025875
109 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3607 | 7284 |  | 3700 | 5.304 | 10.749 | 5.445 |  | 0.452 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 13:53 |  | 1 |  | 5.8622 | 12.4291 | 6.5669 | 0.646
110 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6230 | 7283 |  | 3766 | 8.658 | 10.747 | 2.089 |  | 0.463 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 14:03 |  | 1 |  | 10.0714 | 12.4275 | 2.3561 | 0.65769 | 90.7941 |  | 3.7983 |  |  | 83.0564 |  | 0.42302 |  |  | 52.8591 |  | 15.4175 |  |  | 91.1133 |  | 4.1226 |  |  | 10.2663 | 12.4538 | 2.1875 | 0.64441
111 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6473 | 7315 |  | 3616 | 8.968 | 10.794 | 1.825 |  | 0.437 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-14 | 14:14 |  | 1 |  | 10.4613 | 12.4802 | 2.0189 | 0.63112
112 | 6 | [15 empty] | 1 | [9 empty] | 1 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!