## notes
Adopted a daily calibration so that we can adjust to variations in TIC signal over time.
Stored calibration is probably okay for TC and TNb, but in this version I'm using the DAILY calibration for all analytes.
We can't use repeated check standards over the run because they are no longer isolated from the atmosphere after the first run and the concentration changes over time.
We don't have space in the rack for many covered check standards.
Instead we can use a complex sample that doesn't exchange with the atmosphere as rapidly, so... tap or spiked tap water
We previously did a series of samples analyzed at the begining and end of the run and found no differences
Are results for tap water or spiked tap water flat over a run?
close enough (see rolling tap or rolling spiked tap)
ON 25jul22 adopted a correction for the part of the signal attributable to the Type I water used to make the standards in daily calibrations.
This will change the water blanks.
Previously, they would have been roughly the same as the calibration zero
Now they will be above that because the calibration is adjusted for water signal, but samples are not
ON 04oct21 started using a spiked tap water as run in and as reference instead of plain tap water so it would have some TOC in it
On 23sep22 started using a flush before the first water blank

## rolling spiked blank
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
20 | 9 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1933 | 5883 |  | 3010 | 1.898 | 5.263 | 3.365 |  | 0.199 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 16:00 |  | 1 |  | 3.2854 | 5.5471 | 2.2617 | 0.29433 |  |  | 0.50767 |  |  |  |  | 0.052864 |  |  |  |  | 0.60413 |  |  |  |  | 0.24165 |  |  | 3.2771 | 5.5456 | 2.2686 | 0.29469 |  | 1
21 | 9 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1923 | 5880 |  | 3017 | 1.89 | 5.26 | 3.37 |  | 0.2 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 16:08 |  | 1 |  | 3.2687 | 5.5442 | 2.2754 | 0.29504
23 | 10 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1756 | 5822 |  | 2801 | 1.762 | 5.211 | 3.449 |  | 0.177 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 16:31 |  | 1 |  | 2.9909 | 5.4875 | 2.4966 | 0.27307 |  |  | 0.38861 |  |  |  |  | 1.0452 |  |  |  |  | 1.8261 |  |  |  |  | 2.5743 |  |  | 2.9967 | 5.5163 | 2.5196 | 0.27663 |  | 2
24 | 10 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1763 | 5881 |  | 2871 | 1.767 | 5.261 | 3.494 |  | 0.184 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 16:39 |  | 1 |  | 3.0026 | 5.5451 | 2.5426 | 0.28019
26 | 11 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1748 | 5839 |  | 2844 | 1.756 | 5.225 | 3.469 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 17:02 |  | 1 |  | 2.9776 | 5.5041 | 2.5265 | 0.27744 |  |  | 2.8095 |  |  |  |  | 0.28447 |  |  |  |  | 4.0578 |  |  |  |  | 1.8527 |  |  | 3.02 | 5.4963 | 2.4762 | 0.28004 |  | 3
27 | 11 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1799 | 5823 |  | 2895 | 1.795 | 5.212 | 3.417 |  | 0.187 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 17:10 |  | 1 |  | 3.0625 | 5.4885 | 2.426 | 0.28263
113 | 4 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1140 | 6840 |  | 3222 | 1.289 | 6.074 | 4.784 |  | 0.221 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 6:29 |  | 1 |  | 2.3361 | 5.4713 | 3.1352 | 0.2601
114 | 4 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1139 | 6832 |  | 3156 | 1.288 | 6.067 | 4.778 |  | 0.214 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 6:38 |  | 1 |  | 2.3339 | 5.4651 | 3.1312 | 0.25479 | [20 empty] | 2.3262 | 5.4554 | 3.1292 | 0.24764 |  | 4
116 | 5 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1132 | 6787 |  | 3000 | 1.283 | 6.028 | 4.745 |  | 0.198 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 7:01 |  | 1 |  | 2.3185 | 5.4301 | 3.1116 | 0.24225
117 | 5 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1117 | 6870 |  | 3003 | 1.272 | 6.098 | 4.827 |  | 0.198 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 7:10 |  | 1 |  | 2.2854 | 5.4946 | 3.2092 | 0.24249 | [20 empty] | 2.3019 | 5.4624 | 3.1604 | 0.24237 |  | 5
119 | 6 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1108 | 6818 |  | 3015 | 1.265 | 6.055 | 4.789 |  | 0.199 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 7:34 |  | 1 |  | 2.2656 | 5.4542 | 3.1886 | 0.24346
120 | 6 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1101 | 6841 |  | 2979 | 1.259 | 6.074 | 4.815 |  | 0.196 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 7:43 |  | 1 |  | 2.2502 | 5.4721 | 3.2219 | 0.24056 | [20 empty] | 2.2579 | 5.4631 | 3.2053 | 0.24201 |  | 6
21 | 9 | spiked blank + 150 a | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1000 | 7140 |  | 3455 | 1.182 | 6.327 | 5.145 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 15:57 |  | 1 |  | 2.7334 | 5.5493 | 2.8159 | 0.29195 |  |  | 0.52007 |  |  |  |  | 0.14896 |  |  |  |  | 0.79412 |  |  |  |  | 4.1612 |  |  | 2.7263 | 5.5534 | 2.8272 | 0.286 |  | 7
22 | 9 | spiked blank + 150 a | TIC/TC/TNb |  |  | 0.5 | 0.5 | 995 | 7150 |  | 3327 | 1.178 | 6.336 | 5.158 |  | 0.232 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 16:06 |  | 1 |  | 2.7192 | 5.5576 | 2.8384 | 0.28005
24 | 10 | spiked blank + 150 b | TIC/TC/TNb |  |  | 0.5 | 0.5 | 937 | 6985 |  | 3370 | 1.134 | 6.197 | 5.063 |  | 0.236 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 16:31 |  | 1 |  | 2.5547 | 5.4211 | 2.8664 | 0.28405 |  |  | 1.4535 |  |  |  |  | 0.51749 |  |  |  |  | 2.2419 |  |  |  |  | 0.098246 |  |  | 2.5363 | 5.4351 | 2.8989 | 0.28391 |  | 8
25 | 10 | spiked blank + 150 b | TIC/TC/TNb |  |  | 0.5 | 0.5 | 924 | 7019 |  | 3367 | 1.124 | 6.225 | 5.101 |  | 0.236 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 16:40 |  | 1 |  | 2.5178 | 5.4492 | 2.9313 | 0.28377
27 | 11 | spiked blank + 150 c | TIC/TC/TNb |  |  | 0.5 | 0.5 | 930 | 6877 |  | 3169 | 1.129 | 6.105 | 4.976 |  | 0.215 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 17:04 |  | 1 |  | 2.5349 | 5.3317 | 2.7969 | 0.26536 |  |  | 1.0119 |  |  |  |  | 2.3005 |  |  |  |  | 5.2098 |  |  |  |  | 6.114 |  |  | 2.5221 | 5.3938 | 2.8717 | 0.27373 |  | 9
28 | 11 | spiked blank + 150 c | TIC/TC/TNb |  |  | 0.5 | 0.5 | 921 | 7027 |  | 3349 | 1.121 | 6.232 | 5.11 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 17:14 |  | 1 |  | 2.5093 | 5.4558 | 2.9465 | 0.28209
58 | 19 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 902 | 7007 |  | 3333 | 1.107 | 6.215 | 5.108 |  | 0.233 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-16 | 22:19 |  | 1 |  | 3.0295 | 5.9483 | 2.9188 | 0.29141 |  |  | 0.11639 |  |  |  |  | 1.2262 |  |  |  |  | 2.6389 |  |  |  |  | 2.118 |  |  | 3.0313 | 5.912 | 2.8808 | 0.28836 |  | 10
59 | 19 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 903 | 6923 |  | 3259 | 1.108 | 6.144 | 5.036 |  | 0.225 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-16 | 22:28 |  | 1 |  | 3.033 | 5.8758 | 2.8427 | 0.28531
61 | 20 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 863 | 7065 |  | 3421 | 1.077 | 6.264 | 5.187 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-16 | 22:52 |  | 1 |  | 2.8919 | 5.9983 | 3.1064 | 0.29868 |  |  | 1.4748 |  |  |  |  | 0.77995 |  |  |  |  | 0.13735 |  |  |  |  | 0.44312 |  |  | 2.8707 | 5.975 | 3.1043 | 0.29802 |  | 11
62 | 20 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 851 | 7011 |  | 3405 | 1.067 | 6.218 | 5.151 |  | 0.24 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-16 | 23:00 |  | 1 |  | 2.8496 | 5.9517 | 3.1022 | 0.29736
96 | 28 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 976 | 7106 |  | 3375 | 1.164 | 6.298 | 5.134 |  | 0.237 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-17 | 4:54 |  | 1 |  | 3.2906 | 6.0337 | 2.7431 | 0.29488 |  |  | 2.8272 |  |  |  |  | 0.085855 |  |  |  |  | 3.1064 |  |  |  |  | 0.61388 |  |  | 3.2447 | 6.0311 | 2.7864 | 0.29579 |  | 12
97 | 28 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 950 | 7100 |  | 3397 | 1.143 | 6.294 | 5.15 |  | 0.239 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-17 | 5:03 |  | 1 |  | 3.1988 | 6.0285 | 2.8297 | 0.2967
99 | 29 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 927 | 6988 |  | 3406 | 1.126 | 6.199 | 5.072 |  | 0.24 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-17 | 5:26 |  | 1 |  | 3.1177 | 5.9319 | 2.8142 | 0.29744 |  |  | 0.22608 |  |  |  |  | 0.97953 |  |  |  |  | 2.3323 |  |  |  |  | 0.52862 |  |  | 3.1212 | 5.903 | 2.7817 | 0.29665 |  | 13
100 | 29 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 929 | 6921 |  | 3387 | 1.128 | 6.142 | 5.015 |  | 0.238 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-17 | 5:35 |  | 1 |  | 3.1248 | 5.874 | 2.7493 | 0.29587
27 | 9 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2547 | 6134 |  | 4188 | 2.369 | 5.475 | 3.106 |  | 0.322 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 16:20 |  | 1 |  | 2.4662 | 5.9818 | 3.5157 | 0.42406
28 | 9 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2923 | 6169 |  | 4244 | 2.658 | 5.504 | 2.847 |  | 0.328 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 16:29 |  | 1 |  | 2.8328 | 6.0171 | 3.1842 | 0.42984 |  |  | 0.85688 |  |  |  |  | 0.40227 |  |  |  |  | 1.5359 |  |  |  |  | 1.5222 |  |  | 2.845 | 6.005 | 3.16 | 0.42659 |  | 14
29 | 9 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2948 | 6145 |  | 4181 | 2.676 | 5.484 | 2.808 |  | 0.321 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 16:38 |  | 1 |  | 2.8572 | 5.9929 | 3.1357 | 0.42334
30 | 10 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3134 | 6507 |  | 3152 | 2.819 | 5.791 | 2.972 |  | 0.214 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 16:53 |  | 1 |  | 3.0386 | 6.3573 | 3.3187 | 0.31728
31 | 10 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3230 | 6506 |  | 3085 | 2.893 | 5.79 | 2.897 |  | 0.207 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 17:01 |  | 1 |  | 3.1322 | 6.3563 | 3.2241 | 0.31038 |  |  | 1.5372 |  |  |  |  | 1.2276 |  |  |  |  | 3.8419 |  |  |  |  | 0.099574 |  |  | 3.1083 | 6.3955 | 3.2872 | 0.31053 |  | 15
32 | 10 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3181 | 6584 |  | 3088 | 2.855 | 5.856 | 3.001 |  | 0.207 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 17:10 |  | 1 |  | 3.0844 | 6.4348 | 3.3503 | 0.31069
33 | 11 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2915 | 6044 |  | 3046 | 2.651 | 5.399 | 2.748 |  | 0.203 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 17:25 |  | 1 |  | 2.825 | 5.8912 | 3.0662 | 0.30636
34 | 11 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2825 | 5678 |  | 2731 | 2.582 | 5.089 | 2.506 |  | 0.17 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 17:34 |  | 1 |  | 2.7373 | 5.5229 | 2.7856 | 0.27389 |  |  | 0.82274 |  |  |  |  | 6.5913 |  |  |  |  | 13.3622 |  |  |  |  | 9.4646 |  |  | 2.7261 | 5.7111 | 2.985 | 0.2875 |  | 16
35 | 11 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2802 | 6052 |  | 2995 | 2.565 | 5.406 | 2.841 |  | 0.197 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 17:43 |  | 1 |  | 2.7148 | 5.8993 | 3.1845 | 0.3011
27 | 9 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2995 | 6848 |  | 3199 | 2.713 | 6.08 | 3.367 |  | 0.219 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 16:42 |  | 1 |  | 3.0242 | 6.5175 | 3.4933 | 0.31243
28 | 9 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3408 | 6813 |  | 3183 | 3.029 | 6.05 | 3.021 |  | 0.217 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 16:50 |  | 1 |  | 3.4332 | 6.4831 | 3.05 | 0.31086 |  |  | 0.81082 |  |  |  |  | 0.30316 |  |  |  |  | 0.26522 |  |  |  |  | 1.7426 |  |  | 3.4193 | 6.4733 | 3.054 | 0.30818 |  | 17
29 | 9 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3380 | 6793 |  | 3128 | 3.008 | 6.034 | 3.026 |  | 0.211 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 16:59 |  | 1 |  | 3.4054 | 6.4635 | 3.0581 | 0.30549
30 | 10 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3335 | 6850 |  | 3451 | 2.974 | 6.082 | 3.108 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 17:15 |  | 1 |  | 3.3609 | 6.5194 | 3.1586 | 0.33703
31 | 10 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3264 | 6716 |  | 3425 | 2.919 | 5.968 | 3.049 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 17:23 |  | 1 |  | 3.2906 | 6.388 | 3.0974 | 0.33449 |  |  | 1.4549 |  |  |  |  | 0.84127 |  |  |  |  | 3.224 |  |  |  |  | 0.34967 |  |  | 3.2668 | 6.4149 | 3.1481 | 0.33508 |  | 18
32 | 10 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3216 | 6771 |  | 3437 | 2.882 | 6.015 | 3.133 |  | 0.243 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 17:33 |  | 1 |  | 3.2431 | 6.4419 | 3.1989 | 0.33566
33 | 11 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3252 | 6718 |  | 3107 | 2.91 | 5.97 | 3.061 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 17:48 |  | 1 |  | 3.2787 | 6.3899 | 3.1112 | 0.30344
34 | 11 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3225 | 6609 |  | 3103 | 2.889 | 5.877 | 2.988 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 17:57 |  | 1 |  | 3.252 | 6.283 | 3.031 | 0.30305 |  |  | 1.5032 |  |  |  |  | 0.32742 |  |  |  |  | 2.2548 |  |  |  |  | 0.74379 |  |  | 3.2277 | 6.2933 | 3.0656 | 0.30193 |  | 19
35 | 11 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3176 | 6630 |  | 3080 | 2.851 | 5.895 | 3.044 |  | 0.206 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 18:06 |  | 1 |  | 3.2034 | 6.3036 | 3.1001 | 0.30081
109 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3858 | 7028 |  | 3207 | 3.374 | 6.233 | 2.859 |  | 0.219 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 11:09 |  | 1 |  | 3.3538 | 6.0179 | 2.6641 | 0.26132
110 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3835 | 7260 |  | 3139 | 3.357 | 6.429 | 3.072 |  | 0.212 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 11:20 |  | 1 |  | 3.331 | 6.2397 | 2.9087 | 0.25524 |  |  | 9.9539 |  |  |  |  | 4.4169 |  |  |  |  | 1.5757 |  |  |  |  | 1.9416 |  |  | 3.1731 | 6.1049 | 2.9318 | 0.25775 |  | 20
111 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3516 | 6978 |  | 3195 | 3.112 | 6.191 | 3.078 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 11:28 |  | 1 |  | 3.0152 | 5.9701 | 2.9549 | 0.26025
112 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3894 | 7129 |  | 3183 | 3.402 | 6.318 | 2.916 |  | 0.217 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 17:08 |  | 1 |  | 3.2906 | 6.0252 | 2.7345 | 0.26282
113 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3574 | 7106 |  | 3186 | 3.157 | 6.299 | 3.142 |  | 0.217 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 17:19 |  | 1 |  | 3 | 6.0041 | 3.0041 | 0.26309 |  |  | 0.24252 |  |  |  |  | 0.045725 |  |  |  |  | 0.15041 |  |  |  |  | 1.656 |  |  | 2.9963 | 6.0027 | 3.0064 | 0.26093 |  | 21
114 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3566 | 7103 |  | 3138 | 3.15 | 6.296 | 3.146 |  | 0.212 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 17:32 |  | 1 |  | 2.9927 | 6.0014 | 3.0087 | 0.25877
109 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4087 | 7360 |  | 3224 | 3.55 | 6.513 | 2.963 |  | 0.221 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 12:39 |  | 1 |  | 3.4645 | 6.3208 | 2.8563 | 0.2789
110 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3723 | 7343 |  | 3220 | 3.271 | 6.499 | 3.228 |  | 0.221 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 12:50 |  | 1 |  | 3.1353 | 6.3054 | 3.1701 | 0.27854 |  |  | 1.5739 |  |  |  |  | 0.3886 |  |  |  |  | 2.3682 |  |  |  |  | 0.29708 |  |  | 3.1602 | 6.2932 | 3.133 | 0.27895 |  | 22
111 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3778 | 7316 |  | 3229 | 3.313 | 6.477 | 3.163 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 13:01 |  | 1 |  | 3.185 | 6.2809 | 3.0959 | 0.27937
103 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2703 | 7080 |  | 3178 | 2.489 | 6.277 | 3.788 |  | 0.216 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 13:06 |  | 1 |  | 2.157 | 5.9082 | 3.7512 | 0.28802
104 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3584 | 7146 |  | 3261 | 3.165 | 6.333 | 3.168 |  | 0.225 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 13:16 |  | 1 |  | 2.9422 | 5.9673 | 3.0251 | 0.29539 |  |  | 0.12125 |  |  |  |  | 0.23975 |  |  |  |  | 0.58962 |  |  |  |  | 0.24043 |  |  | 2.9404 | 5.9745 | 3.034 | 0.29575 |  | 23
105 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3580 | 7162 |  | 3269 | 3.162 | 6.346 | 3.184 |  | 0.226 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 13:28 |  | 1 |  | 2.9387 | 5.9816 | 3.043 | 0.2961
109 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4056 | 7447 |  | 3299 | 3.526 | 6.588 | 3.061 |  | 0.229 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 14:34 |  | 1 |  | 3.3323 | 6.1653 | 2.833 | 0.2814
110 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3680 | 7469 |  | 3369 | 3.238 | 6.607 | 3.368 |  | 0.236 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 14:45 |  | 1 |  | 3.0003 | 6.1848 | 3.1845 | 0.28781 |  |  | 0.41283 |  |  |  |  | 1.0493 |  |  |  |  | 1.6527 |  |  |  |  | 1.7332 |  |  | 2.9941 | 6.1525 | 3.1584 | 0.28533 |  | 24
111 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3666 | 7396 |  | 3315 | 3.227 | 6.545 | 3.318 |  | 0.231 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 14:56 |  | 1 |  | 2.9879 | 6.1202 | 3.1323 | 0.28286
103 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4326 | 7658 |  | 3201 | 3.734 | 6.767 | 3.033 |  | 0.219 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 14:56 |  | 1 |  | 3.6357 | 6.4785 | 2.8428 | 0.28576
104 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3951 | 7640 |  | 3231 | 3.446 | 6.751 | 3.306 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 15:07 |  | 1 |  | 3.2899 | 6.4622 | 3.1723 | 0.28859 |  |  | 2.5829 |  |  |  |  | 0.48751 |  |  |  |  | 3.575 |  |  |  |  | 2.1112 |  |  | 3.248 | 6.478 | 3.23 | 0.28557 |  | 25
105 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3860 | 7675 |  | 3167 | 3.377 | 6.781 | 3.404 |  | 0.215 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 15:18 |  | 1 |  | 3.206 | 6.4938 | 3.2878 | 0.28256
163 | 2 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3723 | 7161 |  | 2880 | 3.271 | 6.345 | 3.074 |  | 0.185 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 4:09 |  | 1 |  | 3.1446 | 6.0365 | 2.8919 | 0.26623
164 | 2 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3614 | 7053 |  | 2836 | 3.188 | 6.253 | 3.066 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 4:19 |  | 1 |  | 3.0444 | 5.9368 | 2.8924 | 0.26158 |  |  | 1.6442 |  |  |  |  | 0.18643 |  |  |  |  | 2.0777 |  |  |  |  | 1.9613 |  |  | 3.0195 | 5.9423 | 2.9228 | 0.26417 |  | 26
165 | 2 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3560 | 7065 |  | 2885 | 3.146 | 6.264 | 3.118 |  | 0.186 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 4:30 |  | 1 |  | 2.9947 | 5.9479 | 2.9532 | 0.26676
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2555 | 6964 |  | 3157 | 2.375 | 6.179 | 3.804 |  | 0.214 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:37 |  | 1 |  | 2.5319 | 6.7694 | 4.2375 | 0.34669
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3363 | 7048 |  | 3097 | 2.995 | 6.249 | 3.254 |  | 0.208 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:47 |  | 1 |  | 3.2708 | 6.8477 | 3.5769 | 0.34049 |  |  | 0.67328 |  |  |  |  | 0.12255 |  |  |  |  | 0.37839 |  |  |  |  | 3.2102 |  |  | 3.2598 | 6.8435 | 3.5837 | 0.33511 |  | 27
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3339 | 7039 |  | 2993 | 2.977 | 6.242 | 3.266 |  | 0.197 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:58 |  | 1 |  | 3.2488 | 6.8393 | 3.5905 | 0.32973
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2565 | 7043 |  | 3289 | 2.383 | 6.246 | 3.863 |  | 0.228 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 19:42 |  | 1 |  | 2.4722 | 6.75 | 4.2778 | 0.31385
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3493 | 7131 |  | 3234 | 3.095 | 6.319 | 3.225 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 19:53 |  | 1 |  | 3.3141 | 6.8305 | 3.5163 | 0.30885 |  |  | 1.6561 |  |  |  |  | 0.026771 |  |  |  |  | 1.5874 |  |  |  |  | 0.70844 |  |  | 3.2869 | 6.8314 | 3.5445 | 0.30776 |  | 28
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3433 | 7133 |  | 3210 | 3.049 | 6.321 | 3.272 |  | 0.22 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 20:04 |  | 1 |  | 3.2597 | 6.8323 | 3.5726 | 0.30667
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5721 | 7959 |  | 3684 | 4.804 | 7.022 | 2.218 |  | 0.269 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 17:34 |  | 1 |  | 5.5447 | 7.6343 | 2.0896 | 0.36269
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4089 | 7929 |  | 3587 | 3.552 | 6.996 | 3.444 |  | 0.259 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 17:44 |  | 1 |  | 4.006 | 7.6067 | 3.6007 | 0.3543 |  |  | 0.63745 |  |  |  |  | 0.30192 |  |  |  |  | 1.3368 |  |  |  |  | 0.21995 |  |  | 3.9933 | 7.6182 | 3.6249 | 0.35391 |  | 29
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4062 | 7954 |  | 3578 | 3.531 | 7.017 | 3.486 |  | 0.258 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 17:55 |  | 1 |  | 3.9806 | 7.6297 | 3.6491 | 0.35352
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5043 | 6907 |  | 3345 | 4.284 | 6.13 | 1.847 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:16 |  | 1 |  | 5.0176 | 7.1112 | 2.0937 | 0.33852
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3545 | 7179 |  | 3324 | 3.135 | 6.36 | 3.225 |  | 0.232 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:26 |  | 1 |  | 3.5293 | 7.3933 | 3.864 | 0.33644 |  |  | 1.7607 |  |  |  |  | 3.9625 |  |  |  |  | 6.016 |  |  |  |  | 1.333 |  |  | 3.4985 | 7.2497 | 3.7512 | 0.33421 |  | 30
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3483 | 6902 |  | 3279 | 3.087 | 6.126 | 3.039 |  | 0.227 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:37 |  | 1 |  | 3.4677 | 7.106 | 3.6384 | 0.33199
25 | 2 | Spiked Blank 100ml + 300uL NOT COVERED | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5148 | 7473 |  | 3415 | 4.365 | 6.61 | 2.245 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:27 |  | 1 |  | 5.1966 | 7.9063 | 2.7097 | 0.37189
26 | 2 | Spiked Blank 100ml + 300uL NOT COVERED | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3783 | 7443 |  | 3416 | 3.317 | 6.584 | 3.266 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:37 |  | 1 |  | 3.8591 | 7.8771 | 4.018 | 0.372 |  |  | 1.8192 |  |  |  |  | 0.62093 |  |  |  |  | 0.51663 |  |  |  |  | 0.83047 |  |  | 3.8243 | 7.8527 | 4.0284 | 0.37355 |  | 31
27 | 2 | Spiked Blank 100ml + 300uL NOT COVERED | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3712 | 7393 |  | 3446 | 3.263 | 6.542 | 3.279 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:48 |  | 1 |  | 3.7895 | 7.8283 | 4.0388 | 0.3751
28 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4938 | 7932 |  | 3522 | 4.203 | 6.998 | 2.795 |  | 0.252 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 17:48 |  | 1 |  | 4.7818 | 7.8664 | 3.0846 | 0.36141
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3647 | 7719 |  | 3439 | 3.213 | 6.818 | 3.604 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 17:59 |  | 1 |  | 3.5142 | 7.6572 | 4.143 | 0.35331 |  |  | 0.2798 |  |  |  |  | 1.0447 |  |  |  |  | 1.6981 |  |  |  |  | 0.55453 |  |  | 3.5093 | 7.6174 | 4.1081 | 0.35233 |  | 32
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3637 | 7638 |  | 3419 | 3.205 | 6.749 | 3.544 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 18:10 |  | 1 |  | 3.5044 | 7.5776 | 4.0732 | 0.35135
28 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5005 | 7715 |  | 3111 | 4.255 | 6.815 | 2.56 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 19:34 |  | 1 |  | 4.9222 | 7.5247 | 2.6025 | 0.33478
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3766 | 7635 |  | 3084 | 3.304 | 6.746 | 3.442 |  | 0.207 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 19:45 |  | 1 |  | 3.7104 | 7.4482 | 3.7378 | 0.3319 |  |  | 1.5941 |  |  |  |  | 3.3682 |  |  |  |  | 5.1605 |  |  |  |  | 2.1738 |  |  | 3.6811 | 7.3248 | 3.6437 | 0.32833 |  | 33
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3706 | 7377 |  | 3017 | 3.258 | 6.528 | 3.27 |  | 0.2 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 19:56 |  | 1 |  | 3.6517 | 7.2014 | 3.5497 | 0.32476
136 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5276 | 7426 |  | 3110 | 4.462 | 6.57 | 2.107 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 19:47 |  | 1 |  | 5.5845 | 7.7309 | 2.1465 | 0.35213
137 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4005 | 7293 |  | 3148 | 3.487 | 6.457 | 2.97 |  | 0.213 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 19:57 |  | 1 |  | 4.2973 | 7.5981 | 3.3008 | 0.35617 |  |  | 2.9412 |  |  |  |  | 0.21006 |  |  |  |  | 4.169 |  |  |  |  | 0.47586 |  |  | 4.235 | 7.6061 | 3.3711 | 0.35702 |  | 34
138 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3882 | 7309 |  | 3164 | 3.393 | 6.471 | 3.078 |  | 0.215 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 20:08 |  | 1 |  | 4.1728 | 7.6141 | 3.4413 | 0.35787
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5559 | 7403 |  | 3849 | 4.68 | 6.55 | 1.87 |  | 0.287 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 20:27 |  | 1 |  | 5.3038 | 7.0926 | 1.7888 | 0.376
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3880 | 7368 |  | 3850 | 3.392 | 6.52 | 3.129 |  | 0.287 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 20:38 |  | 1 |  | 3.7101 | 7.0589 | 3.3487 | 0.3761 |  |  | 0.71891 |  |  |  |  | 2.2226 |  |  |  |  | 3.9152 |  |  |  |  | 2.2929 |  |  | 3.6968 | 6.9813 | 3.2844 | 0.37184 |  | 35
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3852 | 7207 |  | 3763 | 3.37 | 6.384 | 3.014 |  | 0.278 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 20:49 |  | 1 |  | 3.6835 | 6.9037 | 3.2202 | 0.36758
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5183 | 7619 |  | 3372 | 4.391 | 6.734 | 2.342 |  | 0.237 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 19:50 |  | 1 |  | 5.1197 | 7.6501 | 2.5303 | 0.3655
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3734 | 7667 |  | 3403 | 3.279 | 6.774 | 3.494 |  | 0.24 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 20:00 |  | 1 |  | 3.7285 | 7.6975 | 3.9689 | 0.36879 |  |  | 1.1775 |  |  |  |  | 0.99245 |  |  |  |  | 3.0746 |  |  |  |  | 0.4881 |  |  | 3.7506 | 7.6595 | 3.9088 | 0.3697 |  | 36
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3780 | 7590 |  | 3420 | 3.315 | 6.709 | 3.394 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 20:11 |  | 1 |  | 3.7727 | 7.6215 | 3.8488 | 0.3706
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5111 | 7033 |  | 3051 | 4.336 | 6.237 | 1.901 |  | 0.203 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 15:50 |  | 1 |  | 4.9138 | 7.0996 | 2.1857 | 0.33817
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3541 | 7051 |  | 3132 | 3.132 | 6.252 | 3.12 |  | 0.212 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 16:01 |  | 1 |  | 3.3962 | 7.1178 | 3.7216 | 0.3469 |  |  | 0.51363 |  |  |  |  | 0.057075 |  |  |  |  | 0.57505 |  |  |  |  | 0.12426 |  |  | 3.3875 | 7.1199 | 3.7324 | 0.34668 |  | 37
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3523 | 7055 |  | 3128 | 3.118 | 6.256 | 3.138 |  | 0.211 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 16:12 |  | 1 |  | 3.3788 | 7.1219 | 3.7431 | 0.34646
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5218 | 8133 |  | 3549 | 4.418 | 7.169 | 2.751 |  | 0.255 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 20:28 |  | 1 |  | 5.2211 | 8.0458 | 2.8247 | 0.37314
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3504 | 7864 |  | 3434 | 3.103 | 6.941 | 3.837 |  | 0.243 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 20:39 |  | 1 |  | 3.5491 | 7.7852 | 4.2361 | 0.36143 |  |  | 0.16478 |  |  |  |  | 0.29915 |  |  |  |  | 0.6895 |  |  |  |  | 0.36555 |  |  | 3.552 | 7.7736 | 4.2215 | 0.36209 |  | 38
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3510 | 7840 |  | 3447 | 3.107 | 6.921 | 3.813 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 20:49 |  | 1 |  | 3.555 | 7.7619 | 4.207 | 0.36276
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5359 | 7204 |  | 3058 | 4.526 | 6.382 | 1.856 |  | 0.204 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 21:53 |  | 1 |  | 5.4623 | 7.204 | 1.7417 | 0.32611
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3567 | 7121 |  | 3051 | 3.152 | 6.311 | 3.16 |  | 0.203 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 22:04 |  | 1 |  | 3.6815 | 7.1222 | 3.4408 | 0.32539 |  |  | 1.3588 |  |  |  |  | 0.47161 |  |  |  |  | 0.46908 |  |  |  |  | 1.4375 |  |  | 3.6566 | 7.1055 | 3.4488 | 0.32307 |  | 39
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3517 | 7087 |  | 3006 | 3.113 | 6.282 | 3.169 |  | 0.198 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 22:15 |  | 1 |  | 3.6318 | 7.0887 | 3.4569 | 0.32074
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5054 | 8053 |  | 3916 | 4.293 | 7.101 | 2.808 |  | 0.294 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 18:59 |  | 1 |  | 4.419 | 7.6745 | 3.2555 | 0.37785
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3666 | 7906 |  | 3981 | 3.227 | 6.977 | 3.749 |  | 0.3 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 19:09 |  | 1 |  | 3.1904 | 7.5367 | 4.3463 | 0.38433 |  |  | 0.027748 |  |  |  |  | 2.8695 |  |  |  |  | 4.9434 |  |  |  |  | 2.7604 |  |  | 3.19 | 7.6464 | 4.4564 | 0.37909 |  | 40
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3665 | 8140 |  | 3876 | 3.227 | 7.174 | 3.948 |  | 0.289 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 19:20 |  | 1 |  | 3.1895 | 7.7561 | 4.5666 | 0.37386
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5049 | 7799 |  | 3421 | 4.288 | 6.885 | 2.597 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 21:11 |  | 1 |  | 4.4619 | 7.4703 | 3.0084 | 0.33331
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4035 | 7485 |  | 3424 | 3.511 | 6.62 | 3.109 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 21:21 |  | 1 |  | 3.6001 | 7.1764 | 3.5762 | 0.3336 |  |  | 1.3308 |  |  |  |  | 0.44246 |  |  |  |  | 2.1963 |  |  |  |  | 0.84241 |  |  | 3.5763 | 7.1923 | 3.616 | 0.33501 |  | 41
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3979 | 7519 |  | 3453 | 3.467 | 6.648 | 3.181 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 21:32 |  | 1 |  | 3.5526 | 7.2082 | 3.6557 | 0.33642
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5882 | 7415 |  | 3376 | 4.928 | 6.561 | 1.633 |  | 0.237 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 17:48 |  | 1 |  | 5.6398 | 7.1992 | 1.5594 | 0.34809
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3911 | 7422 |  | 3312 | 3.416 | 6.566 | 3.151 |  | 0.23 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 17:58 |  | 1 |  | 3.7512 | 7.2059 | 3.4547 | 0.34153 |  |  | 1.0528 |  |  |  |  | 0.3051 |  |  |  |  | 1.7589 |  |  |  |  | 0.81342 |  |  | 3.7316 | 7.2169 | 3.4853 | 0.34014 |  | 42
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3870 | 7445 |  | 3285 | 3.384 | 6.586 | 3.202 |  | 0.228 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 18:10 |  | 1 |  | 3.7119 | 7.2279 | 3.516 | 0.33876
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5905 | 9388 |  | 3351 | 4.945 | 8.232 | 3.287 |  | 0.234 | 1 | 0 | 0 |  | 0 |  | TCpe | 2022-11-22 | 22:24 |  | 1 |  | 5.5267 | 8.9526 | 3.426 | 0.34687
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4102 | 7557 |  | 3280 | 3.562 | 6.681 | 3.119 |  | 0.227 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 22:35 |  | 1 |  | 3.8552 | 7.2229 | 3.3677 | 0.33979 |  |  | 0.38549 |  |  |  |  | 1.5952 |  |  |  |  | 2.9982 |  |  |  |  | 0.20553 |  |  | 3.8478 | 7.1657 | 3.3179 | 0.33945 |  | 43
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4086 | 7436 |  | 3273 | 3.55 | 6.578 | 3.028 |  | 0.226 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 22:46 |  | 1 |  | 3.8404 | 7.1085 | 3.2682 | 0.3391
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | Stored Cal TIC [mg/l] | Stored Cal TC [mg/l] | Stored Cal TOC (Diff.) [mg/l] | Stored Cal NPOC [mg/l] | Stored Cal TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] |  | TIC Absolute value Percent error for check standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes |  | TC Absolute value Percent error for check standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes |  | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes |  | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
[27 empty] | Known | [24 empty] | Known | 2.991 | 5.9642 | 2.991 | 0.2991
[5 empty] | Mean |  |  | 3239.6121 | 7120.0431 |  | 3282.7328 | 2.9003 | 6.3106 | 3.4103 |  | 0.22732 | [10 empty] | Mean |  | 3.436 | 6.5922 | 3.1562 | [20 empty] | Mean | 3.6281 | 7.3415 | 3.7134 | 0.34761 | Min | 1
[5 empty] | Std |  |  | 1310.0791 | 574.2788 |  | 277.0055 | 1.0051 | 0.48655 | 0.84043 |  | 0.028944 | [10 empty] | Std |  | 0.79539 | 0.81175 | 0.55078 | [20 empty] | Std | 0.26971 | 0.32452 | 0.33095 | 0.019809 | Max | 43
[5 empty] | CV |  |  | 40.4394 | 8.0657 |  | 8.4383 | 34.6559 | 7.7101 | 24.6442 |  | 12.7327 | [10 empty] | CV |  | 23.1485 | 12.3138 | 17.451 | [20 empty] | CV | 7.4338 | 4.4203 | 8.9123 | 5.6988
|  |  |  | Warning Limits | Xbar + 2SD |  |  | 5859.7702 | 8268.6007 |  | 3836.7438 | 4.9105 | 7.2837 | 5.0911 |  | 0.28521 | [9 empty] | Warning Limits | Xbar + 2SD |  | 5.0268 | 8.2157 | 4.2577 | [19 empty] | Warning Limits | Xbar + 2SD | 4.1675 | 7.9905 | 4.3753 | 0.38722
[5 empty] | Xbar - 2SD |  |  | 619.4539 | 5971.4855 |  | 2728.7217 | 0.89004 | 5.3375 | 1.7294 |  | 0.16943 | [10 empty] | Xbar - 2SD |  | 1.8452 | 4.9687 | 2.0546 | [20 empty] | Xbar - 2SD | 3.0887 | 6.6924 | 3.0515 | 0.30799
|  |  |  | Control Limits | Xbar + 3SD |  |  | 7169.8493 | 8842.8795 |  | 4113.7493 | 5.9156 | 7.7702 | 5.9315 |  | 0.31415 | [9 empty] | Control Limits | Xbar + 3SD |  | 5.8222 | 9.0274 | 4.8085 | [19 empty] | Control Limits | Xbar + 3SD | 4.4372 | 8.315 | 4.7062 | 0.40703
[5 empty] | Xbar - 3SD |  |  | -690.6252 | 5397.2067 |  | 2451.7162 | -0.11507 | 4.8509 | 0.88896 |  | 0.14049 | [10 empty] | Xbar - 3SD |  | 1.0499 | 4.1569 | 1.5038 | [20 empty] | Xbar - 3SD | 2.819 | 6.3679 | 2.7205 | 0.28818
[26 empty] | Warning Limits as % | Xbar + 2SD |  | 146.2971 | 124.6276 | 134.902 | [19 empty] | Warning Limits as % | Xbar + 2SD | 114.8676 | 108.8407 | 117.8246 | 111.3977
[27 empty] | Xbar - 2SD |  | 53.7029 | 75.3724 | 65.098 | [20 empty] | Xbar - 2SD | 85.1324 | 91.1593 | 82.1754 | 88.6023
[26 empty] | Control Limits as % | Xbar + 3SD |  | 169.4456 | 136.9415 | 152.353 | [19 empty] | Control Limits as % | Xbar + 3SD | 122.3015 | 113.261 | 126.7369 | 117.0965
[27 empty] | Xbar - 3SD |  | 30.5544 | 63.0585 | 47.647 | [20 empty] | Xbar - 3SD | 77.6985 | 86.739 | 73.2631 | 82.9035
[26 empty] | Coefficient of Variation % |  |  | 23.1485 | 12.3138 | 17.451 | [19 empty] | Coefficient of Variation % |  | 7.4338 | 4.4203 | 8.9123 | 5.6988
[26 empty] | 3 x Coefficient of Variation % |  |  | 69.4456 | 36.9415 | 52.353 | [19 empty] | 3 x Coefficient of Variation % |  | 22.3015 | 13.261 | 26.7369 | 17.0965
[51 empty] | Count |  | 17 | 17 | 17 | 17
[51 empty] | 99th percentile.inc |  | 4.1964 | 7.84 | 4.4188 | 0.37821
[51 empty] | Maximum |  | 4.235 | 7.8527 | 4.4564 | 0.37909
[51 empty] | Known |  | 2.991 | 5.9642 | 2.991 | 0.2991
[51 empty] | MDL |  | 0.69678 | 0.83838 | 0.855 | 0.051178
[51 empty] | LOQ |  | 2.6971 | 3.2452 | 3.3095 | 0.19809
[51 empty] | ratio mean/mdl |  | 5.2069 | 8.7567 | 4.3431 | 6.7921
Note: 2022 field season starts on row 111 so updating the above to find 2022 MDL (12Jan2023)

## rolling ref spiked tap
We care most about the variation across just one run.
This is to check on signal drift over a run
And regular check standards don't work
because they exchange with atmosphere and change conc
Over longer terms...
There is no guarantee that one batch of tap water will
be similar to another
Although the water authority tries to keep TOC < 1 mg/L
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
103 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2083 | 8658 |  | 3690 | 3.356 | 12.689 | 9.333 |  | 0.45 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 10:01 |  | 1 |  | 4.0943 | 18.4052 | 14.311 | 0.89281
104 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3516 | 8805 |  | 3737 | 5.188 | 12.897 | 7.71 |  | 0.458 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 10:10 |  | 1 |  | 7.2511 | 18.735 | 11.4839 | 0.90465 |  |  | 6.8368 |  |  |  |  | 0.37054 |  |  |  |  | 3.9402 |  |  |  |  | 0.27905 |  |  | 7.5078 | 18.7698 | 11.262 | 0.90339 |  | 1
105 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3749 | 8836 |  | 3727 | 5.485 | 12.94 | 7.455 |  | 0.456 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 10:20 |  | 1 |  | 7.7644 | 18.8046 | 11.0402 | 0.90213
103 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2097 | 8813 |  | 3595 | 3.373 | 12.907 | 9.535 |  | 0.433 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 11:00 |  | 1 |  | 4.1234 | 18.6598 | 14.5364 | 0.89835
104 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3569 | 9082 |  | 3695 | 5.255 | 13.288 | 8.034 |  | 0.451 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 11:10 |  | 1 |  | 7.3517 | 19.2548 | 11.9031 | 0.9244 |  |  | 4.1768 |  |  |  |  | 0.3452 |  |  |  |  | 3.244 |  |  |  |  | 1.8151 |  |  | 7.5085 | 19.2216 | 11.7131 | 0.93286 |  | 2
105 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3712 | 9052 |  | 3760 | 5.437 | 13.245 | 7.807 |  | 0.462 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 11:20 |  | 1 |  | 7.6653 | 19.1885 | 11.5231 | 0.94133
103 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2183 | 8920 |  | 3894 | 3.483 | 13.059 | 9.576 |  | 0.485 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 9:23 |  | 1 |  | 4.4491 | 18.7805 | 14.3314 | 0.93692
104 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3620 | 9512 |  | 4653 | 5.32 | 13.895 | 8.575 |  | 0.618 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 9:33 |  | 1 |  | 7.5825 | 20.0857 | 12.5032 | 1.1247 |  |  | 4.1674 |  |  |  |  | 0.55035 |  |  |  |  | 3.5238 |  |  |  |  | 0.50471 |  |  | 7.7439 | 20.0306 | 12.2867 | 1.1276 |  | 3
105 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3768 | 9462 |  | 4676 | 5.51 | 13.824 | 8.314 |  | 0.622 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 9:43 |  | 1 |  | 7.9052 | 19.9755 | 12.0703 | 1.1304
103 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2251 | 9162 |  | 3871 | 3.57 | 13.4 | 9.83 |  | 0.481 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 9:56 |  | 1 |  | 4.5063 | 18.4274 | 13.921 | 0.84202
104 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3850 | 9576 |  | 4255 | 5.614 | 13.985 | 8.372 |  | 0.548 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 10:05 |  | 1 |  | 8.0795 | 19.3014 | 11.2219 | 0.92358 |  |  | 3.1845 |  |  |  |  | 2.9952 |  |  |  |  | 2.8587 |  |  |  |  | 4.364 |  |  | 8.2102 | 19.5948 | 11.3846 | 0.94419 |  | 4
105 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3967 | 9854 |  | 4449 | 5.763 | 14.377 | 8.614 |  | 0.582 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 10:15 |  | 1 |  | 8.341 | 19.8883 | 11.5473 | 0.96479
103 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2117 | 8654 |  | 3977 | 3.399 | 12.684 | 9.285 |  | 0.5 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-09 | 7:32 |  | 1 |  | 4.0527 | 16.8335 | 12.7808 | 0.78008
104 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3503 | 8887 |  | 4264 | 5.17 | 13.013 | 7.843 |  | 0.55 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-09 | 7:40 |  | 1 |  | 7.0029 | 17.3108 | 10.3079 | 0.83785 |  |  | 6.6698 |  |  |  |  | 0.48404 |  |  |  |  | 3.9493 |  |  |  |  | 0.096144 |  |  | 7.2445 | 17.3528 | 10.1083 | 0.83744 |  | 5
105 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3730 | 8928 |  | 4260 | 5.461 | 13.07 | 7.609 |  | 0.549 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-09 | 7:50 |  | 1 |  | 7.4861 | 17.3948 | 9.9087 | 0.83704
103 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2967 | 12629 |  | 9290 | 4.485 | 18.296 | 13.811 |  | 1.426 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 10:01 |  | 1 |  | 4.1193 | 18.956 | 14.8367 | 1.3415
104 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5173 | 13104 |  | 9574 | 7.306 | 18.967 | 11.662 |  | 1.475 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 10:11 |  | 1 |  | 7.7596 | 19.713 | 11.9533 | 1.3838 |  |  | 3.284 |  |  |  |  | 2.3143 |  |  |  |  | 6.1223 |  |  |  |  | 0.78261 |  |  | 7.8892 | 19.4875 | 11.5983 | 1.3893 |  | 6
105 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5330 | 12821 |  | 9647 | 7.506 | 18.567 | 11.061 |  | 1.488 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 10:20 |  | 1 |  | 8.0187 | 19.262 | 11.2433 | 1.3947
106 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2964 | 11571 |  | 6269 | 4.481 | 16.802 | 12.32 |  | 0.899 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 15:38 |  | 1 |  | 4.0765 | 16.8153 | 12.7389 | 0.90104
107 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5070 | 11849 |  | 6450 | 7.174 | 17.194 | 10.02 |  | 0.931 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 15:49 |  | 1 |  | 7.2648 | 17.2393 | 9.9745 | 0.92819 |  |  | 4.5615 |  |  |  |  | 0.13259 |  |  |  |  | 3.2216 |  |  |  |  | 1.5803 |  |  | 7.4343 | 17.2507 | 9.8164 | 0.92092 |  | 7
108 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5294 | 11864 |  | 6353 | 7.461 | 17.216 | 9.755 |  | 0.914 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 16:02 |  | 1 |  | 7.6039 | 17.2621 | 9.6582 | 0.91364
103 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3167 | 12156 |  | 5812 | 4.741 | 17.628 | 12.888 |  | 0.82 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 11:16 |  | 1 |  | 4.3875 | 17.7747 | 13.3872 | 0.86201
104 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5419 | 12350 |  | 5978 | 7.621 | 17.903 | 10.282 |  | 0.849 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 11:27 |  | 1 |  | 7.7818 | 18.0676 | 10.2858 | 0.88748 |  |  | 4.7093 |  |  |  |  | 0.79062 |  |  |  |  | 2.2802 |  |  |  |  | 1.7091 |  |  | 7.9695 | 18.1393 | 10.1698 | 0.87996 |  | 8
105 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5668 | 12445 |  | 5880 | 7.938 | 18.036 | 10.098 |  | 0.832 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 11:38 |  | 1 |  | 8.1571 | 18.211 | 10.0539 | 0.87244
103 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3075 | 11805 |  | 5462 | 4.624 | 17.133 | 12.509 |  | 0.759 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 13:12 |  | 1 |  | 4.1103 | 16.6991 | 12.5888 | 0.79914
104 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5197 | 12103 |  | 5925 | 7.336 | 17.553 | 10.217 |  | 0.839 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 13:22 |  | 1 |  | 7.2328 | 17.1383 | 9.9056 | 0.86981 |  |  | 4.9407 |  |  |  |  | 0.034408 |  |  |  |  | 3.8305 |  |  |  |  | 0.52506 |  |  | 7.416 | 17.1354 | 9.7194 | 0.8721 |  | 9
105 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5446 | 12099 |  | 5955 | 7.654 | 17.547 | 9.893 |  | 0.845 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 13:33 |  | 1 |  | 7.5992 | 17.1324 | 9.5333 | 0.87439
4 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6313 | 12857 |  | 7563 | 8.763 | 18.618 | 9.854 |  | 1.125 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-24 | 16:00 |  | 1 |  | 9.1125 | 18.6159 | 9.5034 | 1.1611
5 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6371 | 12875 |  | 7386 | 8.837 | 18.643 | 9.806 |  | 1.094 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-24 | 16:11 |  | 1 |  | 9.2016 | 18.643 | 9.4414 | 1.1333 |  |  | 0.066771 |  |  |  |  | 1.3562 |  |  |  |  | 2.7628 |  |  |  |  | 3.1235 |  |  | 9.2047 | 18.5174 | 9.3127 | 1.1159 |  | 10
6 | 2 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6375 | 12708 |  | 7164 | 8.843 | 18.408 | 9.565 |  | 1.055 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-24 | 16:22 |  | 1 |  | 9.2077 | 18.3918 | 9.1841 | 1.0985
154 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4524 | 11513 |  | 6774 | 6.476 | 16.721 | 10.245 |  | 0.988 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 2:11 |  | 1 |  | 6.4684 | 16.7572 | 10.2888 | 1.13
155 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5233 | 11575 |  | 6746 | 7.382 | 16.808 | 9.426 |  | 0.983 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 2:21 |  | 1 |  | 7.5548 | 16.8526 | 9.2978 | 1.125 |  |  | 0.16213 |  |  |  |  | 0.14619 |  |  |  |  | 0.39742 |  |  |  |  | 1.2452 |  |  | 7.5609 | 16.8403 | 9.2794 | 1.1181 |  | 11
156 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5241 | 11559 |  | 6667 | 7.393 | 16.785 | 9.393 |  | 0.969 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 2:32 |  | 1 |  | 7.567 | 16.828 | 9.2609 | 1.1111
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3530 | 12343 |  | 7085 | 5.206 | 17.892 | 12.686 |  | 1.042 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 11:23 |  | 1 |  | 6.065 | 19.7914 | 13.7264 | 1.2613
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5914 | 12413 |  | 6633 | 8.253 | 17.991 | 9.737 |  | 0.963 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 11:32 |  | 1 |  | 9.8057 | 19.9006 | 10.095 | 1.1831 | 100 |  | 5.0992 |  |  | 100 |  | 0.0078423 |  |  | 100 |  | 5.1989 |  |  | 100 |  | 4.1051 |  |  | 10.0622 | 19.9014 | 9.8392 | 1.1593 |  | 12
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6241 | 12414 |  | 6358 | 8.672 | 17.992 | 9.321 |  | 0.915 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 11:42 |  | 1 |  | 10.3187 | 19.9022 | 9.5835 | 1.1355
109 | 4 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3446 | 12325 |  | 6730 | 5.098 | 17.866 | 12.768 |  | 0.98 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 11:24 |  | 1 |  | 5.9332 | 19.7633 | 13.8301 | 1.1999
110 | 4 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6031 | 12509 |  | 6826 | 8.403 | 18.126 | 9.724 |  | 0.997 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 11:34 |  | 1 |  | 9.9892 | 20.0505 | 10.0612 | 1.2165 | 101.1271 |  | 3.5487 |  |  | 100.9828 |  | 0.4117 |  |  | 88.5093 |  | 2.8035 |  |  | 105.4422 |  | 0.65225 |  |  | 10.1697 | 20.0918 | 9.9221 | 1.2205 |  | 13
111 | 4 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6261 | 12562 |  | 6872 | 8.697 | 18.202 | 9.505 |  | 1.004 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 11:45 |  | 1 |  | 10.3501 | 20.1332 | 9.7831 | 1.2244
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7823 | 12767 |  | 7219 | 10.694 | 18.491 | 7.797 |  | 1.065 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 14:21 |  | 1 |  | 12.1831 | 19.9473 | 7.7642 | 1.1204
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7263 | 12756 |  | 7223 | 9.979 | 18.476 | 8.497 |  | 1.066 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 14:31 |  | 1 |  | 11.3311 | 19.9306 | 8.5994 | 1.121 | 100 |  | 1.2429 |  |  | 100 |  | 0.22895 |  |  | 100 |  | 2.1358 |  |  | 100 |  | 1.7412 |  |  | 11.2611 | 19.9534 | 8.6923 | 1.1308 |  | 14
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7171 | 12786 |  | 7353 | 9.861 | 18.517 | 8.657 |  | 1.088 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 14:42 |  | 1 |  | 11.1912 | 19.9762 | 8.7851 | 1.1406
109 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3057 | 11127 |  | 6946 | 4.601 | 16.175 | 11.574 |  | 1.017 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 14:38 |  | 1 |  | 4.9323 | 17.45 | 12.5177 | 1.079
110 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5242 | 11250 |  | 7023 | 7.394 | 16.348 | 8.954 |  | 1.031 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 14:48 |  | 1 |  | 8.2565 | 17.6373 | 9.3808 | 1.0907 | 74.5254 |  | 4.907 |  |  | 87.8475 |  | 0.53673 |  |  | 82.7389 |  | 5.5856 |  |  | 96.5011 |  | 0.27735 |  |  | 8.4641 | 17.5901 | 9.126 | 1.0922 |  | 15
111 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5515 | 11188 |  | 7043 | 7.743 | 16.262 | 8.519 |  | 1.034 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 14:59 |  | 1 |  | 8.6718 | 17.5429 | 8.8711 | 1.0937
124 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4199 | 11413 |  | 6937 | 6.061 | 16.579 | 10.519 |  | 1.016 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 18:03 |  | 1 |  | 6.6697 | 17.8855 | 11.2158 | 1.0776
125 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5405 | 11339 |  | 6944 | 7.602 | 16.475 | 8.872 |  | 1.017 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 18:14 |  | 1 |  | 8.5044 | 17.7728 | 9.2684 | 1.0787 | 75.6408 |  | 1.9135 |  |  | 88.9515 |  | 0.35921 |  |  | 83.7438 |  | 1.0886 |  |  | 95.6229 |  | 0.69959 |  |  | 8.5866 | 17.8048 | 9.2182 | 1.0825 |  | 16
126 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5513 | 11381 |  | 6994 | 7.741 | 16.533 | 8.793 |  | 1.026 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 18:24 |  | 1 |  | 8.6687 | 17.8368 | 9.168 | 1.0863
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3053 | 10428 |  | 6247 | 4.595 | 15.189 | 10.594 |  | 0.896 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 2:54 |  | 1 |  | 5.0489 | 16.5098 | 11.461 | 0.97394
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5150 | 10481 |  | 6215 | 7.276 | 15.263 | 7.987 |  | 0.89 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 3:04 |  | 1 |  | 8.3439 | 16.5911 | 8.2472 | 0.96933 | 90.3913 |  | 5.53 |  |  | 90.6823 |  | 0.46105 |  |  | 79.6739 |  | 4.9435 |  |  | 95.3826 |  | 0.69651 |  |  | 8.5812 | 16.6295 | 8.0483 | 0.97272 |  | 17
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5452 | 10531 |  | 6262 | 7.662 | 15.333 | 7.671 |  | 0.898 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 3:15 |  | 1 |  | 8.8185 | 16.6678 | 7.8493 | 0.97611
109 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3051 | 9779 |  | 5952 | 4.592 | 14.273 | 9.68 |  | 0.844 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 12:18 |  | 1 |  | 5.0457 | 15.5147 | 10.4689 | 0.93142
110 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5243 | 9802 |  | 5862 | 7.396 | 14.304 | 6.908 |  | 0.828 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 12:29 |  | 1 |  | 8.49 | 15.5499 | 7.0599 | 0.91844 | 90.8517 |  | 3.0976 |  |  | 84.623 |  | 0.039437 |  |  | 74.4932 |  | 3.7664 |  |  | 90.6964 |  | 2.1734 |  |  | 8.6236 | 15.553 | 6.9294 | 0.92854 |  | 18
111 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5413 | 9806 |  | 6002 | 7.612 | 14.31 | 6.698 |  | 0.853 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 12:39 |  | 1 |  | 8.7572 | 15.5561 | 6.7989 | 0.93863
22 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5337 | 11804 |  | 6516 | 7.516 | 17.131 | 9.615 |  | 0.942 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:35 |  | 1 |  | 8.8494 | 20.3163 | 11.4669 | 1.0874
23 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5366 | 12082 |  | 6436 | 7.552 | 17.524 | 9.972 |  | 0.929 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:46 |  | 1 |  | 8.8974 | 20.7968 | 11.8994 | 1.0742 | 100 |  | 0.92621 |  |  | 100 |  | 2.704 |  |  | 100 |  | 5.506 |  |  | 100 |  | 1.4544 |  |  | 8.9388 | 20.5194 | 11.5806 | 1.0665 |  | 19
24 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5416 | 11761 |  | 6342 | 7.617 | 17.07 | 9.453 |  | 0.912 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 20:56 |  | 1 |  | 8.9802 | 20.242 | 11.2618 | 1.0587
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2855 | 10731 |  | 6516 | 4.342 | 15.616 | 11.274 |  | 0.942 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:37 |  | 1 |  | 4.7396 | 18.4617 | 13.7221 | 1.0874
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4729 | 10802 |  | 6479 | 6.738 | 15.716 | 8.978 |  | 0.936 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:47 |  | 1 |  | 7.8426 | 18.5844 | 10.7417 | 1.0813 | 90.3821 |  | 5.8814 |  |  | 90.0726 |  | 1.1976 |  |  | 88.3581 |  | 6.701 |  |  | 101.4165 |  | 0.015258 |  |  | 8.0803 | 18.4738 | 10.3935 | 1.0814 |  | 20
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5016 | 10674 |  | 6480 | 7.106 | 15.536 | 8.43 |  | 0.936 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:58 |  | 1 |  | 8.3179 | 18.3632 | 10.0453 | 1.0815
109 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2859 | 10984 |  | 6743 | 4.347 | 15.974 | 11.627 |  | 0.982 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:53 |  | 1 |  | 4.7462 | 18.899 | 14.1528 | 1.1249
110 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4840 | 11094 |  | 6993 | 6.88 | 16.129 | 9.249 |  | 1.026 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 16:03 |  | 1 |  | 8.0264 | 19.0891 | 11.0627 | 1.1661 | 91.6806 |  | 4.1416 |  |  | 92.8574 |  | 0.43557 |  |  | 85.7445 |  | 3.8927 |  |  | 108.4207 |  | 1.8854 |  |  | 8.1962 | 19.0476 | 10.8514 | 1.1552 |  | 21
111 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5045 | 11046 |  | 6861 | 7.142 | 16.061 | 8.918 |  | 1.003 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 16:14 |  | 1 |  | 8.3659 | 19.0061 | 10.6402 | 1.1444
22 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5436 | 11813 |  | 6235 | 7.643 | 17.144 | 9.501 |  | 0.894 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 20:46 |  | 1 |  | 9.1314 | 20.2312 | 11.0998 | 1.1058
23 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5318 | 11834 |  | 6169 | 7.492 | 17.174 | 9.682 |  | 0.882 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 20:56 |  | 1 |  | 8.9387 | 20.2653 | 11.3266 | 1.0945 | 100 |  | 0.091394 |  |  | 100 |  | 1.0967 |  |  | 100 |  | 1.8973 |  |  | 100 |  | 0.55267 |  |  | 8.9346 | 20.1548 | 11.2202 | 1.0914 |  | 22
24 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5313 | 11698 |  | 6134 | 7.485 | 16.982 | 9.497 |  | 0.876 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:08 |  | 1 |  | 8.9305 | 20.0443 | 11.1137 | 1.0884
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2781 | 11134 |  | 5855 | 4.247 | 16.186 | 11.939 |  | 0.827 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:46 |  | 1 |  | 4.7954 | 19.1276 | 14.3322 | 1.0403
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4683 | 11098 |  | 5984 | 6.68 | 16.134 | 9.455 |  | 0.85 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:56 |  | 1 |  | 7.9017 | 19.0691 | 11.1674 | 1.0626 | 91.0733 |  | 6.3247 |  |  | 93.63 |  | 1.399 |  |  | 80.7706 |  | 7.247 |  |  | 96.4074 |  | 1.7507 |  |  | 8.1597 | 18.9366 | 10.7769 | 1.0534 |  | 23
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4999 | 10935 |  | 5877 | 7.083 | 15.905 | 8.821 |  | 0.831 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 7:07 |  | 1 |  | 8.4177 | 18.8041 | 10.3864 | 1.0441
22 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5436 | 11813 |  | 6235 | 7.643 | 17.144 | 9.501 |  | 0.894 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 20:46 |  | 1 |  | 9.1314 | 20.2312 | 11.0998 | 1.1058
23 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5318 | 11834 |  | 6169 | 7.492 | 17.174 | 9.682 |  | 0.882 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 20:56 |  | 1 |  | 8.9387 | 20.2653 | 11.3266 | 1.0945 | 100 |  | 0.091394 |  |  | 100 |  | 1.0967 |  |  | 100 |  | 1.8973 |  |  | 100 |  | 0.55267 |  |  | 8.9346 | 20.1548 | 11.2202 | 1.0914 |  | 24
24 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5313 | 11698 |  | 6134 | 7.485 | 16.982 | 9.497 |  | 0.876 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:08 |  | 1 |  | 8.9305 | 20.0443 | 11.1137 | 1.0884
67 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2781 | 11134 |  | 5855 | 4.247 | 16.186 | 11.939 |  | 0.827 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:46 |  | 1 |  | 4.7954 | 19.1276 | 14.3322 | 1.0403
68 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4683 | 11098 |  | 5984 | 6.68 | 16.134 | 9.455 |  | 0.85 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:56 |  | 1 |  | 7.9017 | 19.0691 | 11.1674 | 1.0626 | 91.0733 |  | 6.3247 |  |  | 93.63 |  | 1.399 |  |  | 80.7706 |  | 7.247 |  |  | 96.4074 |  | 1.7507 |  |  | 8.1597 | 18.9366 | 10.7769 | 1.0534 |  | 25
69 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4999 | 10935 |  | 5877 | 7.083 | 15.905 | 8.821 |  | 0.831 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 7:07 |  | 1 |  | 8.4177 | 18.8041 | 10.3864 | 1.0441
109 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2842 | 10666 |  | 5450 | 4.326 | 15.525 | 11.199 |  | 0.757 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 16:10 |  | 1 |  | 4.895 | 18.3669 | 13.4719 | 0.97054
110 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4768 | 10742 |  | 5493 | 6.788 | 15.631 | 8.843 |  | 0.764 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 16:20 |  | 1 |  | 8.0405 | 18.4905 | 10.45 | 0.97795 | 91.5624 |  | 3.9424 |  |  | 91.0675 |  | 0.4758 |  |  | 80.9743 |  | 4.0132 |  |  | 89.8805 |  | 1.2609 |  |  | 8.2022 | 18.4466 | 10.2444 | 0.98416 |  | 26
111 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4966 | 10688 |  | 5565 | 7.041 | 15.555 | 8.514 |  | 0.777 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 16:31 |  | 1 |  | 8.3638 | 18.4027 | 10.0388 | 0.99036
25 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5203 | 11811 |  | 5693 | 7.344 | 17.142 | 9.798 |  | 0.799 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 17:07 |  | 1 |  | 8.4034 | 19.4624 | 11.059 | 0.95582
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5027 | 11688 |  | 5556 | 7.119 | 16.967 | 9.849 |  | 0.775 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 17:17 |  | 1 |  | 8.1154 | 19.261 | 11.1456 | 0.93352 | 100 |  | 0.99302 |  |  | 100 |  | 0.3833 |  |  | 100 |  | 0.05834 |  |  | 100 |  | 0.19167 |  |  | 8.0753 | 19.2241 | 11.1488 | 0.93441 |  | 27
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4978 | 11643 |  | 5567 | 7.057 | 16.904 | 9.847 |  | 0.777 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 17:28 |  | 1 |  | 8.0352 | 19.1873 | 11.1521 | 0.93531
70 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2620 | 10445 |  | 4977 | 4.042 | 15.212 | 11.171 |  | 0.674 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 3:17 |  | 1 |  | 4.1763 | 17.2256 | 13.0493 | 0.83925
71 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4291 | 10378 |  | 4868 | 6.177 | 15.118 | 8.94 |  | 0.655 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 3:28 |  | 1 |  | 6.9109 | 17.1159 | 10.205 | 0.8215 | 85.8971 |  | 0.28376 |  |  | 88.5346 |  | 0.96128 |  |  | 90.5148 |  | 1.8133 |  |  | 87.4314 |  | 0.21825 |  |  | 6.9207 | 17.034 | 10.1133 | 0.82061 |  | 28
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4303 | 10278 |  | 4857 | 6.193 | 14.977 | 8.784 |  | 0.653 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 3:39 |  | 1 |  | 6.9305 | 16.9522 | 10.0216 | 0.81971
112 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2686 | 10613 |  | 5260 | 4.126 | 15.45 | 11.324 |  | 0.724 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 12:43 |  | 1 |  | 4.2843 | 17.5007 | 13.2164 | 0.88533
113 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4365 | 10505 |  | 5164 | 6.273 | 15.297 | 9.024 |  | 0.707 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 12:54 |  | 1 |  | 7.032 | 17.3239 | 10.2919 | 0.8697 | 87.1164 |  | 0.32635 |  |  | 89.5718 |  | 1.0643 |  |  | 91.4153 |  | 1.5716 |  |  | 92.3492 |  | 1.0539 |  |  | 7.0205 | 17.2322 | 10.2116 | 0.86514 |  | 29
114 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4351 | 10393 |  | 5108 | 6.254 | 15.139 | 8.885 |  | 0.697 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 13:04 |  | 1 |  | 7.0091 | 17.1405 | 10.1314 | 0.86058
25 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5206 | 11376 |  | 5213 | 7.348 | 16.527 | 9.179 |  | 0.715 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 18:54 |  | 1 |  | 8.5312 | 18.3759 | 9.8447 | 0.93115
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4865 | 11213 |  | 5203 | 6.912 | 16.296 | 9.384 |  | 0.714 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 19:04 |  | 1 |  | 7.9754 | 18.1161 | 10.1407 | 0.92938 | 100 |  | 0.96524 |  |  | 100 |  | 1.4265 |  |  | 100 |  | 1.7908 |  |  | 100 |  | 1.4818 |  |  | 7.9371 | 17.9878 | 10.0507 | 0.92254 |  | 30
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4818 | 11052 |  | 5126 | 6.852 | 16.069 | 9.217 |  | 0.7 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 19:15 |  | 1 |  | 7.8988 | 17.8595 | 9.9607 | 0.91571
70 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 1123 | 1845 |  | 1136 | 2.127 | 3.069 | 0.942 |  | 0.005 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 4:57 |  | 1 |  | 1.8758 | 3.1856 | 1.3099 | 0.20732
71 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 589 | 8876 |  | 3963 | 1.445 | 12.997 | 11.552 |  | 0.497 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 5:06 |  | 1 |  | 1.0053 | 14.3915 | 13.3861 | 0.70923 | 43.881 |  | 142.6871 |  |  | 80.1931 |  | 1.1342 |  |  | 108.1386 |  | 44.1548 |  |  | 76.019 |  | 1.8698 |  |  | 3.5083 | 14.4736 | 10.9653 | 0.70266 |  | 31
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3660 | 8979 |  | 3889 | 5.372 | 13.142 | 7.77 |  | 0.485 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 5:17 |  | 1 |  | 6.0112 | 14.5556 | 8.5444 | 0.69609
112 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2610 | 9893 |  | 4512 | 4.029 | 14.432 | 10.404 |  | 0.593 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 14:20 |  | 1 |  | 4.2997 | 16.0123 | 11.7127 | 0.8067
113 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4149 | 9954 |  | 4539 | 5.996 | 14.519 | 8.522 |  | 0.598 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 14:30 |  | 1 |  | 6.8083 | 16.1096 | 9.3013 | 0.81149 | 87.9893 |  | 100 |  |  | 89.5845 |  | 100 |  |  | 90.8123 |  | 100 |  |  | 88.4694 |  | 100 |  |  | 6.8083 | 16.1096 | 9.3013 | 0.81149 |  | 32
114 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4371 | 9992 |  | 4599 | 6.281 | 14.572 | 8.292 |  | 0.608 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 14:41
133 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5248 | 10787 |  | 4650 | 7.402 | 15.696 | 8.294 |  | 0.617 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 19:06 |  | 1 |  | 9.2602 | 18.4787 | 9.2185 | 0.85942
134 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5149 | 10825 |  | 4629 | 7.275 | 15.749 | 8.473 |  | 0.614 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 19:16 |  | 1 |  | 9.0931 | 18.542 | 9.4489 | 0.8557 | 100 |  | 2.3284 |  |  | 100 |  | 2.4993 |  |  | 100 |  | 2.6639 |  |  | 100 |  | 3.0657 |  |  | 8.9884 | 18.3131 | 9.3247 | 0.84278 |  | 33
135 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5025 | 10550 |  | 4483 | 7.117 | 15.361 | 8.243 |  | 0.588 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 19:27 |  | 1 |  | 8.8838 | 18.0843 | 9.2005 | 0.82986
178 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2886 | 9171 |  | 3914 | 4.382 | 13.414 | 9.032 |  | 0.489 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 5:13 |  | 1 |  | 5.2735 | 15.7891 | 10.5156 | 0.72917
179 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4818 | 9352 |  | 3856 | 6.852 | 13.669 | 6.817 |  | 0.479 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 5:24 |  | 1 |  | 8.5344 | 16.0904 | 7.556 | 0.7189 | 92.127 |  | 5.3402 |  |  | 87.0035 |  | 1.1129 |  |  | 82.3498 |  | 3.4571 |  |  | 84.8112 |  | 0.39309 |  |  | 8.3125 | 16.0013 | 7.6889 | 0.72032 |  | 34
180 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4555 | 9245 |  | 3872 | 6.516 | 13.517 | 7.001 |  | 0.482 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 5:34 |  | 1 |  | 8.0905 | 15.9123 | 7.8218 | 0.72174
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5708 | 8809 |  | 3655 | 7.99 | 12.903 | 4.912 |  | 0.444 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 19:47 |  | 1 |  | 9.0753 | 14.0795 | 5.0041 | 0.59499
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5859 | 8639 |  | 3590 | 8.183 | 12.663 | 4.479 |  | 0.432 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 19:57 |  | 1 |  | 9.3142 | 13.8064 | 4.4922 | 0.58437 | 100 |  | 1.249 |  |  | 100 |  | 0.25629 |  |  | 100 |  | 3.4512 |  |  | 100 |  | 2.4902 |  |  | 9.3728 | 13.7887 | 4.416 | 0.57719 |  | 35
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5933 | 8617 |  | 3502 | 8.277 | 12.631 | 4.354 |  | 0.417 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 20:08 |  | 1 |  | 9.4313 | 13.7711 | 4.3398 | 0.57
71 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3578 | 6707 |  | 2791 | 5.266 | 9.934 | 4.668 |  | 0.293 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 5:54 |  | 1 |  | 5.7058 | 10.703 | 4.9972 | 0.45387
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6107 | 6731 |  | 2792 | 8.5 | 9.968 | 1.468 |  | 0.293 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 6:04 |  | 1 |  | 9.7065 | 10.7416 | 1.035 | 0.45403 | 104.8083 |  | 2.3356 |  |  | 77.8802 |  | 0.34454 |  |  | 19.7657 |  | 29.5314 |  |  | 79.0891 |  | 0.89531 |  |  | 9.8212 | 10.7231 | 0.90185 | 0.45608 |  | 36
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6252 | 6708 |  | 2817 | 8.686 | 9.936 | 1.251 |  | 0.298 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 6:15 |  | 1 |  | 9.9359 | 10.7046 | 0.76868 | 0.45812
113 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 1677 | 7088 |  | 3521 | 2.836 | 10.472 | 7.636 |  | 0.42 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 15:10 |  | 1 |  | 2.6985 | 11.315 | 8.6165 | 0.5731
114 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3850 | 6815 |  | 3128 | 5.614 | 10.087 | 4.473 |  | 0.352 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 15:20 |  | 1 |  | 6.1361 | 10.8765 | 4.7404 | 0.50891 | 85.6089 |  | 47.18 |  |  | 79.213 |  | 0.57432 |  |  | 65.41 |  | 129.499 |  |  | 89.2837 |  | 2.4098 |  |  | 8.0305 | 10.9078 | 2.8773 | 0.51512 |  | 37
115 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6245 | 6854 |  | 3204 | 8.676 | 10.141 | 1.465 |  | 0.365 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 15:31 |  | 1 |  | 9.9249 | 10.9391 | 1.0143 | 0.52133
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5307 | 10377 |  | 4875 | 7.478 | 15.116 | 7.639 |  | 0.656 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 19:09 |  | 1 |  | 8.7313 | 17.2881 | 8.5568 | 0.87507
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5194 | 10386 |  | 4820 | 7.333 | 15.129 | 7.796 |  | 0.647 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 19:19 |  | 1 |  | 8.5505 | 17.3029 | 8.7524 | 0.86534 | 100 |  | 0.97792 |  |  | 100 |  | 0.1237 |  |  | 100 |  | 0.70383 |  |  | 100 |  | 0.24503 |  |  | 8.5089 | 17.2922 | 8.7833 | 0.8664 |  | 38
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5142 | 10373 |  | 4832 | 7.266 | 15.111 | 7.844 |  | 0.649 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 19:30 |  | 1 |  | 8.4673 | 17.2815 | 8.8142 | 0.86746
71 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2934 | 8838 |  | 3944 | 4.443 | 12.943 | 8.5 |  | 0.494 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 5:14 |  | 1 |  | 4.9341 | 14.7559 | 9.8218 | 0.71037
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4914 | 8968 |  | 3968 | 6.974 | 13.127 | 6.153 |  | 0.498 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 5:24 |  | 1 |  | 8.1025 | 14.9698 | 6.8673 | 0.71461 | 97.4845 |  | 4.7808 |  |  | 86.0639 |  | 0.77237 |  |  | 74.7386 |  | 7.7447 |  |  | 81.4961 |  | 1.7481 |  |  | 8.3009 | 14.9122 | 6.6113 | 0.70842 |  | 39
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5162 | 8898 |  | 3898 | 7.292 | 13.029 | 5.737 |  | 0.486 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 5:35 |  | 1 |  | 8.4993 | 14.8546 | 6.3553 | 0.70223
113 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3173 | 8054 |  | 3635 | 4.748 | 11.837 | 7.089 |  | 0.44 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 14:34 |  | 1 |  | 5.3165 | 13.4659 | 8.1494 | 0.6557
114 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5304 | 8131 |  | 3829 | 7.473 | 11.944 | 4.471 |  | 0.474 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 14:45 |  | 1 |  | 8.7265 | 13.5926 | 4.866 | 0.69002 | 103.6184 |  | 1.8531 |  |  | 77.8361 |  | 1.2669 |  |  | 52.269 |  | 7.1153 |  |  | 77.8492 |  | 3.7614 |  |  | 8.8081 | 13.507 | 4.6989 | 0.67728 |  | 40
115 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5406 | 8027 |  | 3685 | 7.604 | 11.798 | 4.194 |  | 0.449 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 14:55 |  | 1 |  | 8.8897 | 13.4215 | 4.5317 | 0.66455
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5421 | 11352 |  | 4633 | 7.623 | 16.492 | 8.87 |  | 0.614 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 15:10 |  | 1 |  | 8.6891 | 19.1455 | 10.4564 | 0.84758
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5220 | 11411 |  | 4811 | 7.366 | 16.577 | 9.211 |  | 0.645 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 15:20 |  | 1 |  | 8.3653 | 19.2454 | 10.8801 | 0.87953 | 100 |  | 0.21207 |  |  | 100 |  | 0.56466 |  |  | 100 |  | 0.83659 |  |  | 100 |  | 2.0618 |  |  | 8.3565 | 19.1912 | 10.8348 | 0.87055 |  | 41
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5209 | 11347 |  | 4711 | 7.352 | 16.486 | 9.134 |  | 0.628 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 15:31 |  | 1 |  | 8.3476 | 19.137 | 10.7894 | 0.86158
71 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2251 | 10167 |  | 4364 | 3.569 | 14.82 | 11.251 |  | 0.567 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 1:02 |  | 1 |  | 3.5821 | 17.1391 | 13.557 | 0.7993
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4743 | 10107 |  | 4384 | 6.756 | 14.736 | 7.979 |  | 0.571 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 1:12 |  | 1 |  | 7.5968 | 17.0375 | 9.4406 | 0.80289 | 91.0921 |  | 0.29646 |  |  | 88.8435 |  | 0.049678 |  |  | 86.9414 |  | 0.14935 |  |  | 92.113 |  | 0.067046 |  |  | 7.6081 | 17.0417 | 9.4336 | 0.80315 |  | 42
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4757 | 10112 |  | 4387 | 6.773 | 14.743 | 7.969 |  | 0.571 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 1:23 |  | 1 |  | 7.6194 | 17.0459 | 9.4265 | 0.80342
113 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2657 | 9147 |  | 3884 | 4.089 | 13.38 | 9.291 |  | 0.484 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 10:17 |  | 1 |  | 4.2362 | 15.412 | 11.1758 | 0.71314
114 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4449 | 9296 |  | 3914 | 6.38 | 13.591 | 7.211 |  | 0.489 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 10:27 |  | 1 |  | 7.1232 | 15.6643 | 8.5411 | 0.71852 | 86.8444 |  | 3.5326 |  |  | 81.2593 |  | 1.0759 |  |  | 76.535 |  | 5.0879 |  |  | 82.6507 |  | 1.0437 |  |  | 7.2513 | 15.5805 | 8.3292 | 0.72229 |  | 43
115 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4608 | 9197 |  | 3956 | 6.583 | 13.451 | 6.868 |  | 0.496 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 10:37 |  | 1 |  | 7.3794 | 15.4967 | 8.1173 | 0.72606
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5649 | 12711 |  | 4870 | 7.915 | 18.412 | 10.498 |  | 0.656 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 19:48 |  | 1 |  | 9.4026 | 20.8027 | 11.4001 | 0.84607
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5476 | 12816 |  | 4785 | 7.693 | 18.56 | 10.867 |  | 0.641 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 19:58 |  | 1 |  | 9.1213 | 20.9722 | 11.8509 | 0.83165 | 100 |  | 1.0285 |  |  | 100 |  | 1.0449 |  |  | 100 |  | 2.6705 |  |  | 100 |  | 0.49092 |  |  | 9.1685 | 20.8632 | 11.6948 | 0.82961 |  | 44
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5534 | 12681 |  | 4761 | 7.767 | 18.37 | 10.603 |  | 0.637 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 20:09 |  | 1 |  | 9.2156 | 20.7542 | 11.5386 | 0.82757
71 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2958 | 12097 |  | 4697 | 4.474 | 17.545 | 13.071 |  | 0.625 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 5:38 |  | 1 |  | 5.0275 | 19.8111 | 14.7837 | 0.81671
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4968 | 12084 |  | 4682 | 7.044 | 17.526 | 10.482 |  | 0.623 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 5:48 |  | 1 |  | 8.2954 | 19.7902 | 11.4948 | 0.81417 | 91.7257 |  | 3.1444 |  |  | 95.0033 |  | 0.4478 |  |  | 97.4943 |  | 1.5446 |  |  | 98.4391 |  | 0.68546 |  |  | 8.4279 | 19.8346 | 11.4067 | 0.81697 |  | 45
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5131 | 12139 |  | 4715 | 7.253 | 17.605 | 10.352 |  | 0.629 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 5:59 |  | 1 |  | 8.5604 | 19.879 | 11.3186 | 0.81977
113 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2935 | 12115 |  | 4677 | 4.445 | 17.57 | 13.125 |  | 0.622 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 14:59 |  | 1 |  | 4.9901 | 19.8402 | 14.8501 | 0.81332
114 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4847 | 12307 |  | 4768 | 6.889 | 17.841 | 10.953 |  | 0.638 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 15:09 |  | 1 |  | 8.0987 | 20.1503 | 12.0516 | 0.82876 | 90.5813 |  | 5.4484 |  |  | 96.4819 |  | 0.11226 |  |  | 100.9664 |  | 4.0311 |  |  | 99.2772 |  | 1.2154 |  |  | 8.3255 | 20.139 | 11.8135 | 0.82376 |  | 46
115 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5126 | 12293 |  | 4709 | 7.246 | 17.821 | 10.575 |  | 0.627 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 15:20 |  | 1 |  | 8.5523 | 20.1277 | 11.5754 | 0.81875
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5556 | 12132 |  | 4635 | 7.796 | 17.594 | 9.798 |  | 0.615 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 21:13 |  | 1 |  | 9.4301 | 20.1017 | 10.6716 | 0.81477
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5485 | 12228 |  | 4644 | 7.705 | 17.73 | 10.025 |  | 0.616 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 21:23 |  | 1 |  | 9.3125 | 20.2594 | 10.9469 | 0.81631 | 100 |  | 0.42594 |  |  | 100 |  | 1.314 |  |  | 100 |  | 2.8182 |  |  | 100 |  | 1.893 |  |  | 9.3324 | 20.1271 | 10.7947 | 0.80866 |  | 47
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5509 | 12067 |  | 4555 | 7.736 | 17.502 | 9.767 |  | 0.601 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 21:34 |  | 1 |  | 9.3523 | 19.9949 | 10.6426 | 0.80101
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 2950 | 10979 |  | 5530 | 4.464 | 15.966 | 11.502 |  | 0.771 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 7:04 |  | 1 |  | 5.1139 | 18.2077 | 13.0938 | 0.96871
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4363 | 10816 |  | 5478 | 6.271 | 15.736 | 9.466 |  | 0.762 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 7:14 |  | 1 |  | 7.4542 | 17.94 | 10.4858 | 0.95976 | 80.9532 |  | 3.7928 |  |  | 89.434 |  | 0.87516 |  |  | 96.4439 |  | 1.2523 |  |  | 118.4911 |  | 1.0087 |  |  | 7.5983 | 18.0188 | 10.4205 | 0.95495 |  | 48
74 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4537 | 10912 |  | 5422 | 6.492 | 15.872 | 9.38 |  | 0.752 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 7:25 |  | 1 |  | 7.7424 | 18.0977 | 10.3553 | 0.95013
115 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3171 | 10830 |  | 5466 | 4.746 | 15.756 | 11.01 |  | 0.759 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 16:19 |  | 1 |  | 5.4799 | 17.963 | 12.483 | 0.9577
116 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4436 | 10609 |  | 5439 | 6.364 | 15.444 | 9.08 |  | 0.755 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 16:29 |  | 1 |  | 7.5751 | 17.5999 | 10.0248 | 0.95306 | 81.9993 |  | 3.0785 |  |  | 87.6518 |  | 0.71609 |  |  | 92.3239 |  | 1.107 |  |  | 117.8824 |  | 0.6155 |  |  | 7.6935 | 17.6632 | 9.9697 | 0.95013 |  | 49
117 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4579 | 10686 |  | 5405 | 6.546 | 15.553 | 9.007 |  | 0.749 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 16:39 |  | 1 |  | 7.8119 | 17.7264 | 9.9145 | 0.94721
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5570 | 11650 |  | 4947 | 7.814 | 16.913 | 9.099 |  | 0.669 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 18:12 |  | 1 |  | 8.1263 | 18.4121 | 10.2858 | 0.801
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6100 | 11617 |  | 4939 | 8.491 | 16.868 | 8.377 |  | 0.668 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 18:24 |  | 1 |  | 8.9081 | 18.3605 | 9.4524 | 0.79967 | 100 |  | 0.66464 |  |  | 100 |  | 0.88973 |  |  | 100 |  | 2.3327 |  |  | 100 |  | 0.99209 |  |  | 8.8786 | 18.4426 | 9.5639 | 0.80366 |  | 50
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6060 | 11722 |  | 4987 | 8.44 | 17.015 | 8.575 |  | 0.676 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 18:35 |  | 1 |  | 8.8491 | 18.5246 | 9.6755 | 0.80764
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3380 | 11043 |  | 4842 | 5.013 | 16.056 | 11.043 |  | 0.651 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 4:39 |  | 1 |  | 4.8954 | 17.4635 | 12.5681 | 0.78356
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5214 | 11151 |  | 4910 | 7.358 | 16.209 | 8.852 |  | 0.663 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 4:49 |  | 1 |  | 7.6011 | 17.6323 | 10.0312 | 0.79485 | 88.4211 |  | 6.0967 |  |  | 95.0898 |  | 0.97076 |  |  | 102.3437 |  | 6.6789 |  |  | 98.8616 |  | 0.14639 |  |  | 7.8401 | 17.5471 | 9.707 | 0.79427 |  | 51
74 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5538 | 11042 |  | 4903 | 7.773 | 16.055 | 8.282 |  | 0.661 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 5:00 |  | 1 |  | 8.079 | 17.4619 | 9.3829 | 0.79369
115 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3591 | 11348 |  | 4879 | 5.283 | 16.487 | 11.204 |  | 0.657 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 14:30 |  | 1 |  | 5.2067 | 17.9401 | 12.7334 | 0.78971
116 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5408 | 11480 |  | 4913 | 7.607 | 16.674 | 9.067 |  | 0.663 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 14:40 |  | 1 |  | 7.8873 | 18.1464 | 10.2592 | 0.79535 | 90.4523 |  | 3.3653 |  |  | 97.3092 |  | 2.1676 |  |  | 104.7679 |  | 6.6377 |  |  | 98.3881 |  | 1.261 |  |  | 8.0222 | 17.9519 | 9.9296 | 0.79037 |  | 52
117 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5591 | 11231 |  | 4853 | 7.841 | 16.322 | 8.482 |  | 0.653 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 14:50 |  | 1 |  | 8.1572 | 17.7573 | 9.6001 | 0.78539
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5746 | 11635 |  | 4982 | 8.039 | 16.893 | 8.855 |  | 0.675 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 20:29 |  | 1 |  | 8.4238 | 18.4344 | 10.0106 | 0.80869
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5223 | 11738 |  | 4925 | 7.37 | 17.038 | 9.668 |  | 0.665 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 20:39 |  | 1 |  | 7.683 | 18.5951 | 10.9121 | 0.79945 | 100 |  | 7.3179 |  |  | 100 |  | 0.13431 |  |  | 100 |  | 5.7368 |  |  | 100 |  | 0.040568 |  |  | 7.9748 | 18.5826 | 10.6078 | 0.79961 |  | 53
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5635 | 11722 |  | 4927 | 7.897 | 17.015 | 9.118 |  | 0.666 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 20:51 |  | 1 |  | 8.2666 | 18.5701 | 10.3035 | 0.79977
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3242 | 11341 |  | 4784 | 4.836 | 16.478 | 11.642 |  | 0.64 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 6:54 |  | 1 |  | 4.877 | 17.9758 | 13.0988 | 0.77658
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4673 | 11200 |  | 4751 | 6.666 | 16.278 | 9.612 |  | 0.635 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 7:05 |  | 1 |  | 6.9039 | 17.7558 | 10.8519 | 0.77123 | 93.5071 |  | 15.2912 |  |  | 95.0043 |  | 0.98884 |  |  | 96.2942 |  | 12.9285 |  |  | 96.8839 |  | 0.90024 |  |  | 7.4755 | 17.6685 | 10.193 | 0.77471 |  | 54
74 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5480 | 11088 |  | 4794 | 7.699 | 16.12 | 8.421 |  | 0.642 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 7:16 |  | 1 |  | 8.047 | 17.5811 | 9.5341 | 0.7782
115 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3683 | 11205 |  | 4815 | 5.401 | 16.285 | 10.885 |  | 0.646 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 16:42 |  | 1 |  | 5.5016 | 17.7636 | 12.262 | 0.78161
116 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4873 | 11348 |  | 4798 | 6.923 | 16.488 | 9.565 |  | 0.643 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 16:53 |  | 1 |  | 7.1872 | 17.9867 | 10.7995 | 0.77885 | 90.2652 |  | 1.0781 |  |  | 96.5516 |  | 0.39104 |  |  | 101.9679 |  | 1.3809 |  |  | 97.838 |  | 0.89147 |  |  | 7.2262 | 17.9516 | 10.7254 | 0.78234 |  | 55
117 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4928 | 11303 |  | 4841 | 6.993 | 16.424 | 9.431 |  | 0.65 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 17:05 |  | 1 |  | 7.2651 | 17.9165 | 10.6514 | 0.78582
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6420 | 9406 |  | 3254 | 8.9 | 13.745 | 4.845 |  | 0.374 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 17:08 |  | 1 |  | 10.2588 | 15.1754 | 4.9166 | 0.55931
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6973 | 9321 |  | 3329 | 9.608 | 13.626 | 4.018 |  | 0.387 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 17:18 |  | 1 |  | 11.1419 | 15.0398 | 3.8979 | 0.57212 | 100 |  | 1.522 |  |  | 100 |  | 0.68129 |  |  | 100 |  | 7.2578 |  |  | 100 |  | 5.8059 |  |  | 11.2274 | 14.9888 | 3.7614 | 0.55598 |  | 56
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7080 | 9257 |  | 3140 | 9.744 | 13.535 | 3.792 |  | 0.354 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 17:29 |  | 1 |  | 11.3128 | 14.9377 | 3.6249 | 0.53984
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4630 | 8147 |  | 3188 | 6.611 | 11.968 | 5.357 |  | 0.362 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 3:22 |  | 1 |  | 7.4002 | 13.1666 | 5.7664 | 0.54803
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6540 | 8134 |  | 3185 | 9.053 | 11.95 | 2.896 |  | 0.362 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 3:32 |  | 1 |  | 10.4504 | 13.1459 | 2.6954 | 0.54752 | 95.7731 |  | 5.6286 |  |  | 87.6198 |  | 0.1213 |  |  | 62.2983 |  | 24.5461 |  |  | 98.9488 |  | 0.96234 |  |  | 10.7531 | 13.1538 | 2.4008 | 0.55017 |  | 57
74 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6919 | 8144 |  | 3216 | 9.539 | 11.963 | 2.424 |  | 0.367 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 3:43 |  | 1 |  | 11.0557 | 13.1618 | 2.1061 | 0.55282
115 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4490 | 8813 |  | 3218 | 6.433 | 12.908 | 6.476 |  | 0.368 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 13:02 |  | 1 |  | 7.1766 | 14.2293 | 7.0526 | 0.55316
116 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6459 | 8830 |  | 3310 | 8.951 | 12.932 | 3.981 |  | 0.384 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 13:12 |  | 1 |  | 10.3211 | 14.2564 | 3.9353 | 0.56887 | 93.8661 |  | 4.1367 |  |  | 95.0425 |  | 0.033582 |  |  | 98.6961 |  | 11.8646 |  |  | 100.9584 |  | 2.7082 |  |  | 10.5391 | 14.254 | 3.7149 | 0.56127 |  | 58
117 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6732 | 8827 |  | 3221 | 9.299 | 12.927 | 3.628 |  | 0.368 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 13:23 |  | 1 |  | 10.7571 | 14.2516 | 3.4945 | 0.55367
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6531 | 8882 |  | 3551 | 9.042 | 13.006 | 3.964 |  | 0.426 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 21:38 |  | 1 |  | 10.1783 | 14.1243 | 3.946 | 0.61134
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7114 | 8758 |  | 3393 | 9.788 | 12.83 | 3.042 |  | 0.398 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 21:49 |  | 1 |  | 11.0791 | 13.9291 | 2.85 | 0.58509 | 100 |  | 3.6152 |  |  | 100 |  | 0.47364 |  |  | 100 |  | 12.7571 |  |  | 100 |  | 0.7981 |  |  | 11.2831 | 13.9622 | 2.6791 | 0.58277 |  | 59
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7378 | 8800 |  | 3365 | 10.125 | 12.89 | 2.765 |  | 0.393 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 21:59 |  | 1 |  | 11.487 | 13.9952 | 2.5082 | 0.58044
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4571 | 8075 |  | 3345 | 6.536 | 11.865 | 5.329 |  | 0.39 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 8:06 |  | 1 |  | 7.15 | 12.8537 | 5.7037 | 0.57712
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6761 | 8150 |  | 3398 | 9.336 | 11.972 | 2.635 |  | 0.399 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 8:16 |  | 1 |  | 10.5337 | 12.9718 | 2.4381 | 0.58592 | 94.3486 |  | 2.1906 |  |  | 93.006 |  | 0.36348 |  |  | 86.6601 |  | 7.9347 |  |  | 100.5031 |  | 0.11346 |  |  | 10.6504 | 12.9954 | 2.345 | 0.58559 |  | 60
74 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6912 | 8180 |  | 3394 | 9.529 | 12.014 | 2.485 |  | 0.398 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 8:27 |  | 1 |  | 10.767 | 13.019 | 2.252 | 0.58526
85 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4611 | 8140 |  | 3508 | 6.588 | 11.957 | 5.37 |  | 0.418 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 10:46 |  | 1 |  | 7.2118 | 12.956 | 5.7443 | 0.6042
86 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6744 | 8219 |  | 3520 | 9.315 | 12.069 | 2.754 |  | 0.42 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 10:56 |  | 1 |  | 10.5074 | 13.0804 | 2.573 | 0.60619 | 94.459 |  | 2.9126 |  |  |  |  | 0.47056 |  |  |  |  | 15.5831 |  |  |  |  | 2.3286 |  |  | 10.6627 | 13.0497 | 2.387 | 0.59921 |  | 61
87 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6945 | 8180 |  | 3436 | 9.571 | 12.014 | 2.443 |  | 0.406 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 11:07 |  | 1 |  | 10.818 | 13.019 | 2.201 | 0.59224
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | Stored Cal TIC [mg/l] | Stored Cal TC [mg/l] | Stored Cal TOC (Diff.) [mg/l] | Stored Cal NPOC [mg/l] | Stored Cal TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] |  | TIC Absolute value Percent error for check standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes |  | TC Absolute value Percent error for check standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes |  | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes |  | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
[27 empty] | Known or added |  |  |  | 9.901
[5 empty] | Mean |  |  | 4746.0383 | 10498.5301 |  | 5069.2295 | 6.76 | 15.2879 | 8.528 |  | 0.6903 | [10 empty] | Mean |  | 7.6928 | 17.3126 | 9.6198 | [20 empty] | Mean | 8.455 | 17.3942 | 8.9393 | 0.87517 | Min | 1
[5 empty] | Std |  |  | 1324.4745 | 1647.6409 |  | 1353.5671 | 1.6935 | 2.3264 | 2.578 |  | 0.23595 | [10 empty] | Std |  | 2.0097 | 2.5909 | 3.1495 | [20 empty] | Std | 1.2977 | 2.406 | 2.9022 | 0.19699 | Max | 61
[5 empty] | CV |  |  | 27.9069 | 15.694 |  | 26.7016 | 25.0517 | 15.2176 | 30.2294 |  | 34.1809 | [10 empty] | CV |  | 26.1241 | 14.9656 | 32.7394 | [20 empty] | CV | 15.3487 | 13.8321 | 32.4657 | 22.5088
|  |  |  | Warning Limits | Xbar + 2SD |  |  | 7394.9872 | 13793.8118 |  | 7776.3636 | 10.147 | 19.9408 | 13.6839 |  | 1.1622 | [9 empty] | Warning Limits | Xbar + 2SD |  | 11.7121 | 22.4945 | 15.9188 | [19 empty] | Warning Limits | Xbar + 2SD | 11.0504 | 22.2062 | 14.7437 | 1.2692
[5 empty] | Xbar - 2SD |  |  | 2097.0893 | 7203.2483 |  | 2362.0954 | 3.373 | 10.635 | 3.3721 |  | 0.2184 | [10 empty] | Xbar - 2SD |  | 3.6734 | 12.1307 | 3.3209 | [20 empty] | Xbar - 2SD | 5.8595 | 12.5823 | 3.1349 | 0.48119
|  |  |  | Control Limits | Xbar + 3SD |  |  | 8719.4616 | 15441.4527 |  | 9129.9307 | 11.8405 | 22.2672 | 16.2619 |  | 1.3981 | [9 empty] | Control Limits | Xbar + 3SD |  | 13.7217 | 25.0854 | 19.0682 | [19 empty] | Control Limits | Xbar + 3SD | 12.3481 | 24.6122 | 17.6459 | 1.4661
[5 empty] | Xbar - 3SD |  |  | 772.6149 | 5555.6075 |  | 1008.5283 | 1.6795 | 8.3086 | 0.7941 |  | -0.017552 | [10 empty] | Xbar - 3SD |  | 1.6638 | 9.5398 | 0.17141 | [20 empty] | Xbar - 3SD | 4.5618 | 10.1763 | 0.23268 | 0.2842
[26 empty] | Warning Limits as % | Xbar + 2SD |  | 152.2481 | 129.9312 | 165.4788 | [19 empty] | Warning Limits as % | Xbar + 2SD | 130.6975 | 127.6642 | 164.9314 | 145.0176
[27 empty] | Xbar - 2SD |  | 47.7519 | 70.0688 | 34.5212 | [20 empty] | Xbar - 2SD | 69.3025 | 72.3358 | 35.0686 | 54.9824
[26 empty] | Control Limits as % | Xbar + 3SD |  | 178.3722 | 144.8968 | 198.2181 | [19 empty] | Control Limits as % | Xbar + 3SD | 146.0462 | 141.4963 | 197.3971 | 167.5265
[27 empty] | Xbar - 3SD |  | 21.6278 | 55.1032 | 1.7819 | [20 empty] | Xbar - 3SD | 53.9538 | 58.5037 | 2.6029 | 32.4735
[26 empty] | Coefficient of Variation % |  |  | 26.1241 | 14.9656 | 32.7394 | [19 empty] | Coefficient of Variation % |  | 15.3487 | 13.8321 | 32.4657 | 22.5088
[26 empty] | 3 x Coefficient of Variation % |  |  | 78.3722 | 44.8968 | 98.2181 | [19 empty] | 3 x Coefficient of Variation % |  | 46.0462 | 41.4963 | 97.3971 | 67.5265

## rolling ref plain tap
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
23 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11457 | 13719 |  | 17038 | 9.204 | 11.901 | 2.697 |  | 1.666 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 16:01 |  | 1 |  | 11.4235 | 13.2792 | 1.8557 | 1.633
24 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12394 | 13603 |  | 17085 | 9.924 | 11.803 | 1.879 |  | 1.671 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 16:10 |  | 1 |  | 12.3641 | 13.1631 | 0.79903 | 1.6374 |  |  | 1.8101 |  |  |  |  | 0.13676 |  |  |  |  | 29.8996 |  |  |  |  | 0.35627 |  |  | 12.477 | 13.1721 | 0.69512 | 1.6403 |  | 1
25 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12619 | 13621 |  | 17147 | 10.096 | 11.818 | 1.722 |  | 1.677 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 16:19 |  | 1 |  | 12.5899 | 13.1811 | 0.5912 | 1.6432
67 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8904 | 13189 |  | 16839 | 7.246 | 11.452 | 4.206 |  | 1.645 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 23:43 |  | 1 |  | 8.8609 | 12.7488 | 3.8879 | 1.6142
68 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12188 | 13207 |  | 16885 | 9.765 | 11.468 | 1.703 |  | 1.65 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 23:52 |  | 1 |  | 12.1573 | 12.7668 | 0.6095 | 1.6185 |  |  | 1.2472 |  |  |  |  | 0.68434 |  |  |  |  | 48.9398 |  |  |  |  | 0.74842 |  |  | 12.2336 | 12.7233 | 0.48967 | 1.6246 |  | 2
69 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12340 | 13120 |  | 17014 | 9.882 | 11.393 | 1.511 |  | 1.663 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-16 | 0:01 |  | 1 |  | 12.3099 | 12.6797 | 0.36985 | 1.6307
114 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9008 | 11477 |  | 14208 | 7.326 | 10.002 | 2.676 |  | 1.37 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-16 | 7:55 |  | 1 |  | 8.9653 | 11.0354 | 2.0701 | 1.3662
115 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12309 | 11410 |  | 14297 | 9.858 | 9.945 | 0.086 |  | 1.379 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-16 | 8:04 |  | 1 |  | 12.2788 | 10.9684 | -1.3104 | 1.3746 |  |  | 2.327 |  |  |  |  | 0.91734 |  |  |  |  | 13.387 |  |  |  |  | 2.0227 |  |  | 12.4233 | 11.0189 | -1.4044 | 1.3886 |  | 3
116 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12597 | 11511 |  | 14595 | 10.079 | 10.031 | 0 |  | 1.411 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-16 | 8:13 |  | 1 |  | 12.5679 | 11.0694 | -1.4984 | 1.4027
23 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10221 | 11834 |  | 14425 | 8.256 | 10.304 | 2.048 |  | 1.393 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 17:04 |  | 1 |  | 10.2179 | 11.6474 | 1.4295 | 1.5682
24 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10833 | 11690 |  | 14441 | 8.725 | 10.182 | 1.457 |  | 1.394 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 17:12 |  | 1 |  | 10.823 | 11.505 | 0.68202 | 1.57 |  |  | 0.41934 |  |  |  |  | 0.29175 |  |  |  |  | 1.755 |  |  |  |  | 0.31364 |  |  | 10.8457 | 11.5218 | 0.67609 | 1.5675 |  | 4
25 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10879 | 11724 |  | 14396 | 8.761 | 10.211 | 1.45 |  | 1.39 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 17:21 |  | 1 |  | 10.8685 | 11.5386 | 0.67016 | 1.565
67 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7741 | 9807 |  | 11899 | 6.354 | 8.587 | 2.233 |  | 1.129 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 0:45 |  | 1 |  | 7.7659 | 9.6433 | 1.8774 | 1.2922
68 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10657 | 9910 |  | 12253 | 8.591 | 8.674 | 0.084 |  | 1.166 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 0:54 |  | 1 |  | 10.649 | 9.7452 | -0.90383 | 1.3309 |  |  | 1.9582 |  |  |  |  | 15.0401 |  |  |  |  | 634.2 |  |  |  |  | 15.0951 |  |  | 10.7543 | 10.5376 | -0.21669 | 1.4396 |  | 5
69 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10870 | 11513 |  | 14242 | 8.754 | 10.032 | 1.278 |  | 1.374 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 1:04 |  | 1 |  | 10.8596 | 11.33 | 0.47044 | 1.5482
114 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7809 | 11336 |  | 13364 | 6.406 | 9.883 | 3.477 |  | 1.282 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 9:00 |  | 1 |  | 7.8331 | 11.155 | 3.3219 | 1.4523
115 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10575 | 11370 |  | 13626 | 8.528 | 9.911 | 1.383 |  | 1.309 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 9:08 |  | 1 |  | 10.5679 | 11.1886 | 0.62073 | 1.4809 |  |  | 0.77353 |  |  |  |  | 0.36164 |  |  |  |  | 6.9215 |  |  |  |  | 0.63245 |  |  | 10.609 | 11.2089 | 0.59997 | 1.4856 |  | 6
116 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10658 | 11411 |  | 13712 | 8.592 | 9.946 | 1.354 |  | 1.318 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 9:17 |  | 1 |  | 10.65 | 11.2292 | 0.5792 | 1.4903
23 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10215 | 9846 |  | 11524 | 8.252 | 8.62 | 0.368 |  | 1.089 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 15:05 |  | 1 |  | 10.0307 | 10.0356 | 0.0049155 | 1.307
24 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10692 | 11526 |  | 13706 | 8.617 | 10.043 | 1.426 |  | 1.317 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 15:14 |  | 1 |  | 10.5005 | 11.7299 | 1.2293 | 1.5559 |  |  | 0.86671 |  |  |  |  | 0.18932 |  |  |  |  | 5.4158 |  |  |  |  | 0.095272 |  |  | 10.4552 | 11.7188 | 1.2635 | 1.5567 |  | 7
25 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10600 | 11504 |  | 13719 | 8.547 | 10.025 | 1.478 |  | 1.319 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 15:23 |  | 1 |  | 10.4099 | 11.7077 | 1.2978 | 1.5574
67 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7522 | 10982 |  | 13199 | 6.185 | 9.582 | 3.397 |  | 1.264 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 22:44 |  | 1 |  | 7.3782 | 11.1813 | 3.8031 | 1.4981
68 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10289 | 11075 |  | 13317 | 8.308 | 9.661 | 1.353 |  | 1.277 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 22:52 |  | 1 |  | 10.1036 | 11.2751 | 1.1714 | 1.5116 |  |  | 0 |  |  |  |  | 0.1698 |  |  |  |  | 1.6224 |  |  |  |  | 2.0098 |  |  | 10.1036 | 11.2846 | 1.181 | 1.5269 |  | 8
69 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10289 | 11094 |  | 13586 | 8.308 | 9.677 | 1.369 |  | 1.305 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 23:02 |  | 1 |  | 10.1036 | 11.2942 | 1.1906 | 1.5423
114 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7622 | 11087 |  | 13027 | 6.262 | 9.672 | 3.409 |  | 1.246 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-19 | 6:58 |  | 1 |  | 7.4767 | 11.2872 | 3.8105 | 1.4785
115 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10366 | 11083 |  | 13203 | 8.367 | 9.668 | 1.3 |  | 1.265 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-19 | 7:07 |  | 1 |  | 10.1794 | 11.2831 | 1.1037 | 1.4986 |  |  | 0.048368 |  |  |  |  | 0.1429 |  |  |  |  | 1.0106 |  |  |  |  | 0.6677 |  |  | 10.1819 | 11.2912 | 1.1093 | 1.5036 |  | 9
116 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10371 | 11099 |  | 13291 | 8.371 | 9.681 | 1.31 |  | 1.274 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-19 | 7:16 |  | 1 |  | 10.1844 | 11.2993 | 1.1149 | 1.5086
23 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10060 | 11412 |  | 13665 | 8.133 | 9.946 | 1.813 |  | 1.313 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-21 | 18:34 |  | 1 |  | 9.7711 | 10.9524 | 1.1813 | 1.4281
24 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10424 | 11269 |  | 13780 | 8.412 | 9.825 | 1.413 |  | 1.325 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-21 | 18:42 |  | 1 |  | 10.1251 | 10.8136 | 0.68841 | 1.4401 |  |  | 0.076882 |  |  |  |  | 0.65344 |  |  |  |  | 10.8105 |  |  |  |  | 0.24665 |  |  | 10.1213 | 10.849 | 0.72774 | 1.4419 |  | 10
25 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10416 | 11342 |  | 13814 | 8.406 | 9.887 | 1.482 |  | 1.329 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-21 | 18:51 |  | 1 |  | 10.1174 | 10.8844 | 0.76708 | 1.4437
67 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7458 | 10717 |  | 12813 | 6.136 | 9.358 | 3.221 |  | 1.224 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 2:13 |  | 1 |  | 7.2402 | 10.2775 | 3.0373 | 1.3389
68 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10100 | 10762 |  | 13112 | 8.164 | 9.396 | 1.232 |  | 1.255 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 2:21 |  | 1 |  | 9.81 | 10.3212 | 0.5112 | 1.3702 |  |  | 0.19811 |  |  |  |  | 0.41314 |  |  |  |  | 4.4517 |  |  |  |  | 0.083938 |  |  | 9.8197 | 10.3426 | 0.52284 | 1.3708 |  | 11
69 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10120 | 10806 |  | 13123 | 8.179 | 9.433 | 1.254 |  | 1.256 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 2:31 |  | 1 |  | 9.8294 | 10.3639 | 0.53448 | 1.3714
114 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7439 | 10728 |  | 12547 | 6.122 | 9.367 | 3.245 |  | 1.196 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 10:27 |  | 1 |  | 7.2217 | 10.2882 | 3.0665 | 1.3111
115 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10057 | 9474 |  | 11001 | 8.13 | 8.305 | 0.174 |  | 1.035 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 10:35 |  | 1 |  | 9.7682 | 9.0704 | -0.69777 | 1.1494 |  |  | 2.5074 |  |  |  |  | 1.1073 |  |  |  |  | 19.0639 |  |  |  |  | 1.8126 |  |  | 9.8922 | 9.1209 | -0.77129 | 1.1599 |  | 12
116 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10312 | 9578 |  | 11202 | 8.326 | 8.393 | 0.067 |  | 1.056 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 10:45 |  | 1 |  | 10.0162 | 9.1714 | -0.84481 | 1.1704
23 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9853 | 11384 |  | 12787 | 7.974 | 9.923 | 1.949 |  | 1.221 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 15:42 |  | 1 |  | 9.5906 | 11.266 | 1.6755 | 1.3104
24 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10362 | 11261 |  | 12909 | 8.365 | 9.819 | 1.454 |  | 1.234 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 15:50 |  | 1 |  | 10.0869 | 11.1422 | 1.0553 | 1.3229 |  |  | 0.37632 |  |  |  |  | 0.51358 |  |  |  |  | 1.8161 |  |  |  |  | 0.946 |  |  | 10.1059 | 11.1709 | 1.065 | 1.3292 |  | 13
25 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10401 | 11318 |  | 13031 | 8.394 | 9.867 | 1.473 |  | 1.247 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 15:59 |  | 1 |  | 10.1249 | 11.1996 | 1.0747 | 1.3355
67 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7449 | 9528 |  | 11020 | 6.13 | 8.351 | 2.221 |  | 1.037 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 23:24 |  | 1 |  | 7.2463 | 9.3979 | 2.1516 | 1.1282
68 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10209 | 9638 |  | 11267 | 8.247 | 8.444 | 0.197 |  | 1.062 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 23:33 |  | 1 |  | 9.9377 | 9.5087 | -0.42906 | 1.1537 |  |  | 0.34403 |  |  |  |  | 1.6274 |  |  |  |  | 56.9304 |  |  |  |  | 1.8674 |  |  | 9.9206 | 9.5867 | -0.33399 | 1.1646 |  | 14
69 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10174 | 9793 |  | 11478 | 8.22 | 8.575 | 0.355 |  | 1.084 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 23:42 |  | 1 |  | 9.9036 | 9.6647 | -0.23892 | 1.1754
84 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7296 | 9460 |  | 10888 | 6.012 | 8.293 | 2.281 |  | 1.023 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-23 | 2:19 |  | 1 |  | 7.0971 | 9.3295 | 2.2324 | 1.1146
85 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10004 | 9572 |  | 11083 | 8.09 | 8.388 | 0.297 |  | 1.043 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-23 | 2:28 |  | 1 |  | 9.7378 | 9.4422 | -0.29558 | 1.1347 |  |  | 1.6289 |  |  |  |  | 0.51298 |  |  |  |  | 52.0981 |  |  |  |  | 1.2098 |  |  | 9.8178 | 9.4181 | -0.3997 | 1.1416 |  | 15
86 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10168 | 9524 |  | 11217 | 8.216 | 8.347 | 0.131 |  | 1.057 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-23 | 2:37 |  | 1 |  | 9.8977 | 9.3939 | -0.50382 | 1.1485
23 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12522 | 14538 |  | 81175 | 10.021 | 12.595 | 2.573 |  | 8.374 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 15:17 |  | 1 |  | 10.4585 | 11.6426 | 1.1841 | 0.70756
24 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 13316 | 14497 |  | 80916 | 10.63 | 12.56 | 1.93 |  | 8.347 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 15:27 |  | 1 |  | 11.1263 | 11.6091 | 0.48282 | 0.70529 |  |  | 1.4111 |  |  |  |  | 0.232 |  |  |  |  | 47.4045 |  |  |  |  | 0.090667 |  |  | 11.2054 | 11.5957 | 0.39031 | 0.70561 |  | 16
25 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 13504 | 14464 |  | 80989 | 10.775 | 12.532 | 1.758 |  | 8.355 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 15:38 |  | 1 |  | 11.2844 | 11.5822 | 0.2978 | 0.70593
67 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9581 | 12727 |  | 70433 | 7.765 | 11.061 | 3.296 |  | 7.25 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 23:53 |  | 1 |  | 7.9849 | 10.1662 | 2.1813 | 0.61342
68 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 13224 | 12966 |  | 71889 | 10.56 | 11.263 | 0.703 |  | 7.403 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 0:04 |  | 1 |  | 11.0489 | 10.3611 | -0.68788 | 0.62618 |  |  | 1.466 |  |  |  |  | 0.82274 |  |  |  |  | 10.6751 |  |  |  |  | 0.060163 |  |  | 11.1305 | 10.4039 | -0.72666 | 0.62637 |  | 17
69 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 13418 | 13071 |  | 71932 | 10.709 | 11.352 | 0.643 |  | 7.407 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 0:14 |  | 1 |  | 11.2121 | 10.4467 | -0.76545 | 0.62655
114 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9585 | 12632 |  | 67879 | 7.768 | 10.98 | 3.212 |  | 6.983 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 9:20 |  | 1 |  | 7.9883 | 10.0888 | 2.1005 | 0.59103
115 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 13252 | 12924 |  | 69891 | 10.582 | 11.227 | 0.646 |  | 7.194 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 9:30 |  | 1 |  | 11.0725 | 10.3268 | -0.74567 | 0.60867 |  |  | 0.80196 |  |  |  |  | 0.078972 |  |  |  |  | 12.2502 |  |  |  |  | 0.118 |  |  | 11.1171 | 10.3228 | -0.79432 | 0.60903 |  | 18
116 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 13358 | 12914 |  | 69973 | 10.663 | 11.219 | 0.557 |  | 7.202 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 9:41 |  | 1 |  | 11.1617 | 10.3187 | -0.84297 | 0.60939
23 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10009 | 12350 |  | 70820 | 8.094 | 10.742 | 2.648 |  | 7.291 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 15:33 |  | 1 |  | 8.3449 | 9.8589 | 1.514 | 0.61681
24 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12568 | 12503 |  | 71028 | 10.057 | 10.871 | 0.814 |  | 7.313 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 15:43 |  | 1 |  | 10.4972 | 9.9836 | -0.51357 | 0.61863 |  |  | 1.2501 |  |  |  |  | 1.547 |  |  |  |  | 43.477 |  |  |  |  | 0.06656 |  |  | 10.5632 | 9.907 | -0.65623 | 0.61884 |  | 19
25 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12725 | 12315 |  | 71075 | 10.177 | 10.712 | 0.535 |  | 7.318 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 15:53 |  | 1 |  | 10.6292 | 9.8304 | -0.79888 | 0.61904
67 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9097 | 12279 |  | 69055 | 7.394 | 10.682 | 3.288 |  | 7.106 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 0:14 |  | 1 |  | 7.5778 | 9.801 | 2.2232 | 0.60134
68 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12521 | 12470 |  | 69605 | 10.02 | 10.843 | 0.823 |  | 7.164 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 0:24 |  | 1 |  | 10.4577 | 9.9567 | -0.50094 | 0.60616 |  |  | 0.62537 |  |  |  |  | 2.5036 |  |  |  |  | 47.4689 |  |  |  |  | 0.41408 |  |  | 10.4905 | 9.8336 | -0.65684 | 0.60742 |  | 20
69 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12599 | 12168 |  | 69892 | 10.08 | 10.587 | 0.507 |  | 7.194 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 0:35 |  | 1 |  | 10.5233 | 9.7105 | -0.81274 | 0.60868
114 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9141 | 12196 |  | 67511 | 7.428 | 10.611 | 3.183 |  | 6.945 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 9:32 |  | 1 |  | 7.6148 | 9.7334 | 2.1185 | 0.58781
115 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12567 | 12186 |  | 68764 | 10.056 | 10.602 | 0.546 |  | 7.076 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 9:43 |  | 1 |  | 10.4964 | 9.7252 | -0.77115 | 0.59879 |  |  | 0.83784 |  |  |  |  | 0.64337 |  |  |  |  | 3.2583 |  |  |  |  | 0.62736 |  |  | 10.5405 | 9.7566 | -0.78392 | 0.60067 |  | 21
116 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12672 | 12263 |  | 69194 | 10.136 | 10.668 | 0.531 |  | 7.121 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 9:53 |  | 1 |  | 10.5847 | 9.788 | -0.79669 | 0.60256
22 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10435 | 11750 |  | 14369 | 8.42 | 10.233 | 1.813 |  | 1.387 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 15:37 |  | 1 |  | 9.9524 | 11.0921 | 1.1397 | 1.3988
23 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10857 | 11770 |  | 14429 | 8.744 | 10.25 | 1.506 |  | 1.393 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 15:46 |  | 1 |  | 10.3627 | 11.1116 | 0.74889 | 1.4047 |  |  | 0.40433 |  |  |  |  | 11.8075 |  |  |  |  | 796.088 |  |  |  |  | 11.8221 |  |  | 10.3418 | 10.4922 | 0.15037 | 1.3263 |  | 22
24 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10814 | 10501 |  | 12840 | 8.712 | 9.175 | 0.464 |  | 1.227 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 15:55 |  | 1 |  | 10.3209 | 9.8728 | -0.44816 | 1.2479
64 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7718 | 9853 |  | 12177 | 6.336 | 8.626 | 2.291 |  | 1.158 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 23:15 |  | 1 |  | 7.3103 | 9.2401 | 1.9299 | 1.1825
65 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10533 | 9967 |  | 12402 | 8.496 | 8.723 | 0.227 |  | 1.181 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 23:24 |  | 1 |  | 10.0477 | 9.3514 | -0.69622 | 1.2047 |  |  | 0.71363 |  |  |  |  | 0.55177 |  |  |  |  | 2.8625 |  |  |  |  | 2.7148 |  |  | 10.0836 | 9.3773 | -0.70633 | 1.2213 |  | 23
66 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10607 | 10020 |  | 12738 | 8.552 | 8.767 | 0.215 |  | 1.216 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 23:34 |  | 1 |  | 10.1196 | 9.4032 | -0.71644 | 1.2378
109 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7652 | 9847 |  | 12151 | 6.285 | 8.62 | 2.335 |  | 1.155 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-03 | 7:29 |  | 1 |  | 7.2461 | 9.2343 | 1.9882 | 1.1799
110 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10441 | 9940 |  | 12367 | 8.425 | 8.699 | 0.274 |  | 1.177 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-03 | 7:38 |  | 1 |  | 9.9582 | 9.3251 | -0.63312 | 1.2012 |  |  | 0.38985 |  |  |  |  | 0.49084 |  |  |  |  | 1.1096 |  |  |  |  | 1.2165 |  |  | 9.9777 | 9.348 | -0.62963 | 1.2086 |  | 24
111 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10481 | 9987 |  | 12516 | 8.456 | 8.74 | 0.284 |  | 1.193 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-03 | 7:48 |  | 1 |  | 9.9971 | 9.371 | -0.62613 | 1.2159
22 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10619 | 10588 |  | 12347 | 8.561 | 9.248 | 0.687 |  | 1.175 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 16:08 |  | 1 |  | 9.8815 | 10.376 | 0.4945 | 1.2516
23 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11001 | 10581 |  | 12557 | 8.855 | 9.242 | 0.388 |  | 1.197 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 16:17 |  | 1 |  | 10.2401 | 10.369 | 0.12883 | 1.2731 |  |  | 0.32037 |  |  |  |  | 0.71495 |  |  |  |  | 27.7671 |  |  |  |  | 1.4209 |  |  | 10.2566 | 10.4062 | 0.1496 | 1.2822 |  | 25
24 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11036 | 10655 |  | 12735 | 8.881 | 9.305 | 0.424 |  | 1.216 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 16:26 |  | 1 |  | 10.273 | 10.4434 | 0.17037 | 1.2913
64 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7828 | 10166 |  | 12285 | 6.421 | 8.891 | 2.47 |  | 1.169 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 23:51 |  | 1 |  | 7.2612 | 9.9517 | 2.6905 | 1.2452
65 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10630 | 10260 |  | 12473 | 8.57 | 8.97 | 0.4 |  | 1.188 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 24:00 |  | 1 |  | 9.8918 | 10.0462 | 0.15441 | 1.2645 |  |  | 2.3727 |  |  |  |  | 0.71797 |  |  |  |  | 229.8704 |  |  |  |  | 1.7888 |  |  | 10.0106 | 10.0824 | 0.071838 | 1.2759 |  | 26
66 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10883 | 10332 |  | 12696 | 8.764 | 9.031 | 0.267 |  | 1.212 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 0:09 |  | 1 |  | 10.1294 | 10.1186 | -0.010729 | 1.2873
109 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7783 | 10056 |  | 11948 | 6.386 | 8.798 | 2.412 |  | 1.134 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 8:08 |  | 1 |  | 7.219 | 9.8411 | 2.6221 | 1.2107
110 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10576 | 10099 |  | 12216 | 8.528 | 8.834 | 0.305 |  | 1.162 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 8:17 |  | 1 |  | 9.8411 | 9.8844 | 0.043233 | 1.2382 |  |  | 2.6545 |  |  |  |  | 1.2835 |  |  |  |  | 541.7826 |  |  |  |  | 1.6153 |  |  | 9.9735 | 9.9482 | -0.025299 | 1.2483 |  | 27
111 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10858 | 10226 |  | 12413 | 8.745 | 8.942 | 0.197 |  | 1.182 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 8:27 |  | 1 |  | 10.1059 | 10.0121 | -0.093831 | 1.2583
22 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10655 | 11002 |  | 12295 | 8.589 | 9.6 | 1.011 |  | 1.17 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 17:26 |  | 1 |  | 9.8919 | 10.5794 | 0.68749 | 1.2167
23 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11070 | 10901 |  | 12475 | 8.907 | 9.514 | 0.607 |  | 1.189 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 17:35 |  | 1 |  | 10.2804 | 10.4809 | 0.20048 | 1.2346 |  |  | 1.4106 |  |  |  |  | 0.59368 |  |  |  |  | 52.717 |  |  |  |  | 1.1455 |  |  | 10.3534 | 10.5121 | 0.15866 | 1.2418 |  | 28
24 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11226 | 10965 |  | 12618 | 9.028 | 9.568 | 0.54 |  | 1.204 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 17:45 |  | 1 |  | 10.4265 | 10.5433 | 0.11684 | 1.2489
64 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7876 | 10167 |  | 11868 | 6.457 | 8.892 | 2.435 |  | 1.125 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 1:12 |  | 1 |  | 7.2902 | 9.7652 | 2.475 | 1.1743
65 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10678 | 10316 |  | 12082 | 8.607 | 9.018 | 0.411 |  | 1.148 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 1:21 |  | 1 |  | 9.9134 | 9.9105 | -0.0029772 | 1.1955 |  |  | 0.58381 |  |  |  |  | 0.31536 |  |  |  |  | 187.4911 |  |  |  |  | 1.3306 |  |  | 9.9425 | 9.8948 | -0.047601 | 1.2036 |  | 29
66 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10740 | 10284 |  | 12243 | 8.655 | 8.991 | 0.336 |  | 1.164 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 1:31 |  | 1 |  | 9.9715 | 9.8792 | -0.092226 | 1.2116
109 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7939 | 10282 |  | 11790 | 6.506 | 8.989 | 2.484 |  | 1.117 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 9:31 |  | 1 |  | 7.3492 | 9.8773 | 2.5281 | 1.1665
110 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10793 | 10273 |  | 11979 | 8.695 | 8.982 | 0.287 |  | 1.137 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 9:40 |  | 1 |  | 10.0211 | 9.8685 | -0.15257 | 1.1853 |  |  | 0.39315 |  |  |  |  | 0.92453 |  |  |  |  | 150.4161 |  |  |  |  | 1.2592 |  |  | 10.0014 | 9.9144 | -0.08708 | 1.1928 |  | 30
111 | 1 | Reference | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10751 | 10367 |  | 12130 | 8.662 | 9.062 | 0.399 |  | 1.153 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 9:50 |  | 1 |  | 9.9818 | 9.9602 | -0.021589 | 1.2003
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | Stored Cal TIC [mg/l] | Stored Cal TC [mg/l] | Stored Cal TOC (Diff.) [mg/l] | Stored Cal NPOC [mg/l] | Stored Cal TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] |  | TIC Absolute value Percent error for check standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes |  | TC Absolute value Percent error for check standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes |  | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes |  | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
[5 empty] | Mean |  |  | 10471.8444 | 11277.0222 |  | 24843.7222 | 8.4486 | 9.8323 | 1.3842 |  | 2.4824 | [10 empty] | Mean |  | 9.811 | 10.5366 | 0.72561 | [20 empty] | Mean | 10.525 | 10.5587 | 0.033701 | 1.2103 | Min | 1
[5 empty] | Std |  |  | 1612.827 | 1264.4943 |  | 23716.4036 | 1.2373 | 1.0713 | 1.0476 |  | 2.4805 | [10 empty] | Std |  | 1.3877 | 0.9809 | 1.2665 | [20 empty] | Std | 0.74028 | 0.97205 | 0.70269 | 0.32808 | Max | 30
[5 empty] | CV |  |  | 15.4016 | 11.213 |  | 95.4624 | 14.6448 | 10.8953 | 75.6818 |  | 99.9246 | [10 empty] | CV |  | 14.1443 | 9.3094 | 174.543 | [20 empty] | CV | 7.0336 | 9.2062 | 2085.0601 | 27.1065
|  |  |  | Warning Limits | Xbar + 2SD |  |  | 13697.4985 | 13806.0108 |  | 72276.5295 | 10.9232 | 11.9748 | 3.4793 |  | 7.4433 | [9 empty] | Warning Limits | Xbar + 2SD |  | 12.5864 | 12.4984 | 3.2586 | [19 empty] | Warning Limits | Xbar + 2SD | 12.0055 | 12.5028 | 1.4391 | 1.8665
[5 empty] | Xbar - 2SD |  |  | 7246.1904 | 8748.0337 |  | -22589.085 | 5.9741 | 7.6898 | -0.71096 |  | -2.4786 | [10 empty] | Xbar - 2SD |  | 7.0356 | 8.5748 | -1.8074 | [20 empty] | Xbar - 2SD | 9.0444 | 8.6146 | -1.3717 | 0.55417
|  |  |  | Control Limits | Xbar + 3SD |  |  | 15310.3255 | 15070.5051 |  | 95992.9331 | 12.1605 | 13.0461 | 4.5269 |  | 9.9238 | [9 empty] | Control Limits | Xbar + 3SD |  | 13.9741 | 13.4793 | 4.5251 | [19 empty] | Control Limits | Xbar + 3SD | 12.7458 | 13.4748 | 2.1418 | 2.1946
[5 empty] | Xbar - 3SD |  |  | 5633.3634 | 7483.5394 |  | -46305.4887 | 4.7368 | 6.6185 | -1.7585 |  | -4.9591 | [10 empty] | Xbar - 3SD |  | 5.6479 | 7.5939 | -3.0739 | [20 empty] | Xbar - 3SD | 8.3041 | 7.6425 | -2.0744 | 0.2261
[26 empty] | Warning Limits as % | Xbar + 2SD |  | 128.2886 | 118.6189 | 449.086 | [19 empty] | Warning Limits as % | Xbar + 2SD | 114.0671 | 118.4124 | 4270.1203 | 154.213
[27 empty] | Xbar - 2SD |  | 71.7114 | 81.3811 | -249.086 | [20 empty] | Xbar - 2SD | 85.9329 | 81.5876 | -4070.1203 | 45.787
[26 empty] | Control Limits as % | Xbar + 3SD |  | 142.4329 | 127.9283 | 623.629 | [19 empty] | Control Limits as % | Xbar + 3SD | 121.1007 | 127.6186 | 6355.1804 | 181.3195
[27 empty] | Xbar - 3SD |  | 57.5671 | 72.0717 | -423.629 | [20 empty] | Xbar - 3SD | 78.8993 | 72.3814 | -6155.1804 | 18.6805
[26 empty] | Coefficient of Variation % |  |  | 14.1443 | 9.3094 | 174.543 | [19 empty] | Coefficient of Variation % |  | 7.0336 | 9.2062 | 2085.0601 | 27.1065
[26 empty] | 3 x Coefficient of Variation % |  |  | 42.4329 | 27.9283 | 523.629 | [19 empty] | 3 x Coefficient of Variation % |  | 21.1007 | 27.6186 | 6255.1804 | 81.3195

## rolling spikes
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes |  | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
51 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3524 | 18588 |  | 29080 | 3.119 | 16.026 | 12.908 |  | 2.925 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-09 | 20:26 |  | 1 |  | 10.2679 | 16.3551 | 6.0872 | 2.5904 |  |  | 0.59708 |  | 156.8982 |  |  | 0.059769 |  | 116.9449 |  |  | 0.85315 |  | 76.9916 |  |  | 0.33603 |  | 87.5661 |  | 10.2986 | 16.36 | 6.0614 | 2.5948 |  | 1
98 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2565 | 14051 |  | 6285 | 2.383 | 12.182 | 9.8 |  | 0.541 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-10 | 4:49 |  | 1 |  | 7.4598 | 12.322 | 4.8622 | 0.54141 |  |  | 0.43271 |  | 106.5834 |  |  | 1.5631 |  | 83.3511 |  |  | 3.3225 |  | 60.1188 |  |  | 5.0557 |  | 81.3188 |  | 7.4437 | 12.2265 | 4.7828 | 0.52806 |  | 2
51 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5354 | 14780 |  | 8968 | 4.522 | 12.8 | 8.278 |  | 0.822 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-11 | 19:58 |  | 1 |  | 8.187 | 14.5018 | 6.3148 | 0.73599 |  |  | 0.32013 |  | 113.1842 |  |  | 0.49052 |  | 118.8987 |  |  | 0.71186 |  | 124.6132 |  |  | 1.18 |  | 77.1841 |  | 8.1739 | 14.4663 | 6.2924 | 0.74036 |  | 3
98 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2962 | 11929 |  | 4497 | 2.687 | 10.385 | 7.697 |  | 0.354 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-12 | 4:02 |  | 1 |  | 4.5052 | 11.6524 | 7.1472 | 0.3675 |  |  | 0.6131 |  | 63.7238 |  |  | 0.41255 |  | 107.1888 |  |  | 1.0645 |  | 150.6537 |  |  | 1.037 |  | 76.4699 |  | 4.519 | 11.6284 | 7.1094 | 0.3656 |  | 4
51 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4464 | 10919 |  | 4327 | 3.84 | 9.529 | 5.689 |  | 0.336 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-12 | 22:10 |  | 1 |  | 8.8328 | 11.2527 | 2.4198 | 0.38875 |  |  | 1.4105 |  | 126.7206 |  |  | 0.79029 |  | 111.5934 |  |  | 1.5067 |  | 96.4661 |  |  | 1.1487 |  | 94.2202 |  | 8.8955 | 11.2973 | 2.4017 | 0.38653 |  | 5
98 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4055 | 11379 |  | 5213 | 3.526 | 9.919 | 6.393 |  | 0.429 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-13 | 6:17 |  | 1 |  | 8.2959 | 11.7475 | 3.4516 | 0.46743 |  |  | 0.8066 |  | 115.5982 |  |  | 0.10077 |  | 105.8384 |  |  | 2.3158 |  | 96.0787 |  |  | 1.2041 |  | 91.7485 |  | 8.3295 | 11.7415 | 3.4121 | 0.46463 |  | 6
51 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1808 | 23433 |  | 13378 | 1.802 | 20.131 | 18.329 |  | 1.283 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-17 | 22:35 |  | 1 |  | 6.7423 | 15.4097 | 8.6673 | 0.79059 |  |  | 2.1413 |  | 103.0898 |  |  | 0.64602 |  | 119.0532 |  |  | 2.8692 |  | 135.0166 |  |  | 1.4703 |  | 99.5934 |  | 6.8153 | 15.36 | 8.5447 | 0.78482 |  | 7
98 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1523 | 21834 |  | 8398 | 1.583 | 18.776 | 17.192 |  | 0.762 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-18 | 7:58 |  | 1 |  | 5.6759 | 14.3794 | 8.7035 | 0.50469 |  |  | 3.0793 |  | 89.9307 |  |  | 1.0608 |  | 106.4026 |  |  | 3.671 |  | 122.8745 |  |  | 0.20496 |  | 85.5093 |  | 5.5898 | 14.456 | 8.8662 | 0.50418 |  | 8
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1571 | 16797 |  | 5821 | 1.621 | 14.509 | 12.888 |  | 0.493 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-09 | 23:57 |  | 1 |  | 6.8339 | 11.725 | 4.8912 | 0.39412 |  |  | 1.8341 |  | 107.961 |  |  | 0.2074 |  | 77.0625 |  |  | 2.1106 |  | 46.164 |  |  | 0.43919 |  | 79.46 |  | 6.8971 | 11.7372 | 4.8401 | 0.39326 |  | 9
100 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1572 | 13146 |  | 5568 | 1.621 | 11.415 | 9.795 |  | 0.466 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-10 | 8:09 |  | 1 |  | 6.8382 | 9.1857 | 2.3475 | 0.37664 |  |  | 1.0786 |  | 108.6153 |  |  | 0.41731 |  | 73.6835 |  |  | 4.9059 |  | 38.7517 |  |  | 2.5637 |  | 75.0615 |  | 6.8753 | 9.1666 | 2.2913 | 0.37188 |  | 10
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3070 | 19194 |  | 7245 | 2.77 | 16.54 | 13.769 |  | 0.642 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-10 | 21:43 |  | 1 |  | 12.6334 | 14.1259 | 1.4924 | 0.56344 |  |  | 1.7063 |  | 104.3038 |  |  | 0.28844 |  | 61.2651 |  |  | 12.5782 |  | 18.2263 |  |  | 0.51969 |  | 76.1212 |  | 12.7422 | 14.1463 | 1.4041 | 0.5649 |  | 11
100 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1587 | 11507 |  | 5967 | 1.632 | 10.027 | 8.395 |  | 0.508 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-11 | 5:50 |  | 1 |  | 6.3113 | 8.525 | 2.2137 | 0.46964 |  |  | 5.3896 |  | 110.6274 |  |  | 1.1346 |  | 68.0048 |  |  | 22.3904 |  | 25.3823 |  |  | 0.3444 |  | 92.488 |  | 6.4861 | 8.4769 | 1.9908 | 0.46883 |  | 12
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1800 | 16139 |  | 5405 | 1.796 | 13.951 | 12.155 |  | 0.449 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-11 | 21:16 |  | 1 |  | 5.484 | 12.3332 | 6.8492 | 0.41543 |  |  | 5.5901 |  | 105.0363 |  |  | 0.30611 |  | 85.8326 |  |  | 5.2913 |  | 66.6289 |  |  | 1.7732 |  | 89.8352 |  | 5.6417 | 12.3144 | 6.6727 | 0.41915 |  | 13
100 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2256 | 15390 |  | 10443 | 2.146 | 13.317 | 11.171 |  | 0.976 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-12 | 5:31 |  | 1 |  | 7.5057 | 11.7685 | 4.2628 | 0.77896 |  |  | 2.49 |  | 102.1546 |  |  | 0.33368 |  | 83.596 |  |  | 5.5067 |  | 65.0374 |  |  | 1.023 |  | 77.8812 |  | 7.6003 | 11.7489 | 4.1486 | 0.78296 |  | 14
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4872 | 10694 |  | 5689 | 4.152 | 9.339 | 5.186 |  | 0.479 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 22:10 |  | 1 |  | 4.8136 | 10.2518 | 5.4382 | 0.56325 |  |  | 1.0579 |  | 87.1953 |  |  | 0.2048 |  | 94.0505 |  |  | 0.55644 |  | 100.9057 |  |  | 0.83323 |  | 116.3901 |  | 4.8392 | 10.2623 | 5.4231 | 0.5656 |  | 15
105 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5148 | 10515 |  | 5420 | 4.365 | 9.187 | 4.822 |  | 0.451 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-16 | 6:24 |  | 1 |  | 5.0907 | 10.0726 | 4.982 | 0.53789 |  |  | 0.66817 |  | 87.7641 |  |  | 0.089383 |  | 103.016 |  |  | 0.50552 |  | 118.2679 |  |  | 0.56232 |  | 124.7474 |  | 5.1077 | 10.0771 | 4.9694 | 0.53638 |  | 16
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5932 | 10572 |  | 4866 | 4.966 | 9.235 | 4.269 |  | 0.393 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 23:13 |  | 2 | the headspace was compromised in this spike-big bubble under foil | 5.9773 | 10.3997 | 4.4223 | 0.52386 |  |  | 0 |  | 103.9464 |  |  | 13.2035 |  | 104.7184 |  |  | 28.506 |  | 105.4905 |  |  | 16.1251 |  | 138.5867 |  | 5.9773 | 11.1348 | 5.1574 | 0.5698 |  | 17
105 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5593 | 10538 |  | 5402 | 4.706 | 9.207 | 4.501 |  | 0.449 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 7:28 |  | 1 |  | 5.6422 | 10.3661 | 4.7239 | 0.58242 |  |  | 0.36867 |  | 87.9128 |  |  | 1.4581 |  | 101.637 |  |  | 3.5968 |  | 115.3612 |  |  | 0.13122 |  | 153.9003 |  | 5.6318 | 10.4422 | 4.8104 | 0.5828 |  | 18
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5798 | 10596 |  | 3902 | 4.863 | 9.255 | 4.392 |  | 0.292 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 21:13 |  | 1 |  | 5.6801 | 10.792 | 5.1119 | 0.43747 |  |  | 0.32893 |  | 87.2024 |  |  | 0.20579 |  | 94.7295 |  |  | 0.80333 |  | 102.2567 |  |  | 0.39194 |  | 98.3211 |  | 5.6895 | 10.7809 | 5.0914 | 0.43662 |  | 19
105 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5715 | 10618 |  | 4145 | 4.799 | 9.274 | 4.475 |  | 0.318 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-19 | 5:26 |  | 1 |  | 5.5984 | 10.8142 | 5.2158 | 0.46519 |  |  | 0.5469 |  | 89.1231 |  |  | 0.037295 |  | 94.8556 |  |  | 0.66056 |  | 100.5881 |  |  | 0.19638 |  | 100.3176 |  | 5.5831 | 10.8162 | 5.2331 | 0.46474 |  | 20
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5571 | 10530 |  | 3769 | 4.689 | 9.2 | 4.511 |  | 0.278 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 0:42 |  | 1 |  | 5.4047 | 10.0959 | 4.6912 | 0.39293 |  |  | 1.0857 |  | 88.2871 |  |  | 0.23059 |  | 91.4871 |  |  | 1.7259 |  | 94.687 |  |  | 2.9636 |  | 87.3769 |  | 5.3755 | 10.1076 | 4.732 | 0.39884 |  | 21
105 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5881 | 10471 |  | 4036 | 4.927 | 9.149 | 4.222 |  | 0.306 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 8:55 |  | 1 |  | 5.7062 | 10.0386 | 4.3324 | 0.42086 |  |  | 0.28936 |  | 88.1899 |  |  | 0.45364 |  | 93.4131 |  |  | 0.6696 |  | 98.6364 |  |  | 0.19864 |  | 101.7944 |  | 5.7145 | 10.0614 | 4.3469 | 0.42128 |  | 22
58 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6558 | 11247 |  | 3621 | 5.446 | 9.807 | 4.361 |  | 0.263 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 21:53 |  | 1 |  | 6.3775 | 11.1281 | 4.7507 | 0.36562 |  |  | 1.9232 |  | 89.4855 |  |  | 1.1736 |  | 80.4288 |  |  | 5.4864 |  | 71.3722 |  |  | 0.36714 |  | 76.4947 |  | 6.4394 | 11.0632 | 4.6238 | 0.36495 |  | 23
58 | 19 | SPIKE_C_IM_06jun21 S site5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8358 | 14593 |  | 38260 | 6.827 | 12.641 | 5.814 |  | 3.886 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 22:12 |  | 1 |  | 6.9563 | 11.6874 | 4.7311 | 0.33146 |  |  | 0.2906 |  | 95.7096 |  |  | 6.6814 |  | 96.0174 |  |  | 16.0938 |  | 96.3252 |  |  | 14.0032 |  |  | 9.6178 | 6.9462 | 12.0913 | 5.1452 | 0.35642 |  | 24
105 | 31 | SPIKE_F21jun21S1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7830 | 15900 |  | 50407 | 6.422 | 13.749 | 7.327 |  | 5.156 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 7:33 |  | 1 |  | 6.5122 | 12.7529 | 6.2407 | 0.43791 |  |  | 1.079 |  | 98.7861 |  |  | 0.16606 |  | 93.4552 |  |  | 0.79561 |  | 88.1243 |  |  | 10.9991 |  |  | 10.2029 | 6.5475 | 12.7635 | 6.216 | 0.4634 |  | 25
58 | 19 | SPIKE_C01jul21S9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8578 | 14310 |  | 36318 | 6.996 | 12.402 | 5.406 |  | 3.682 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 22:33 |  | 1 |  | 7.1413 | 11.4567 | 4.3154 | 0.31444 |  |  | 0.48405 |  | 84.03 |  |  | 0.49685 |  | 86.6603 |  |  | 2.0992 |  | 89.2906 |  |  | 0.65559 |  |  | 8.66 | 7.1241 | 11.4852 | 4.3612 | 0.31548 |  | 26
105 | 31 | SPIKE_F31may21S1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7272 | 14415 |  | 40199 | 5.994 | 12.491 | 6.496 |  | 4.088 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 7:51 |  | 1 |  | 6.0429 | 11.5423 | 5.4994 | 0.34845 |  |  | 0.055689 |  | 90.644 |  |  | 0.31124 |  | 84.8523 |  |  | 0.59281 |  | 79.0607 |  |  | 1.3415 |  |  | 7.8446 | 6.0412 | 11.5244 | 5.4832 | 0.35081 |  | 27
56 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6453 | 10665 |  | 4961 | 5.366 | 9.314 | 3.948 |  | 0.403 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 21:48 |  | 1 |  | 6.0801 | 10.0329 | 3.9527 | 0.47043 |  |  | 0.81236 |  | 97.4056 |  |  | 0.33999 |  | 83.7161 |  |  | 0.39101 |  | 70.0266 |  |  | 2.322 |  | 106.9612 |  | 6.1049 | 10.05 | 3.945 | 0.47596 |  | 28
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7079 | 12200 |  | 4566 | 5.846 | 10.614 | 4.768 |  | 0.361 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-03 | 6:03 |  | 1 |  | 6.6889 | 11.5314 | 4.8425 | 0.43145 |  |  | 1.8294 |  | 94.0499 |  |  | 0.80752 |  | 92.075 |  |  | 4.5675 |  | 90.1 |  |  | 0.77458 |  | 103.1982 |  | 6.7506 | 11.485 | 4.7344 | 0.43313 |  | 29
56 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6761 | 10967 |  | 5214 | 5.601 | 9.569 | 3.968 |  | 0.429 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 22:23 |  | 1 |  | 6.2595 | 10.7571 | 4.4975 | 0.52154 |  |  | 1.769 |  | 94.1501 |  |  | 0.7913 |  | 97.0993 |  |  | 0.58562 |  | 100.0484 |  |  | 1.4997 |  | 101.3149 |  | 6.3154 | 10.7998 | 4.4844 | 0.52548 |  | 30
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6571 | 11271 |  | 8764 | 5.456 | 9.827 | 4.371 |  | 0.801 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 6:41 |  | 1 |  | 6.0811 | 11.0627 | 4.9816 | 0.88487 |  |  | 0.63098 |  | 96.6674 |  |  | 0.29946 |  | 94.7115 |  |  | 0.10673 |  | 92.7556 |  |  | 0.85227 |  | 98.3787 |  | 6.1004 | 11.0793 | 4.9789 | 0.88866 |  | 31
56 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7226 | 10876 |  | 5097 | 5.958 | 9.493 | 3.534 |  | 0.417 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 23:44 |  | 1 |  | 6.6817 | 10.4565 | 3.7749 | 0.50075 |  |  | 0.66072 |  | 94.8537 |  |  | 1.6188 |  | 94.0879 |  |  | 5.529 |  | 93.3222 |  |  | 2.2974 |  | 104.8686 |  | 6.6597 | 10.5418 | 3.8822 | 0.50657 |  | 32
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5819 | 7809 |  | 6925 | 4.879 | 6.894 | 2.015 |  | 0.608 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 8:03 |  | 1 |  | 5.3644 | 7.4658 | 2.1014 | 0.68258 |  |  | 0.12224 |  | 102.0942 |  |  | 0.15661 |  | 41.5071 |  |  | 0.86494 |  | -19.08 |  |  | 0.8272 |  | 170.4749 |  | 5.3611 | 7.4717 | 2.1105 | 0.68542 |  | 33
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4973 | 10152 |  | 4222 | 4.23 | 8.879 | 4.649 |  | 0.326 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 0:04 |  | 1 |  | 6.2765 | 13.0543 | 6.7778 | 0.61615 |  |  | 0.33638 |  | 97.7816 |  |  | 0.082461 |  | 96.4239 |  |  | 0.15326 |  | 95.0662 |  |  | 0.31862 |  | 99.9647 |  | 6.2871 | 13.0596 | 6.7726 | 0.61713 |  | 34
95 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5090 | 9897 |  | 3957 | 4.32 | 8.663 | 4.343 |  | 0.298 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 8:26 |  | 1 |  | 6.4312 | 12.711 | 6.2798 | 0.57607 |  |  | 1.2256 |  | 102.6832 |  |  | 0.20101 |  | 106.9855 |  |  | 0.85928 |  | 111.2878 |  |  | 1.2265 |  | 115.759 |  | 6.4708 | 12.7238 | 6.253 | 0.57962 |  | 35
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6377 | 10820 |  | 7764 | 5.307 | 9.445 | 4.138 |  | 0.696 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 1:08 |  | 1 |  | 8.106 | 13.8593 | 5.7533 | 1.1906 |  |  | 0.19461 |  | 102.8141 |  |  | 0.95296 |  | 97.6156 |  |  | 2.0117 |  | 92.417 |  |  | 0.75108 |  | 84.7548 |  | 8.1139 | 13.9257 | 5.8118 | 1.1862 |  | 36
95 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4708 | 8525 |  | 2887 | 4.026 | 7.501 | 3.475 |  | 0.186 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 9:24 |  | 1 |  | 5.9098 | 10.8137 | 4.9039 | 0.42835 |  |  | 1.4578 |  | 101.9645 |  |  | 0.25738 |  | 100.1367 |  |  | 2.286 |  | 98.3089 |  |  | 0.036495 |  | 104.6232 |  | 5.8671 | 10.8276 | 4.9606 | 0.42827 |  | 37
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5804 | 9193 |  | 5857 | 4.868 | 8.067 | 3.199 |  | 0.497 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 23:26 |  | 1 |  | 7.4069 | 11.6294 | 4.2226 | 0.85358 |  |  | 0.37025 |  | 95.9012 |  |  | 0.056859 |  | 90.9941 |  |  | 0.49524 |  | 86.087 |  |  | 1.3306 |  | 98.7203 |  | 7.4206 | 11.6327 | 4.2122 | 0.84794 |  | 38
95 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5763 | 9637 |  | 8998 | 4.836 | 8.443 | 3.607 |  | 0.825 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 7:47 |  | 1 |  | 7.3532 | 12.2168 | 4.8636 | 1.3199 |  |  | 0.51731 |  | 99.8656 |  |  | 0.94837 |  | 87.64 |  |  | 3.1241 |  | 75.4143 |  |  | 0.022498 |  | 71.2357 |  | 7.3342 | 12.275 | 4.9408 | 1.3197 |  | 39
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6197 | 11186 |  | 4865 | 5.169 | 9.756 | 4.586 |  | 0.393 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 0:02 |  | 1 |  | 7.9945 | 13.6202 | 5.6257 | 0.63189 |  |  | 2.4274 |  | 122.9056 |  |  | 0.027896 |  | 110.2715 |  |  | 3.4163 |  | 97.6375 |  |  | 0.34226 |  | 128.2942 |  | 7.8987 | 13.6221 | 5.7235 | 0.63297 |  | 40
95 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5961 | 10732 |  | 7796 | 4.988 | 9.371 | 4.383 |  | 0.699 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 8:20 |  | 1 |  | 7.6781 | 13.0451 | 5.367 | 1.0054 |  |  | 0.5225 |  | 110.4879 |  |  | 0.96594 |  | 108.3389 |  |  | 3.1343 |  | 106.1898 |  |  | 0.75064 |  | 121.478 |  | 7.6982 | 12.9824 | 5.2842 | 1.0016 |  | 41
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5873 | 10258 |  | 4122 | 4.921 | 8.969 | 4.048 |  | 0.315 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 21:48 |  | 1 |  | 7.2287 | 12.0717 | 4.843 | 0.48556 |  |  | 1.072 |  | 94.5627 |  |  | 1.0534 |  | 92.1967 |  |  | 1.0256 |  | 89.8306 |  |  | 0.52097 |  | 97.0315 |  | 7.2676 | 12.1356 | 4.868 | 0.48683 |  | 42
95 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5283 | 9213 |  | 8194 | 4.468 | 8.084 | 3.616 |  | 0.741 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-09 | 5:58 |  | 1 |  | 6.4751 | 10.7872 | 4.3121 | 0.97734 |  |  | 0.90322 |  | 101.8478 |  |  | 0.29671 |  | 90.1175 |  |  | 2.126 |  | 78.3871 |  |  | 0.46847 |  | 123.711 |  | 6.5045 | 10.7712 | 4.2667 | 0.97964 |  | 43
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8165 | 15661 |  | 6597 | 6.679 | 13.547 | 6.868 |  | 0.574 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 0:31 |  | 1 |  | 7.6182 | 14.2728 | 6.6546 | 0.56427 |  |  | 0.71738 |  | 119.6551 |  |  | 0.14059 |  | 110.5197 |  |  | 1.1138 |  | 101.3843 |  |  | 0.86727 |  | 101.4165 |  | 7.591 | 14.2828 | 6.6919 | 0.56673 |  | 44
95 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8102 | 14177 |  | 4697 | 6.63 | 12.289 | 5.659 |  | 0.375 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 8:31 |  | 1 |  | 7.5558 | 12.8538 | 5.298 | 0.39448 |  |  | 7.7831 |  | 116.6835 |  |  | 3.0186 |  | 109.8727 |  |  | 4.1999 |  | 103.0619 |  |  | 0.54222 |  | 104.6265 |  | 7.8618 | 13.0508 | 5.189 | 0.39555 |  | 45
56 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8473 | 16028 |  | 7351 | 6.915 | 13.857 | 6.942 |  | 0.653 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 3:13 |  | 1 |  | 7.45 | 14.167 | 6.717 | 0.63802 |  |  | 0.35296 |  | 113.7386 |  |  | 0.18711 |  | 112.1455 |  |  | 0.002835 |  | 110.5523 |  |  | 0.52341 |  | 139.3404 |  | 7.4631 | 14.1802 | 6.7171 | 0.63636 |  | 46
98 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11360 | 15859 |  | 15581 | 9.13 | 13.714 | 4.584 |  | 1.514 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 13:38 |  | 1 |  | 10.0724 | 14.0124 | 3.94 | 1.3789 |  |  | 0.92458 |  | 122.2563 |  |  | 0.01959 |  | 107.7285 |  |  | 2.4745 |  | 93.2006 |  |  | 1.323 |  | 79.1063 |  | 10.1191 | 14.011 | 3.8918 | 1.3881 |  | 47
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8079 | 14778 |  | 5351 | 6.613 | 12.798 | 6.185 |  | 0.444 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 0:10 |  | 1 |  | 7.0747 | 13.0397 | 5.9651 | 0.47476 |  |  | 0.1788 |  | 110.4806 |  |  | 0.38974 |  | 101.2236 |  |  | 1.0683 |  | 91.9666 |  |  | 4.1778 |  | 94.4402 |  | 7.081 | 13.0144 | 5.9334 | 0.46504 |  | 48
95 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7339 | 14126 |  | 6113 | 6.045 | 12.246 | 6.201 |  | 0.523 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 9:30 |  | 1 |  | 6.4054 | 12.4492 | 6.0437 | 0.54492 |  |  | 0.25381 |  | 119.9117 |  |  | 0.44283 |  | 110.4119 |  |  | 0.64278 |  | 100.9122 |  |  | 1.1592 |  | 107.9947 |  | 6.4136 | 12.4768 | 6.0632 | 0.5481 |  | 49
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8938 | 15212 |  | 5782 | 7.272 | 13.166 | 5.894 |  | 0.489 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 1:46 |  | 1 |  | 7.7142 | 13.1883 | 5.4741 | 0.51947 |  |  | 0.31147 |  | 120.9043 |  |  | 0.82475 |  | 104.9308 |  |  | 2.4481 |  | 88.9572 |  |  | 0.8247 |  | 87.045 |  | 7.7262 | 13.1342 | 5.4079 | 0.51734 |  | 50
95 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10128 | 16295 |  | 4452 | 8.185 | 14.084 | 5.899 |  | 0.35 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 11:16 |  | 1 |  | 8.7748 | 14.1579 | 5.383 | 0.40126 |  |  | 0.18267 |  | 110.8081 |  |  | 0.35348 |  | 102.4919 |  |  | 0.63129 |  | 94.1757 |  |  | 0.71138 |  | 85.7591 |  | 8.7829 | 14.183 | 5.4001 | 0.39983 |  | 51
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8448 | 15323 |  | 5128 | 6.896 | 13.26 | 6.364 |  | 0.42 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 1:59 |  | 1 |  | 7.21 | 13.1307 | 5.9207 | 0.4489 |  |  | 0.061246 |  | 109.0372 |  |  | 0.49288 |  | 98.6205 |  |  | 1.021 |  | 88.2037 |  |  | 0.91386 |  | 89.7589 |  | 7.2078 | 13.0984 | 5.8907 | 0.45096 |  | 52
95 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8508 | 14659 |  | 4717 | 6.942 | 12.698 | 5.755 |  | 0.377 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 11:26 |  | 1 |  | 7.2629 | 12.5435 | 5.2805 | 0.41126 |  |  | 0.67844 |  | 109.7381 |  |  | 0.58691 |  | 101.2217 |  |  | 2.3538 |  | 92.7053 |  |  | 1.437 |  | 91.8009 |  | 7.2877 | 12.5068 | 5.2191 | 0.41423 |  | 53
56 | 19 | SPIKE 25 + 80 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11401 | 16990 |  | 15081 | 9.162 | 14.672 | 5.51 |  | 1.461 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 4:13 |  | 1 |  | 10.1582 | 14.8988 | 4.7406 | 1.4049 |  |  | 0.53403 |  | 125.0213 |  |  | 0.46134 |  | 108.4338 |  |  | 2.628 |  | 91.8463 |  |  | 0.46828 |  | 86.608 |  | 10.1854 | 14.8645 | 4.6791 | 1.4082 |  | 54
98 | 31 | SPIKE 25 + 80 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11434 | 16466 |  | 13183 | 9.187 | 14.229 | 5.042 |  | 1.263 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 13:43 |  | 1 |  | 10.1886 | 14.426 | 4.2374 | 1.2261 |  |  | 2.9772 |  | 137.2405 |  |  | 0.84166 |  | 111.6862 |  |  | 10.6595 |  | 86.1318 |  |  | 1.2218 |  | 69.1131 |  | 10.3425 | 14.3655 | 4.023 | 1.2336 |  | 55
53 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7436 | 13909 |  | 4999 | 6.119 | 12.062 | 5.943 |  | 0.407 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 3:28 |  | 1 |  | 6.5583 | 12.2663 | 5.708 | 0.49029 |  |  | 0.18208 |  | 99.6403 |  |  | 0.015054 |  | 95.7691 |  |  | 0.24203 |  | 91.8978 |  |  | 2.2243 |  | 103.3519 |  | 6.5643 | 12.2654 | 5.7011 | 0.4849 |  | 56
95 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7917 | 14515 |  | 4955 | 6.489 | 12.575 | 6.086 |  | 0.402 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 12:55 |  | 1 |  | 7.0005 | 12.8258 | 5.8253 | 0.48564 |  |  | 3.927 |  | 113.1906 |  |  | 0.83122 |  | 103.5232 |  |  | 6.8642 |  | 93.8559 |  |  | 1.0506 |  | 101.5489 |  | 7.1407 | 12.7727 | 5.632 | 0.4831 |  | 57
137 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11050 | 15778 |  | 14452 | 8.892 | 13.646 | 4.753 |  | 1.395 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 22:21 |  | 1 |  | 9.881 | 13.9918 | 4.1108 | 1.4898 |  |  | 3.0602 |  | 122.5632 |  |  | 0.56584 |  | 103.7121 |  |  | 5.6964 |  | 84.8609 |  |  | 1.6262 |  | 101.839 |  | 10.0346 | 14.0315 | 3.997 | 1.5021 |  | 58
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7396 | 17777 |  | 8520 | 6.089 | 15.339 | 9.25 |  | 0.775 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 0:12 |  | 1 |  | 6.9588 | 16.8459 | 9.887 | 0.90143 |  |  | 0.81145 |  | 110.0954 |  |  | 0.19895 |  | 102.9478 |  |  | 0.23441 |  | 95.8002 |  |  | 0.21779 |  | 95.3879 |  | 6.9872 | 16.8626 | 9.8755 | 0.90241 |  | 59
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11060 | 16474 |  | 8508 | 8.9 | 14.235 | 5.335 |  | 0.774 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 9:38 |  | 1 |  | 10.3095 | 15.6316 | 5.3222 | 0.90019 |  |  | 0.81275 |  | 148.339 |  |  | 0.61808 |  | 127.0849 |  |  | 0.23993 |  | 105.8307 |  |  | 1.6976 |  | 107.1101 |  | 10.3515 | 15.6801 | 5.3285 | 0.9079 |  | 60
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7641 | 15886 |  | 10361 | 6.277 | 13.737 | 7.46 |  | 0.968 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 3:21 |  | 1 |  | 7.0774 | 14.8361 | 7.7587 | 0.95632 |  |  | 2.1246 |  | 131.4899 |  |  | 0.2901 |  | 104.3463 |  |  | 1.3545 |  | 77.2027 |  |  | 0.84191 |  | 88.9279 |  | 7.003 | 14.8146 | 7.8116 | 0.96036 |  | 61
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9897 | 16506 |  | 8446 | 8.008 | 14.262 | 6.255 |  | 0.767 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 12:51 |  | 1 |  | 9.203 | 15.3845 | 6.1815 | 0.78372 |  |  | 0.58349 |  | 131.5218 |  |  | 0.36753 |  | 113.6335 |  |  | 0.045134 |  | 95.7452 |  |  | 1.6559 |  | 125.1293 |  | 9.23 | 15.4129 | 6.1829 | 0.79026 |  | 62
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8931 | 13900 |  | 6816 | 7.267 | 12.055 | 4.788 |  | 0.597 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 1:07 |  | 1 |  | 8.571 | 13.1003 | 4.5293 | 0.63358 |  |  | 0.71755 |  | 113.8742 |  |  | 0.1054 |  | 104.5645 |  |  | 1.0428 |  | 95.2548 |  |  | 0.43779 |  | 105.0366 |  | 8.5404 | 13.0934 | 4.553 | 0.6322 |  | 63
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9423 | 15710 |  | 8313 | 7.644 | 13.588 | 5.944 |  | 0.753 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 10:32 |  | 1 |  | 9.0349 | 14.7656 | 5.7307 | 0.76306 |  |  | 1.1103 |  | 137.2932 |  |  | 0.21785 |  | 115.1836 |  |  | 1.2056 |  | 93.074 |  |  | 0.03401 |  | 144.4361 |  | 9.0853 | 14.7817 | 5.6964 | 0.76293 |  | 64
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12270 | 17005 |  | 4945 | 9.829 | 14.685 | 4.857 |  | 0.401 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 4:50 |  | 1 |  | 12.1977 | 17.5836 | 5.386 | 0.49692 |  |  | 0.30092 |  | 98.3962 |  |  | 1.3001 |  | 95.1018 |  |  | 5.0224 |  | 91.8074 |  |  | 0.19903 |  | 101.1552 |  | 12.216 | 17.4701 | 5.254 | 0.49742 |  | 65
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8101 | 16869 |  | 6008 | 6.63 | 14.57 | 7.94 |  | 0.512 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:06 |  | 1 |  | 8.0557 | 17.4426 | 9.3869 | 0.60216 |  |  | 0.57798 |  | 138.0878 |  |  | 0.5485 |  | 120.7149 |  |  | 1.5254 |  | 103.3421 |  |  | 0.082172 |  | 105.1498 |  | 8.0791 | 17.3949 | 9.3158 | 0.60241 |  | 66
59 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8615 | 13793 |  | 4154 | 7.024 | 11.964 | 4.94 |  | 0.318 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 4:59 |  | 1 |  | 8.5939 | 14.0696 | 5.4757 | 0.44831 |  |  | 1.9868 |  | 122.8851 |  |  | 0.013863 |  | 104.8476 |  |  | 3.2367 |  | 86.8101 |  |  | 0.41605 |  | 88.6693 |  | 8.6801 | 14.0686 | 5.3885 | 0.44738 |  | 67
101 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6954 | 12801 |  | 4472 | 5.75 | 11.123 | 5.373 |  | 0.352 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 14:23 |  | 1 |  | 6.9663 | 13.1022 | 6.1359 | 0.48119 |  |  | 0.30897 |  | 85.6284 |  |  | 0.10415 |  | 89.8481 |  |  | 0.12891 |  | 94.0677 |  |  | 1.0475 |  | 96.1202 |  | 6.9771 | 13.109 | 6.1319 | 0.48373 |  | 68
62 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10737 | 16747 |  | 4130 | 8.652 | 14.467 | 5.814 |  | 0.316 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 1:28 |  | 1 |  | 10.4759 | 16.5269 | 6.0509 | 0.42081 |  |  | 5.3369 |  | 96.5269 |  |  | 1.2984 |  | 91.2783 |  |  | 13.9225 |  | 86.0297 |  |  | 1.9929 |  | 75.6721 |  | 10.7631 | 16.4203 | 5.6571 | 0.41666 |  | 69
104 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7635 | 14024 |  | 4824 | 6.272 | 12.16 | 5.888 |  | 0.389 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 10:54 |  | 1 |  | 7.43 | 13.8516 | 6.4216 | 0.4886 |  |  | 0.21167 |  | 116.5375 |  |  | 1.1485 |  | 114.0756 |  |  | 2.2435 |  | 111.6137 |  |  | 0.41898 |  | 109.8052 |  | 7.4222 | 13.7725 | 6.3503 | 0.48963 |  | 70
62 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6558 | 15761 |  | 7959 | 5.446 | 13.631 | 8.185 |  | 0.716 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 3:13 |  | 1 |  | 6.441 | 15.2188 | 8.7777 | 0.85121 |  |  | 4.5979 |  | 92.1117 |  |  | 0.99774 |  | 95.571 |  |  | 1.5644 |  | 99.0304 |  |  | 0.45154 |  | 102.4856 |  | 6.2963 | 15.1432 | 8.8469 | 0.84929 |  | 71
104 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8671 | 14505 |  | 5250 | 7.067 | 12.567 | 5.5 |  | 0.433 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 12:33 |  | 1 |  | 8.5076 | 14.0177 | 5.5101 | 0.56263 |  |  | 1.1544 |  | 119.7345 |  |  | 0.21806 |  | 101.3015 |  |  | 1.2451 |  | 82.8685 |  |  | 1.764 |  | 95.6042 |  | 8.557 | 14.033 | 5.476 | 0.56764 |  | 72
170 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5202 | 12873 |  | 5835 | 4.405 | 11.184 | 6.779 |  | 0.494 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 3:26 |  | 1 |  | 5.5095 | 13.1703 | 7.6608 | 0.64147 |  |  | 3.9742 |  | 56.2727 |  |  | 0.19733 |  | 79.815 |  |  | 2.4336 |  | 103.3573 |  |  | 0.37999 |  | 102.4099 |  | 5.4022 | 13.1573 | 7.7552 | 0.64269 |  | 73
212 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10864 | 15513 |  | 7374 | 8.749 | 13.421 | 4.672 |  | 0.655 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 12:50 |  | 1 |  | 11.2434 | 15.8066 | 4.5632 | 0.80488 |  |  | 0.55688 |  | 162.1906 |  |  | 0.41149 |  | 125.1614 |  |  | 2.8381 |  | 88.1323 |  |  | 0.93226 |  | 100.1117 |  | 11.2748 | 15.7742 | 4.4994 | 0.80865 |  | 74
63 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 15545 | 20359 |  | 5347 | 12.341 | 17.527 | 5.186 |  | 0.443 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 4:06 |  | 1 |  | 14.7822 | 19.5794 | 4.7972 | 0.52281 |  |  | 0.10279 |  | 133.6146 |  |  | 0.23162 |  | 108.7812 |  |  | 0.62966 |  | 83.9477 |  |  | 0.28157 |  | 79.4097 |  | 14.7746 | 19.5567 | 4.7821 | 0.52207 |  | 75
105 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6745 | 14289 |  | 6174 | 5.589 | 12.384 | 6.794 |  | 0.53 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 13:25 |  | 1 |  | 6.4295 | 13.7292 | 7.2997 | 0.60385 |  |  | 2.9512 |  | 124.2648 |  |  | 0.54606 |  | 104.4591 |  |  | 3.528 |  | 84.6534 |  |  | 1.482 |  | 93.8784 |  | 6.336 | 13.7668 | 7.4308 | 0.60836 |  | 76
63 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9165 | 14689 |  | 10809 | 7.446 | 12.723 | 5.277 |  | 1.015 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 3:27 |  | 1 |  | 8.9429 | 14.6298 | 5.6869 | 1.1549 |  |  | 2.4811 |  | 158.607 |  |  | 0.68594 |  | 132.0061 |  |  | 2.2039 |  | 105.4051 |  |  | 1.8124 |  | 103.1504 |  | 9.0552 | 14.6801 | 5.6249 | 1.1655 |  | 77
105 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9607 | 14827 |  | 14967 | 7.785 | 12.84 | 5.055 |  | 1.449 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 12:47 |  | 1 |  | 9.3673 | 14.766 | 5.3987 | 1.5963 |  |  | 3.7119 |  | 163.3463 |  |  | 1.2689 |  | 127.5988 |  |  | 3.1175 |  | 91.8512 |  |  | 2.2746 |  | 85.5507 |  | 9.5444 | 14.8603 | 5.3159 | 1.6146 |  | 78
63 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7507 | 21306 |  | 5828 | 6.174 | 18.329 | 12.155 |  | 0.494 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 23:16 |  | 1 |  | 7.2299 | 21.5997 | 14.3699 | 0.63724 |  |  | 4.4831 |  | 115.4754 |  |  | 1.8369 |  | 94.3212 |  |  | 5.1738 |  | 73.1669 |  |  | 1.1559 |  | 107.9818 |  | 7.3956 | 21.4031 | 14.0075 | 0.63358 |  | 79
105 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11466 | 15473 |  | 3613 | 9.212 | 13.387 | 4.176 |  | 0.262 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 8:33 |  | 1 |  | 11.0568 | 15.6739 | 4.6171 | 0.3987 |  |  | 0.55229 |  | 116.468 |  |  | 0.12323 |  | 104.6862 |  |  | 0.89685 |  | 92.9044 |  |  | 1.8466 |  | 100.1265 |  | 11.0263 | 15.6642 | 4.6379 | 0.40241 |  | 80
64 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8205 | 14921 |  | 7209 | 6.709 | 12.919 | 6.21 |  | 0.638 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 2:54 |  | 1 |  | 7.2081 | 14.1143 | 6.9062 | 0.70604 |  |  | 0.30747 |  | 100.2937 |  |  | 2.6931 |  | 90.7286 |  |  | 5.2444 |  | 81.1635 |  |  | 1.1369 |  | 110.9949 |  | 7.1971 | 13.9268 | 6.7297 | 0.71007 |  | 81
107 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8930 | 15177 |  | 9275 | 7.266 | 13.136 | 5.871 |  | 0.854 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 12:43 |  | 1 |  | 7.8499 | 14.3544 | 6.5045 | 0.91194 |  |  | 2.8566 |  | 122.4166 |  |  | 0.92329 |  | 118.6745 |  |  | 1.461 |  | 114.9324 |  |  | 2.7485 |  | 205.6204 |  | 7.9636 | 14.4209 | 6.4573 | 0.92465 |  | 82
64 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12856 | 18198 |  | 4702 | 10.278 | 15.696 | 5.418 |  | 0.376 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 4:57 |  | 1 |  | 11.097 | 17.2034 | 6.1064 | 0.45797 |  |  | 3.2102 |  | 92.0827 |  |  | 0.19024 |  | 101.6472 |  |  | 5.5419 |  | 111.2117 |  |  | 0.19143 |  | 108.2742 |  | 11.278 | 17.2198 | 5.9418 | 0.45753 |  | 83
107 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9098 | 14559 |  | 6586 | 7.394 | 12.613 | 5.218 |  | 0.573 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 14:54 |  | 1 |  | 7.9031 | 13.7974 | 5.8943 | 0.64131 |  |  | 0.32313 |  | 108.8998 |  |  | 0.25811 |  | 99.3919 |  |  | 0.17099 |  | 89.8839 |  |  | 1.4674 |  | 89.6727 |  | 7.8903 | 13.7796 | 5.8893 | 0.63664 |  | 84
64 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12643 | 18477 |  | 4180 | 10.114 | 15.932 | 5.818 |  | 0.321 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 1:34 |  | 1 |  | 12.1181 | 17.7894 | 5.6713 | 0.43048 |  |  | 1.4834 |  | 117.0156 |  |  | 0.25325 |  | 109.3696 |  |  | 4.0677 |  | 101.7235 |  |  | 2.6288 |  | 85.5943 |  | 12.2086 | 17.7669 | 5.5583 | 0.42489 |  | 85
107 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9092 | 14836 |  | 6156 | 7.39 | 12.848 | 5.457 |  | 0.528 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 11:11 |  | 1 |  | 8.7155 | 14.3037 | 5.5881 | 0.63297 |  |  | 0.6065 |  | 136.2413 |  |  | 0.33409 |  | 114.7813 |  |  | 1.7836 |  | 93.3212 |  |  | 0.29099 |  | 101.917 |  | 8.6892 | 14.3276 | 5.6384 | 0.63389 |  | 86
64 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8419 | 14430 |  | 5117 | 6.874 | 12.503 | 5.629 |  | 0.419 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 6:13 |  | 1 |  | 7.8573 | 13.7159 | 5.8586 | 0.52288 |  |  | 0.059011 |  | 129.1308 |  |  | 0.60429 |  | 106.675 |  |  | 1.487 |  | 84.2192 |  |  | 3.0577 |  | 86.9922 |  | 7.855 | 13.7574 | 5.9025 | 0.51501 |  | 87
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | Stored Cal TIC [mg/l] | Stored Cal TC [mg/l] | Stored Cal TOC (Diff.) [mg/l] | Stored Cal NPOC [mg/l] | Stored Cal TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] |  | TIC Absolute value Percent error for check standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes |  | TC Absolute value Percent error for check standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes |  | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes |  | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes |  | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order | Tap Water Conc TIC | Tap Water Conc TC | Tap Water Conc TNb | TIC Signal per mass injected | TC Signal per mass injected | TNb Signal per mass injected | TIC signal per ml
| Mean | [32 empty] | Mean |  |  | 109.5684 |  |  |  |  | 100.0327 |  |  |  |  | 90.4969 |  |  |  |  | 100.7142 | [5 empty] | Min | 1
| Std | [32 empty] | Std |  |  | 19.6683 |  |  |  |  | 14.4485 |  |  |  |  | 22.8426 |  |  |  |  | 21.6425 | [5 empty] | Max | 87
Warning Limit | Xbar + 2SD | [30 empty] | Warning Limit |  | Xbar + 2SD |  |  | 148.9051 |  |  |  |  | 128.9297 |  |  |  |  | 136.1821 |  |  |  |  | 143.9993
| Xbar - 2SD | [32 empty] | Xbar - 2SD |  |  | 70.2317 |  |  |  |  | 71.1357 |  |  |  |  | 44.8118 |  |  |  |  | 57.4291
Control Limit | Xbar + 3SD | [30 empty] | Control Limit |  | Xbar + 3SD |  |  | 168.5734 |  |  |  |  | 143.3782 |  |  |  |  | 159.0246 |  |  |  |  | 165.6419
| Xbar - 3SD | [32 empty] | Xbar - 3SD |  |  | 50.5634 |  |  |  |  | 56.6872 |  |  |  |  | 21.9692 |  |  |  |  | 35.7866
[32 empty] | Warning Limits as % |  | Xbar + 2SD |  |  | 135.9015 |  |  |  |  | 128.8876 |  |  |  |  | 150.4825 |  |  |  |  | 142.9781
[34 empty] | Xbar - 2SD |  |  | 64.0985 |  |  |  |  | 71.1124 |  |  |  |  | 49.5175 |  |  |  |  | 57.0219
[32 empty] | Control Limits as % |  | Xbar + 3SD |  |  | 153.8522 |  |  |  |  | 143.3314 |  |  |  |  | 175.7238 |  |  |  |  | 164.4672
[34 empty] | Xbar - 3SD |  |  | 46.1478 |  |  |  |  | 56.6686 |  |  |  |  | 24.2762 |  |  |  |  | 35.5328
Coefficient of Variation % | [31 empty] | Coefficient of Variation % |  |  |  |  | 17.9507 |  |  |  |  | 14.4438 |  |  |  |  | 25.2413 |  |  |  |  | 21.4891
[32 empty] | 3 x Coefficient of Variation % |  |  |  |  | 53.8522 |  |  |  |  | 43.3314 |  |  |  |  | 75.7238 |  |  |  |  | 64.4672

## rolling dups
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
50 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3940 | 13155 |  | 4816 | 3.438 | 11.423 | 7.986 |  | 0.388 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-14 | 18:43 |  | 1 |  | 6.86 | 12.2552 | 5.3952 | 0.39733 |  |  | 0.075598 |  | 95.5302 |  |  | 1.9056 |  | 93.4609 |  |  | 4.2813 |  | 91.3915 |  |  | 3.9136 |  | 65.6924 | 6.8574 | 12.1395 | 5.2822 | 0.38971 |  | 1
97 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3957 | 12405 |  | 4632 | 3.451 | 10.788 | 7.337 |  | 0.368 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-15 | 2:42 |  | 1 |  | 6.8893 | 11.5584 | 4.6691 | 0.38316 |  |  | 3.5236 |  | 113.7607 |  |  | 0.25754 |  | 95.8454 |  |  | 6.1103 |  | 77.93 |  |  | 0.30201 |  | 72.4067 | 7.0129 | 11.5435 | 4.5307 | 0.38258 |  | 2
50 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1816 | 12659 |  | 4214 | 1.808 | 11.003 | 9.195 |  | 0.325 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-16 | 19:21 |  | 1 |  | 3.8258 | 9.9935 | 6.1677 | 0.33984 |  |  | 1.1454 |  | 62.437 |  |  | 1.3781 |  | 74.716 |  |  | 2.9758 |  | 86.9949 |  |  | 2.4179 |  | 81.4963 | 3.8478 | 9.9251 | 6.0773 | 0.33578 |  | 3
4
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2884 | 5716 |  | 2113 | 2.627 | 5.121 | 2.494 |  | 0.105 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 22:43 |  | 1 |  | 2.8181 | 5.2698 | 2.4517 | 0.22619 |  |  | 0.035613 | 0.73273 |  |  |  | 1.3956 | 0.71921 |  |  |  | 2.9361 | 0.7039 |  |  |  | 1.9354 | 4.1444 |  | 2.8186 | 5.3068 | 2.4882 | 0.22402 |  | 5
108 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2414 | 4278 |  | 1292 | 2.267 | 3.903 | 1.637 |  | 0.019 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-16 | 6:56 |  | 1 |  | 2.3463 | 3.8306 | 1.4843 | 0.1488 |  |  | 1.1185 | 1.1031 |  |  |  | 0.41715 | 11.4189 |  |  |  | 2.7974 | 28.0174 |  |  |  | 0.12677 | 16.1401 |  | 2.3333 | 3.8386 | 1.5054 | 0.14871 |  | 6
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3348 | 6532 |  | 3289 | 2.983 | 5.812 | 2.829 |  | 0.228 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 23:46 |  | 1 |  | 3.4225 | 6.4054 | 2.9829 | 0.35157 |  |  | 1.0346 | 2.4583 |  |  |  | 12.4617 | 2.6173 |  |  |  | 24.056 | 2.779 |  |  |  | 13.1786 | 1.964 |  | 3.4403 | 6.831 | 3.3907 | 0.37637 |  | 7
108 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2920 | 5101 |  | 1599 | 2.655 | 4.6 | 1.945 |  | 0.051 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 8:01 |  | 1 |  | 2.9993 | 4.9906 | 1.9912 | 0.16693 |  |  | 0.92728 | 3.4874 |  |  |  | 0.4766 | 6.595 |  |  |  | 0.19844 | 11.4314 |  |  |  | 2.8544 | 26.6697 |  | 2.9855 | 4.9787 | 1.9932 | 0.16459 |  | 8
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2479 | 4532 |  | 1540 | 2.317 | 4.118 | 1.801 |  | 0.045 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 21:45 |  | 1 |  | 2.411 | 4.6766 | 2.2656 | 0.16801 |  |  | 1.1918 | 0.68039 |  |  |  | 4.8821 | 2.5278 |  |  |  | 8.9609 | 4.6097 |  |  |  | 19.9378 | 2.0748 |  | 2.3967 | 4.5652 | 2.1684 | 0.15278 |  | 9
108 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3096 | 5370 |  | 1324 | 2.79 | 4.828 | 2.038 |  | 0.022 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-19 | 5:59 |  | 1 |  | 3.0187 | 5.5217 | 2.503 | 0.14336 |  |  | 2.1103 | 3.694 |  |  |  | 0.16451 | 4.2944 |  |  |  | 2.1329 | 5.008 |  |  |  | 0.39867 | 5.5309 |  | 2.9872 | 5.5172 | 2.5299 | 0.14308 |  | 10
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2884 | 5076 |  | 1344 | 2.628 | 4.579 | 1.951 |  | 0.025 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 1:14 |  | 1 |  | 2.7911 | 4.7994 | 2.0083 | 0.13927 |  |  | 2.4696 | 1.9594 |  |  |  | 0.18227 | 3.0012 |  |  |  | 2.9121 | 4.4279 |  |  |  | 1.1329 | 0.11323 |  | 2.7571 | 4.7951 | 2.038 | 0.13849 |  | 11
108 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3822 | 5284 |  | 2316 | 3.347 | 4.755 | 1.408 |  | 0.126 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 9:28 |  | 1 |  | 3.7035 | 5.0014 | 1.2979 | 0.24095 |  |  | 0.84399 | 1.756 |  |  |  | 12.0209 | 3.7468 |  |  |  | 52.0428 | 21.2081 |  |  |  | 11.9624 | 3.2809 |  | 3.6879 | 4.7179 | 1.0299 | 0.22735 |  | 12
61 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3305 | 3924 |  | 1817 | 2.951 | 3.603 | 0.652 |  | 0.074 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 22:26 |  | 1 |  | 3.2053 | 3.7574 | 0.5521 | 0.17969 |  |  | 1.132 | 2.8077 |  |  |  | 1.4361 | 8.9064 |  |  |  | 15.1397 | 53.927 |  |  |  | 3.6802 | 4.3432 |  | 3.1873 | 3.7846 | 0.59732 | 0.17644 |  | 13
61 | 20 | DUP_F07jul21S1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4641 | 8721 |  | 21492 | 3.976 | 7.667 | 3.692 |  | 2.132 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 22:50 |  | 1 |  | 3.83 | 6.9005 | 3.0705 | 0.18451 |  |  | 0.94877 | 1.5901 |  |  |  | 0.81848 | 3.2327 |  |  |  | 0.6562 | 8.9262 |  |  |  | 1.644 | 0.63393 |  | 3.8119 | 6.8724 | 3.0605 | 0.18604 |  | 14
108 | 32 | DUP_F28jun21S8.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4967 | 7138 |  | 22238 | 4.225 | 6.326 | 2.1 |  | 2.21 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 8:15 |  | 1 |  | 4.1042 | 5.6101 | 1.5059 | 0.19105 |  |  | 1.1816 | 3.0691 |  |  |  | 10.1698 | 2.5047 |  |  |  | 48.9008 | 24.1513 |  |  |  | 10.3385 | 9.678 |  | 4.1286 | 5.3386 | 1.21 | 0.18166 |  | 15
61 | 20 | DUP_F31may21S9.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5207 | 7155 |  | 25168 | 4.409 | 6.34 | 1.93 |  | 2.516 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 23:10 |  | 1 |  | 4.306 | 5.6239 | 1.3179 | 0.21673 |  |  | 0.88284 | 0.83728 |  |  |  | 0.24612 | 3.0275 |  |  |  | 3.848 | 10.3462 |  |  |  | 1.3056 | 0.72169 |  | 4.2871 | 5.6308 | 1.3437 | 0.21815 |  | 16
108 | 32 | DUP_C_IM_01jun21 S site6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5177 | 8158 |  | 9913 | 4.387 | 7.19 | 2.803 |  | 0.921 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 8:29 |  | 1 |  | 4.2808 | 6.4416 | 2.1608 | 0.083035 |  |  | 0.33345 | 1.2136 |  |  |  | 0.89451 | 2.1521 |  |  |  | 1.9968 | 4.0216 |  |  |  | 0.72027 | 3.9314 |  | 4.288 | 6.4705 | 2.1825 | 0.082737 |  | 17
59 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5056 | 7177 |  | 1233 | 4.294 | 6.358 | 2.065 |  | 0.013 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 22:21 |  | 1 |  | 4.7216 | 6.6277 | 1.9061 | 0.10257 |  |  | 0.41106 | 5.762 |  |  |  | 12.2559 | 10.9048 |  |  |  | 51.8506 | 28.7781 |  |  |  | 9.5771 | 8.2397 |  | 4.7314 | 6.245 | 1.5136 | 0.097881 |  | 18
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2574 | 3652 |  | 2199 | 2.39 | 3.372 | 0.983 |  | 0.114 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-03 | 6:36 |  | 1 |  | 2.3081 | 3.1864 | 0.87835 | 0.19789 |  |  | 2.9066 | 28.9685 |  |  |  | 0.18366 | 15.6974 |  |  |  | 7.8727 | 28.398 |  |  |  | 0.99233 | 7.2762 |  | 2.275 | 3.1893 | 0.91434 | 0.19888 |  | 19
59 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4489 | 5574 |  | 3968 | 3.859 | 5.001 | 1.142 |  | 0.299 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 22:57 |  | 1 |  | 4.1265 | 5.335 | 1.2085 | 0.39401 |  |  | 0.49928 | 2.0288 |  |  |  | 0.33864 | 2.5532 |  |  |  | 0.21182 | 4.371 |  |  |  | 3.0693 | 3.4077 |  | 4.1368 | 5.344 | 1.2072 | 0.40015 |  | 20
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4372 | 5401 |  | 2513 | 3.769 | 4.854 | 1.085 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 7:14 |  | 1 |  | 4.0167 | 5.161 | 1.1444 | 0.24509 |  |  | 1.081 | 1.9662 |  |  |  | 0.058425 | 0.37071 |  |  |  | 3.9575 | 8.8025 |  |  |  | 0.04175 | 1.4085 |  | 3.9951 | 5.1625 | 1.1675 | 0.24514 |  | 21
59 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3528 | 5413 |  | 1593 | 3.122 | 4.864 | 1.742 |  | 0.051 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 0:18 |  | 1 |  | 3.2196 | 5.1294 | 1.9098 | 0.15221 |  |  | 0.75891 | 2.662 |  |  |  | 1.3204 | 1.3582 |  |  |  | 2.274 | 0.81809 |  |  |  | 2.8501 | 0.86537 |  | 3.2074 | 5.0958 | 1.8884 | 0.15007 |  | 22
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4354 | 6902 |  | 1426 | 3.756 | 6.126 | 2.37 |  | 0.033 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 8:37 |  | 1 |  | 3.9929 | 6.5814 | 2.5885 | 0.1356 |  |  | 0.56431 | 1.6357 |  |  |  | 0.60563 | 1.0991 |  |  |  | 2.3838 | 0.28904 |  |  |  | 0.88414 | 6.2456 |  | 3.9816 | 6.6014 | 2.6197 | 0.135 |  | 23
56 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3459 | 5270 |  | 8618 | 3.068 | 4.744 | 1.675 |  | 0.785 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 0:40 |  | 1 |  | 4.2753 | 6.4824 | 2.2071 | 1.2811 |  |  | 2.251 | 7.2574 |  |  |  | 0.12452 | 4.4877 |  |  |  | 4.5709 | 0.50222 |  |  |  | 0.25942 | 1.1023 |  | 4.2277 | 6.4865 | 2.2587 | 1.2827 |  | 24
98 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2063 | 4647 |  | 1481 | 1.998 | 4.216 | 2.218 |  | 0.039 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 9:02 |  | 1 |  | 2.4301 | 5.6438 | 3.2137 | 0.20156 |  |  | 5.3043 | 8.1347 |  |  |  | 1.2562 | 7.2199 |  |  |  | 5.9452 | 6.5709 |  |  |  | 2.8844 | 4.693 |  | 2.3673 | 5.6795 | 3.3121 | 0.20451 |  | 25
56 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3107 | 5975 |  | 1852 | 2.798 | 5.341 | 2.542 |  | 0.078 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 1:44 |  | 1 |  | 3.8031 | 7.4297 | 3.6266 | 0.26658 |  |  | 1.4638 | 1.7403 |  |  |  | 0.3388 | 2.081 |  |  |  | 2.1948 | 2.4329 |  |  |  | 5.6674 | 2.2062 |  | 3.7755 | 7.4423 | 3.6668 | 0.25924 |  | 26
98 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4257 | 5900 |  | 4892 | 3.681 | 5.277 | 1.596 |  | 0.396 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 10:01 |  | 1 |  | 5.3164 | 7.3302 | 2.0138 | 0.74173 |  |  | 0.2719 | 4.9001 |  |  |  | 1.7052 | 4.6844 |  |  |  | 5.3922 | 4.1317 |  |  |  | 2.0852 | 5.8595 |  | 5.3236 | 7.3932 | 2.0696 | 0.74954 |  | 27
56 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2856 | 4165 |  | 885 | 2.606 | 3.807 | 1.201 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 0:02 |  | 1 |  | 3.55 | 4.9781 | 1.4281 | 0.11544 |  |  | 0.51729 | 22.2777 |  |  |  | 0.53007 | 16.499 |  |  |  | 3.0867 | 0.67846 |  |  |  | 4.2016 | 3.7366 |  | 3.5408 | 4.9913 | 1.4505 | 0.11307 |  | 28
98 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3677 | 5400 |  | 2486 | 3.236 | 4.854 | 1.618 |  | 0.144 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 8:24 |  | 1 |  | 4.6241 | 6.6118 | 1.9877 | 0.35312 |  |  | 0.50799 | 5.6888 |  |  |  | 0.38087 | 3.0218 |  |  |  | 2.4796 | 3.0091 |  |  |  | 0.37909 | 8.657 |  | 4.6359 | 6.5992 | 1.9634 | 0.35246 |  | 29
56 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4666 | 6656 |  | 5272 | 3.994 | 5.918 | 1.924 |  | 0.435 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 0:38 |  | 1 |  | 5.9418 | 7.8821 | 1.9403 | 0.68375 |  |  | 0.81566 | 6.8059 |  |  |  | 2.6013 | 3.7324 |  |  |  | 12.3775 | 4.5773 |  |  |  | 1.0752 | 1.0623 |  | 5.9177 | 7.986 | 2.0683 | 0.68745 |  | 30
98 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3252 | 7659 |  | 2324 | 2.91 | 6.768 | 3.858 |  | 0.127 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 8:57 |  | 1 |  | 4.0459 | 9.1526 | 5.1067 | 0.30807 |  |  | 0.99914 | 4.8438 |  |  |  | 2.8782 | 2.9834 |  |  |  | 5.8455 | 1.5823 |  |  |  | 1.4374 | 2.7478 |  | 4.0258 | 9.2862 | 5.2604 | 0.3103 |  | 31
56 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3384 | 5648 |  | 1114 | 3.011 | 5.063 | 2.052 |  | 0.001 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 22:24 |  | 1 |  | 4.0498 | 6.4051 | 2.3553 | 0.12228 |  |  | 1.6855 | 6.383 |  |  |  | 1.2965 | 4.4636 |  |  |  | 6.2228 | 1.3698 |  |  |  | 5.5692 | 0.57447 |  | 4.0159 | 6.4469 | 2.431 | 0.12578 |  | 32
98 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4635 | 6118 |  | 7362 | 3.971 | 5.461 | 1.49 |  | 0.654 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-09 | 6:33 |  | 1 |  | 5.6475 | 6.9828 | 1.3353 | 0.87686 |  |  | 1.3662 | 1.5257 |  |  |  | 1.0332 | 2.1774 |  |  |  | 10.579 | 4.8126 |  |  |  | 2.1258 | 1.1857 |  | 5.6092 | 7.0191 | 1.4099 | 0.88628 |  | 33
56 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4735 | 8399 |  | 2477 | 4.048 | 7.394 | 3.346 |  | 0.143 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 1:12 |  | 1 |  | 4.2221 | 7.3289 | 3.1068 | 0.19609 |  |  | 7.2162 | 2.1868 |  |  |  | 3.6941 | 1.7357 |  |  |  | 0.90487 | 7.098 |  |  |  | 2.305 | 7.4102 |  | 4.0751 | 7.1959 | 3.1209 | 0.19385 |  | 34
98 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2842 | 6751 |  | 2555 | 2.596 | 5.998 | 3.402 |  | 0.151 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 9:06 |  | 1 |  | 2.3478 | 5.753 | 3.4052 | 0.20306 |  |  | 1.9589 | 6.17 |  |  |  | 0.16635 | 1.0199 |  |  |  | 1.0512 | 6.2153 |  |  |  | 3.3749 | 9.1158 |  | 2.325 | 5.7483 | 3.4232 | 0.20654 |  | 35
59 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6361 | 8675 |  | 8544 | 5.295 | 7.628 | 2.333 |  | 0.778 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 3:58 |  | 1 |  | 5.5315 | 7.4396 | 1.9081 | 0.74542 |  |  | 0.08214 | 2.4024 |  |  |  | 0.28245 | 4.2772 |  |  |  | 1.3319 | 9.8732 |  |  |  | 0.72197 | 3.8011 |  | 5.5293 | 7.4501 | 1.9209 | 0.74812 |  | 36
101 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3883 | 7194 |  | 3043 | 3.394 | 6.373 | 2.979 |  | 0.202 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 14:23 |  | 1 |  | 3.2806 | 6.0846 | 2.804 | 0.25022 |  |  | 3.2356 | 9.0891 |  |  |  | 0.13524 | 5.8221 |  |  |  | 3.9399 | 2.2552 |  |  |  | 1.6322 | 1.1124 |  | 3.2284 | 6.0887 | 2.8603 | 0.24819 |  | 37
56 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5090 | 8171 |  | 2792 | 4.32 | 7.201 | 2.881 |  | 0.176 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 0:52 |  | 1 |  | 4.3716 | 7.0554 | 2.6838 | 0.23913 |  |  | 1.9637 | 2.4316 |  |  |  | 0.12846 | 3.0585 |  |  |  | 2.7905 | 4.0636 |  |  |  | 0.038499 | 5.9669 |  | 4.329 | 7.0508 | 2.7218 | 0.23917 |  | 38
98 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5111 | 8252 |  | 2789 | 4.336 | 7.269 | 2.933 |  | 0.176 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 10:11 |  | 1 |  | 4.3905 | 7.1287 | 2.7382 | 0.23885 |  |  | 1.7034 | 7.065 |  |  |  | 0.26647 | 5.7971 |  |  |  | 3.346 | 3.8459 |  |  |  | 0.89061 | 4.3319 |  | 4.3535 | 7.1382 | 2.7848 | 0.23779 |  | 39
56 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5268 | 8402 |  | 2392 | 4.456 | 7.397 | 2.94 |  | 0.134 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 2:28 |  | 1 |  | 4.4432 | 7.0917 | 2.6486 | 0.21815 |  |  | 0.36173 | 4.8684 |  |  |  | 0.25216 | 5.313 |  |  |  | 1.2736 | 6.0571 |  |  |  | 0.16311 | 5.3809 |  | 4.4351 | 7.1007 | 2.6655 | 0.21797 |  | 40
98 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8317 | 11172 |  | 2026 | 6.796 | 9.743 | 2.947 |  | 0.096 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 11:59 |  | 1 |  | 7.1607 | 9.5716 | 2.4108 | 0.18562 |  |  | 0.037334 | 3.714 |  |  |  | 1.072 | 2.8904 |  |  |  | 4.4407 | 0.42961 |  |  |  | 2.0381 | 1.1311 |  | 7.1621 | 9.5205 | 2.3585 | 0.18753 |  | 41
56 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4976 | 6362 |  | 2042 | 4.233 | 5.669 | 1.436 |  | 0.098 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 2:41 |  | 1 |  | 4.1445 | 5.2058 | 1.0612 | 0.16628 |  |  | 1.3943 | 3.2481 |  |  |  | 1.1654 | 6.9649 |  |  |  | 10.5684 | 21.8692 |  |  |  | 13.2191 | 8.3873 |  | 4.1159 | 5.2363 | 1.1204 | 0.17805 |  | 42
98 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4735 | 8347 |  | 2217 | 4.047 | 7.35 | 3.303 |  | 0.116 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 12:07 |  | 1 |  | 3.9318 | 6.9613 | 3.0295 | 0.18231 |  |  | 0.51781 | 2.9818 |  |  |  | 0.89796 | 3.5658 |  |  |  | 2.7059 | 4.3163 |  |  |  | 1.1488 | 2.1737 |  | 3.9216 | 6.9927 | 3.071 | 0.18336 |  | 43
56 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4504 | 6207 |  | 728 | 3.871 | 5.537 | 1.666 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 4:09 |  | 1 |  | 3.8626 | 5.1557 | 1.2931 | 0.038678 |  |  | 0.38157 | 6.2592 |  |  |  | 0.26896 | 6.4599 |  |  |  | 0.066655 | 7.0602 |  |  |  | 1.0995 | 3.6434 |  | 3.8553 | 5.1488 | 1.2936 | 0.038467 |  | 44
98 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4738 | 8187 |  | 2026 | 4.05 | 7.214 | 3.164 |  | 0.096 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 13:37 |  | 1 |  | 4.0778 | 6.9837 | 2.9059 | 0.17593 |  |  | 2.8354 | 17.9614 |  |  |  | 0.052864 | 10.5473 |  |  |  | 3.9699 | 0.52878 |  |  |  | 2.3756 | 8.1741 |  | 4.0208 | 6.9856 | 2.9648 | 0.17804 |  | 45
140 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4870 | 8541 |  | 2431 | 4.151 | 7.515 | 3.363 |  | 0.138 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 23:02 |  | 1 |  | 4.1991 | 7.3105 | 3.1114 | 0.21875 |  |  | 2.1444 | 2.0898 |  |  |  | 0.26555 | 0.15186 |  |  |  | 3.6126 | 3.3508 |  |  |  | 1.2158 | 3.2871 |  | 4.2446 | 7.3008 | 3.0562 | 0.21743 |  | 46
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4320 | 10940 |  | 4805 | 3.729 | 9.547 | 5.818 |  | 0.387 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 0:53 |  | 1 |  | 4.1459 | 10.4746 | 6.3287 | 0.51716 |  |  | 2.1402 | 1.6068 |  |  |  | 0.7322 | 3.1288 |  |  |  | 0.17957 | 4.1267 |  |  |  | 1.3289 | 1.1543 |  | 4.102 | 10.4364 | 6.3343 | 0.51375 |  | 47
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2660 | 8441 |  | 6159 | 2.456 | 7.43 | 4.974 |  | 0.528 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 10:19 |  | 1 |  | 2.6279 | 8.1458 | 5.5179 | 0.65721 |  |  | 1.1198 | 22.9679 |  |  |  | 0.82031 | 5.0367 |  |  |  | 1.7312 | 4.6976 |  |  |  | 0.11023 | 7.3975 |  | 2.6133 | 8.1793 | 5.5661 | 0.65685 |  | 48
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6842 | 9265 |  | 7204 | 5.664 | 8.128 | 2.464 |  | 0.637 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 4:03 |  | 1 |  | 6.3525 | 8.7818 | 2.4293 | 0.66951 |  |  | 1.5689 | 2.7259 |  |  |  | 0.83648 | 0.85581 |  |  |  | 1.0539 | 10.7307 |  |  |  | 1.7521 | 4.2495 |  | 6.3031 | 8.7453 | 2.4422 | 0.6637 |  | 49
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4092 | 8245 |  | 3915 | 3.554 | 7.263 | 3.709 |  | 0.293 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 13:32 |  | 1 |  | 3.9041 | 7.8368 | 3.9327 | 0.37196 |  |  | 0.28096 | 2.3571 |  |  |  | 0.069927 | 2.5614 |  |  |  | 0.41705 | 2.7639 |  |  |  | 0.17117 | 5.4122 |  | 3.8987 | 7.8396 | 3.9409 | 0.37164 |  | 50
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5125 | 7841 |  | 9917 | 4.347 | 6.921 | 2.575 |  | 0.921 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 1:48 |  | 1 |  | 4.9828 | 7.5258 | 2.543 | 0.90179 |  |  | 1.2758 | 7.3948 |  |  |  | 0.47793 | 3.847 |  |  |  | 1.0672 | 2.6896 |  |  |  | 0.30738 | 1.9692 |  | 4.9512 | 7.5078 | 2.5566 | 0.9004 |  | 51
104 | 32 | DUP of B50 18jul22 8.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5194 | 8441 |  | 18316 | 4.4 | 7.429 | 3.03 |  | 1.8 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 11:13 |  | 1 |  | 5.0478 | 8.0778 | 3.0299 | 1.6282 |  |  | 0.69345 | 4.0579 |  |  |  | 0.11396 | 0.11403 |  |  |  | 0.84405 | 7.415 |  |  |  | 0.58603 | 2.0642 |  | 5.0304 | 8.0732 | 3.0428 | 1.6235 |  | 52
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4965 | 9537 |  | 2599 | 4.224 | 8.358 | 4.134 |  | 0.156 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 5:32 |  | 1 |  | 4.9401 | 9.8387 | 4.8987 | 0.26466 |  |  | 0.66588 | 8.949 |  |  |  | 14.226 | 12.3981 |  |  |  | 29.8925 | 16.5275 |  |  |  | 7.4899 | 6.2637 |  | 4.9237 | 9.1854 | 4.2617 | 0.25511 |  | 53
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4821 | 8318 |  | 2841 | 4.113 | 7.325 | 3.212 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 14:47 |  | 1 |  | 4.797 | 8.5745 | 3.7775 | 0.28862 |  |  | 0.1035 | 4.9967 |  |  |  | 0.32603 | 7.5954 |  |  |  | 0.60789 | 10.9832 |  |  |  | 0.82665 | 7.5788 |  | 4.7995 | 8.5885 | 3.7891 | 0.28743 |  | 54
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3039 | 6460 |  | 1953 | 2.746 | 5.751 | 3.005 |  | 0.088 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 6:05 |  | 1 |  | 3.2155 | 7.4167 | 4.2012 | 0.21559 |  |  | 0.06132 | 2.3419 |  |  |  | 1.1895 | 4.9329 |  |  |  | 2.1575 | 6.984 |  |  |  | 3.3386 | 1.233 |  | 3.2165 | 7.3729 | 4.1564 | 0.21205 |  | 55
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3491 | 5799 |  | 1932 | 3.093 | 5.192 | 2.099 |  | 0.086 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 15:22 |  | 1 |  | 3.6612 | 6.7347 | 3.0734 | 0.21358 |  |  | 0.8385 | 5.5458 |  |  |  | 0.44332 | 5.9828 |  |  |  | 1.9491 | 6.4985 |  |  |  | 0.584 | 9.1613 |  | 3.646 | 6.7496 | 3.1037 | 0.21296 |  | 56
62 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4495 | 7820 |  | 2279 | 3.864 | 6.903 | 3.04 |  | 0.122 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 5:40 |  | 1 |  | 4.5568 | 8.2447 | 3.688 | 0.25442 |  |  | 0.8205 | 7.1451 |  |  |  | 0.84334 | 5.5797 |  |  |  | 0.87157 | 3.6773 |  |  |  | 0.57064 | 0.53129 |  | 4.5382 | 8.2101 | 3.672 | 0.2537 |  | 57
104 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5891 | 10646 |  | 2380 | 4.934 | 9.298 | 4.363 |  | 0.133 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 15:04 |  | 1 |  | 5.9247 | 11.0007 | 5.0759 | 0.26487 |  |  | 0.54728 | 5.811 |  |  |  | 0.8815 | 3.0005 |  |  |  | 1.273 | 0.19579 |  |  |  | 1.4154 | 4.8527 |  | 5.9085 | 10.9524 | 5.0438 | 0.263 |  | 58
65 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5218 | 7785 |  | 2344 | 4.418 | 6.874 | 2.456 |  | 0.129 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 2:10 |  | 1 |  | 5.0568 | 7.722 | 2.6652 | 0.24634 |  |  | 0.77974 | 18.9153 |  |  |  | 0.25478 | 7.4869 |  |  |  | 0.73371 | 11.0839 |  |  |  | 1.6796 | 0.75692 |  | 5.0371 | 7.7122 | 2.675 | 0.24429 |  | 59
107 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8400 | 11402 |  | 1625 | 6.859 | 9.939 | 3.08 |  | 0.054 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 11:36 |  | 1 |  | 8.1812 | 11.2756 | 3.0944 | 0.1761 |  |  | 0.14392 | 1.3225 |  |  |  | 0.21759 | 1.4421 |  |  |  | 0.41212 | 9.1253 |  |  |  | 1.5413 | 5.5134 |  | 8.1871 | 11.2879 | 3.1008 | 0.17747 |  | 60
65 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4574 | 7874 |  | 2647 | 3.924 | 6.95 | 3.025 |  | 0.161 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 3:54 |  | 1 |  | 4.5006 | 7.6767 | 3.1761 | 0.28535 |  |  | 4.1479 | 3.0286 |  |  |  | 0.5496 | 3.1207 |  |  |  | 4.3375 | 3.246 |  |  |  | 1.9222 | 9.3088 |  | 4.4092 | 7.6557 | 3.2465 | 0.28263 |  | 61
107 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4323 | 7574 |  | 2106 | 3.731 | 6.695 | 2.964 |  | 0.104 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 13:14 |  | 1 |  | 4.2552 | 7.3898 | 3.1347 | 0.22772 |  |  | 1.3243 | 13.7767 |  |  |  | 1.6439 | 1.983 |  |  |  | 5.8178 | 12.5587 |  |  |  | 2.6064 | 4.855 |  | 4.2835 | 7.3296 | 3.0461 | 0.22479 |  | 62
173 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4377 | 7152 |  | 7963 | 3.773 | 6.338 | 2.565 |  | 0.717 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 4:07 |  | 1 |  | 4.674 | 7.4573 | 2.7833 | 0.86742 |  |  | 0.58329 | 3.1712 |  |  |  | 0.90124 | 3.316 |  |  |  | 3.4446 | 3.5647 |  |  |  | 0.22009 | 3.9463 |  | 4.6877 | 7.4238 | 2.7361 | 0.86838 |  | 63
215 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6239 | 8892 |  | 4017 | 5.202 | 7.811 | 2.61 |  | 0.304 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 13:31 |  | 1 |  | 6.5597 | 9.1949 | 2.6352 | 0.44844 |  |  | 4.8505 | 5.074 |  |  |  | 0.55542 | 5.3678 |  |  |  | 15.4094 | 6.1794 |  |  |  | 0.023681 | 4.4295 |  | 6.7227 | 9.1694 | 2.4467 | 0.44838 |  | 64
66 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5995 | 8142 |  | 1652 | 5.014 | 7.176 | 2.162 |  | 0.057 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 4:48 |  | 1 |  | 5.7176 | 7.8048 | 2.0872 | 0.1607 |  |  | 0.62885 | 9.927 |  |  |  | 0.2349 | 6.4314 |  |  |  | 2.6398 | 2.7152 |  |  |  | 1.4743 | 1.2673 |  | 5.7357 | 7.7957 | 2.06 | 0.15953 |  | 65
108 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5931 | 8456 |  | 2920 | 4.965 | 7.443 | 2.477 |  | 0.189 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 14:07 |  | 1 |  | 5.6569 | 8.1075 | 2.4506 | 0.28496 |  |  | 0.56887 | 4.2111 |  |  |  | 0.45071 | 5.3836 |  |  |  | 0.17743 | 8.1477 |  |  |  | 0.7248 | 1.0984 |  | 5.673 | 8.1258 | 2.4528 | 0.28394 |  | 66
66 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3380 | 11562 |  | 7754 | 3.008 | 10.073 | 7.066 |  | 0.695 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 4:09 |  | 1 |  | 3.3887 | 11.5427 | 8.1541 | 0.83064 |  |  | 4.6763 | 6.0434 |  |  |  | 0.2998 | 3.588 |  |  |  | 2.4432 | 2.5483 |  |  |  | 1.7142 | 3.9227 |  | 3.4698 | 11.5254 | 8.0557 | 0.82358 |  | 67
108 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6171 | 8902 |  | 3390 | 5.149 | 7.82 | 2.671 |  | 0.239 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 13:29 |  | 1 |  | 6.0683 | 8.9167 | 2.8484 | 0.36741 |  |  | 0.12649 | 3.6385 |  |  |  | 0.28745 | 4.4183 |  |  |  | 0.6295 | 6.0956 |  |  |  | 1.3964 | 8.2705 |  | 6.0722 | 8.9295 | 2.8574 | 0.36487 |  | 68
66 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6037 | 8398 |  | 1733 | 5.046 | 7.393 | 2.347 |  | 0.065 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 23:57 |  | 1 |  | 5.8089 | 8.4863 | 2.6774 | 0.19623 |  |  | 2.1699 | 35.3642 |  |  |  | 0.29883 | 11.3696 |  |  |  | 5.4532 | 25.7246 |  |  |  | 1.471 | 4.4839 |  | 5.7466 | 8.499 | 2.7524 | 0.19768 |  | 69
108 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2793 | 7025 |  | 1324 | 2.557 | 6.23 | 3.673 |  | 0.022 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 9:12 |  | 1 |  | 2.6732 | 7.0914 | 4.4183 | 0.15218 |  |  | 57.1176 | 10.8363 |  |  |  | 1.5212 | 13.4915 |  |  |  | 59.5921 | 48.6418 |  |  |  | 7.0329 | 33.2924 |  | 3.7418 | 7.1458 | 3.404 | 0.15773 |  | 70
67 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3022 | 7039 |  | 3114 | 2.733 | 6.242 | 3.509 |  | 0.21 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 3:34 |  | 1 |  | 2.6204 | 6.7237 | 4.1034 | 0.29792 |  |  | 20.9566 | 2.3903 |  |  |  | 2.3157 | 5.845 |  |  |  | 11.7636 | 12.5348 |  |  |  | 0.30062 | 8.4306 |  | 2.9271 | 6.8025 | 3.8754 | 0.29837 |  | 71
110 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3855 | 6348 |  | 1707 | 3.372 | 5.656 | 2.284 |  | 0.063 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 13:23 |  | 1 |  | 3.3577 | 6.0758 | 2.7181 | 0.15769 |  |  | 13.4044 | 3.9421 |  |  |  | 1.055 | 4.8708 |  |  |  | 22.3382 | 19.3864 |  |  |  | 4.1466 | 2.9836 |  | 3.5989 | 6.0439 | 2.445 | 0.16103 |  | 72
67 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7962 | 9797 |  | 14979 | 6.523 | 8.578 | 2.055 |  | 1.451 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 5:41 |  | 1 |  | 6.9376 | 9.3403 | 2.4027 | 1.4581 |  |  | 3.224 | 1.6567 |  |  |  | 1.3239 | 4.0989 |  |  |  | 13.3826 | 10.8701 |  |  |  | 2.1929 | 1.2639 |  | 6.8276 | 9.4026 | 2.575 | 1.4423 |  | 73
110 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5401 | 8359 |  | 4137 | 4.559 | 7.361 | 2.802 |  | 0.317 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 15:35 |  | 1 |  | 4.7611 | 7.9944 | 3.2333 | 0.40298 |  |  | 0.69861 | 5.2012 |  |  |  | 1.2912 | 5.9224 |  |  |  | 4.1503 | 6.9677 |  |  |  | 1.0569 | 4.4696 |  | 4.7445 | 8.0464 | 3.3019 | 0.40512 |  | 74
67 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10013 | 12800 |  | 1576 | 8.097 | 11.122 | 3.025 |  | 0.049 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 2:16 |  | 1 |  | 9.598 | 12.3545 | 2.7565 | 0.16363 |  |  | 0.37864 | 1.3442 |  |  |  | 0.038737 | 1.9163 |  |  |  | 1.1539 | 3.9498 |  |  |  | 5.866 | 2.5791 |  | 9.6162 | 12.3569 | 2.7407 | 0.15897 |  | 75
110 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8538 | 10320 |  | 17671 | 6.965 | 9.021 | 2.056 |  | 1.732 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 11:52 |  | 1 |  | 8.1847 | 9.9803 | 1.7956 | 1.813 |  |  | 0.24615 | 3.0283 |  |  |  | 0.24971 | 3.156 |  |  |  | 0.26597 | 3.7406 |  |  |  | 1.0235 | 8.9752 |  | 8.1747 | 9.9678 | 1.7932 | 1.8223 |  | 76
67 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10107 | 12450 |  | 1576 | 8.168 | 10.826 | 2.657 |  | 0.049 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 6:55 |  | 1 |  | 9.4221 | 11.8453 | 2.4232 | 0.16996 |  |  | 14.433 | 7.836 |  |  |  | 1.1551 | 1.105 |  |  |  | 98.7279 | 32.7914 |  |  |  | 7.9887 | 17.4499 |  | 10.155 | 11.7773 | 1.6223 | 0.16343 |  | 77
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | Stored Cal TIC [mg/l] | Stored Cal TC [mg/l] | Stored Cal TOC (Diff.) [mg/l] | Stored Cal NPOC [mg/l] | Stored Cal TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] |  | TIC Absolute value Percent error for check standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes |  | TC Absolute value Percent error for check standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes |  | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes |  | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order | Tap Water Conc TIC | Tap Water Conc TC | Tap Water Conc TNb | TIC Signal per mass injected | TC Signal per mass injected | TNb Signal per mass injected | TIC signal per ml
| Mean | [32 empty] | Mean |  | 5.8225 | 90.576 |  |  |  | 4.7264 | 88.0074 |  |  |  | 9.1069 | 85.4388 |  |  |  | 5.2271 | 73.1985 |  |  |  |  | Min | 1
| Std | [32 empty] | Std |  | 6.4205 | 26.0181 |  |  |  | 3.4957 | 11.5723 |  |  |  | 10.4425 | 6.8643 |  |  |  | 5.4203 | 7.9316 |  |  |  |  | Max | 77
Warning Limit | Xbar + 2SD | [30 empty] | Warning Limit |  | Xbar + 2SD |  | 18.6635 | 142.6121 |  |  |  | 11.7179 | 111.152 |  |  |  | 29.992 | 99.1675 |  |  |  | 16.0676 | 89.0617
| Xbar - 2SD | [32 empty] | Xbar - 2SD |  | -7.0184 | 38.5398 |  |  |  | -2.265 | 64.8628 |  |  |  | -11.7782 | 71.7102 |  |  |  | -5.6135 | 57.3352
Control Limit | Xbar + 3SD | [30 empty] | Control Limit |  | Xbar + 3SD |  | 25.084 | 168.6302 |  |  |  | 15.2136 | 122.7243 |  |  |  | 40.4346 | 106.0318 |  |  |  | 21.4879 | 96.9934
| Xbar - 3SD | [32 empty] | Xbar - 3SD |  | -13.4389 | 12.5218 |  |  |  | -5.7608 | 53.2905 |  |  |  | -22.2207 | 64.8459 |  |  |  | -11.0338 | 49.4036
[32 empty] | Warning Limits as % |  | Xbar + 2SD |  | 320.5388 | 157.4503 |  |  |  | 247.9232 | 126.2985 |  |  |  | 329.3316 | 116.0684 |  |  |  | 307.3936 | 121.6716
[34 empty] | Xbar - 2SD |  | -120.5388 | 42.5497 |  |  |  | -47.9232 | 73.7015 |  |  |  | -129.3316 | 83.9316 |  |  |  | -107.3936 | 78.3284
[32 empty] | Control Limits as % |  | Xbar + 3SD |  | 430.8081 | 186.1754 |  |  |  | 321.8848 | 139.4477 |  |  |  | 443.9974 | 124.1026 |  |  |  | 411.0904 | 132.5073
[34 empty] | Xbar - 3SD |  | -230.8081 | 13.8246 |  |  |  | -121.8848 | 60.5523 |  |  |  | -243.9974 | 75.8974 |  |  |  | -211.0904 | 67.4927
Coefficient of Variation % | [31 empty] | Coefficient of Variation % |  |  |  | 110.2694 | 28.7251 |  |  |  | 73.9616 | 13.1492 |  |  |  | 114.6658 | 8.0342 |  |  |  | 103.6968 | 10.8358
[32 empty] | 3 x Coefficient of Variation % |  |  |  | 330.8081 | 86.1754 |  |  |  | 221.8848 | 39.4477 |  |  |  | 343.9974 | 24.1026 |  |  |  | 311.0904 | 32.5073

## rolling blanks
We know that our Type I water has TOC in it.
We know we have carry over so the first run is thrown out
On 25JUL22 (Analysis Order = 70) we adopted a correction for
the signal from the water in the daily calibration curve
by subtracting the typical signal for the water
from the signal for the standard prepared in that water
This will correspond to a shift upward in concentration, but not necessarily signal
We also have an upward trend in signal to account for
Changed E-pure cartridges on 14jul22
This could make things worse...
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 0 | 1.041 | 0.7 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 12:42 |  | 1 |  | 0.74229 | 0.047595 | -0.69469 | 0.027025
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 138 | 474 |  | 0 | 0.521 | 0.68 | 0.159 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 12:50 |  | 1 |  | 0.061721 | 0.023576 | -0.038145 | 0.027025 | [20 empty] | 0.034117 | 0.013568 | -0.020549 | 0.027025 |  | 1
7 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 83 | 454 |  | 0 | 0.479 | 0.663 | 0.185 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 12:58 |  | 1 |  | 0.006513 | 0.0035596 | -0.0029534 | 0.027025
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 164 | 0.534 | 0.444 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 23:12 |  | 1 |  | 0.079789 | -0.25565 | -0.33544 | 0.042483
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 27 | 212 |  | 161 | 0.435 | 0.458 | 0.023 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 23:19 |  | 1 |  | -0.049699 | -0.23863 | -0.18894 | 0.0422 |  |  | 17.5779 |  |  |  |  | 5.704 |  |  |  |  | 11.0231 |  |  |  |  | 3.1766 |  |  | -0.045684 | -0.24564 | -0.19996 | 0.04154 |  | 2
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 35 | 198 |  | 147 | 0.442 | 0.446 | 0.004 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 23:27 |  | 1 |  | -0.041669 | -0.25265 | -0.21098 | 0.04088
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 196 | 0.5 | 0.445 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-16 | 7:25 |  | 1 |  | 0.034619 | -0.25465 | -0.28927 | 0.045499
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 20 | 154 |  | 163 | 0.43 | 0.409 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-16 | 7:32 |  | 1 |  | -0.056725 | -0.29668 | -0.23996 | 0.042389 |  |  | 25.994 |  |  |  |  | 0 |  |  |  |  | 5.2942 |  |  |  |  | 10.0416 |  |  | -0.050201 | -0.29668 | -0.24648 | 0.040362 |  | 3
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 33 | 154 |  | 120 | 0.44 | 0.409 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-16 | 7:40 |  | 1 |  | -0.043676 | -0.29668 | -0.25301 | 0.038336
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 141 | 0.846 | 0.47 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 13:50 |  | 1 |  | 0.66896 | 0.17072 | -0.49823 | 0.0076447
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 51 | 191 |  | 155 | 0.454 | 0.44 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 13:57 |  | 1 |  | 0.16274 | 0.13612 | -0.026618 | 0.0091742 | [20 empty] | 0.15878 | 0.12821 | -0.030573 | 0.0097205 |  | 2
7 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 43 | 175 |  | 165 | 0.448 | 0.427 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 14:05 |  | 1 |  | 0.15483 | 0.1203 | -0.034528 | 0.010267
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 120 | 0.569 | 0.431 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 0:15 |  | 1 |  | 0.31104 | 0.12524 | -0.1858 | 0.0053504
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 35 | 183 |  | 102 | 0.442 | 0.434 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 0:22 |  | 1 |  | 0.14692 | 0.12821 | -0.018708 | 0.0033838 |  |  | 7.1385 |  |  |  |  | 12.3032 |  |  |  |  | 37.6808 |  |  |  |  | 33.9975 |  |  | 0.15236 | 0.13661 | -0.015742 | 0.0028922 |  | 5
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 46 | 200 |  | 93 | 0.45 | 0.448 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 0:30 |  | 1 |  | 0.15779 | 0.14502 | -0.012776 | 0.0024006
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 319 | 0.559 | 0.551 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 8:29 |  | 1 |  | 0.29819 | 0.26564 | -0.032553 | 0.027092
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 55 | 192 |  | 147 | 0.457 | 0.441 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 8:36 |  | 1 |  | 0.16669 | 0.13711 | -0.029584 | 0.0083002 |  |  | 4.6351 |  |  |  |  | 24.257 |  |  |  |  | 77.6735 |  |  |  |  | 5.13 |  |  | 0.17065 | 0.12228 | -0.04837 | 0.0085187 |  | 6
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 63 | 162 |  | 151 | 0.463 | 0.416 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 8:44 |  | 1 |  | 0.1746 | 0.10745 | -0.067155 | 0.0087372
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 10 | 0.816 | 0.415 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 11:48 |  | 1 |  | 0.48442 | 0.26955 | -0.21487 | -0.006541
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 55 | 119 |  | 0 | 0.457 | 0.379 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 11:55 |  | 1 |  | 0.023456 | 0.22618 | 0.20273 | -0.0076819 | [20 empty] | 0.016561 | 0.23324 | 0.21668 | -0.0076819 |  | 7
7 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 41 | 133 |  | 0 | 0.446 | 0.391 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 12:03 |  | 1 |  | 0.0096663 | 0.2403 | 0.23064 | -0.0076819
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 93 | 0.53 | 0.48 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 22:13 |  | 1 |  | 0.11703 | 0.34619 | 0.22916 | 0.0029279
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 51 | 210 |  | 100 | 0.454 | 0.457 | 0.003 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 22:20 |  | 1 |  | 0.019516 | 0.31795 | 0.29844 | 0.0037265 |  |  | 26.2998 |  |  |  |  | 1.2768 |  |  |  |  | 3.3884 |  |  |  |  | 59.6075 |  |  | 0.022471 | 0.31594 | 0.29347 | 0.0028709 |  | 8
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 57 | 206 |  | 85 | 0.459 | 0.453 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 22:28 |  | 1 |  | 0.025426 | 0.31392 | 0.28849 | 0.0020152
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 138 | 0.54 | 0.452 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-19 | 6:28 |  | 1 |  | 0.12983 | 0.3119 | 0.18207 | 0.0080617
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 50 | 168 |  | 64 | 0.453 | 0.421 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-19 | 6:35 |  | 1 |  | 0.018531 | 0.2756 | 0.25707 | -0.00038051 |  |  | 105.5721 |  |  |  |  | 0.73453 |  |  |  |  | 4.1102 |  |  |  |  | 1203.6891 |  |  | 0.012129 | 0.27459 | 0.26246 | 7.5823e-05 |  | 9
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 37 | 166 |  | 72 | 0.443 | 0.419 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-19 | 6:42 |  | 1 |  | 0.0057265 | 0.27358 | 0.26786 | 0.00053216
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 21 | 0.793 | 0.465 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-21 | 15:21 |  | 1 |  | 0.46544 | 0.083705 | -0.38173 | 0.00088675
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 58 | 190 |  | 6 | 0.459 | 0.44 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-21 | 15:28 |  | 1 |  | 0.042318 | 0.054572 | 0.012254 | -0.00068226 | [20 empty] | 0.030646 | 0.05263 | 0.021984 | -0.00099607 |  | 10
7 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 34 | 186 |  | 0 | 0.441 | 0.436 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-21 | 15:36 |  | 1 |  | 0.018974 | 0.050687 | 0.031714 | -0.0013099
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 119 | 0.531 | 0.454 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 1:42 |  | 1 |  | 0.13278 | 0.071081 | -0.061697 | 0.011138
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 46 | 159 |  | 33 | 0.45 | 0.414 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 1:49 |  | 1 |  | 0.030646 | 0.024467 | -0.0061788 | 0.002142 |  |  | 9.9977 |  |  |  |  | 38.4675 |  |  |  |  | 1316.3611 |  |  |  |  | 119.6607 |  |  | 0.029187 | 0.030294 | 0.0011069 | 0.0053323 |  | 11
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 43 | 171 |  | 94 | 0.448 | 0.423 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 1:57 |  | 1 |  | 0.027728 | 0.036121 | 0.0083927 | 0.0085226
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 182 | 0.568 | 0.408 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 9:56 |  | 1 |  | 0.18044 | 0.018641 | -0.1618 | 0.017728
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 59 | 122 |  | 183 | 0.46 | 0.381 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 10:03 |  | 1 |  | 0.043291 | -0.011464 | -0.054755 | 0.017832 |  |  | 50.6945 |  |  |  |  | 291.372 |  |  |  |  | 79.3407 |  |  |  |  | 16.5103 |  |  | 0.034537 | -0.0046661 | -0.039203 | 0.016472 |  | 12
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 41 | 136 |  | 157 | 0.447 | 0.394 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 10:11 |  | 1 |  | 0.025783 | 0.0021317 | -0.023651 | 0.015113
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 85 | 0.821 | 0.506 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 12:24 |  | 1 |  | 0.49834 | 0.078628 | -0.41971 | 0.0011696
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 41 | 219 |  | 35 | 0.446 | 0.464 | 0.018 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 12:31 |  | 1 |  | 0.022473 | 0.028303 | 0.0058301 | -0.0039838 | [20 empty] | 0.030761 | 0.050949 | 0.020188 | -0.0016132 |  | 13
7 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 58 | 264 |  | 81 | 0.459 | 0.502 | 0.043 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 12:39 |  | 1 |  | 0.03905 | 0.073596 | 0.034546 | 0.00075736
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 118 | 0.562 | 0.491 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 22:54 |  | 1 |  | 0.16972 | 0.060511 | -0.10921 | 0.0045709
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 87 | 197 |  | 114 | 0.481 | 0.445 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 23:01 |  | 1 |  | 0.067329 | 0.0061593 | -0.06117 | 0.0041587 |  |  | 106.5592 |  |  |  |  | 39.373 |  |  |  |  | 137.4271 |  |  |  |  | 53.3784 |  |  | 0.043926 | 0.0076691 | -0.036257 | 0.0032826 |  | 14
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 39 | 200 |  | 97 | 0.445 | 0.448 | 0.003 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 23:09 |  | 1 |  | 0.020522 | 0.0091788 | -0.011343 | 0.0024065
81 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 129 | 0.54 | 0.455 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-23 | 1:48 |  | 1 |  | 0.14144 | 0.018237 | -0.1232 | 0.0057047
82 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 38 | 189 |  | 113 | 0.444 | 0.439 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-23 | 1:55 |  | 1 |  | 0.019547 | -0.0018928 | -0.02144 | 0.0040556 |  |  | 180.1733 |  |  |  |  | 114.1407 |  |  |  |  | 91.8513 |  |  |  |  | 43.2792 |  |  | 0.010283 | -0.0044091 | -0.014692 | 0.0033341 |  | 15
83 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 19 | 184 |  | 99 | 0.429 | 0.434 | 0.005 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-23 | 2:03 |  | 1 |  | 0.0010194 | -0.0069254 | -0.0079448 | 0.0026126
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 543 | 1.134 | 0.436 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 11:40 |  | 1 |  | 0.71549 | -0.057246 | -0.77274 | 0.00091914
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 142 | 186 |  | 509 | 0.524 | 0.436 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 11:47 |  | 1 |  | 0.045994 | -0.057246 | -0.10324 | 0.00062118 | [20 empty] | 0.029593 | -0.051539 | -0.081133 | 0.00051163 |  | 16
7 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 103 | 200 |  | 484 | 0.494 | 0.448 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 11:55 |  | 1 |  | 0.013192 | -0.045833 | -0.059025 | 0.00040208
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 574 | 0.539 | 0.433 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 23:21 |  | 1 |  | 0.062816 | -0.060506 | -0.12332 | 0.0011908
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6 | 135 |  | 484 | 0.42 | 0.393 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 23:28 |  | 1 |  | -0.068392 | -0.098821 | -0.030429 | 0.00040208 |  |  | 1.2374 |  |  |  |  | 10.9185 |  |  |  |  | 33.5209 |  |  |  |  | 74.9519 |  |  | -0.067971 | -0.10453 | -0.036556 | 0.00064309 |  | 17
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7 | 121 |  | 539 | 0.42 | 0.381 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 23:36 |  | 1 |  | -0.067551 | -0.11023 | -0.042683 | 0.00088409
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 412 | 0.574 | 0.419 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 8:48 |  | 1 |  | 0.10151 | -0.07355 | -0.17506 | -0.00022891
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 28 | 114 |  | 158 | 0.436 | 0.375 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 8:55 |  | 1 |  | -0.049888 | -0.11594 | -0.066052 | -0.0024549 |  |  | 21.1114 |  |  |  |  | 4.3096 |  |  |  |  | 28.8749 |  |  |  |  | 41.879 |  |  | -0.055776 | -0.11349 | -0.057719 | -0.0020299 |  | 18
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 14 | 120 |  | 255 | 0.426 | 0.38 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 9:03 |  | 1 |  | -0.061663 | -0.11105 | -0.049386 | -0.0016048
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 735 | 0.759 | 0.429 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 12:04 |  | 1 |  | 0.30336 | -0.064583 | -0.36795 | 0.0026018
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 20 | 157 |  | 586 | 0.43 | 0.412 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 12:11 |  | 1 |  | -0.056617 | -0.080887 | -0.02427 | 0.001296 | [20 empty] | -0.062084 | -0.087408 | -0.025324 | 0.0011601 |  | 19
7 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7 | 141 |  | 555 | 0.42 | 0.398 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 12:19 |  | 1 |  | -0.067551 | -0.09393 | -0.026379 | 0.0010243
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 538 | 0.55 | 0.442 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 23:41 |  | 1 |  | 0.074591 | -0.051539 | -0.12613 | 0.00087533
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 0 | 146 |  | 573 | 0 | 0.402 | 0.402 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 23:49 |  | 1 |  | -0.073438 | -0.089854 | -0.016415 | 0.0011821 |  |  | 8.3517 |  |  |  |  | 0.91139 |  |  |  |  | 26.7646 |  |  |  |  | 21.7994 |  |  | -0.070495 | -0.089446 | -0.018952 | 0.0013267 |  | 20
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7 | 147 |  | 606 | 0.42 | 0.403 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-22 | 23:56 |  | 1 |  | -0.067551 | -0.089039 | -0.021488 | 0.0014713
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 535 | 0.565 | 0.431 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 9:00 |  | 1 |  | 0.091412 | -0.062137 | -0.15355 | 0.00084903
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 21 | 163 |  | 577 | 0.431 | 0.416 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 9:07 |  | 1 |  | -0.055776 | -0.075995 | -0.02022 | 0.0012171 |  |  | 22.7228 |  |  |  |  | 6.6502 |  |  |  |  | 180.6106 |  |  |  |  | 107.3242 |  |  | -0.062925 | -0.07355 | -0.010625 | 0.00079207 |  | 21
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4 | 169 |  | 480 | 0.418 | 0.422 | 0.003 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-23 | 9:15 |  | 1 |  | -0.070074 | -0.071104 | -0.00103 | 0.00036703
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 170 | 0.832 | 0.529 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 12:42 |  | 1 |  | 0.59637 | 0.088523 | -0.50785 | 0.016675
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 50 | 587 |  | 113 | 0.454 | 0.776 | 0.322 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 12:49 |  | 1 |  | 0.10822 | 0.37406 | 0.26584 | 0.011109 | [20 empty] | 0.10426 | 0.22688 | 0.12261 | 0.0075943 |  | 22
7 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 42 | 287 |  | 41 | 0.447 | 0.521 | 0.074 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 12:57 |  | 1 |  | 0.1003 | 0.079692 | -0.020609 | 0.0040792
37 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 134 | 0.557 | 0.525 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 18:26 |  | 1 |  | 0.24189 | 0.083617 | -0.15828 | 0.01316
38 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 57 | 223 |  | 62 | 0.459 | 0.468 | 0.009 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 18:33 |  | 1 |  | 0.11515 | 0.016894 | -0.098259 | 0.0061297 |  |  | 4.3937 |  |  |  |  | 75.7653 |  |  |  |  | 29.8972 |  |  |  |  | 14.7539 |  |  | 0.11268 | 0.027197 | -0.085481 | 0.0066179 |  | 23
39 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 52 | 244 |  | 72 | 0.455 | 0.485 | 0.031 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 18:41 |  | 1 |  | 0.1102 | 0.037499 | -0.072703 | 0.0071061
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 336 | 3.193 | 0.776 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 12:11 |  | 1 |  | 3.3262 | 0.19419 | -3.132 | 0.014056
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 576 | 459 |  | 178 | 0.856 | 0.667 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 12:19 |  | 1 |  | 0.36513 | 0.069229 | -0.2959 | -0.0015343 | [20 empty] | 0.25914 | 0.069229 | -0.18991 | 4.4482e-05 |  | 24
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 358 | 459 |  | 210 | 0.69 | 0.668 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 12:27 |  | 1 |  | 0.15314 | 0.069229 | -0.083913 | 0.0016233
61 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 122 | 1.51 | 0.492 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 22:45 |  | 1 |  | 1.1927 | -0.13286 | -1.3255 | -0.0070601
62 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 91 | 252 |  | 143 | 0.485 | 0.492 | 0.008 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 22:52 |  | 1 |  | -0.1065 | -0.13286 | -0.026357 | -0.004988 |  |  | 15.9641 |  |  |  |  | 16.8252 |  |  |  |  | 20.224 |  |  |  |  | 72.9108 |  |  | -0.11574 | -0.14506 | -0.029322 | -0.0078496 |  | 25
63 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 72 | 227 |  | 85 | 0.47 | 0.471 | 0.001 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 23:00 |  | 1 |  | -0.12497 | -0.15726 | -0.032287 | -0.010711
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 234 | 0.982 | 0.437 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-03 | 6:59 |  | 1 |  | 0.52461 | -0.19631 | -0.72092 | 0.0039915
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 40 | 188 |  | 213 | 0.446 | 0.437 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-03 | 7:06 |  | 1 |  | -0.15609 | -0.19534 | -0.039243 | 0.0019193 |  |  | 0 |  |  |  |  | 13.9479 |  |  |  |  | 54.3504 |  |  |  |  | 7283.1859 |  |  | -0.15609 | -0.20998 | -0.053887 | -5.4194e-05 |  | 26
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 40 | 158 |  | 173 | 0.445 | 0.412 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-03 | 7:14 |  | 1 |  | -0.15609 | -0.22462 | -0.06853 | -0.0020277
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 234 | 3.03 | 0.588 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 12:46 |  | 1 |  | 3.1116 | 0.097833 | -3.0138 | 0.011841
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 241 | 365 |  | 164 | 0.6 | 0.588 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 12:53 |  | 1 |  | 0.13836 | 0.097833 | -0.040523 | 0.0046766 | [20 empty] | 0.1285 | 0.079736 | -0.048763 | 0.003039 |  | 27
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 220 | 329 |  | 132 | 0.584 | 0.557 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 13:01 |  | 1 |  | 0.11864 | 0.061639 | -0.057002 | 0.0014015
61 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 166 | 1.409 | 0.518 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 23:20 |  | 1 |  | 1.1288 | 0.01539 | -1.1134 | 0.0048813
62 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 125 | 262 |  | 141 | 0.511 | 0.5 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 23:27 |  | 1 |  | 0.029452 | -0.0057229 | -0.035175 | 0.0023226 |  |  | 42.5182 |  |  |  |  | 140.6826 |  |  |  |  | 38.5886 |  |  |  |  | 24.7609 |  |  | 0.024289 | -0.019296 | -0.043584 | 0.0020667 |  | 28
63 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 114 | 235 |  | 136 | 0.502 | 0.478 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 23:35 |  | 1 |  | 0.019125 | -0.032869 | -0.051994 | 0.0018109
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 217 | 1.392 | 0.513 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 7:38 |  | 1 |  | 1.1082 | 0.0093581 | -1.0988 | 0.010101
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 135 | 244 |  | 165 | 0.519 | 0.485 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 7:45 |  | 1 |  | 0.038841 | -0.02382 | -0.062661 | 0.004779 |  |  | 76.997 |  |  |  |  | 70.8599 |  |  |  |  | 7.0029 |  |  |  |  | 17.5805 |  |  | 0.028044 | -0.03689 | -0.064934 | 0.0052395 |  | 29
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 112 | 218 |  | 174 | 0.501 | 0.464 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 7:53 |  | 1 |  | 0.017247 | -0.04996 | -0.067208 | 0.0057001
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 201 | 3.056 | 0.582 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 14:04 |  | 1 |  | 3.1391 | 0.20115 | -2.9379 | 0.013751
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 266 | 344 |  | 98 | 0.619 | 0.57 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 14:12 |  | 1 |  | 0.16568 | 0.18653 | 0.020846 | 0.0035057 | [20 empty] | 0.15164 | 0.1797 | 0.028063 | 0.00091952 |  | 30
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 236 | 330 |  | 46 | 0.596 | 0.558 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 14:19 |  | 1 |  | 0.13759 | 0.17287 | 0.03528 | -0.0016667
61 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 139 | 1.207 | 0.558 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 0:42 |  | 1 |  | 0.88375 | 0.17287 | -0.71087 | 0.007584
62 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 97 | 281 |  | 115 | 0.489 | 0.517 | 0.027 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 0:49 |  | 1 |  | 0.0074624 | 0.12509 | 0.11763 | 0.0051967 |  |  | 91.3926 |  |  |  |  | 0.77649 |  |  |  |  | 4.6955 |  |  |  |  | 16.5823 |  |  | 0.0051219 | 0.12558 | 0.12046 | 0.0047988 |  | 31
63 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 92 | 282 |  | 107 | 0.485 | 0.518 | 0.032 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 0:57 |  | 1 |  | 0.0027814 | 0.12607 | 0.12329 | 0.004401
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 116 | 1.418 | 0.583 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 9:01 |  | 1 |  | 1.1403 | 0.20115 | -0.93911 | 0.0052962
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 157 | 317 |  | 144 | 0.536 | 0.547 | 0.012 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 9:08 |  | 1 |  | 0.063635 | 0.1602 | 0.096564 | 0.0080813 |  |  | 98.9644 |  |  |  |  | 10.9125 |  |  |  |  | 23.3694 |  |  |  |  | 34.6599 |  |  | 0.04257 | 0.15191 | 0.10934 | 0.0068877 |  | 32
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 112 | 300 |  | 120 | 0.501 | 0.533 | 0.032 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-06 | 9:16 |  | 1 |  | 0.021505 | 0.14362 | 0.12212 | 0.0056941
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 334 | 2.323 | 1.132 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-04 | 14:20 |  | 1 |  | 2.9906 | 0.74522 | -2.2453 | 0.028076
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 520 | 853 |  | 304 | 0.814 | 1.001 | 0.187 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-04 | 14:28 |  | 1 |  | 0.39061 | 0.53657 | 0.14596 | 0.023538 | [20 empty] | 0.30932 | 0.52445 | 0.21513 | 0.023765 |  | 33
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 397 | 835 |  | 307 | 0.72 | 0.986 | 0.266 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-04 | 14:37 |  | 1 |  | 0.22803 | 0.51234 | 0.28431 | 0.023992
58 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 379 | 1.318 | 0.585 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 1:06 |  | 1 |  | 1.259 | -0.12438 | -1.3834 | 0.034882
59 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 226 | 366 |  | 359 | 0.588 | 0.588 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 1:14 |  | 1 |  | 0.0020042 | -0.119 | -0.121 | 0.031857 |  |  | 216.4071 |  |  |  |  | 47.4544 |  |  |  |  | 16.0852 |  |  |  |  | 12.0834 |  |  | -0.024431 | -0.15602 | -0.13159 | 0.030042 |  | 34
60 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 186 | 311 |  | 335 | 0.557 | 0.542 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 1:22 |  | 1 |  | -0.050867 | -0.19304 | -0.14217 | 0.028227
100 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 263 | 0.967 | 0.586 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 9:28 |  | 1 |  | 0.65497 | -0.12304 | -0.778 | 0.017337
101 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 148 | 347 |  | 237 | 0.529 | 0.573 | 0.044 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 9:36 |  | 1 |  | -0.10109 | -0.14458 | -0.043481 | 0.013404 |  |  | 30.9448 |  |  |  |  | 12.2377 |  |  |  |  | 52.8042 |  |  |  |  | 17.1528 |  |  | -0.1196 | -0.154 | -0.034399 | 0.012345 |  | 35
102 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 120 | 333 |  | 223 | 0.507 | 0.561 | 0.054 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 9:44 |  | 1 |  | -0.1381 | -0.16342 | -0.025317 | 0.011286
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 368 | 2.25 | 1.029 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 15:20 |  | 1 |  | 2.8622 | 0.67625 | -2.186 | 0.034638
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 482 | 669 |  | 250 | 0.785 | 0.845 | 0.06 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 15:28 |  | 1 |  | 0.3489 | 0.38828 | 0.039371 | 0.016195 | [20 empty] | 0.23903 | 0.40088 | 0.16185 | 0.018149 |  | 36
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 315 | 688 |  | 275 | 0.657 | 0.861 | 0.205 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 15:37 |  | 1 |  | 0.12915 | 0.41349 | 0.28434 | 0.020103
58 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 188 | 1.195 | 0.631 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 2:10 |  | 1 |  | 1.0529 | 0.052528 | -1.0004 | 0.006505
59 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 194 | 349 |  | 126 | 0.564 | 0.574 | 0.01 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 2:18 |  | 1 |  | -0.030069 | -0.036386 | -0.0063164 | -0.0031854 |  |  | 29.7939 |  |  |  |  | 3.715 |  |  |  |  | 3044.9454 |  |  |  |  | 352.2685 |  |  | -0.035333 | -0.035722 | -0.00038931 | -0.0011536 |  | 37
60 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 186 | 350 |  | 152 | 0.558 | 0.575 | 0.017 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 2:26 |  | 1 |  | -0.040596 | -0.035059 | 0.0055378 | 0.00087827
100 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 187 | 1.499 | 0.62 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 10:27 |  | 1 |  | 1.574 | 0.035276 | -1.5387 | 0.0063487
101 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 296 | 360 |  | 141 | 0.642 | 0.584 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 10:35 |  | 1 |  | 0.10415 | -0.021788 | -0.12594 | -0.00084099 |  |  | 417.1725 |  |  |  |  | 64.4216 |  |  |  |  | 301.4413 |  |  |  |  | 432.9693 |  |  | 0.033751 | -0.01648 | -0.050231 | 0.00072198 |  | 38
102 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 189 | 368 |  | 161 | 0.56 | 0.591 | 0.031 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 10:44 |  | 1 |  | -0.036649 | -0.011172 | 0.025477 | 0.0022849
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 311 | 2.217 | 1.025 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 13:37 |  | 1 |  | 2.8866 | 0.6338 | -2.2528 | 0.030227
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 387 | 747 |  | 260 | 0.712 | 0.911 | 0.199 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 13:45 |  | 1 |  | 0.31974 | 0.45653 | 0.13679 | 0.022656 | [20 empty] | 0.24582 | 0.44661 | 0.20079 | 0.022804 |  | 39
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 274 | 732 |  | 262 | 0.625 | 0.899 | 0.274 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 13:54 |  | 1 |  | 0.1719 | 0.43669 | 0.26479 | 0.022953
58 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 123 | 1.133 | 0.576 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 0:27 |  | 1 |  | 1.038 | -0.067322 | -1.1053 | 0.002317
59 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 186 | 312 |  | 133 | 0.557 | 0.543 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 0:35 |  | 1 |  | 0.056769 | -0.11891 | -0.17568 | 0.0038016 |  |  | 141.7803 |  |  |  |  | 53.6035 |  |  |  |  | 76.6685 |  |  |  |  | 42.6395 |  |  | 0.03322 | -0.093779 | -0.127 | 0.0031335 |  | 40
60 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 150 | 350 |  | 124 | 0.53 | 0.575 | 0.045 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 0:44 |  | 1 |  | 0.0096702 | -0.068645 | -0.078315 | 0.0024655
100 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 201 | 1.339 | 0.639 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 8:50 |  | 1 |  | 1.3886 | 0.03057 | -1.3581 | 0.013897
101 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 264 | 369 |  | 127 | 0.617 | 0.591 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 8:58 |  | 1 |  | 0.15882 | -0.04351 | -0.20233 | 0.0029109 |  |  | 77.8109 |  |  |  |  | 122.604 |  |  |  |  | 86.3614 |  |  |  |  | 62.9212 |  |  | 0.11434 | -0.026974 | -0.14131 | 0.004247 |  | 41
102 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 196 | 394 |  | 145 | 0.565 | 0.612 | 0.046 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 9:07 |  | 1 |  | 0.069852 | -0.010439 | -0.080291 | 0.0055831
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 459 | 2.301 | 1.363 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 14:23 |  | 1 |  | 2.9813 | 1.0724 | -1.909 | 0.070402
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 526 | 900 |  | 386 | 0.818 | 1.041 | 0.223 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 14:31 |  | 1 |  | 0.39096 | 0.59102 | 0.20006 | 0.0611 | [20 empty] | 0.26627 | 0.71009 | 0.44382 | 0.059443 |  | 42
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 340 | 1088 |  | 360 | 0.675 | 1.2 | 0.525 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 14:40 |  | 1 |  | 0.14158 | 0.82916 | 0.68758 | 0.057786
58 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 253 | 1.598 | 0.683 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 1:04 |  | 1 |  | 1.7532 | 0.056474 | -1.6967 | 0.044151
59 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 287 | 453 |  | 243 | 0.635 | 0.663 | 0.028 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 1:12 |  | 1 |  | 0.070516 | 0.024807 | -0.045709 | 0.042876 |  |  | 91.429 |  |  |  |  | 33.9211 |  |  |  |  | 293.6303 |  |  |  |  | 22.1156 |  |  | 0.048393 | 0.029874 | -0.01852 | 0.038607 |  | 43
60 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 254 | 461 |  | 176 | 0.61 | 0.669 | 0.059 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 1:21 |  | 1 |  | 0.026271 | 0.034941 | 0.00867 | 0.034338
100 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 182 | 1.245 | 0.724 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 9:23 |  | 1 |  | 1.1364 | 0.11728 | -1.0192 | 0.035103
101 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 210 | 550 |  | 242 | 0.576 | 0.745 | 0.169 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 9:31 |  | 1 |  | -0.032724 | 0.14768 | 0.1804 | 0.042749 |  |  | 62.1356 |  |  |  |  | 41.4147 |  |  |  |  | 12.4669 |  |  |  |  | 0.59444 |  |  | -0.047472 | 0.12234 | 0.16981 | 0.042876 |  | 44
102 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 188 | 510 |  | 244 | 0.559 | 0.711 | 0.152 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 9:40 |  | 1 |  | -0.062221 | 0.097008 | 0.15923 | 0.043004
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 386 | 2.337 | 1.064 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 12:16 |  | 1 |  | 2.9271 | 0.60334 | -2.3238 | 0.034353
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 421 | 772 |  | 281 | 0.738 | 0.933 | 0.195 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 12:24 |  | 1 |  | 0.26551 | 0.41159 | 0.14608 | 0.021672 | [20 empty] | 0.19399 | 0.47059 | 0.2766 | 0.023786 |  | 45
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 309 | 868 |  | 316 | 0.652 | 1.014 | 0.362 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 12:33 |  | 1 |  | 0.12246 | 0.52959 | 0.40713 | 0.025899
58 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 114 | 1.239 | 0.596 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 22:49 |  | 1 |  | 1.1008 | -0.0764 | -1.1772 | 0.0015031
59 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 177 | 399 |  | 128 | 0.551 | 0.616 | 0.065 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 22:57 |  | 1 |  | -0.046123 | -0.0469 | -0.00077704 | 0.0031939 |  |  | 18.1198 |  |  |  |  | 46.3198 |  |  |  |  | 191.709 |  |  |  |  | 283.2102 |  |  | -0.042291 | -0.061035 | -0.018744 | 0.001322 |  | 46
60 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 183 | 376 |  | 97 | 0.555 | 0.597 | 0.042 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 23:06 |  | 1 |  | -0.03846 | -0.075171 | -0.036712 | -0.00055
100 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 325 | 1.577 | 0.607 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-09 | 6:59 |  | 1 |  | 1.6627 | -0.060421 | -1.7232 | 0.026986
101 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 264 | 417 |  | 245 | 0.618 | 0.632 | 0.014 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-09 | 7:07 |  | 1 |  | 0.064992 | -0.024774 | -0.089766 | 0.017324 |  |  | 114.2431 |  |  |  |  | 0 |  |  |  |  | 71.4497 |  |  |  |  | 14.1848 |  |  | 0.041364 | -0.024774 | -0.066138 | 0.016177 |  | 47
102 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 227 | 417 |  | 226 | 0.589 | 0.632 | 0.043 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-09 | 7:15 |  | 1 |  | 0.017736 | -0.024774 | -0.04251 | 0.01503
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 570 | 2.913 | 1.4 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-09 | 13:55 |  | 1 |  | 2.7577 | 0.56279 | -2.1949 | 0.025666
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 883 | 1063 |  | 468 | 1.093 | 1.179 | 0.087 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-09 | 14:05 |  | 1 |  | 0.40819 | 0.31417 | -0.094014 | 0.016551 | [20 empty] | 0.31165 | 0.40979 | 0.098142 | 0.022002 |  | 48
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 688 | 1263 |  | 590 | 0.942 | 1.349 | 0.406 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-09 | 14:15 |  | 1 |  | 0.21512 | 0.50541 | 0.2903 | 0.027454
58 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 142 | 1.459 | 0.63 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 1:36 |  | 1 |  | 0.88146 | -0.30544 | -1.1869 | -0.012582
59 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 222 | 456 |  | 173 | 0.585 | 0.665 | 0.08 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 1:43 |  | 1 |  | -0.24628 | -0.26624 | -0.019962 | -0.0098116 |  |  | 16.5907 |  |  |  |  | 8.2635 |  |  |  |  | 235.8948 |  |  |  |  | 10.5473 |  |  | -0.26855 | -0.27771 | -0.0091593 | -0.0093201 |  | 49
60 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 177 | 432 |  | 184 | 0.551 | 0.644 | 0.094 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 1:51 |  | 1 |  | -0.29083 | -0.28919 | 0.0016439 | -0.0088286
100 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 293 | 1.114 | 0.654 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 9:30 |  | 1 |  | 0.43591 | -0.27867 | -0.71458 | 0.00091222
101 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 154 | 400 |  | 268 | 0.533 | 0.617 | 0.085 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 9:37 |  | 1 |  | -0.3136 | -0.31979 | -0.006182 | -0.0013219 |  |  | 8.176 |  |  |  |  | 6.4828 |  |  |  |  | 271.7802 |  |  |  |  | 127.4753 |  |  | -0.32697 | -0.30975 | 0.017225 | -0.0036454 |  | 50
102 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 127 | 421 |  | 216 | 0.513 | 0.635 | 0.122 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 9:45 |  | 1 |  | -0.34034 | -0.29971 | 0.040631 | -0.0059689
7 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 592 | 2.983 | 1.25 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 15:01 |  | 1 |  | 2.7938 | 0.55214 | -2.2416 | 0.029575
8 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 576 | 1018 |  | 474 | 0.857 | 1.141 | 0.284 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 15:12 |  | 1 |  | 0.27674 | 0.43412 | 0.15737 | 0.018953 | [20 empty] | 0.22088 | 0.45058 | 0.2297 | 0.018278 |  | 51
9 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 453 | 1054 |  | 459 | 0.762 | 1.172 | 0.409 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 15:24 |  | 1 |  | 0.16502 | 0.46705 | 0.30203 | 0.017602
61 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 479 | 2.051 | 0.743 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 4:31 |  | 1 |  | 1.6901 | 0.0041065 | -1.686 | 0.019403
62 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 418 | 494 |  | 350 | 0.736 | 0.697 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 4:42 |  | 1 |  | 0.13323 | -0.045299 | -0.17852 | 0.0077902 |  |  | 128.9694 |  |  |  |  | 230.3188 |  |  |  |  | 149.8785 |  |  |  |  | 10.9706 |  |  | 0.080996 | -0.021054 | -0.10205 | 0.0073851 |  | 52
63 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 303 | 547 |  | 341 | 0.647 | 0.742 | 0.095 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 4:54 |  | 1 |  | 0.028766 | 0.0031916 | -0.025574 | 0.00698
103 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 246 | 1.386 | 0.691 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 14:55 |  | 1 |  | 0.9035 | -0.051703 | -0.95521 | -0.001572
104 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 286 | 464 |  | 202 | 0.634 | 0.671 | 0.037 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 15:06 |  | 1 |  | 0.013324 | -0.072746 | -0.08607 | -0.0055329 |  |  | 514.3848 |  |  |  |  | 23.9398 |  |  |  |  | 104.7103 |  |  |  |  | 56.2941 |  |  | -0.0084763 | -0.064969 | -0.056493 | -0.0043176 |  | 53
105 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 238 | 481 |  | 229 | 0.598 | 0.686 | 0.088 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 15:18 |  | 1 |  | -0.030277 | -0.057193 | -0.026916 | -0.0031023
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 485 | 2.946 | 1.25 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-13 | 13:12 |  | 1 |  | 2.7519 | 0.6933 | -2.0586 | 0.026702
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 623 | 1031 |  | 443 | 0.893 | 1.152 | 0.259 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-13 | 13:21 |  | 1 |  | 0.33182 | 0.58823 | 0.25642 | 0.022835 | [20 empty] | 0.28344 | 0.58506 | 0.30163 | 0.022513 |  | 54
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 516 | 1024 |  | 436 | 0.811 | 1.146 | 0.335 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-13 | 13:31 |  | 1 |  | 0.23505 | 0.58189 | 0.34684 | 0.02219
58 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 265 | 1.694 | 0.744 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 1:21 |  | 1 |  | 1.276 | 0.15166 | -1.1243 | 0.006445
59 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 274 | 468 |  | 324 | 0.625 | 0.675 | 0.05 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 1:29 |  | 1 |  | 0.0162 | 0.07829 | 0.06209 | 0.011878 |  |  | 715.4896 |  |  |  |  | 19.0027 |  |  |  |  | 17.2686 |  |  |  |  | 35.6181 |  |  | 0.0035391 | 0.071497 | 0.067958 | 0.010082 |  | 55
60 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 246 | 453 |  | 285 | 0.604 | 0.662 | 0.058 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 1:39 |  | 1 |  | -0.0091218 | 0.064704 | 0.073826 | 0.0082866
100 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 378 | 1.694 | 0.726 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 10:40 |  | 1 |  | 1.276 | 0.13264 | -1.1433 | 0.01685
101 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 315 | 493 |  | 377 | 0.657 | 0.696 | 0.039 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 10:49 |  | 1 |  | 0.053278 | 0.10093 | 0.047656 | 0.016758 |  |  | 175.0845 |  |  |  |  | 32.3464 |  |  |  |  | 96.3994 |  |  |  |  | 6.2325 |  |  | 0.028409 | 0.12041 | 0.092 | 0.016251 |  | 56
102 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 260 | 536 |  | 366 | 0.614 | 0.733 | 0.118 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 10:58 |  | 1 |  | 0.0035391 | 0.13988 | 0.13634 | 0.015745
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 417 | 3.098 | 1.241 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 14:45 |  | 1 |  | 2.8647 | 0.58687 | -2.2778 | 0.042603
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 602 | 1007 |  | 384 | 0.877 | 1.132 | 0.255 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 14:54 |  | 1 |  | 0.28439 | 0.47139 | 0.18699 | 0.03967 | [20 empty] | 0.22691 | 0.47855 | 0.25164 | 0.038114 |  | 57
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 473 | 1023 |  | 349 | 0.778 | 1.145 | 0.368 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 15:04 |  | 1 |  | 0.16942 | 0.48571 | 0.31629 | 0.036559
58 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 232 | 1.757 | 0.78 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 2:57 |  | 1 |  | 1.3076 | 0.099858 | -1.2077 | 0.026159
59 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 340 | 533 |  | 248 | 0.675 | 0.73 | 0.055 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 3:06 |  | 1 |  | 0.050875 | 0.047039 | -0.0038365 | 0.027582 |  |  | 284.1326 |  |  |  |  | 41.3383 |  |  |  |  | 242.7113 |  |  |  |  | 7.0046 |  |  | 0.021017 | 0.038982 | 0.017965 | 0.026648 |  | 58
60 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 273 | 515 |  | 227 | 0.624 | 0.715 | 0.09 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 3:15 |  | 1 |  | -0.0088411 | 0.030925 | 0.039766 | 0.025715
100 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 203 | 2.469 | 0.842 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 12:28 |  | 1 |  | 2.1338 | 0.16521 | -1.9686 | 0.023582
101 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 436 | 610 |  | 227 | 0.749 | 0.795 | 0.046 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 12:38 |  | 1 |  | 0.13644 | 0.11597 | -0.020467 | 0.025715 |  |  | 21.7265 |  |  |  |  | 18.8754 |  |  |  |  | 1020.5516 |  |  |  |  | 11.3219 |  |  | 0.12307 | 0.12806 | 0.0049885 | 0.024337 |  | 59
102 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 406 | 637 |  | 196 | 0.726 | 0.819 | 0.092 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 12:48 |  | 1 |  | 0.1097 | 0.14014 | 0.030444 | 0.022959
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 489 | 3.177 | 1.219 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 14:55 |  | 1 |  | 2.9306 | 0.56098 | -2.3696 | 0.024054
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 602 | 1080 |  | 402 | 0.877 | 1.193 | 0.317 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 15:04 |  | 1 |  | 0.28275 | 0.53445 | 0.2517 | 0.016087 | [20 empty] | 0.23419 | 0.52428 | 0.29009 | 0.019796 |  | 60
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 492 | 1057 |  | 483 | 0.793 | 1.174 | 0.381 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 15:14 |  | 1 |  | 0.18563 | 0.51411 | 0.32848 | 0.023505
58 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 131 | 1.64 | 0.667 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 3:09 |  | 1 |  | 1.1612 | -0.014754 | -1.176 | -0.0087317
59 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 285 | 440 |  | 153 | 0.634 | 0.651 | 0.018 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 3:18 |  | 1 |  | 0.0028668 | -0.031557 | -0.034424 | -0.0067169 |  |  | 236.5797 |  |  |  |  | 56.358 |  |  |  |  | 43.5853 |  |  |  |  | 38.5274 |  |  | -0.015674 | -0.043938 | -0.028264 | -0.0083196 |  | 61
60 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 243 | 412 |  | 118 | 0.602 | 0.627 | 0.026 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 3:28 |  | 1 |  | -0.034215 | -0.05632 | -0.022105 | -0.0099223
100 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 145 | 1.624 | 0.76 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 12:36 |  | 1 |  | 1.1427 | 0.081644 | -1.061 | -0.0074496
101 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 280 | 530 |  | 137 | 0.63 | 0.727 | 0.097 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 12:44 |  | 1 |  | -0.0015477 | 0.048038 | 0.049585 | -0.0081822 |  |  | 183.106 |  |  |  |  | 1.8581 |  |  |  |  | 49.555 |  |  |  |  | 53.3467 |  |  | -0.018323 | 0.047595 | 0.065918 | -0.011159 |  | 62
102 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 242 | 529 |  | 72 | 0.6 | 0.726 | 0.126 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 12:54 |  | 1 |  | -0.035098 | 0.047153 | 0.082251 | -0.014135
7 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.3 | 0.3 |  |  |  | 791 | 4.611 | 1.881 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-24 | 16:45 |  | 1 |  | 4.1219 | 0.79087 | -3.331 | 0.097883
8 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.3 | 0.3 | 747 | 989 |  | 591 | 1.647 | 1.86 | 0.213 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-24 | 16:55 |  | 1 |  | 0.56023 | 0.76831 | 0.20807 | 0.066482 | [20 empty] | 0.44039 | 0.75402 | 0.31364 | 0.063734 |  | 63
9 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.3 | 0.3 | 591 | 970 |  | 556 | 1.447 | 1.833 | 0.387 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-24 | 17:07 |  | 1 |  | 0.32054 | 0.73973 | 0.4192 | 0.060987
61 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 482 | 1.206 | 0.811 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 5:23 |  | 1 |  | 0.59796 | 0.13615 | -0.46181 | 0.029621
62 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 231 | 604 |  | 440 | 0.592 | 0.79 | 0.198 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 5:32 |  | 1 |  | -0.13956 | 0.1136 | 0.25316 | 0.025665 |  |  | 32.1532 |  |  |  |  | 24.0221 |  |  |  |  | 10.8729 |  |  |  |  | 14.5704 |  |  | -0.1663 | 0.10142 | 0.26772 | 0.023922 |  | 64
63 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 173 | 577 |  | 403 | 0.547 | 0.767 | 0.22 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 5:41 |  | 1 |  | -0.19304 | 0.089235 | 0.28227 | 0.022179
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 635 | 3.028 | 1.383 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 16:22 |  | 1 |  | 2.854 | 0.62923 | -2.2248 | 0.028844
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 541 | 1037 |  | 587 | 0.83 | 1.157 | 0.327 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 16:31 |  | 1 |  | 0.21904 | 0.38273 | 0.1637 | 0.023769 | [20 empty] | 0.1864 | 0.39473 | 0.20834 | 0.024139 |  | 65
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 470 | 1063 |  | 594 | 0.775 | 1.179 | 0.404 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 16:42 |  | 1 |  | 0.15376 | 0.40674 | 0.25298 | 0.024509
58 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 354 | 1.494 | 0.657 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 4:37 |  | 1 |  | 1.0152 | -0.16196 | -1.1772 | -0.00086833
59 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 222 | 415 |  | 302 | 0.585 | 0.63 | 0.045 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 4:46 |  | 1 |  | -0.074252 | -0.19151 | -0.11725 | -0.0063668 |  |  | 16.995 |  |  |  |  | 4.4349 |  |  |  |  | 20.8089 |  |  |  |  | 36.629 |  |  | -0.081147 | -0.18735 | -0.1062 | -0.0077942 |  | 66
60 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 207 | 424 |  | 275 | 0.573 | 0.638 | 0.064 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 4:55 |  | 1 |  | -0.088043 | -0.1832 | -0.095154 | -0.0092217
100 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 407 | 1.55 | 0.755 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 14:06 |  | 1 |  | 1.0824 | -0.055794 | -1.1381 | 0.0047358
101 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 293 | 590 |  | 413 | 0.639 | 0.778 | 0.139 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 14:15 |  | 1 |  | -0.0089744 | -0.029944 | -0.020969 | 0.0053703 |  |  | 150.276 |  |  |  |  | 94.4099 |  |  |  |  | 3.3865 |  |  |  |  | 60.4495 |  |  | -0.036097 | -0.056717 | -0.02062 | 0.0076965 |  | 67
102 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 234 | 532 |  | 457 | 0.595 | 0.729 | 0.134 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 14:24 |  | 1 |  | -0.063219 | -0.08349 | -0.020271 | 0.010023
142 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 381 | 1.63 | 0.804 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 23:31 |  | 1 |  | 1.178 | -0.0022473 | -1.1802 | 0.0019866
143 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 323 | 520 |  | 383 | 0.663 | 0.719 | 0.056 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 23:40 |  | 1 |  | 0.018608 | -0.094569 | -0.11318 | 0.0021981 |  |  | 448.7963 |  |  |  |  | 18.0978 |  |  |  |  | 44.8183 |  |  |  |  | 14235.5126 |  |  | 0.0057361 | -0.086721 | -0.092457 | 3.0454e-05 |  | 68
144 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 295 | 537 |  | 342 | 0.641 | 0.733 | 0.092 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-26 | 23:50 |  | 1 |  | -0.0071356 | -0.078874 | -0.071738 | -0.0021372
157 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 738 | 2.525 | 1.346 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 2:50 |  | 1 |  | 2.25 | 0.58861 | -1.6614 | 0.039735
158 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 473 | 979 |  | 611 | 0.778 | 1.108 | 0.33 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 2:59 |  | 1 |  | 0.15652 | 0.32919 | 0.17267 | 0.026307 |  |  | 22.1876 |  |  |  |  | 13.5924 |  |  |  |  | 37.3368 |  |  |  |  | 11.3728 |  |  | 0.14089 | 0.35319 | 0.2123 | 0.027893 |  | 69
159 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 439 | 1031 |  | 641 | 0.751 | 1.152 | 0.401 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 3:09 |  | 1 |  | 0.12526 | 0.37719 | 0.25194 | 0.029479
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 793 | 3.269 | 1.42 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 12:01 |  | 1 |  | 3.5981 | 1.535 | -2.0632 | 0.10217
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1013 | 1242 |  | 577 | 1.192 | 1.331 | 0.139 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 12:10 |  | 1 |  | 1.1217 | 1.4371 | 0.31539 | 0.079822 |  |  | 27.9698 |  |  |  |  | 5.7316 |  |  |  |  | 72.6775 |  |  |  |  | 26.9042 |  |  | 0.98412 | 1.4795 | 0.49542 | 0.070358 |  | 70
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 712 | 1333 |  | 394 | 0.961 | 1.408 | 0.447 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 12:20 |  | 1 |  | 0.84649 | 1.5219 | 0.67545 | 0.060893
22 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 618 | 2.953 | 1.249 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:00 |  | 1 |  | 3.2214 | 1.3477 | -1.8737 | 0.084063
23 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 518 | 1113 |  | 491 | 0.812 | 1.221 | 0.409 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:09 |  | 1 |  | 0.66909 | 1.3169 | 0.64784 | 0.070927 |  |  | 10.0479 |  |  |  |  | 7.0981 |  |  |  |  | 22.0964 |  |  |  |  | 5.8542 |  |  | 0.63708 | 1.3654 | 0.72831 | 0.06891 |  | 71
24 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 448 | 1217 |  | 452 | 0.758 | 1.309 | 0.551 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:19 |  | 1 |  | 0.60507 | 1.4138 | 0.80877 | 0.066893
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 499 | 1.502 | 0.857 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 1:22 |  | 1 |  | 1.4912 | 0.91622 | -0.57498 | 0.071754
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 210 | 629 |  | 382 | 0.576 | 0.811 | 0.235 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 1:31 |  | 1 |  | 0.38743 | 0.86589 | 0.47847 | 0.059652 |  |  | 2.6306 |  |  |  |  | 1.0705 |  |  |  |  | 3.9697 |  |  |  |  | 5.1584 |  |  | 0.3824 | 0.87055 | 0.48816 | 0.058152 |  | 72
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 199 | 639 |  | 353 | 0.568 | 0.82 | 0.252 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 1:40 |  | 1 |  | 0.37737 | 0.87521 | 0.49784 | 0.056652
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 577 | 1.111 | 0.838 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 10:47 |  | 1 |  | 1.0257 | 0.89478 | -0.13095 | 0.079822
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 171 | 588 |  | 474 | 0.546 | 0.776 | 0.231 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 10:56 |  | 1 |  | 0.35176 | 0.82769 | 0.47592 | 0.069168 |  |  | 2.3127 |  |  |  |  | 4.6195 |  |  |  |  | 6.2903 |  |  |  |  | 4.2448 |  |  | 0.35588 | 0.84726 | 0.49138 | 0.070668 |  | 73
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 180 | 630 |  | 503 | 0.553 | 0.812 | 0.259 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 11:05 |  | 1 |  | 0.35999 | 0.86683 | 0.50683 | 0.072168
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 587 | 3.448 | 1.189 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 15:00 |  | 1 |  | 3.7324 | 1.2928 | -2.4396 | 0.068382
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 630 | 1078 |  | 515 | 0.898 | 1.192 | 0.294 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 15:09 |  | 1 |  | 0.71669 | 1.2956 | 0.57888 | 0.061841 |  |  | 10.8078 |  |  |  |  | 4.2146 |  |  |  |  | 20.0875 |  |  |  |  | 2.8955 |  |  | 0.67995 | 1.3235 | 0.64351 | 0.06275 |  | 74
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 549 | 1139 |  | 535 | 0.836 | 1.243 | 0.408 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 15:19 |  | 1 |  | 0.6432 | 1.3514 | 0.70815 | 0.063658
22 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 582 | 2.972 | 1.247 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 19:05 |  | 1 |  | 3.169 | 1.355 | -1.814 | 0.067928
23 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 543 | 1140 |  | 502 | 0.831 | 1.244 | 0.413 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 19:14 |  | 1 |  | 0.63776 | 1.3523 | 0.71451 | 0.06066 |  |  | 10.0081 |  |  |  |  | 1.4302 |  |  |  |  | 5.6553 |  |  |  |  | 2.1189 |  |  | 0.60737 | 1.3427 | 0.7353 | 0.060025 |  | 75
24 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 476 | 1119 |  | 488 | 0.78 | 1.227 | 0.446 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 19:24 |  | 1 |  | 0.57698 | 1.3331 | 0.75609 | 0.059389
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 |  |  |  | 665 | 2.163 | 0.822 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 4:32 |  | 1 |  | 2.2127 | 0.89598 | -1.3168 | 0.075468
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 408 | 552 |  | 508 | 0.728 | 0.746 | 0.018 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 4:41 |  | 1 |  | 0.51528 | 0.81459 | 0.29931 | 0.061206 |  |  | 15.5609 |  |  |  |  | 1.6697 |  |  |  |  | 25.6607 |  |  |  |  | 0.59196 |  |  | 0.47809 | 0.82145 | 0.34337 | 0.061387 |  | 76
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 326 | 567 |  | 512 | 0.665 | 0.759 | 0.094 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 4:50 |  | 1 |  | 0.44089 | 0.82831 | 0.38742 | 0.061569
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1402 | 563 |  | 413 | 1.49 | 0.756 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 14:01 |  | 1 |  | 1.4487 | 0.81811 | -0.63055 | 0.053707
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 274 | 560 |  | 487 | 0.625 | 0.753 | 0.128 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 14:10 |  | 1 |  | 0.419 | 0.81537 | 0.39637 | 0.060432 |  |  | 4.9105 |  |  |  |  | 1.5566 |  |  |  |  | 7.9634 |  |  |  |  | 1.821 |  |  | 0.40896 | 0.82176 | 0.4128 | 0.059887 |  | 77
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 252 | 574 |  | 475 | 0.608 | 0.765 | 0.156 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-17 | 14:19 |  | 1 |  | 0.39892 | 0.82816 | 0.42924 | 0.059342
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3505 | 1099 |  | 647 | 3.104 | 1.21 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 12:49 |  | 1 |  | 3.4555 | 1.3228 | -2.1327 | 0.10002
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 723 | 1142 |  | 564 | 0.969 | 1.246 | 0.276 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 12:58 |  | 1 |  | 0.83262 | 1.3624 | 0.52974 | 0.09284 |  |  | 5.7067 |  |  |  |  | 0.53881 |  |  |  |  | 9.6235 |  |  |  |  | 2.2611 |  |  | 0.80952 | 1.366 | 0.55652 | 0.091802 |  | 78
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 674 | 1150 |  | 540 | 0.932 | 1.252 | 0.32 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 13:08 |  | 1 |  | 0.78642 | 1.3697 | 0.5833 | 0.090764
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1749 | 491 |  | 565 | 1.757 | 0.695 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 2:17 |  | 1 |  | 1.7999 | 0.76341 | -1.0365 | 0.092926
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 278 | 534 |  | 496 | 0.628 | 0.731 | 0.103 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 2:26 |  | 1 |  | 0.41307 | 0.80297 | 0.3899 | 0.086958 |  |  | 1.6105 |  |  |  |  | 1.0366 |  |  |  |  | 0.43204 |  |  |  |  | 4.0592 |  |  | 0.40977 | 0.79883 | 0.38906 | 0.085228 |  | 79
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 271 | 525 |  | 456 | 0.623 | 0.724 | 0.101 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 2:35 |  | 1 |  | 0.40648 | 0.79469 | 0.38822 | 0.083499
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1722 | 535 |  | 788 | 1.736 | 0.732 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 11:42 |  | 1 |  | 1.7745 | 0.80389 | -0.97057 | 0.11221
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 295 | 515 |  | 789 | 0.641 | 0.714 | 0.073 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 11:51 |  | 1 |  | 0.4291 | 0.78549 | 0.35639 | 0.1123 |  |  | 0.875 |  |  |  |  | 3.792 |  |  |  |  | 7.1927 |  |  |  |  | 0.61804 |  |  | 0.43099 | 0.80067 | 0.36968 | 0.11195 |  | 80
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 299 | 548 |  | 781 | 0.644 | 0.743 | 0.099 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-19 | 12:00 |  | 1 |  | 0.43287 | 0.81585 | 0.38298 | 0.11161
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3519 | 2119 |  | 1165 | 3.114 | 2.074 | 0 |  | 0.006 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 16:16 |  | 1 |  | 3.5034 | 2.1457 | -1.3577 | 0.12269
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 837 | 3048 |  | 858 | 1.057 | 2.861 | 1.804 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 16:27 |  | 1 |  | 0.83884 | 3.1091 | 2.2703 | 0.092299 |  |  | 6.9869 |  |  |  |  | 50.6691 |  |  |  |  | 71.8681 |  |  |  |  | 51.7018 |  |  | 0.81053 | 2.4807 | 1.6702 | 0.07334 |  | 81
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 780 | 1836 |  | 475 | 1.013 | 1.834 | 0.82 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 16:38 |  | 1 |  | 0.78221 | 1.8522 | 1.07 | 0.054381
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1623 | 587 |  | 422 | 1.66 | 0.776 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:01 |  | 1 |  | 1.6197 | 0.5569 | -1.0628 | 0.049134
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 326 | 530 |  | 321 | 0.665 | 0.727 | 0.062 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:10 |  | 1 |  | 0.33116 | 0.49779 | 0.16663 | 0.039135 |  |  | 3.8256 |  |  |  |  | 6.0607 |  |  |  |  | 10.3549 |  |  |  |  | 3.3437 |  |  | 0.33761 | 0.51334 | 0.17573 | 0.038491 |  | 82
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 339 | 560 |  | 308 | 0.675 | 0.753 | 0.078 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 6:19 |  | 1 |  | 0.34407 | 0.5289 | 0.18483 | 0.037847
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1598 | 502 |  | 500 | 1.641 | 0.704 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:17 |  | 1 |  | 1.5949 | 0.46875 | -1.1262 | 0.056856
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 323 | 554 |  | 472 | 0.663 | 0.748 | 0.086 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:26 |  | 1 |  | 0.32817 | 0.52268 | 0.1945 | 0.054084 |  |  | 1.8001 |  |  |  |  | 6.1539 |  |  |  |  | 13.0819 |  |  |  |  | 3.5953 |  |  | 0.33116 | 0.53927 | 0.20812 | 0.055074 |  | 83
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 329 | 586 |  | 492 | 0.667 | 0.775 | 0.108 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 15:35 |  | 1 |  | 0.33414 | 0.55586 | 0.22173 | 0.056064
4 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2997 | 998 |  | 561 | 2.714 | 1.124 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 16:43 |  | 1 |  | 3.0889 | 1.5919 | -1.497 | 0.076768
5 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 519 | 996 |  | 556 | 0.813 | 1.122 | 0.309 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 16:53 |  | 1 |  | 0.66075 | 1.59 | 0.92922 | 0.076251 |  |  | 5.3287 |  |  |  |  | 0.67241 |  |  |  |  | 4.7306 |  |  |  |  | 8.3346 |  |  | 0.64361 | 1.5953 | 0.95174 | 0.073201 |  | 84
6 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 484 | 1007 |  | 497 | 0.787 | 1.132 | 0.345 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 17:02 |  | 1 |  | 0.62646 | 1.6007 | 0.97425 | 0.07015
64 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1474 | 495 |  | 354 | 1.546 | 0.698 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:09 |  | 1 |  | 1.5965 | 1.1014 | -0.49514 | 0.055363
65 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 322 | 480 |  | 240 | 0.662 | 0.685 | 0.023 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:18 |  | 1 |  | 0.46772 | 1.0868 | 0.61906 | 0.043575 |  |  | 1.6902 |  |  |  |  | 4.0443 |  |  |  |  | 8.1651 |  |  |  |  | 9.2783 |  |  | 0.4638 | 1.1092 | 0.64541 | 0.045695 |  | 85
66 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 314 | 526 |  | 281 | 0.656 | 0.724 | 0.068 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 6:27 |  | 1 |  | 0.45988 | 1.1316 | 0.67176 | 0.047815
106 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2025 | 591 |  | 312 | 1.968 | 0.78 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 15:33 |  | 1 |  | 2.1365 | 1.195 | -0.94144 | 0.05102
107 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 439 | 606 |  | 307 | 0.751 | 0.792 | 0.04 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 15:42 |  | 1 |  | 0.58236 | 1.2096 | 0.62729 | 0.050503 |  |  | 12.8959 |  |  |  |  | 4.0287 |  |  |  |  | 17.4981 |  |  |  |  | 2.0687 |  |  | 0.54709 | 1.2345 | 0.68743 | 0.049986 |  | 86
108 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 367 | 657 |  | 297 | 0.697 | 0.835 | 0.138 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-09 | 15:52 |  | 1 |  | 0.51181 | 1.2594 | 0.74757 | 0.049469
7 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 274 | 521 |  | 385 | 0.625 | 0.72 | 0.095 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 13:00 |  | 1 |  | 0.2022 | 0.58534 | 0.38314 | 0.054966
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 178 | 433 |  | 381 | 0.551 | 0.645 | 0.094 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 13:09 |  | 1 |  | 0.10794 | 0.49888 | 0.39094 | 0.054575 |  |  | 112.172 |  |  |  |  | 2.3356 |  |  |  |  | 20.5132 |  |  |  |  | 2.3544 |  |  | 0.069154 | 0.50478 | 0.43562 | 0.05394 |  | 87
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 99 | 445 |  | 368 | 0.491 | 0.656 | 0.165 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 13:18 |  | 1 |  | 0.030368 | 0.51067 | 0.4803 | 0.053305
67 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1585 | 721 |  | 290 | 1.631 | 0.889 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 2:40 |  | 1 |  | 1.4895 | 0.78183 | -0.70765 | 0.045685
68 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 349 | 683 |  | 292 | 0.682 | 0.857 | 0.175 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 2:49 |  | 1 |  | 0.27585 | 0.7445 | 0.46865 | 0.045881 |  |  | 1.4138 |  |  |  |  | 6.9645 |  |  |  |  | 12.2348 |  |  |  |  | 1.5016 |  |  | 0.27781 | 0.71945 | 0.44164 | 0.045539 |  | 88
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 353 | 632 |  | 285 | 0.686 | 0.814 | 0.128 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 2:58 |  | 1 |  | 0.27977 | 0.69439 | 0.41462 | 0.045197
109 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2746 | 738 |  | 240 | 2.522 | 0.904 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 12:06 |  | 1 |  | 2.6295 | 0.79854 | -1.831 | 0.040801
110 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 506 | 672 |  | 219 | 0.803 | 0.848 | 0.044 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 12:15 |  | 1 |  | 0.43001 | 0.73369 | 0.30369 | 0.038749 |  |  | 10.3259 |  |  |  |  | 6.9787 |  |  |  |  | 27.1189 |  |  |  |  | 1.7804 |  |  | 0.40889 | 0.76022 | 0.35133 | 0.038407 |  | 89
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 463 | 726 |  | 212 | 0.77 | 0.893 | 0.123 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 12:25 |  | 1 |  | 0.38778 | 0.78675 | 0.39896 | 0.038066
7 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 252 | 243 |  | 240 | 0.608 | 0.485 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 14:44 |  | 1 |  | 0.27361 | 0.37945 | 0.10584 | 0.028945
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 100 | 244 |  | 289 | 0.491 | 0.485 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 14:52 |  | 1 |  | 0.12495 | 0.38041 | 0.25545 | 0.034165 |  |  | 16.9835 |  |  |  |  | 7.0256 |  |  |  |  | 2.4804 |  |  |  |  | 11.5426 |  |  | 0.11517 | 0.3675 | 0.25232 | 0.032301 |  | 90
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 80 | 217 |  | 254 | 0.476 | 0.462 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 15:01 |  | 1 |  | 0.10539 | 0.35459 | 0.2492 | 0.030437
67 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1456 | 538 |  | 237 | 1.532 | 0.734 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 4:22 |  | 1 |  | 1.4512 | 0.66155 | -0.78961 | 0.028626
68 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 324 | 478 |  | 206 | 0.663 | 0.683 | 0.02 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 4:31 |  | 1 |  | 0.34403 | 0.60417 | 0.26014 | 0.025324 |  |  | 22.0986 |  |  |  |  | 2.9627 |  |  |  |  | 28.5479 |  |  |  |  | 16.8776 |  |  | 0.3098 | 0.61326 | 0.30346 | 0.023353 |  | 91
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 254 | 497 |  | 169 | 0.609 | 0.699 | 0.09 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 4:41 |  | 1 |  | 0.27557 | 0.62234 | 0.34677 | 0.021382
109 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1414 | 457 |  | 292 | 1.5 | 0.665 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 13:43 |  | 1 |  | 1.4101 | 0.58409 | -0.82599 | 0.034485
110 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 266 | 463 |  | 258 | 0.619 | 0.671 | 0.052 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 13:52 |  | 1 |  | 0.28731 | 0.58983 | 0.30252 | 0.030863 |  |  | 2.0635 |  |  |  |  | 5.4972 |  |  |  |  | 8.8679 |  |  |  |  | 10.92 |  |  | 0.28437 | 0.57405 | 0.28968 | 0.029265 |  | 92
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 260 | 430 |  | 228 | 0.614 | 0.643 | 0.029 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 14:01 |  | 1 |  | 0.28144 | 0.55827 | 0.27684 | 0.027667
115 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2415 | 872 |  | 423 | 2.268 | 1.017 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 14:59 |  | 1 |  | 2.6871 | 1.1861 | -1.5011 | 0.066826
116 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 434 | 837 |  | 386 | 0.748 | 0.988 | 0.239 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 15:08 |  | 1 |  | 0.68097 | 1.1511 | 0.47013 | 0.062897 |  |  | 3.075 |  |  |  |  | 1.662 |  |  |  |  | 8.9421 |  |  |  |  | 5.5521 |  |  | 0.69161 | 1.1416 | 0.45001 | 0.061198 |  | 93
117 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 455 | 818 |  | 354 | 0.764 | 0.972 | 0.208 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 15:18 |  | 1 |  | 0.70224 | 1.1321 | 0.42989 | 0.059499
175 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1474 | 535 |  | 319 | 1.545 | 0.732 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 4:36 |  | 1 |  | 1.7342 | 0.84952 | -0.88466 | 0.055783
176 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 326 | 509 |  | 307 | 0.665 | 0.71 | 0.045 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 4:45 |  | 1 |  | 0.5716 | 0.82356 | 0.25196 | 0.054509 |  |  | 12.6193 |  |  |  |  | 3.705 |  |  |  |  | 13.9876 |  |  |  |  | 3.9733 |  |  | 0.53768 | 0.80858 | 0.2709 | 0.053447 |  | 94
177 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 259 | 479 |  | 287 | 0.614 | 0.684 | 0.07 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 4:54 |  | 1 |  | 0.50375 | 0.7936 | 0.28985 | 0.052385
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 63 | 162 |  | 165 | 0.463 | 0.416 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 15:44 |  | 1 |  | 0.087136 | 0.11383 | 0.026699 | 0.014978
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 63 | 200 |  | 182 | 0.463 | 0.448 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 15:53 |  | 1 |  | 0.087136 | 0.15046 | 0.063323 | 0.016644 |  |  | 22.2649 |  |  |  |  | 7.994 |  |  |  |  | 71.6282 |  |  |  |  | 5.7195 |  |  | 0.098051 | 0.14468 | 0.046625 | 0.017134 |  | 95
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 86 | 188 |  | 192 | 0.481 | 0.438 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 16:02 |  | 1 |  | 0.10897 | 0.13889 | 0.029926 | 0.017624
68 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1916 | 455 |  | 231 | 1.885 | 0.664 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 5:17 |  | 1 |  | 1.846 | 0.39622 | -1.4497 | 0.021446
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 303 | 555 |  | 206 | 0.647 | 0.748 | 0.101 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 5:27 |  | 1 |  | 0.31494 | 0.4926 | 0.17767 | 0.018996 |  |  | 2.9691 |  |  |  |  | 15.1543 |  |  |  |  | 57.0702 |  |  |  |  | 7.4931 |  |  | 0.31968 | 0.45791 | 0.13822 | 0.01831 |  | 96
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 313 | 483 |  | 192 | 0.655 | 0.688 | 0.033 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 5:36 |  | 1 |  | 0.32443 | 0.42321 | 0.098781 | 0.017624
110 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1966 | 451 |  | 240 | 1.923 | 0.661 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 14:35 |  | 1 |  | 1.8934 | 0.39237 | -1.501 | 0.022328
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 315 | 488 |  | 249 | 0.657 | 0.692 | 0.035 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 14:44 |  | 1 |  | 0.32633 | 0.42803 | 0.1017 | 0.02321 |  |  | 5.3763 |  |  |  |  | 1.1322 |  |  |  |  | 11.375 |  |  |  |  | 4.3133 |  |  | 0.31778 | 0.42562 | 0.10783 | 0.02272 |  | 97
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 297 | 483 |  | 239 | 0.643 | 0.687 | 0.045 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-24 | 14:53 |  | 1 |  | 0.30924 | 0.42321 | 0.11397 | 0.02223
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 62 | 158 |  | 176 | 0.462 | 0.412 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 15:09 |  | 1 |  | 0.203 | 0.28438 | 0.081376 | 0.026259
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 125 | 164 |  | 193 | 0.511 | 0.417 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 15:17 |  | 1 |  | 0.26349 | 0.2903 | 0.026812 | 0.028064 |  |  | 20.4874 |  |  |  |  | 6.5778 |  |  |  |  | 112.332 |  |  |  |  | 1.5245 |  |  | 0.239 | 0.30017 | 0.061167 | 0.027851 |  | 98
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 74 | 184 |  | 189 | 0.471 | 0.434 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 15:26 |  | 1 |  | 0.21452 | 0.31004 | 0.095522 | 0.027639
68 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1083 | 666 |  | 322 | 1.246 | 0.843 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 4:37 |  | 1 |  | 1.1833 | 0.78589 | -0.39739 | 0.041757
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 257 | 605 |  | 233 | 0.612 | 0.791 | 0.179 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 4:46 |  | 1 |  | 0.39022 | 0.72567 | 0.33544 | 0.03231 |  |  | 5.8238 |  |  |  |  | 3.4757 |  |  |  |  | 13.2891 |  |  |  |  | 2.2736 |  |  | 0.37918 | 0.7385 | 0.35932 | 0.032681 |  | 99
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 234 | 631 |  | 240 | 0.595 | 0.813 | 0.218 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 4:56 |  | 1 |  | 0.36814 | 0.75133 | 0.3832 | 0.033053
110 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2004 | 538 |  | 392 | 1.952 | 0.734 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 13:58 |  | 1 |  | 2.0675 | 0.65952 | -1.408 | 0.049187
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 402 | 548 |  | 311 | 0.723 | 0.743 | 0.02 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 14:07 |  | 1 |  | 0.52944 | 0.66939 | 0.13996 | 0.040589 |  |  | 9.8966 |  |  |  |  | 1.6092 |  |  |  |  | 35.6829 |  |  |  |  | 26.5375 |  |  | 0.50448 | 0.67482 | 0.17035 | 0.046799 |  | 100
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 350 | 559 |  | 428 | 0.683 | 0.752 | 0.069 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-27 | 14:16 |  | 1 |  | 0.47951 | 0.68025 | 0.20074 | 0.053008
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 92 | 233 |  | 122 | 0.485 | 0.476 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 11:09 |  | 1 |  | 0.062312 | 0.19133 | 0.12902 | 0.022729
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 120 | 212 |  | 160 | 0.507 | 0.458 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 11:18 |  | 1 |  | 0.089378 | 0.17 | 0.080623 | 0.026821 |  |  | 8.2933 |  |  |  |  | 8.7316 |  |  |  |  | 31.5254 |  |  |  |  | 10.5689 |  |  | 0.093244 | 0.16289 | 0.069645 | 0.025475 |  | 101
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 128 | 198 |  | 135 | 0.513 | 0.446 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 11:27 |  | 1 |  | 0.097111 | 0.15578 | 0.058667 | 0.024129
68 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1952 | 451 |  | 189 | 1.912 | 0.661 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 0:26 |  | 1 |  | 1.8602 | 0.4128 | -1.4474 | 0.029944
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 355 | 448 |  | 166 | 0.687 | 0.658 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 0:35 |  | 1 |  | 0.31654 | 0.40976 | 0.093221 | 0.027467 |  |  | 24.3181 |  |  |  |  | 8.7712 |  |  |  |  | 72.5886 |  |  |  |  | 2.3805 |  |  | 0.28222 | 0.42855 | 0.14633 | 0.027144 |  | 102
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 284 | 485 |  | 160 | 0.633 | 0.689 | 0.057 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 0:44 |  | 1 |  | 0.2479 | 0.44735 | 0.19944 | 0.026821
110 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1682 | 462 |  | 266 | 1.705 | 0.67 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 9:41 |  | 1 |  | 1.5993 | 0.42398 | -1.1753 | 0.038237
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 408 | 405 |  | 253 | 0.728 | 0.622 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 9:50 |  | 1 |  | 0.36777 | 0.36607 | -0.0016946 | 0.036837 |  |  | 41.9763 |  |  |  |  | 14.6591 |  |  |  |  | 203.7221 |  |  |  |  | 19.2845 |  |  | 0.30397 | 0.39503 | 0.091057 | 0.040768 |  | 103
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 276 | 462 |  | 326 | 0.627 | 0.67 | 0.044 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-07 | 9:59 |  | 1 |  | 0.24017 | 0.42398 | 0.18381 | 0.044699
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 169 | 373 |  | 131 | 0.545 | 0.594 | 0.049 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 15:49 |  | 1 |  | 0.29583 | 0.52694 | 0.23111 | 0.025127
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 91 | 333 |  | 127 | 0.485 | 0.561 | 0.076 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 15:57 |  | 1 |  | 0.21974 | 0.48818 | 0.26844 | 0.02472 |  |  | 1.323 |  |  |  |  | 11.1034 |  |  |  |  | 22.4948 |  |  |  |  | 13.0876 |  |  | 0.2212 | 0.4625 | 0.2413 | 0.026451 |  | 104
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 94 | 280 |  | 161 | 0.487 | 0.516 | 0.029 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 16:06 |  | 1 |  | 0.22267 | 0.43683 | 0.21416 | 0.028182
68 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1464 | 572 |  | 328 | 1.538 | 0.763 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 5:01 |  | 1 |  | 1.5591 | 0.71975 | -0.83934 | 0.045186
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 273 | 538 |  | 293 | 0.624 | 0.734 | 0.11 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 5:10 |  | 1 |  | 0.39728 | 0.68681 | 0.28953 | 0.041622 |  |  | 9.5175 |  |  |  |  | 0.56272 |  |  |  |  | 12.9135 |  |  |  |  | 9.7489 |  |  | 0.37923 | 0.68875 | 0.30951 | 0.039687 |  | 105
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 236 | 542 |  | 255 | 0.596 | 0.738 | 0.142 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 5:20 |  | 1 |  | 0.36119 | 0.69069 | 0.3295 | 0.037753
110 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1273 | 453 |  | 181 | 1.391 | 0.662 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 14:23 |  | 1 |  | 1.3728 | 0.60445 | -0.76833 | 0.030218
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 209 | 435 |  | 223 | 0.575 | 0.647 | 0.071 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 14:32 |  | 1 |  | 0.33485 | 0.58701 | 0.25216 | 0.034495 |  |  | 4.7724 |  |  |  |  | 1.1621 |  |  |  |  | 3.4397 |  |  |  |  | 3.4804 |  |  | 0.32704 | 0.58362 | 0.25657 | 0.035106 |  | 106
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 193 | 428 |  | 235 | 0.563 | 0.641 | 0.078 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 14:41 |  | 1 |  | 0.31924 | 0.58023 | 0.26099 | 0.035716
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 68 | 207 |  | 137 | 0.467 | 0.453 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 17:14 |  | 1 |  | 0.20432 | 0.3079 | 0.10358 | 0.024661
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 90 | 191 |  | 130 | 0.484 | 0.44 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 17:22 |  | 1 |  | 0.22619 | 0.29213 | 0.065946 | 0.023938 |  |  | 6.3461 |  |  |  |  | 8.0931 |  |  |  |  | 45.2374 |  |  |  |  | 2.1326 |  |  | 0.21923 | 0.30445 | 0.085222 | 0.024196 |  | 107
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 76 | 216 |  | 135 | 0.474 | 0.461 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 17:31 |  | 1 |  | 0.21227 | 0.31677 | 0.1045 | 0.024454
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 535 | 97 |  | 129 | 0.826 | 0.361 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 6:29 |  | 1 |  | 0.66841 | 0.19949 | -0.46892 | 0.023835
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 365 | 142 |  | 62 | 0.695 | 0.399 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 6:37 |  | 1 |  | 0.49947 | 0.24384 | -0.25563 | 0.016921 |  |  | 8.2883 |  |  |  |  | 157.9567 |  |  |  |  | 275.1474 |  |  |  |  | 155.7739 |  |  | 0.47959 | 1.1599 | 0.68035 | 0.076519 |  | 108
71 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 325 | 2001 |  | 1217 | 0.664 | 1.974 | 1.31 |  | 0.011 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 6:46 |  | 1 |  | 0.45972 | 2.076 | 1.6163 | 0.13612
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 898 | 153 |  | 125 | 1.104 | 0.408 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 15:43 |  | 1 |  | 1.0291 | 0.25468 | -0.77446 | 0.023422
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 673 | 187 |  | 28 | 0.931 | 0.437 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 15:51 |  | 1 |  | 0.80555 | 0.28819 | -0.51736 | 0.013412 |  |  | 40.1418 |  |  |  |  | 174.5059 |  |  |  |  | 265.0786 |  |  |  |  | 173.2583 |  |  | 0.67089 | 2.2608 | 1.5899 | 0.10031 |  | 109
114 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 402 | 4190 |  | 1712 | 0.724 | 3.829 | 3.105 |  | 0.063 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-10 | 16:00 |  | 1 |  | 0.53624 | 4.2335 | 3.6972 | 0.1872
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 469 | 274 |  | 217 | 0.775 | 0.511 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 14:00 |  | 1 |  | 0.36056 | 0.38049 | 0.019937 | 0.0091964
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 38 | 260 |  | 186 | 0.444 | 0.499 | 0.055 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 14:08 |  | 1 |  | -0.020946 | 0.36737 | 0.38831 | 0.0061068 |  |  | 63.4681 |  |  |  |  | 9.4823 |  |  |  |  | 13.4608 |  |  |  |  | 8.507 |  |  | -0.030682 | 0.38565 | 0.41633 | 0.0058577 |  | 110
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 16 | 299 |  | 181 | 0.427 | 0.532 | 0.104 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 14:17 |  | 1 |  | -0.040419 | 0.40394 | 0.44436 | 0.0056085
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 869 | 966 |  | 339 | 1.082 | 1.097 | 0.015 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 4:03 |  | 1 |  | 0.71462 | 1.0294 | 0.31473 | 0.021355
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 237 | 453 |  | 334 | 0.597 | 0.662 | 0.065 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 4:12 |  | 1 |  | 0.1552 | 0.54834 | 0.39313 | 0.020857 |  |  | 97.2646 |  |  |  |  | 0.51169 |  |  |  |  | 117.5504 |  |  |  |  | 17.1184 |  |  | 0.30214 | 0.54974 | 0.2476 | 0.019212 |  | 111
71 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 569 | 456 |  | 301 | 0.852 | 0.665 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 4:22 |  | 1 |  | 0.44907 | 0.55115 | 0.10207 | 0.017568
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1179 | 547 |  | 243 | 1.32 | 0.742 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 13:54 |  | 1 |  | 0.98902 | 0.63647 | -0.35255 | 0.011788
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 303 | 329 |  | 237 | 0.647 | 0.558 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 14:02 |  | 1 |  | 0.21362 | 0.43207 | 0.21844 | 0.01119 |  |  | 54.9285 |  |  |  |  | 28.7988 |  |  |  |  | 7.9544 |  |  |  |  | 35.6861 |  |  | 0.16759 | 0.37768 | 0.21009 | 0.0094954 |  | 112
114 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 199 | 213 |  | 203 | 0.567 | 0.459 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 14:11 |  | 1 |  | 0.12156 | 0.3233 | 0.20173 | 0.0078011
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 51 | 216 |  | 209 | 0.454 | 0.461 | 0.008 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 16:31 |  | 1 |  | 0.21422 | 0.37285 | 0.15864 | 0.020727
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 57 | 197 |  | 89 | 0.459 | 0.446 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 16:39 |  | 1 |  | 0.21931 | 0.35507 | 0.13575 | 0.009049 |  |  | 16.0397 |  |  |  |  | 5.4148 |  |  |  |  | 53.1032 |  |  |  |  | 22.0126 |  |  | 0.23844 | 0.34571 | 0.10727 | 0.010168 |  | 113
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 102 | 177 |  | 112 | 0.493 | 0.429 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 16:48 |  | 1 |  | 0.25756 | 0.33635 | 0.078788 | 0.011287
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2318 | 591 |  | 580 | 2.194 | 0.779 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 6:10 |  | 1 |  | 2.1409 | 0.72384 | -1.4171 | 0.056831
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 639 | 587 |  | 527 | 0.905 | 0.776 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 6:19 |  | 1 |  | 0.71395 | 0.72009 | 0.0061481 | 0.051673 |  |  | 29.0368 |  |  |  |  | 29.3654 |  |  |  |  | 75.2171 |  |  |  |  | 7.828 |  |  | 0.62343 | 0.6279 | 0.0044678 | 0.049727 |  | 114
71 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 426 | 390 |  | 487 | 0.742 | 0.609 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 6:28 |  | 1 |  | 0.53292 | 0.53571 | 0.0027875 | 0.047781
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 34 | 0 |  | 112 | 0.441 | 0 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  | TCnp | 2022-11-18 | 16:04 |  | 1 |  | 0.19977 | 0.17068 | -0.029086 | 0.011287
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 219 | 1159 |  | 496 | 0.583 | 1.26 | 0.677 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 16:13 |  | 1 |  | 0.357 | 1.2555 | 0.89847 | 0.048657 |  |  | 79.9082 |  |  |  |  | 12.5163 |  |  |  |  | 106.1294 |  |  |  |  | 2.8398 |  |  | 0.59454 | 1.1815 | 0.58699 | 0.047975 |  | 115
114 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 778 | 1001 |  | 482 | 1.012 | 1.126 | 0.114 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-18 | 16:23 |  | 1 |  | 0.83208 | 1.1076 | 0.2755 | 0.047294
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 93 | 187 |  | 276 | 0.486 | 0.437 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 13:06 |  | 1 |  | 0.092859 | 0.27946 | 0.1866 | 0.030415
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 78 | 167 |  | 319 | 0.475 | 0.42 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 13:14 |  | 1 |  | 0.078486 | 0.26031 | 0.18183 | 0.034821 |  |  | 0 |  |  |  |  | 2.8995 |  |  |  |  | 4.1252 |  |  |  |  | 22.6006 |  |  | 0.078486 | 0.26414 | 0.18566 | 0.031286 |  | 116
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 78 | 175 |  | 250 | 0.475 | 0.427 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 13:23 |  | 1 |  | 0.078486 | 0.26797 | 0.18949 | 0.02775
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2741 | 744 |  | 156 | 2.518 | 0.909 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 2:46 |  | 1 |  | 2.6301 | 0.8127 | -1.8174 | 0.018118
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 531 | 729 |  | 151 | 0.822 | 0.896 | 0.074 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 2:55 |  | 1 |  | 0.51254 | 0.79834 | 0.2858 | 0.017605 |  |  | 5.3752 |  |  |  |  | 1.5469 |  |  |  |  | 12.8584 |  |  |  |  | 20.8549 |  |  | 0.49913 | 0.80457 | 0.30544 | 0.019655 |  | 117
71 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 503 | 742 |  | 191 | 0.801 | 0.907 | 0.106 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 3:05 |  | 1 |  | 0.48571 | 0.81079 | 0.32508 | 0.021704
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2168 | 701 |  | 580 | 2.078 | 0.872 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 12:25 |  | 1 |  | 2.0811 | 0.77154 | -1.3095 | 0.061567
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 466 | 619 |  | 544 | 0.773 | 0.803 | 0.03 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 12:34 |  | 1 |  | 0.45026 | 0.69304 | 0.24277 | 0.057878 |  |  | 19.6302 |  |  |  |  | 1.9528 |  |  |  |  | 24.2781 |  |  |  |  | 16.897 |  |  | 0.41002 | 0.68633 | 0.27632 | 0.053369 |  | 118
114 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 382 | 605 |  | 456 | 0.708 | 0.791 | 0.083 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 12:44 |  | 1 |  | 0.36977 | 0.67963 | 0.30986 | 0.04886
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 106 | 214 |  | 253 | 0.496 | 0.46 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 17:32 |  | 1 |  | 0.1507 | 0.28582 | 0.13512 | 0.038105
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 74 | 210 |  | 254 | 0.472 | 0.456 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 17:40 |  | 1 |  | 0.12104 | 0.28204 | 0.161 | 0.038204 |  |  | 5.9453 |  |  |  |  | 1.6889 |  |  |  |  | 7.8357 |  |  |  |  | 4.8087 |  |  | 0.12474 | 0.27968 | 0.15493 | 0.037307 |  | 119
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 82 | 205 |  | 236 | 0.477 | 0.453 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 17:49 |  | 1 |  | 0.12845 | 0.27732 | 0.14886 | 0.03641
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2703 | 794 |  | 227 | 2.489 | 0.951 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 7:29 |  | 1 |  | 2.5583 | 0.83375 | -1.7245 | 0.035513
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 561 | 765 |  | 272 | 0.845 | 0.926 | 0.081 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 7:38 |  | 1 |  | 0.57251 | 0.80636 | 0.23384 | 0.039998 |  |  | 5.8327 |  |  |  |  | 2.2015 |  |  |  |  | 19.4549 |  |  |  |  | 1.2537 |  |  | 0.55629 | 0.81533 | 0.25904 | 0.039749 |  | 120
71 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 526 | 784 |  | 267 | 0.818 | 0.943 | 0.124 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 7:48 |  | 1 |  | 0.54006 | 0.8243 | 0.28424 | 0.0395
82 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 137 | 0 |  | 137 | 0.52 | 0 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  | TCnp | 2022-11-23 | 10:10 |  | 1 |  | 0.17944 | 0.083649 | -0.095792 | 0.026543
83 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 68 | 173 |  | 199 | 0.467 | 0.425 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 10:19 |  | 1 |  | 0.11547 | 0.24708 | 0.13161 | 0.032723 |  |  | 125.2023 |  |  |  |  | 133.9572 |  |  |  |  | 140.108 |  |  |  |  | 54.6676 |  |  | 0.30876 | 0.74826 | 0.43949 | 0.045031 |  | 121
84 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 485 | 1234 |  | 446 | 0.787 | 1.324 | 0.537 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-23 | 10:28 |  | 1 |  | 0.50205 | 1.2494 | 0.74737 | 0.05734
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | Stored Cal TIC [mg/l] | Stored Cal TC [mg/l] | Stored Cal TOC (Diff.) [mg/l] | Stored Cal NPOC [mg/l] | Stored Cal TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] |  | TIC Absolute value Percent error for check standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes |  | TC Absolute value Percent error for check standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes |  | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes |  | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
[52 empty] | Known | NA | NA | NA | NA
[5 empty] | Mean |  |  | 437.6098 | 527.8885 |  | 302.6088 | 0.93308 | 0.73217 | 0.085377 |  | 0.00022039 | [10 empty] | Mean |  | 0.60076 | 0.3921 | -0.20866 | [20 empty] | Mean | 0.20089 | 0.39127 | 0.19038 | 0.026378 | Min | 1
[5 empty] | Std |  |  | 587.0978 | 420.0444 |  | 200.8533 | 0.68487 | 0.36025 | 0.23002 |  | 0.0033688 | [10 empty] | Std |  | 0.85808 | 0.51496 | 0.80099 | [20 empty] | Std | 0.24357 | 0.50098 | 0.29544 | 0.026079 | Max | 121
[5 empty] | CV |  |  | 134.1601 | 79.5707 |  | 66.3739 | 73.3981 | 49.2035 | 269.4145 |  | 1528.6114 | [10 empty] | CV |  | 142.8323 | 131.3325 | -383.8744 | [20 empty] | CV | 121.2473 | 128.0381 | 155.1814 | 98.8685
|  |  |  | Warning Limits | Xbar + 2SD |  |  | 1611.8054 | 1367.9772 |  | 704.3154 | 2.3028 | 1.4527 | 0.54542 |  | 0.0069581 | [9 empty] | Warning Limits | Xbar + 2SD |  | 2.3169 | 1.422 | 1.3933 | [19 empty] | Warning Limits | Xbar + 2SD | 0.68804 | 1.3932 | 0.78126 | 0.078536
[5 empty] | Xbar - 2SD |  |  | -736.5859 | -312.2002 |  | -99.0977 | -0.43665 | 0.011664 | -0.37466 |  | -0.0065173 | [10 empty] | Xbar - 2SD |  | -1.1154 | -0.63781 | -1.8106 | [20 empty] | Xbar - 2SD | -0.28626 | -0.61068 | -0.40049 | -0.025781
|  |  |  | Control Limits | Xbar + 3SD |  |  | 2198.9032 | 1788.0216 |  | 905.1687 | 2.9877 | 1.8129 | 0.77543 |  | 0.010327 | [9 empty] | Control Limits | Xbar + 3SD |  | 3.175 | 1.937 | 2.1943 | [19 empty] | Control Limits | Xbar + 3SD | 0.93161 | 1.8942 | 1.0767 | 0.10462
[5 empty] | Xbar - 3SD |  |  | -1323.6837 | -732.2446 |  | -299.951 | -1.1215 | -0.34859 | -0.60468 |  | -0.0098861 | [10 empty] | Xbar - 3SD |  | -1.9735 | -1.1528 | -2.6116 | [20 empty] | Xbar - 3SD | -0.52983 | -1.1117 | -0.69593 | -0.05186
[26 empty] | Warning Limits as % | Xbar + 2SD |  | 385.6646 | 362.665 | -667.7488 | [19 empty] | Warning Limits as % | Xbar + 2SD | 342.4946 | 356.0761 | 410.3628 | 297.7371
[27 empty] | Xbar - 2SD |  | -185.6646 | -162.665 | 867.7488 | [20 empty] | Xbar - 2SD | -142.4946 | -156.0761 | -210.3628 | -97.7371
[26 empty] | Control Limits as % | Xbar + 3SD |  | 528.497 | 493.9975 | -1051.6232 | [19 empty] | Control Limits as % | Xbar + 3SD | 463.742 | 484.1142 | 565.5441 | 396.6056
[27 empty] | Xbar - 3SD |  | -328.497 | -293.9975 | 1251.6232 | [20 empty] | Xbar - 3SD | -263.742 | -284.1142 | -365.5441 | -196.6056
[26 empty] | Coefficient of Variation % |  |  | 142.8323 | 131.3325 | -383.8744 | [19 empty] | Coefficient of Variation % |  | 121.2473 | 128.0381 | 155.1814 | 98.8685
[26 empty] | 3 x Coefficient of Variation % |  |  | 428.497 | 393.9975 | -1151.6232 | [19 empty] | 3 x Coefficient of Variation % |  | 363.742 | 384.1142 | 465.5441 | 296.6056
[51 empty] | Count |  | 121 | 80 | 80 | 80
[51 empty] | 99th percentile.inc |  | 0.81032 | 2.307 | 1.6068 | 0.10275
[51 empty] | Maximum |  | 0.98412 | 2.4807 | 1.6702 | 0.11195
[51 empty] | Known |  | NA | NA | NA | NA
[51 empty] | MDL |  | 0.5743 | 1.1896 | 0.70151 | 0.061925
[51 empty] | LOQ |  | 2.4357 | 5.0098 | 2.9544 | 0.26079
[51 empty] | ratio mean/mdl |  | 0.3498 | 0.32892 | 0.27139 | 0.42596

## 1st run mixed std 0.6
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3828 | 8388 |  | 3974 | 2.793 | 6.154 | 3.361 |  | 0.25 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 13:45 |  | 1 |  | 3.1381 | 6.6199 | 3.4819 | 0.33466 |  | 4.6023 |  |  |  |  | 10.3324 |  |  |  |  | 16.0625 |  |  |  |  | 11.5547
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3821 | 8287 |  | 3906 | 2.788 | 6.083 | 3.294 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 13:53 |  | 1 |  | 3.1322 | 6.5357 | 3.4035 | 0.32932 |  | 4.4071 |  |  |  |  | 8.9285 |  |  |  |  | 13.4499 |  |  |  |  | 9.7743 |  |  |  | 3.1443 | 6.5132 | 3.3689 | 0.32563 |  | 1
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3850 | 8233 |  | 3812 | 2.807 | 6.044 | 3.237 |  | 0.236 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-15 | 14:02 |  | 1 |  | 3.1565 | 6.4907 | 3.3342 | 0.32194 |  | 5.2157 |  |  |  |  | 8.1779 |  |  |  |  | 11.1401 |  |  |  |  | 7.3132
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3369 | 7479 |  | 3451 | 2.5 | 5.512 | 3.013 |  | 0.204 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 14:50 |  | 1 |  | 2.8694 | 6.1181 | 3.2487 | 0.30773 |  | 4.3534 |  |  |  |  | 1.9676 |  |  |  |  | 8.2887 |  |  |  |  | 2.5751
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3381 | 7411 |  | 3497 | 2.507 | 5.464 | 2.956 |  | 0.208 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 14:58 |  | 1 |  | 2.8793 | 6.062 | 3.1827 | 0.31191 |  | 4.0239 |  |  |  |  | 1.0339 |  |  |  |  | 6.0916 |  |  |  |  | 3.9711 |  |  |  | 2.8805 | 6.0847 | 3.2042 | 0.31096 |  | 2
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3384 | 7466 |  | 3476 | 2.509 | 5.503 | 2.994 |  | 0.206 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-17 | 15:07 |  | 1 |  | 2.8818 | 6.1073 | 3.2256 | 0.31 |  | 3.9415 |  |  |  |  | 1.7891 |  |  |  |  | 7.5197 |  |  |  |  | 3.3338
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3803 | 6305 |  | 2983 | 2.777 | 4.683 | 1.906 |  | 0.163 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 12:48 |  | 1 |  | 3.0959 | 5.3872 | 2.2913 | 0.27719 |  | 3.1975 |  |  |  |  | 10.2136 |  |  |  |  | 23.6248 |  |  |  |  | 7.6029
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3840 | 7126 |  | 3483 | 2.801 | 5.263 | 2.462 |  | 0.207 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 12:56 |  | 1 |  | 3.1263 | 6.0771 | 2.9509 | 0.32473 |  | 4.2099 |  |  |  |  | 1.2858 |  |  |  |  | 1.6383 |  |  |  |  | 8.2421 |  |  |  | 3.139 | 6.1788 | 3.0398 | 0.33918 |  | 3
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3871 | 7368 |  | 3787 | 2.82 | 5.434 | 2.614 |  | 0.233 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-18 | 13:05 |  | 1 |  | 3.1517 | 6.2805 | 3.1288 | 0.35363 |  | 5.058 |  |  |  |  | 4.6754 |  |  |  |  | 4.2928 |  |  |  |  | 17.8758
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3835 | 7863 |  | 3490 | 2.798 | 5.783 | 2.985 |  | 0.208 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-21 | 16:20 |  | 1 |  | 3.0968 | 6.2549 | 3.1582 | 0.30312 |  | 3.2261 |  |  |  |  | 4.2491 |  |  |  |  | 5.2721 |  |  |  |  | 1.0412
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3800 | 7696 |  | 3537 | 2.775 | 5.665 | 2.89 |  | 0.212 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-21 | 16:29 |  | 1 |  | 3.0684 | 6.1198 | 3.0514 | 0.30722 |  | 2.2805 |  |  |  |  | 1.9966 |  |  |  |  | 1.7128 |  |  |  |  | 2.4068 |  |  |  | 3.0765 | 6.1858 | 3.1092 | 0.31036 |  | 4
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3820 | 7859 |  | 3609 | 2.788 | 5.78 | 2.992 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-21 | 16:37 |  | 1 |  | 3.0846 | 6.2517 | 3.1671 | 0.3135 |  | 2.8208 |  |  |  |  | 4.1951 |  |  |  |  | 5.5695 |  |  |  |  | 4.4988
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3776 | 7957 |  | 3793 | 2.76 | 5.85 | 3.09 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 13:24 |  | 1 |  | 3.0539 | 6.5139 | 3.4601 | 0.31946 |  | 1.7951 |  |  |  |  | 8.5653 |  |  |  |  | 15.3355 |  |  |  |  | 6.4864
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3826 | 7217 |  | 3346 | 2.792 | 5.327 | 2.536 |  | 0.195 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 13:33 |  | 1 |  | 3.0945 | 5.8932 | 2.7988 | 0.28107 |  | 3.1495 |  |  |  |  | 1.7794 |  |  |  |  | 6.7083 |  |  |  |  | 6.3114 |  |  |  | 3.0738 | 5.9176 | 2.8438 | 0.28253 |  | 5
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3775 | 7275 |  | 3380 | 2.759 | 5.368 | 2.609 |  | 0.198 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 13:42 |  | 1 |  | 3.053 | 5.9419 | 2.8888 | 0.28399 |  | 1.7681 |  |  |  |  | 0.96858 |  |  |  |  | 3.7052 |  |  |  |  | 5.3379
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4521 | 9500 |  | 43983 | 3.236 | 6.939 | 3.703 |  | 3.737 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 12:44 |  | 1 |  | 3.1076 | 6.2796 | 3.1721 | 0.31801 |  | 3.5851 |  |  |  |  | 4.6607 |  |  |  |  | 5.7362 |  |  |  |  | 6.0044
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4540 | 9612 |  | 44070 | 3.248 | 7.018 | 3.77 |  | 3.744 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 12:54 |  | 1 |  | 3.1209 | 6.3557 | 3.2349 | 0.31865 |  | 4.029 |  |  |  |  | 5.9288 |  |  |  |  | 7.8285 |  |  |  |  | 6.2162 |  |  |  | 3.1398 | 6.3758 | 3.236 | 0.31886 |  | 6
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4594 | 9671 |  | 44127 | 3.283 | 7.06 | 3.777 |  | 3.749 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-21 | 13:04 |  | 1 |  | 3.1587 | 6.3958 | 3.2371 | 0.31906 |  | 5.2906 |  |  |  |  | 6.5968 |  |  |  |  | 7.9029 |  |  |  |  | 6.3549
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 0 | 4552 |  | 39072 | 0 | 3.446 | 3.446 |  | 3.309 | 1 | 0 | 0 |  | 0 |  | TICnp | 2021-07-22 | 13:05 |  | 1 |  | -0.061199 | 2.9183 | 2.9795 | 0.28215 |  | 102.04 |  |  |  |  | 51.362 |  |  |  |  | 0.68401 |  |  |  |  | 5.9508
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 0 | 4717 |  | 39616 | 0 | 3.562 | 3.562 |  | 3.356 | 1 | 0 | 0 |  | 0 |  | TICnp | 2021-07-22 | 13:14 |  | 1 |  | -0.061199 | 3.0304 | 3.0916 | 0.28612 |  | 102.04 |  |  |  |  | 49.4938 |  |  |  |  | 3.0523 |  |  |  |  | 4.6265 | [5 empty] | 3.0573 | 0.28493 |  | 7
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 0 | 4616 |  | 39290 | 0 | 3.491 | 3.491 |  | 3.328 | 1 | 0 | 0 |  | 0 |  | TICnp | 2021-07-22 | 13:23 |  | 1 |  | -0.061199 | 2.9618 | 3.023 | 0.28374 |  | 102.04 |  |  |  |  | 50.6374 |  |  |  |  | 0.76524 |  |  |  |  | 5.4201
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3700 | 7943 |  | 3878 | 2.711 | 5.84 | 3.129 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 13:42 |  | 1 |  | 3.1019 | 6.3266 | 3.2247 | 0.3156 |  | 3.3969 |  |  |  |  | 5.4433 |  |  |  |  | 7.4898 |  |  |  |  | 5.2009
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3712 | 7891 |  | 3916 | 2.719 | 5.803 | 3.084 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 13:51 |  | 1 |  | 3.1118 | 6.2841 | 3.1723 | 0.31869 |  | 3.7269 |  |  |  |  | 4.7347 |  |  |  |  | 5.7424 |  |  |  |  | 6.2315 |  |  |  | 3.0965 | 6.2992 | 3.2027 | 0.31825 |  | 8
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3675 | 7928 |  | 3905 | 2.695 | 5.829 | 3.134 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 14:01 |  | 1 |  | 3.0813 | 6.3143 | 3.2331 | 0.3178 |  | 2.7093 |  |  |  |  | 5.2389 |  |  |  |  | 7.7686 |  |  |  |  | 5.9332
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3870 | 7615 |  | 3647 | 2.82 | 5.608 | 2.789 |  | 0.221 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 13:13 |  | 1 |  | 2.9736 | 5.8794 | 2.9058 | 0.28398 |  | 0.87964 |  |  |  |  | 2.0098 |  |  |  |  | 3.1399 |  |  |  |  | 5.3415
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3897 | 8420 |  | 4284 | 2.837 | 6.177 | 3.34 |  | 0.277 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 13:22 |  | 1 |  | 2.9955 | 6.5343 | 3.5388 | 0.33636 |  | 0.15032 |  |  |  |  | 8.9053 |  |  |  |  | 17.9609 |  |  |  |  | 12.1186 |  |  |  | 3.0076 | 6.5445 | 3.5368 | 0.33759 |  | 9
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3927 | 8445 |  | 4314 | 2.856 | 6.194 | 3.338 |  | 0.279 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-02 | 13:32 |  | 1 |  | 3.0198 | 6.5547 | 3.5349 | 0.33882 |  | 0.66004 |  |  |  |  | 9.2442 |  |  |  |  | 17.8284 |  |  |  |  | 12.9409
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3899 | 7631 |  | 3717 | 2.838 | 5.619 | 2.781 |  | 0.227 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 13:47 |  | 1 |  | 2.9772 | 6.1692 | 3.1921 | 0.30693 |  | 0.76139 |  |  |  |  | 2.8202 |  |  |  |  | 6.4018 |  |  |  |  | 2.3112
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3911 | 7738 |  | 3764 | 2.846 | 5.695 | 2.849 |  | 0.231 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 13:56 |  | 1 |  | 2.9865 | 6.2589 | 3.2723 | 0.31094 |  | 0.44845 |  |  |  |  | 4.3144 |  |  |  |  | 9.0772 |  |  |  |  | 3.6474 |  |  |  | 3.0178 | 6.2283 | 3.2104 | 0.31158 |  | 10
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3991 | 7665 |  | 3779 | 2.897 | 5.644 | 2.746 |  | 0.233 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-04 | 14:05 |  | 1 |  | 3.0491 | 6.1977 | 3.1486 | 0.31222 |  | 1.6378 |  |  |  |  | 3.295 |  |  |  |  | 4.9522 |  |  |  |  | 4.0739
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3970 | 7746 |  | 3934 | 2.884 | 5.701 | 2.817 |  | 0.246 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 15:05 |  | 1 |  | 3.0278 | 6.1703 | 3.1425 | 0.32089 |  | 0.92715 |  |  |  |  | 2.8387 |  |  |  |  | 4.7502 |  |  |  |  | 6.9642
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3941 | 7809 |  | 3906 | 2.866 | 5.746 | 2.88 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 15:14 |  | 1 |  | 3.0052 | 6.2215 | 3.2163 | 0.31857 |  | 0.17299 |  |  |  |  | 3.6919 |  |  |  |  | 7.2109 |  |  |  |  | 6.1906 |  |  |  | 3.0185 | 6.2207 | 3.2023 | 0.32516 |  | 11
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3975 | 7807 |  | 4065 | 2.887 | 5.744 | 2.856 |  | 0.258 | 1 | 0 | 0 |  | 0 |  |  | 2021-08-05 | 15:23 |  | 1 |  | 3.0317 | 6.2199 | 3.1882 | 0.33175 |  | 1.0572 |  |  |  |  | 3.6648 |  |  |  |  | 6.2725 |  |  |  |  | 10.5838
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 2958 | 5971 |  | 2515 | 2.237 | 4.447 | 2.211 |  | 0.123 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-04 | 15:27 |  | 1 |  | 3.0109 | 6.1884 | 3.1775 | 0.2983 |  | 0.36439 |  |  |  |  | 3.14 |  |  |  |  | 5.9156 |  |  |  |  | 0.56664
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3026 | 5948 |  | 2500 | 2.28 | 4.431 | 2.151 |  | 0.121 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-04 | 15:36 |  | 1 |  | 3.0858 | 6.1626 | 3.0768 | 0.29641 |  | 2.8611 |  |  |  |  | 2.71 |  |  |  |  | 2.5589 |  |  |  |  | 1.1969 |  |  |  | 3.0605 | 6.1643 | 3.1038 | 0.30088 |  | 12
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 2980 | 5951 |  | 2571 | 2.251 | 4.433 | 2.183 |  | 0.127 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-04 | 15:46 |  | 1 |  | 3.0352 | 6.166 | 3.1308 | 0.30536 |  | 1.1721 |  |  |  |  | 2.7661 |  |  |  |  | 4.36 |  |  |  |  | 1.7862
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 2976 | 5891 |  | 2415 | 2.248 | 4.391 | 2.143 |  | 0.114 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 16:27 |  | 1 |  | 3.0256 | 6.0985 | 3.0729 | 0.29548 |  | 0.85319 |  |  |  |  | 1.6419 |  |  |  |  | 2.4307 |  |  |  |  | 1.5059
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 2989 | 5904 |  | 2420 | 2.257 | 4.4 | 2.143 |  | 0.114 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 16:37 |  | 1 |  | 3.0399 | 6.1129 | 3.073 | 0.29613 |  | 1.3284 |  |  |  |  | 1.8815 |  |  |  |  | 2.4347 |  |  |  |  | 1.2889 |  |  |  | 3.0717 | 6.0985 | 3.0269 | 0.29607 |  | 13
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3047 | 5878 |  | 2419 | 2.294 | 4.382 | 2.088 |  | 0.114 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-05 | 16:47 |  | 1 |  | 3.1035 | 6.0841 | 2.9807 | 0.296 |  | 3.4484 |  |  |  |  | 1.4023 |  |  |  |  | 0.64376 |  |  |  |  | 1.3323
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 2937 | 5839 |  | 2459 | 2.224 | 4.354 | 2.13 |  | 0.118 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 14:43 |  | 1 |  | 3.0466 | 5.9938 | 2.9472 | 0.29093 |  | 1.5536 |  |  |  |  | 0.10343 |  |  |  |  | 1.7605 |  |  |  |  | 3.0232
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 2918 | 5926 |  | 2527 | 2.211 | 4.416 | 2.205 |  | 0.124 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 14:52 |  | 1 |  | 3.0259 | 6.0897 | 3.0638 | 0.29934 |  | 0.86311 |  |  |  |  | 1.495 |  |  |  |  | 2.127 |  |  |  |  | 0.21898 |  |  |  | 3.0439 | 6.0654 | 3.0216 | 0.29879 |  | 14
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 2951 | 5882 |  | 2518 | 2.233 | 4.385 | 2.152 |  | 0.123 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-06 | 15:02 |  | 1 |  | 3.0619 | 6.0412 | 2.9793 | 0.29823 |  | 2.0624 |  |  |  |  | 0.68661 |  |  |  |  | 0.68916 |  |  |  |  | 0.59012
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 2958 | 6372 |  | 2790 | 2.237 | 4.73 | 2.493 |  | 0.146 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 15:31 |  | 1 |  | 3.0431 | 6.2686 | 3.2255 | 0.30621 |  | 1.4371 |  |  |  |  | 4.4771 |  |  |  |  | 7.5171 |  |  |  |  | 2.0711
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 2998 | 6375 |  | 2806 | 2.262 | 4.732 | 2.47 |  | 0.148 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 15:41 |  | 1 |  | 3.0878 | 6.2718 | 3.184 | 0.30791 |  | 2.9269 |  |  |  |  | 4.5299 |  |  |  |  | 6.1329 |  |  |  |  | 2.6375 |  |  |  | 3.0755 | 6.2491 | 3.1736 | 0.30749 |  | 15
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 2976 | 6332 |  | 2798 | 2.248 | 4.702 | 2.454 |  | 0.147 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-07 | 15:50 |  | 1 |  | 3.0632 | 6.2264 | 3.1632 | 0.30706 |  | 2.1075 |  |  |  |  | 3.7734 |  |  |  |  | 5.4393 |  |  |  |  | 2.3543
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3094 | 6519 |  | 3182 | 2.324 | 4.834 | 2.511 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 13:22 |  | 1 |  | 3.0662 | 6.2298 | 3.1636 | 0.31003 |  | 2.2056 |  |  |  |  | 3.8299 |  |  |  |  | 5.4542 |  |  |  |  | 3.3422
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3105 | 6509 |  | 3275 | 2.331 | 4.828 | 2.497 |  | 0.189 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 13:31 |  | 1 |  | 3.0779 | 6.2196 | 3.1417 | 0.31939 |  | 2.5958 |  |  |  |  | 3.6592 |  |  |  |  | 4.7226 |  |  |  |  | 6.4621 |  |  |  | 3.0726 | 6.2073 | 3.1347 | 0.31908 |  | 16
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3095 | 6485 |  | 3269 | 2.325 | 4.81 | 2.486 |  | 0.188 | 1 | 0 | 0 |  | 0 |  |  | 2021-10-08 | 13:41 |  | 1 |  | 3.0672 | 6.195 | 3.1277 | 0.31878 |  | 2.241 |  |  |  |  | 3.2494 |  |  |  |  | 4.2579 |  |  |  |  | 6.2609
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4101 | 8628 |  | 4322 | 2.968 | 6.324 | 3.356 |  | 0.28 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-09 | 15:09 |  | 1 |  | 2.9953 | 6.2899 | 3.2945 | 0.3008 |  | 0.15638 |  |  |  |  | 4.8309 |  |  |  |  | 9.8182 |  |  |  |  | 0.26755
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4161 | 8685 |  | 4435 | 3.006 | 6.364 | 3.357 |  | 0.29 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-09 | 15:19 |  | 1 |  | 3.0448 | 6.3353 | 3.2905 | 0.30922 |  | 1.4938 |  |  |  |  | 5.5879 |  |  |  |  | 9.682 |  |  |  |  | 3.0726 |  |  |  | 3.0329 | 6.362 | 3.3291 | 0.3157 |  | 17
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4132 | 8752 |  | 4609 | 2.988 | 6.411 | 3.423 |  | 0.305 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-09 | 15:30 |  | 1 |  | 3.0209 | 6.3887 | 3.3678 | 0.32218 |  | 0.69621 |  |  |  |  | 6.4777 |  |  |  |  | 12.2592 |  |  |  |  | 7.3919
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4265 | 8953 |  | 4728 | 3.073 | 6.553 | 3.48 |  | 0.315 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 16:27 |  | 1 |  | 3.023 | 6.4116 | 3.3886 | 0.33492 |  | 0.76773 |  |  |  |  | 6.8605 |  |  |  |  | 12.9534 |  |  |  |  | 11.6389
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4292 | 8920 |  | 4649 | 3.09 | 6.53 | 3.44 |  | 0.309 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 16:39 |  | 1 |  | 3.0435 | 6.3865 | 3.343 | 0.32899 |  | 1.449 |  |  |  |  | 6.4412 |  |  |  |  | 11.4334 |  |  |  |  | 9.6635 |  |  |  | 3.0454 | 6.3937 | 3.3484 | 0.33143 |  | 18
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4297 | 8939 |  | 4714 | 3.093 | 6.543 | 3.45 |  | 0.314 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 16:52 |  | 1 |  | 3.0473 | 6.401 | 3.3537 | 0.33387 |  | 1.5752 |  |  |  |  | 6.6826 |  |  |  |  | 11.7901 |  |  |  |  | 11.2888
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4278 | 8581 |  | 4102 | 3.081 | 6.29 | 3.209 |  | 0.261 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-13 | 14:26 |  | 1 |  | 3.031 | 6.1889 | 3.1579 | 0.29979 |  | 1.0339 |  |  |  |  | 3.1491 |  |  |  |  | 5.2642 |  |  |  |  | 0.069607
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4296 | 8574 |  | 4159 | 3.092 | 6.285 | 3.193 |  | 0.266 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-13 | 14:37 |  | 1 |  | 3.0446 | 6.1837 | 3.1391 | 0.30416 |  | 1.486 |  |  |  |  | 3.061 |  |  |  |  | 4.6359 |  |  |  |  | 1.3883 |  |  |  | 3.031 | 6.1818 | 3.1508 | 0.30424 |  | 19
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4260 | 8569 |  | 4161 | 3.069 | 6.282 | 3.212 |  | 0.266 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-13 | 14:48 |  | 1 |  | 3.0175 | 6.1799 | 3.1624 | 0.30432 |  | 0.58169 |  |  |  |  | 2.9981 |  |  |  |  | 5.4145 |  |  |  |  | 1.4395
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4425 | 8847 |  | 3647 | 3.175 | 6.478 | 3.303 |  | 0.221 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 15:58 |  | 1 |  | 3.0765 | 6.2418 | 3.1653 | 0.27475 |  | 2.5502 |  |  |  |  | 4.0293 |  |  |  |  | 5.5084 |  |  |  |  | 8.4158
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4503 | 8882 |  | 3728 | 3.225 | 6.503 | 3.278 |  | 0.228 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 16:08 |  | 1 |  | 3.1344 | 6.2679 | 3.1334 | 0.28075 |  | 4.4813 |  |  |  |  | 4.4644 |  |  |  |  | 4.4476 |  |  |  |  | 6.4159 |  |  |  | 3.1374 | 6.2828 | 3.1454 | 0.27901 |  | 20
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4511 | 8922 |  | 3681 | 3.23 | 6.531 | 3.301 |  | 0.224 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 16:20 |  | 1 |  | 3.1404 | 6.2977 | 3.1573 | 0.27727 |  | 4.6793 |  |  |  |  | 4.9618 |  |  |  |  | 5.2442 |  |  |  |  | 7.5764
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4411 | 8979 |  | 4259 | 3.166 | 6.572 | 3.406 |  | 0.274 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 16:10 |  | 1 |  | 3.0381 | 6.2668 | 3.2287 | 0.30776 |  | 1.2696 |  |  |  |  | 4.4471 |  |  |  |  | 7.6245 |  |  |  |  | 2.5873
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4474 | 8931 |  | 4230 | 3.206 | 6.538 | 3.332 |  | 0.272 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 16:21 |  | 1 |  | 3.0844 | 6.2314 | 3.147 | 0.30555 |  | 2.8147 |  |  |  |  | 3.8575 |  |  |  |  | 4.9003 |  |  |  |  | 1.8495 |  |  |  | 3.0624 | 6.2388 | 3.1765 | 0.30372 |  | 21
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4414 | 8951 |  | 4182 | 3.167 | 6.551 | 3.384 |  | 0.268 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 16:32 |  | 1 |  | 3.0403 | 6.2462 | 3.2059 | 0.30189 |  | 1.3432 |  |  |  |  | 4.1032 |  |  |  |  | 6.8631 |  |  |  |  | 0.62844
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4222 | 8663 |  | 4077 | 3.045 | 6.348 | 3.303 |  | 0.259 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-24 | 18:16 |  | 1 |  | 2.9498 | 6.1544 | 3.2046 | 0.3069 |  | 1.673 |  |  |  |  | 2.5732 |  |  |  |  | 6.8194 |  |  |  |  | 2.3002
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4288 | 8753 |  | 4058 | 3.087 | 6.412 | 3.325 |  | 0.257 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-24 | 18:26 |  | 1 |  | 3.0005 | 6.2221 | 3.2215 | 0.30541 |  | 0.017202 |  |  |  |  | 3.7011 |  |  |  |  | 7.385 |  |  |  |  | 1.803 |  |  |  | 3.0009 | 6.2657 | 3.2648 | 0.30443 |  | 22
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4289 | 8869 |  | 4033 | 3.088 | 6.494 | 3.406 |  | 0.255 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-24 | 18:38 |  | 1 |  | 3.0013 | 6.3093 | 3.308 | 0.30345 |  | 0.042811 |  |  |  |  | 5.1548 |  |  |  |  | 10.2668 |  |  |  |  | 1.1488
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4311 | 8311 |  | 3512 | 3.102 | 6.1 | 2.998 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 17:37 |  | 1 |  | 3.071 | 5.9152 | 2.8442 | 0.27755 |  | 2.3661 |  |  |  |  | 1.4141 |  |  |  |  | 5.1943 |  |  |  |  | 7.4845
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4333 | 8524 |  | 3497 | 3.116 | 6.25 | 3.134 |  | 0.208 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 17:48 |  | 1 |  | 3.0878 | 6.079 | 2.9912 | 0.27622 |  | 2.928 |  |  |  |  | 1.3171 |  |  |  |  | 0.29376 |  |  |  |  | 7.9251 |  |  |  | 3.0901 | 6.089 | 2.9989 | 0.27371 |  | 23
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4339 | 8550 |  | 3440 | 3.12 | 6.268 | 3.148 |  | 0.203 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 18:00 |  | 1 |  | 3.0924 | 6.099 | 3.0066 | 0.2712 |  | 3.0812 |  |  |  |  | 1.6505 |  |  |  |  | 0.21977 |  |  |  |  | 9.5993
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4098 | 8602 |  | 3738 | 2.966 | 6.305 | 3.34 |  | 0.229 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 13:15 |  | 1 |  | 3.2857 | 6.9132 | 3.6274 | 0.33899
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4207 | 8679 |  | 3707 | 3.036 | 6.359 | 3.324 |  | 0.226 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 13:25 |  | 1 |  | 3.3688 | 6.973 | 3.6042 | 0.33632 |  | 11.1888 | 1.9876 |  |  |  | 16.0674 | 0.25648 |  |  |  | 20.946 | 1.335 |  |  |  | 10.6844 | 2.57 |  |  | 3.3357 | 6.964 | 3.6284 | 0.33205 |  | 24
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4120 | 8656 |  | 3608 | 2.979 | 6.343 | 3.364 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 13:36 |  | 1 |  | 3.3025 | 6.9551 | 3.6526 | 0.32779
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4220 | 8691 |  | 4212 | 3.044 | 6.368 | 3.324 |  | 0.27 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 16:15 |  | 1 |  | 3.3114 | 6.8808 | 3.5694 | 0.33142
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4341 | 8814 |  | 4222 | 3.121 | 6.454 | 3.334 |  | 0.271 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 16:25 |  | 1 |  | 3.4029 | 6.9745 | 3.5716 | 0.33217 |  | 12.2208 | 2.1559 |  |  |  | 15.4864 | 1.3086 |  |  |  | 18.752 | 0.50803 |  |  |  | 9.7407 | 1.7936 |  |  | 3.3666 | 6.9292 | 3.5626 | 0.32922 |  | 25
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4245 | 8695 |  | 4144 | 3.06 | 6.37 | 3.311 |  | 0.264 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 16:37 |  | 1 |  | 3.3303 | 6.8838 | 3.5535 | 0.32627
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4325 | 8620 |  | 3980 | 3.111 | 6.318 | 3.207 |  | 0.25 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 14:04 |  | 1 |  | 3.5238 | 6.8687 | 3.3449 | 0.32358
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4309 | 8600 |  | 4028 | 3.101 | 6.303 | 3.203 |  | 0.254 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 14:14 |  | 1 |  | 3.5112 | 6.8534 | 3.3422 | 0.32704 |  | 16.7918 | 0.42604 |  |  |  | 14.7982 | 1.0018 |  |  |  | 12.8045 | 2.4801 |  |  |  | 8.6996 | 0.57466 |  |  | 3.5038 | 6.8879 | 3.3841 | 0.3261 |  | 26
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4290 | 8690 |  | 4002 | 3.088 | 6.367 | 3.279 |  | 0.252 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 14:25 |  | 1 |  | 3.4963 | 6.9224 | 3.4261 | 0.32516
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3792 | 8006 |  | 4043 | 2.77 | 5.885 | 3.114 |  | 0.256 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:42 |  | 1 |  | 3.1456 | 6.8758 | 3.7303 | 0.33969
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3900 | 8043 |  | 3960 | 2.839 | 5.911 | 3.072 |  | 0.248 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 17:52 |  | 1 |  | 3.235 | 6.9078 | 3.6728 | 0.33284 |  | 8.8949 | 1.9514 |  |  |  | 14.8417 | 0.50169 |  |  |  | 20.7884 | 2.7133 |  |  |  | 10.0249 | 1.6747 |  |  | 3.2668 | 6.8905 | 3.6237 | 0.33007 |  | 27
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3977 | 8003 |  | 3893 | 2.888 | 5.882 | 2.994 |  | 0.243 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 18:03 |  | 1 |  | 3.2987 | 6.8732 | 3.5745 | 0.32731
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3881 | 7652 |  | 3905 | 2.827 | 5.635 | 2.808 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 18:31 |  | 1 |  | 3.3715 | 7.2055 | 3.834 | 0.33528
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3965 | 7657 |  | 3822 | 2.881 | 5.638 | 2.757 |  | 0.236 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 18:41 |  | 1 |  | 3.4406 | 7.2098 | 3.7693 | 0.32866 |  | 14.261 | 0.74321 |  |  |  | 19.7839 | 0.63408 |  |  |  | 25.3069 | 0.53457 |  |  |  | 9.195 | 0.65709 |  |  | 3.4278 | 7.187 | 3.7592 | 0.32758 |  | 28
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3934 | 7604 |  | 3795 | 2.861 | 5.601 | 2.74 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-07 | 18:52 |  | 1 |  | 3.4151 | 7.1643 | 3.7492 | 0.32651
10 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3965 | 7983 |  | 3917 | 2.88 | 5.868 | 2.988 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 17:58 |  | 1 |  | 3.3645 | 7.0031 | 3.6385 | 0.35317
11 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4059 | 8066 |  | 3875 | 2.941 | 5.926 | 2.986 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 18:08 |  | 1 |  | 3.4413 | 7.0705 | 3.6292 | 0.34955 |  | 15.6896 | 1.694 |  |  |  | 16.867 | 1.6689 |  |  |  | 18.0443 | 4.9647 |  |  |  | 15.5828 | 1.6154 |  |  | 3.4707 | 7.012 | 3.5413 | 0.34675 |  | 29
12 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4131 | 7922 |  | 3810 | 2.987 | 5.825 | 2.838 |  | 0.235 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 18:20 |  | 1 |  | 3.5001 | 6.9535 | 3.4534 | 0.34395
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3922 | 7885 |  | 3884 | 2.853 | 5.799 | 2.946 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 14:14 |  | 1 |  | 3.1535 | 6.5169 | 3.3634 | 0.33065
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3968 | 7866 |  | 3826 | 2.883 | 5.785 | 2.903 |  | 0.237 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 14:25 |  | 1 |  | 3.1912 | 6.5013 | 3.3102 | 0.32593 |  | 6.1126 | 0.48838 |  |  |  | 8.3621 | 0.012592 |  |  |  | 10.6115 | 0.49319 |  |  |  | 10.555 | 3.4608 |  |  | 3.1834 | 6.5017 | 3.3183 | 0.33166 |  | 30
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3949 | 7867 |  | 3967 | 2.87 | 5.786 | 2.916 |  | 0.249 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 14:37 |  | 1 |  | 3.1756 | 6.5021 | 3.3265 | 0.3374
13 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3603 | 7872 |  | 3400 | 2.649 | 5.79 | 3.141 |  | 0.2 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 15:57 |  | 1 |  | 2.9591 | 6.3957 | 3.4365 | 0.30464
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3919 | 7941 |  | 3419 | 2.851 | 5.839 | 2.987 |  | 0.201 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 16:07 |  | 1 |  | 3.2167 | 6.4506 | 3.234 | 0.30632 |  | 7.2231 | 0 |  |  |  | 7.2252 | 0.53262 |  |  |  | 7.2274 | 1.0652 |  |  |  | 3.365 | 2.4333 |  |  | 3.2167 | 6.4335 | 3.2168 | 0.31009 |  | 31
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3919 | 7898 |  | 3504 | 2.851 | 5.808 | 2.957 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 16:19 |  | 1 |  | 3.2167 | 6.4164 | 3.1997 | 0.31387
121 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3798 | 7878 |  | 3616 | 2.774 | 5.794 | 3.02 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 16:13 |  | 1 |  | 3.4064 | 6.8186 | 3.4122 | 0.33822
122 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3836 | 7962 |  | 3591 | 2.798 | 5.853 | 3.055 |  | 0.216 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 16:24 |  | 1 |  | 3.4385 | 6.8885 | 3.45 | 0.336 |  | 17.2884 | 4.557 |  |  |  | 14.7663 | 0.07251 |  |  |  | 12.2442 | 4.91 |  |  |  | 12.2962 | 0.5253 |  |  | 3.5187 | 6.886 | 3.3673 | 0.33689 |  | 32
123 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4026 | 7956 |  | 3611 | 2.92 | 5.849 | 2.929 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 16:35 |  | 1 |  | 3.5988 | 6.8835 | 3.2847 | 0.33777
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3939 | 8224 |  | 4316 | 2.864 | 6.038 | 3.174 |  | 0.279 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 16:56 |  | 1 |  | 3.1384 | 6.5699 | 3.4315 | 0.35148
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4028 | 8218 |  | 4385 | 2.921 | 6.034 | 3.113 |  | 0.286 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 17:06 |  | 1 |  | 3.2088 | 6.5651 | 3.3562 | 0.35711 |  | 6.8425 | 0.2221 |  |  |  | 10.2678 | 1.5417 |  |  |  | 13.6932 | 3.1992 |  |  |  | 18.4242 | 1.0344 |  |  | 3.2053 | 6.6161 | 3.4108 | 0.35527 |  | 33
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4019 | 8345 |  | 4340 | 2.915 | 6.124 | 3.209 |  | 0.282 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 17:18 |  | 1 |  | 3.2017 | 6.6671 | 3.4654 | 0.35343
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3750 | 7701 |  | 3208 | 2.743 | 5.669 | 2.926 |  | 0.183 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 16:20 |  | 1 |  | 3.1199 | 6.4425 | 3.3226 | 0.29008
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3889 | 7827 |  | 3245 | 2.832 | 5.758 | 2.925 |  | 0.186 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 16:30 |  | 1 |  | 3.2311 | 6.5462 | 3.3151 | 0.29335 |  | 7.531 | 0.32243 |  |  |  | 9.0415 | 0.11317 |  |  |  | 10.552 | 0.090365 |  |  |  | 2.9677 | 1.5497 |  |  | 3.2259 | 6.5425 | 3.3166 | 0.2911 |  | 34
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3876 | 7818 |  | 3194 | 2.824 | 5.752 | 2.928 |  | 0.182 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 16:41 |  | 1 |  | 3.2207 | 6.5388 | 3.3181 | 0.28884
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4053 | 7695 |  | 3274 | 2.937 | 5.665 | 2.728 |  | 0.189 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 12:22 |  | 1 |  | 3.2426 | 6.4767 | 3.2341 | 0.30182
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4116 | 7691 |  | 3279 | 2.977 | 5.662 | 2.684 |  | 0.189 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 12:33 |  | 1 |  | 3.2934 | 6.4734 | 3.18 | 0.30227 |  | 9.5638 | 0.39211 |  |  |  | 7.9387 | 0.091505 |  |  |  | 6.3137 | 0.58991 |  |  |  | 1.4604 | 1.3858 |  |  | 3.2869 | 6.4763 | 3.1894 | 0.30438 |  | 35
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4100 | 7698 |  | 3326 | 2.967 | 5.667 | 2.7 |  | 0.193 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 12:44 |  | 1 |  | 3.2805 | 6.4793 | 3.1988 | 0.30649
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3830 | 8073 |  | 3580 | 2.794 | 5.932 | 3.137 |  | 0.215 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 17:00 |  | 1 |  | 3.2226 | 6.6564 | 3.4338 | 0.31358
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3753 | 8006 |  | 3545 | 2.745 | 5.884 | 3.139 |  | 0.212 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 17:11 |  | 1 |  | 3.16 | 6.6023 | 3.4423 | 0.31061 |  | 5.9568 | 1.1764 |  |  |  | 10.1195 | 0.14665 |  |  |  | 14.2821 | 0.80808 |  |  |  | 5.0503 | 2.8808 |  |  | 3.1787 | 6.6072 | 3.4285 | 0.31515 |  | 36
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3799 | 8018 |  | 3652 | 2.774 | 5.893 | 3.118 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 17:22 |  | 1 |  | 3.1974 | 6.612 | 3.4146 | 0.31969
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3772 | 7880 |  | 3735 | 2.757 | 5.795 | 3.038 |  | 0.229 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 18:25 |  | 1 |  | 3.2377 | 6.5586 | 3.3209 | 0.32998
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3737 | 7917 |  | 3731 | 2.735 | 5.821 | 3.086 |  | 0.228 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 18:35 |  | 1 |  | 3.2087 | 6.589 | 3.3803 | 0.32964 |  | 6.7077 | 0.46564 |  |  |  | 9.1797 | 1.166 |  |  |  | 11.6516 | 1.8354 |  |  |  | 9.6931 | 0.33974 |  |  | 3.2012 | 6.5508 | 3.3495 | 0.32908 |  | 37
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3719 | 7824 |  | 3718 | 2.723 | 5.756 | 3.032 |  | 0.227 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 18:46 |  | 1 |  | 3.1938 | 6.5126 | 3.3188 | 0.32852
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4084 | 8111 |  | 4100 | 2.957 | 5.959 | 3.002 |  | 0.261 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 15:11 |  | 1 |  | 2.967 | 6.4407 | 3.4737 | 0.33015
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4216 | 8412 |  | 4189 | 3.041 | 6.171 | 3.13 |  | 0.268 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 15:21 |  | 1 |  | 3.0644 | 6.6759 | 3.6116 | 0.33755 |  | 7.2477 | 9.5143 |  |  |  | 11.8777 | 1.0942 |  |  |  | 16.5077 | 6.6567 |  |  |  | 12.6539 | 0.24574 |  |  | 3.2174 | 6.7127 | 3.4952 | 0.33796 |  | 38
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4631 | 8506 |  | 4199 | 3.307 | 6.237 | 2.93 |  | 0.269 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 15:35 |  | 1 |  | 3.3705 | 6.7494 | 3.3789 | 0.33838
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4165 | 8134 |  | 3942 | 3.009 | 5.975 | 2.966 |  | 0.247 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 17:42 |  | 1 |  | 3.0922 | 6.4865 | 3.3943 | 0.32001
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4258 | 8168 |  | 3972 | 3.068 | 5.999 | 2.931 |  | 0.25 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 17:53 |  | 1 |  | 3.1581 | 6.513 | 3.355 | 0.32244 |  | 5.0443 | 0.42701 |  |  |  | 7.9918 | 1.0352 |  |  |  | 10.9393 | 1.6111 |  |  |  | 7.3445 | 0.25183 |  |  | 3.1513 | 6.4795 | 3.3282 | 0.32203 |  | 39
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4239 | 8082 |  | 3962 | 3.056 | 5.938 | 2.882 |  | 0.249 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 18:04 |  | 1 |  | 3.1446 | 6.446 | 3.3014 | 0.32163
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4089 | 7962 |  | 3579 | 2.96 | 5.853 | 2.893 |  | 0.215 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 14:17 |  | 1 |  | 3.2681 | 6.4357 | 3.1676 | 0.30741
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4001 | 8020 |  | 3593 | 2.904 | 5.894 | 2.99 |  | 0.216 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 14:27 |  | 1 |  | 3.1979 | 6.482 | 3.2841 | 0.3086 |  | 7.1943 | 1.1173 |  |  |  | 7.9599 | 0.13548 |  |  |  | 8.7254 | 1.3707 |  |  |  | 3.6219 | 1.4559 |  |  | 3.2158 | 6.4776 | 3.2618 | 0.31087 |  | 40
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4046 | 8009 |  | 3646 | 2.932 | 5.887 | 2.954 |  | 0.221 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 14:39 |  | 1 |  | 3.2338 | 6.4732 | 3.2394 | 0.31313
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4078 | 7871 |  | 3579 | 2.953 | 5.789 | 2.836 |  | 0.215 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 18:42 |  | 1 |  | 3.1941 | 6.2662 | 3.0721 | 0.308
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4085 | 7982 |  | 3558 | 2.957 | 5.867 | 2.91 |  | 0.213 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 18:53 |  | 1 |  | 3.1995 | 6.3536 | 3.1541 | 0.30625 |  | 6.7799 | 0.24116 |  |  |  | 7.9539 | 3.8164 |  |  |  | 9.1279 | 7.3148 |  |  |  | 2.7897 | 1.3736 |  |  | 3.2034 | 6.4772 | 3.2738 | 0.30837 |  | 41
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4095 | 8296 |  | 3609 | 2.964 | 6.089 | 3.125 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 19:05 |  | 1 |  | 3.2073 | 6.6008 | 3.3936 | 0.31049
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | Stored Cal TIC [mg/l] | Stored Cal TC [mg/l] | Stored Cal TOC (Diff.) [mg/l] | Stored Cal NPOC [mg/l] | Stored Cal TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] |  | TIC Absolute value Percent error for check standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes |  | TC Absolute value Percent error for check standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes |  | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes |  | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
[5 empty] | Mean |  |  | 3811.2033 | 7832.4959 |  | 5539.9431 | 2.7739 | 5.7618 | 2.9879 |  | 0.38615 | [10 empty] | Mean |  | 3.0702 | 6.3435 | 3.2733 | [20 empty] | Mean | 3.1624 | 6.4445 | 3.2766 | 0.31571 | Min | 1
[5 empty] | Std |  |  | 734.3596 | 985.0986 |  | 8252.5345 | 0.51453 | 0.69555 | 0.37493 |  | 0.71927 | [10 empty] | Std |  | 0.51786 | 0.61995 | 0.22109 | [20 empty] | Std | 0.14586 | 0.30165 | 0.1933 | 0.018558 | Max | 41
[5 empty] | CV |  |  | 19.2684 | 12.5771 |  | 148.9642 | 18.5487 | 12.0718 | 12.5485 |  | 186.2689 | [10 empty] | CV |  | 16.8676 | 9.7731 | 6.7545 | [20 empty] | CV | 4.6124 | 4.6808 | 5.8995 | 5.878
|  |  |  | Warning Limits | Xbar + 2SD |  |  | 5279.9224 | 9802.6931 |  | 22045.012 | 3.803 | 7.1529 | 3.7377 |  | 1.8247 | [9 empty] | Warning Limits | Xbar + 2SD |  | 4.1059 | 7.5834 | 3.7155 | [19 empty] | Warning Limits | Xbar + 2SD | 3.4541 | 7.0478 | 3.6632 | 0.35283
[5 empty] | Xbar - 2SD |  |  | 2342.4842 | 5862.2988 |  | -10965.1258 | 1.7449 | 4.3707 | 2.238 |  | -1.0524 | [10 empty] | Xbar - 2SD |  | 2.0344 | 5.1036 | 2.8311 | [20 empty] | Xbar - 2SD | 2.8706 | 5.8412 | 2.89 | 0.2786
|  |  |  | Control Limits | Xbar + 3SD |  |  | 6014.2819 | 10787.7916 |  | 30297.5465 | 4.3175 | 7.8484 | 4.1127 |  | 2.544 | [9 empty] | Control Limits | Xbar + 3SD |  | 4.6237 | 8.2033 | 3.9366 | [19 empty] | Control Limits | Xbar + 3SD | 3.6 | 7.3494 | 3.8565 | 0.37139
[5 empty] | Xbar - 3SD |  |  | 1608.1246 | 4877.2003 |  | -19217.6603 | 1.2303 | 3.6751 | 1.8631 |  | -1.7717 | [10 empty] | Xbar - 3SD |  | 1.5166 | 4.4836 | 2.61 | [20 empty] | Xbar - 3SD | 2.7248 | 5.5395 | 2.6967 | 0.26004
[26 empty] | Warning Limits as % | Xbar + 2SD |  | 133.7352 | 119.5462 | 113.5089 | [19 empty] | Warning Limits as % | Xbar + 2SD | 109.2248 | 109.3617 | 111.7989 | 111.756
[27 empty] | Xbar - 2SD |  | 66.2648 | 80.4538 | 86.4911 | [20 empty] | Xbar - 2SD | 90.7752 | 90.6383 | 88.2011 | 88.244
[26 empty] | Control Limits as % | Xbar + 3SD |  | 150.6028 | 129.3193 | 120.2634 | [19 empty] | Control Limits as % | Xbar + 3SD | 113.8372 | 114.0425 | 117.6984 | 117.634
[27 empty] | Xbar - 3SD |  | 49.3972 | 70.6807 | 79.7366 | [20 empty] | Xbar - 3SD | 86.1628 | 85.9575 | 82.3016 | 82.366
[26 empty] | Coefficient of Variation % |  |  | 16.8676 | 9.7731 | 6.7545 | [19 empty] | Coefficient of Variation % |  | 4.6124 | 4.6808 | 5.8995 | 5.878
[26 empty] | 3 x Coefficient of Variation % |  |  | 50.6028 | 29.3193 | 20.2634 | [19 empty] | 3 x Coefficient of Variation % |  | 13.8372 | 14.0425 | 17.6984 | 17.634