## notes
MDL from spiked blanks okay
Used repeated measures of 5 ppb standard to calculate a similar value over days (and preparations) just for comparison and they are very similar.
MDLb is a problem for FIA
Blanks frequently are misintegrated since the software picks up any air spike or small valve artifact. When there is a peak, there is not as much of a problem so spiked blanks just make more sense.
We could try to fix the integrations for blanks. In the past we have not spent the time to do so. So bounces high are often just misintegrations we did not fix.
General observations…
Some blanks are no numerical result (no peak, ND)
Some blanks do have a deviation from flat baseline or even a peak
EPA says when you have a mix of nondetects and peaks:
if you have <100 observations, use the highest one
If you have > 100 observations, use the 99th percentile
Need data that spans 3 batches on 3 different days
A minimum of 7 reps that have been prepared independently (spikes into a tube are independent, but pouring from a flask into several vials is not)
For the spikes only, if any result is <= 0, a higher spike level is needed. (BUT would this apply only to analytical result BEFORE correction? I think so.)
For blanks only, use numerical results (both + and -) even those below MDL but do not use nondetects (no peak)
NO use of outlier detection
But you can eliminate values for "gross failure" (which could includes misintegration), but document why
Don't use data over 24 months old
Spiking level from 2x to 10x the estimated MDL is okay
Use sample standard deviation (stdev.s)

## rolling spiked blank
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Reduction Efficiency | Analyst Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | NH4-N NOTES: | NH4-N concentration | NH4-N- % Error in check stds | NH4-N- DQO (<=20%) | NH4-N- RPD of duplicates | NH4-N- DQO (<=20% if conc>LOQ) | NH4-N- PR of spikes | NH4-N- DQO (80-120%) if conc>LOQ | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BN PO4-P NOTES: | PO4-P concentration | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BN NO3-N NOTES: | NO3-N concentration | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spiked blank | 1 | 1 | 1 | NH4-N | ug N/L | 0.0275 | 0.629 | 13.1 | PO4-P | ug P/L | 0.0544 | 0.884 | 17.8 | NO3-N+NO2-N | ug N/L | 0.0194 | 0.335 | 18.4 |  | 1 |  | 13.1 | [6 empty] | 1 |  | 17.8 | [6 empty] | 1 |  | 18.4 | [7 empty] | 1
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spiked blank | 2 | 1 | 1 | NH4-N | ug N/L | 0.027 | 0.692 | 14.5 | PO4-P | ug P/L | 0.0545 | 0.858 | 17.1 | NO3-N+NO2-N | ug N/L | 0.02 | 0.342 | 18.8 |  | 1 |  | 14.5 | [6 empty] | 1 |  | 17.1 | [6 empty] | 1 |  | 18.8 | [7 empty] | 2
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spiked blank | 3 | 1 | 1 | NH4-N | ug N/L | 0.027 | 0.574 | 11.9 | PO4-P | ug P/L | 0.0558 | 0.877 | 17.6 | NO3-N+NO2-N | ug N/L | 0.021 | 0.337 | 18.5 |  | 1 |  | 11.9 | [6 empty] | 1 |  | 17.6 | [6 empty] | 1 |  | 18.5 | [7 empty] | 3
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | spiked blank | 1 | 1 | 1 | NH4-N | ug N/L | 0.0257 | 0.597 | 11.5 | PO4-P | ug P/L | 0.0526 | 0.848 | 17.7 | NO3-N+NO2-N | ug N/L | 0.0187 | 0.312 | 17 |  | 1 |  | 11.5 | [6 empty] | 1 |  | 17.7 | [6 empty] | 1 |  | 17 | [7 empty] | 4
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | spiked blank | 2 | 1 | 1 | NH4-N | ug N/L | 0.027 | 0.634 | 12.2 | PO4-P | ug P/L | 0.0564 | 0.92 | 19.6 | NO3-N+NO2-N | ug N/L | 0.0199 | 0.325 | 17.7 |  | 1 |  | 12.2 | [6 empty] | 1 |  | 19.6 | [6 empty] | 1 |  | 17.7 | [7 empty] | 5
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | spiked blank | 3 | 1 | 1 | NH4-N | ug N/L | 0.0291 | 0.522 | 10 | PO4-P | ug P/L | 0.0556 | 0.916 | 19.5 | NO3-N+NO2-N | ug N/L | 0.0201 | 0.336 | 18.3 |  | 1 |  | 10 | [6 empty] | 1 |  | 19.5 | [6 empty] | 1 |  | 18.3 | [7 empty] | 6
2020-11-10 | OM_11-10-2020_02-55-20PM.OMN | spiked blank | 1 | 1 | 1 | NH4-N | ug N/L | 0.0429 | 0.132 | 4.82 | PO4-P | ug P/L | 0.0432 | 0.769 | 18 | NO3-N | ug N/L | 0.00999 | 0.138 | 25.1 |  | 1 | [8 empty] | 1 |  | 18 | [6 empty] | 3 | loop fill incomplete | [8 empty] | 7
2020-11-10 | OM_11-10-2020_02-55-20PM.OMN | spiked blank | 2 | 1 | 1 | NH4-N | ug N/L | 0.0415 | 1.01 | 19.2 | PO4-P | ug P/L | 0.045 | 0.793 | 18.6 | NO3-N | ug N/L | 0.0112 | 0.155 | 28.1 |  | 1 |  | 19.2 | [6 empty] | 1 |  | 18.6 | [6 empty] | 3 | loop fill incomplete | [8 empty] | 8
2020-11-10 | OM_11-10-2020_02-55-20PM.OMN | spiked blank | 3 | 1 | 1 | NH4-N | ug N/L | 0.0435 | 1.03 | 19.6 | PO4-P | ug P/L | 0.0472 | 0.804 | 18.9 | NO3-N | ug N/L | 0.0101 | 0.124 | 22.7 |  | 1 |  | 19.6 | [6 empty] | 1 |  | 18.9 | [6 empty] | 3 | loop fill incomplete | [8 empty] | 9
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spiked blank | 1 | 1 | 1 | NH4-N | ug N/L | 0.0337 | 0.825 | 19.3 | PO4-P | ug P/L | 0.0604 | 0.991 | 19.6 | NO3-N | ug N/L | 0.058 | 0.745 | 21.1 |  | 1 |  | 19.3 | [6 empty] | 1 |  | 19.6 | [6 empty] | 1 |  | 21.1 | [7 empty] | 10
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spiked blank | 2 | 1 | 1 | NH4-N | ug N/L | 0.0313 | 0.735 | 17.3 | PO4-P | ug P/L | 0.0611 | 1.02 | 20.3 | NO3-N | ug N/L | 0.0621 | 0.794 | 22.5 |  | 1 |  | 17.3 | [6 empty] | 1 |  | 20.3 | [6 empty] | 1 |  | 22.5 | [7 empty] | 11
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spiked blank | 3 | 1 | 1 | NH4-N | ug N/L | 0.0331 | 0.827 | 19.3 | PO4-P | ug P/L | 0.0646 | 1.06 | 21.2 | NO3-N | ug N/L | 0.0615 | 0.753 | 21.3 |  | 1 |  | 19.3 | [6 empty] | 1 |  | 21.2 | [6 empty] | 1 |  | 21.3 | [7 empty] | 12
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | spiked blank | 143 | 1 | 1 | NH4-N | ug N/L | 0.0309 | 0.836 | 16 | PO4-P | ug P/L | 0.0586 | 0.942 | 19.2 | NO3-N | ug N/L | 0.061 | 0.737 | 18.3 |  | 1 |  | 16 | [6 empty] | 1 |  | 19.2 | [6 empty] | 1 |  | 18.3 | [7 empty] | 13
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | spiked blank | 144 | 1 | 1 | NH4-N | ug N/L | 0.0393 | 1.08 | 21.1 | PO4-P | ug P/L | 0.0584 | 0.958 | 19.5 | NO3-N | ug N/L | 0.0616 | 0.745 | 18.5 |  | 1 |  | 21.1 | [6 empty] | 1 |  | 19.5 | [6 empty] | 1 |  | 18.5 | [7 empty] | 14
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | spiked blank | 145 | 1 | 1 | NH4-N | ug N/L | 0.0342 | 0.846 | 16.2 | PO4-P | ug P/L | 0.0596 | 0.954 | 19.4 | NO3-N | ug N/L | 0.0654 | 0.791 | 19.7 |  | 1 |  | 16.2 | [6 empty] | 1 |  | 19.4 | [6 empty] | 1 |  | 19.7 | [7 empty] | 15
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | spiked blank | 1 | 1 | 1 | NH4-N | ug N/L | 0.0352 | 0.824 | 16.5 | PO4-P | ug P/L | 0.0843 | 1.08 | 19.5 | NO3-N | ug N/L | 0.0331 | 0.428 | 20.4 |  | 1 |  | 16.5 | [6 empty] | 1 |  | 19.5 | [6 empty] | 1 |  | 20.4 | [7 empty] | 16
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | spiked blank | 2 | 1 | 1 | NH4-N | ug N/L | 0.034 | 0.719 | 14.4 | PO4-P | ug P/L | 0.0799 | 1.07 | 19.4 | NO3-N | ug N/L | 0.0311 | 0.389 | 18.3 |  | 1 |  | 14.4 | [6 empty] | 1 |  | 19.4 | [6 empty] | 1 |  | 18.3 | [7 empty] | 17
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | spiked blank | 3 | 1 | 1 | NH4-N | ug N/L | 0.0394 | 0.709 | 14.2 | PO4-P | ug P/L | 0.0852 | 1 | 17.9 | NO3-N | ug N/L | 0.0278 | 0.38 | 17.8 |  | 1 |  | 14.2 | [6 empty] | 1 |  | 17.9 | [6 empty] | 1 |  | 17.8 | [7 empty] | 18
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | spiked blank + 80 | 1 | 1 | 1 | NH4-N | ug N/L | 0.0265 | 0.648 | 13.9 | PO4-P | ug P/L | 0.0587 | 0.996 | 20.3 | NO3-N | ug N/L | 0.0279 | 0.353 | 18.6 |  | 1 |  | 13.9 | [6 empty] | 1 |  | 20.3 | [6 empty] | 1 |  | 18.6 | [7 empty] | 19
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | spiked blank + 80 | 2 | 1 | 1 | NH4-N | ug N/L | 0.0274 | 0.614 | 13.2 | PO4-P | ug P/L | 0.059 | 0.995 | 20.3 | NO3-N | ug N/L | 0.0284 | 0.352 | 18.5 |  | 1 |  | 13.2 | [6 empty] | 1 |  | 20.3 | [6 empty] | 1 |  | 18.5 | [7 empty] | 20
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | spiked blank + 80 | 3 | 1 | 1 | NH4-N | ug N/L | 0.0338 | 0.606 | 13.1 | PO4-P | ug P/L | 0.06 | 1.02 | 20.9 | NO3-N | ug N/L | 0.0285 | 0.362 | 19.1 |  | 1 |  | 13.1 | [6 empty] | 1 |  | 20.9 | [6 empty] | 1 |  | 19.1 | [7 empty] | 21
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | spiked blank 4+80 | 1 | 1 | 1 | NH4-N | ug N/L | 0.0281 | 0.65 | 14.8 | PO4-P | ug P/L | 0.061 | 1.05 | 21 | NO3-N | ug N/L | 0.0319 | 0.406 | 21.7 |  | 1 |  | 14.8 | [6 empty] | 1 |  | 21 | [6 empty] | 1 |  | 21.7 | [7 empty] | 22
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | spiked blank 4+80 | 2 | 1 | 1 | NH4-N | ug N/L | 0.0277 | 0.714 | 16.2 | PO4-P | ug P/L | 0.0601 | 1.02 | 20.5 | NO3-N | ug N/L | 0.0272 | 0.361 | 19.3 |  | 1 |  | 16.2 | [6 empty] | 1 |  | 20.5 | [6 empty] | 1 |  | 19.3 | [7 empty] | 23
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | spiked blank 4+80 | 3 | 1 | 1 | NH4-N | ug N/L | 0.0313 | 0.735 | 16.7 | PO4-P | ug P/L | 0.0596 | 1 | 20 | NO3-N | ug N/L | 0.0268 | 0.354 | 19 |  | 1 |  | 16.7 | [6 empty] | 1 |  | 20 | [6 empty] | 1 |  | 19 | [7 empty] | 24
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | spiked blank 4 + 80 | 1 | 1 | 1 | NH4-N | ug N/L | 0.0307 | 0.71 | 15.3 | PO4-P | ug P/L | 0.0599 | 1.02 | 19.1 | NO3-N | ug N/L | 0.0345 | 0.409 | 20 |  | 1 |  | 15.3 | [6 empty] | 1 |  | 19.1 | [6 empty] | 1 |  | 20 | [7 empty] | 25
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | spiked blank 4 + 80 | 2 | 1 | 1 | NH4-N | ug N/L | 0.0328 | 0.756 | 16.2 | PO4-P | ug P/L | 0.0622 | 1.05 | 19.7 | NO3-N | ug N/L | 0.0321 | 0.425 | 20.8 |  | 1 |  | 16.2 | [6 empty] | 1 |  | 19.7 | [6 empty] | 1 |  | 20.8 | [7 empty] | 26
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | spiked blank 4 + 80 | 3 | 1 | 1 | NH4-N | ug N/L | 0.0314 | 0.761 | 16.3 | PO4-P | ug P/L | 0.0609 | 1.05 | 19.6 | NO3-N | ug N/L | 0.0321 | 0.414 | 20.3 |  | 1 |  | 16.3 | [6 empty] | 1 |  | 19.6 | [6 empty] | 1 |  | 20.3 | [7 empty] | 27
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | spiked blank + 80 | 1 | 1 | 1 | NH4-N | ug N/L | 0.0279 | 0.668 | 20.6 | PO4-P | ug P/L | 0.0576 | 0.979 | 19.1 | NO3-N | ug N/L | 0.0274 | 0.384 | 15.9 |  | 1 |  | 20.6 | [6 empty] | 1 |  | 19.1 | [6 empty] | 1 |  | 15.9 | [7 empty] | 28
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | spiked blank + 80 | 2 | 1 | 1 | NH4-N | ug N/L | 0.0613 | 0.672 | 20.6 | PO4-P | ug P/L | 0.0571 | 0.979 | 19 | NO3-N | ug N/L | 0.031 | 0.402 | 16.8 |  | 1 |  | 20.6 | [6 empty] | 1 |  | 19 | [6 empty] | 1 |  | 16.8 | [7 empty] | 29
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | spiked blank + 80 | 3 | 1 | 1 | NH4-N | ug N/L | 0.0406 | 0.93 | 25 | PO4-P | ug P/L | 0.0572 | 0.947 | 18.3 | NO3-N | ug N/L | 0.0329 | 0.424 | 17.8 |  | 1 |  | 25 | [6 empty] | 1 |  | 18.3 | [6 empty] | 1 |  | 17.8 | [7 empty] | 30
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | spiked blank 4+80 | 1 | 1 | 1 | NH4-N | ug N/L | 0.033 | 0.735 | 16.3 | PO4-P | ug P/L | 0.0574 | 0.963 | 18.4 | NO3-N | ug N/L | 0.024 | 0.371 | 16.7 |  | 1 |  | 16.3 | [6 empty] | 1 |  | 18.4 | [6 empty] | 1 |  | 16.7 | [7 empty] | 31
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | spiked blank 4+80 | 2 | 1 | 1 | NH4-N | ug N/L | 0.0332 | 0.857 | 19 | PO4-P | ug P/L | 0.0608 | 0.964 | 18.5 | NO3-N | ug N/L | 0.0302 | 0.436 | 19.8 |  | 1 |  | 19 | [6 empty] | 1 |  | 18.5 | [6 empty] | 1 |  | 19.8 | [7 empty] | 32
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | spiked blank 4+80 | 3 | 1 | 1 | NH4-N | ug N/L | 0.0329 | 0.882 | 19.6 | PO4-P | ug P/L | 0.058 | 0.967 | 18.5 | NO3-N | ug N/L | 0.0289 | 0.389 | 17.6 |  | 1 |  | 19.6 | [6 empty] | 1 |  | 18.5 | [6 empty] | 1 |  | 17.6 | [7 empty] | 33
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | spiked blank | 1 | 1 | 1 | NH4-N | ug N/L | 0.0293 | 0.795 | 15.4 | PO4-P | ug P/L | 0.0564 | 1.03 | 21.1 | NO3-N | ug N/L | 0.0291 | 0.394 | 18.5 |  | 1 |  | 15.4 | [6 empty] | 1 |  | 21.1 | [6 empty] | 1 |  | 18.5 | [7 empty] | 34
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | spiked blank | 2 | 1 | 1 | NH4-N | ug N/L | 0.0272 | 0.663 | 12.6 | PO4-P | ug P/L | 0.0574 | 1.02 | 20.9 | NO3-N | ug N/L | 0.0293 | 0.392 | 18.4 |  | 1 |  | 12.6 | [6 empty] | 1 |  | 20.9 | [6 empty] | 1 |  | 18.4 | [7 empty] | 35
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | spiked blank | 3 | 1 | 1 | NH4-N | ug N/L | 0.035 | 0.866 | 16.8 | PO4-P | ug P/L | 0.0575 | 1.01 | 20.8 | NO3-N | ug N/L | 0.0296 | 0.428 | 20.1 |  | 1 |  | 16.8 | [6 empty] | 1 |  | 20.8 | [6 empty] | 1 |  | 20.1 | [7 empty] | 36
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | spiked blank | 1 | 1 | 1 | NH4-N | ug N/L | 0.0339 | 0.779 | 17.7 | PO4-P | ug P/L | 0.0561 | 1 | 19.2 | NO3-N | ug N/L | 0.0289 | 0.426 | 20.5 |  | 1 |  | 17.7 | [6 empty] | 1 |  | 19.2 | [6 empty] | 1 |  | 20.5 | [7 empty] | 37
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | spiked blank | 2 | 1 | 1 | NH4-N | ug N/L | 0.0259 | 0.589 | 13.7 | PO4-P | ug P/L | 0.0568 | 0.976 | 18.6 | NO3-N | ug N/L | 0.0298 | 0.399 | 19.2 |  | 1 |  | 13.7 | [6 empty] | 1 |  | 18.6 | [6 empty] | 1 |  | 19.2 | [7 empty] | 38
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | spiked blank | 3 | 1 | 1 | NH4-N | ug N/L | 0.0289 | 0.769 | 17.5 | PO4-P | ug P/L | 0.0581 | 1.02 | 19.6 | NO3-N | ug N/L | 0.0304 | 0.414 | 19.9 |  | 1 |  | 17.5 | [6 empty] | 1 |  | 19.6 | [6 empty] | 1 |  | 19.9 | [7 empty] | 39
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | spiked blank 4+80 | 1 | 1 | 1 | NH4-N | ug N/L | 0.0275 | 0.675 | 10.5 | PO4-P | ug P/L | 0.0573 | 0.969 | 19.2 | NO3-N | ug N/L | 0.0253 | 0.366 | 17.9 |  | 2 | restricted recal begin to middle | 14.4125 | [6 empty] | 1 |  | 19.2 | [6 empty] | 1 |  | 17.9 | [7 empty] | 40
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | spiked blank 4+80 | 2 | 1 | 1 | NH4-N | ug N/L | 0.0346 | 0.748 | 11.9 | PO4-P | ug P/L | 0.0561 | 0.923 | 18.1 | NO3-N | ug N/L | 0.0265 | 0.369 | 18.1 |  | 2 | restricted recal begin to middle | 16.2446 | [6 empty] | 1 |  | 18.1 | [6 empty] | 1 |  | 18.1 | [7 empty] | 41
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | spiked blank 4+80 | 3 | 1 | 1 | NH4-N | ug N/L | 0.0331 | 0.824 | 13.4 | PO4-P | ug P/L | 0.0564 | 0.952 | 18.8 | NO3-N | ug N/L | 0.027 | 0.371 | 18.1 |  | 2 | restricted recal begin to middle | 18.1472 | [6 empty] | 1 |  | 18.8 | [6 empty] | 1 |  | 18.1 | [7 empty] | 42
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | spike blank | 1 | 1 | 1 | NH4-N | ug N/L | 0.0315 | 0.678 | 12.9 | PO4-P | ug P/L | 0.0746 | 1.17 | 22.3 | NO3-N | ug N/L | 0.029 | 0.428 | 20.6 |  | 1 |  | 14.4879 | [6 empty] | 1 |  | 22.3 | [6 empty] | 1 |  | 20.6 | [7 empty] | 43
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | spike blank | 2 | 1 | 1 | NH4-N | ug N/L | 0.0324 | 0.853 | 16.7 | PO4-P | ug P/L | 0.0752 | 1.21 | 23 | NO3-N | ug N/L | 0.03 | 0.461 | 22.3 |  | 1 |  | 18.8719 | [6 empty] | 1 |  | 23 | [6 empty] | 1 |  | 22.3 | [7 empty] | 44
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | spike blank | 3 | 1 | 1 | NH4-N | ug N/L | 0.0342 | 0.892 | 17.5 | PO4-P | ug P/L | 0.0756 | 1.19 | 22.6 | NO3-N | ug N/L | 0.0291 | 0.459 | 22.2 |  | 1 |  | 19.8453 | [6 empty] | 1 |  | 22.6 | [6 empty] | 1 |  | 22.2 | [7 empty] | 45
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | spiked blank | 1 | 1 | 1 | NH4-N | ug N/L | 0.0341 | 0.847 | 17.9 | PO4-P | ug P/L | 0.0624 | 0.964 | 18.7 | NO3-N | ug N/L | 0.0269 | 0.378 | 20.6 |  | 1 |  | 17.9 | [6 empty] | 1 |  | 18.7 | [6 empty] | 1 |  | 20.6 | [7 empty] | 46
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | spiked blank | 2 | 1 | 1 | NH4-N | ug N/L | 0.0346 | 0.882 | 18.5 | PO4-P | ug P/L | 0.0622 | 0.955 | 18.5 | NO3-N | ug N/L | 0.0264 | 0.368 | 20 |  | 1 |  | 18.5 | [6 empty] | 1 |  | 18.5 | [6 empty] | 1 |  | 20 | [7 empty] | 47
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | spiked blank | 3 | 1 | 1 | NH4-N | ug N/L | 0.0343 | 0.725 | 15.5 | PO4-P | ug P/L | 0.0622 | 0.971 | 18.8 | NO3-N | ug N/L | 0.0256 | 0.36 | 19.6 |  | 1 |  | 15.5 | [6 empty] | 1 |  | 18.8 | [6 empty] | 1 |  | 19.6 | [7 empty] | 48
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | spiked blank 4+ 80 | 1 | 1 | 1 | NH4-N | ug N/L | 0.033 | 0.897 | 18.4 | PO4-P | ug P/L | 0.062 | 0.97 | 19.3 | NO3-N | ug N/L | 0.0311 | 0.405 | 17.5 |  | 1 |  | 18.4 | [6 empty] | 1 |  | 19.3 | [6 empty] | 1 |  | 17.5 | [7 empty] | 49
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | spiked blank 4 + 80 | 2 | 1 | 1 | NH4-N | ug N/L | 0.0448 | 0.936 | 19.3 | PO4-P | ug P/L | 0.0633 | 1.01 | 20.2 | NO3-N | ug N/L | 0.0326 | 0.427 | 18.5 |  | 1 |  | 19.3 | [6 empty] | 1 |  | 20.2 | [6 empty] | 1 |  | 18.5 | [7 empty] | 50
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | spiked blank 4 + 80 | 3 | 1 | 1 | NH4-N | ug N/L | 0.029 | 0.681 | 13.8 | PO4-P | ug P/L | 0.063 | 1 | 20.1 | NO3-N | ug N/L | 0.0358 | 0.462 | 20.2 |  | 1 |  | 13.8 | [6 empty] | 1 |  | 20.1 | [6 empty] | 1 |  | 20.2 | [7 empty] | 51
2022-02-15 | OM_2-15-2022_11-16-18AM.OMN | spiked blank 4 + 80 | 1 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0297 | 0.654 | 12.6 | [11 empty] | 1 |  | 12.6 | [25 empty] | 52
2022-02-15 | OM_2-15-2022_11-16-18AM.OMN | spiked blank 4 + 80 | 2 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0254 | 0.656 | 12.6 | [11 empty] | 1 |  | 12.6 | [25 empty] | 53
2022-02-15 | OM_2-15-2022_11-16-18AM.OMN | spiked blank 4 + 80 | 3 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0333 | 0.636 | 12.2 | [11 empty] | 1 |  | 12.2 | [25 empty] | 54
2022-02-15 | OM_2-15-2022_03-54-18PM.OMN | spiked blank | 1 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0264 | 0.572 | 11.6 | [11 empty] | 1 |  | 11.6 | [25 empty] | 55
2022-02-15 | OM_2-15-2022_03-54-18PM.OMN | spiked blank | 2 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0296 | 0.614 | 12.7 | [11 empty] | 1 |  | 12.7 | [25 empty] | 56
2022-02-15 | OM_2-15-2022_03-54-18PM.OMN | spiked blank | 3 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0308 | 0.691 | 14.6 | [11 empty] | 1 |  | 14.6 | [25 empty] | 57
2022-02-17 | OM_2-17-2022_10-56-00AM.OMN | spiked blank 4 + 80 | 1 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0238 | 0.572 | 11.9 | [11 empty] | 1 |  | 11.9 | [25 empty] | 58
2022-02-17 | OM_2-17-2022_10-56-00AM.OMN | spiked blank 4 + 80 | 2 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0291 | 0.691 | 14.6 | [11 empty] | 1 |  | 14.6 | [25 empty] | 59
2022-02-17 | OM_2-17-2022_10-56-00AM.OMN | spiked blank 4 + 80 | 3 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0257 | 0.62 | 13 | [11 empty] | 1 |  | 13 | [25 empty] | 60
2022-02-18 | OM_2-18-2022_11-06-19AM edit2.omn | spiked blank 4 + 80 | 1 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0314 | 0.617 | 12.7 | [11 empty] | 1 |  | 12.7 | [25 empty] | 61
2022-02-18 | OM_2-18-2022_11-06-19AM edit2.omn | spiked blank 4 + 80 | 2 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0232 | 0.539 | 10.9 | [11 empty] | 1 |  | 10.9 | [25 empty] | 62
2022-02-18 | OM_2-18-2022_11-06-19AM edit2.omn | spiked blank 4 + 80 | 3 | 1 | 1 | NH4-N phenolate | ug N/L | 0.0257 | 0.482 | 9.61 | [11 empty] | 1 |  | 9.61 | [25 empty] | 63
[22 empty] | Known |  | 19.6078 | [8 empty] | 19.6078 | [8 empty] | 19.6078
| Mean | [20 empty] | Mean |  | 15.6254 | [8 empty] | 19.7167 | [8 empty] | 19.1833 | [6 empty] | Min | 1
| Std | [20 empty] | Std |  | 3.0258 | [8 empty] | 1.2566 | [8 empty] | 1.507 | [6 empty] | Max | 63
| CV | [20 empty] | CV |  | 19.3646 | [8 empty] | 6.3735 | [8 empty] | 7.8559
Warning Limits | Xbar + 2SD | [19 empty] | Warning Limits | Xbar + 2SD |  | 21.677 | [8 empty] | 22.2299 | [8 empty] | 22.1974
| Xbar - 2SD | [20 empty] | Xbar - 2SD |  | 9.5738 | [8 empty] | 17.2034 | [8 empty] | 16.1693
Control Limits | Xbar + 3SD | [19 empty] | Control Limits | Xbar + 3SD |  | 24.7028 | [8 empty] | 23.4866 | [8 empty] | 23.7044
| Xbar - 3SD | [20 empty] | Xbar - 3SD |  | 6.548 | [8 empty] | 15.9467 | [8 empty] | 14.6622
Warning Limits as % | Xbar + 2SD | [19 empty] | Warning Limits as % | Xbar + 2SD |  | 138.7293 | [8 empty] | 112.747 | [8 empty] | 115.7119
| Xbar - 2SD | [20 empty] | Xbar - 2SD |  | 61.2707 | [8 empty] | 87.253 | [8 empty] | 84.2881
Control Limits as % | Xbar + 3SD | [19 empty] | Control Limits as % | Xbar + 3SD |  | 158.0939 | [8 empty] | 119.1205 | [8 empty] | 123.5678
| Xbar - 3SD | [20 empty] | Xbar - 3SD |  | 41.9061 | [8 empty] | 80.8795 | [8 empty] | 76.4322
Coefficient of Variation % | 100* std/mean | [19 empty] | Coefficient of Variation % | 100* std/mean |  | 19.3646 | [8 empty] | 6.3735 | [8 empty] | 7.8559
3 x Coefficient of Variation % | [20 empty] | 3 x Coefficient of Variation % |  |  | 58.0939 | [8 empty] | 19.1205 | [8 empty] | 23.5678
Count | [20 empty] | Count |  |  | 48 | [8 empty] | 36 | [8 empty] | 36
99th percentile.inc | [20 empty] | 99th percentile.inc |  |  | 22.855 | [8 empty] | 22.824 | [8 empty] | 22.418
Maximum | [20 empty] | Maximum |  |  | 25 | [8 empty] | 23 | [8 empty] | 22.5
[21 empty] | Known |  |  | 19.6078 | [8 empty] | 19.6078 | [8 empty] | 19.6078
[21 empty] | MDL |  |  | 7.2872 | [8 empty] | 3.0633 | [8 empty] | 3.6737
[21 empty] | LOQ |  |  | 30.258 | [8 empty] | 12.5664 | [8 empty] | 15.0703
[21 empty] | ratio mean/mdl |  |  | 2.1442 | [8 empty] | 6.4363 | [8 empty] | 5.2218

## rolling 5 ppb
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Reduction Efficiency | Analyst Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | NH4-N NOTES: | NH4-N concentration | NH4-N- % Error in check stds | NH4-N- DQO (<=20%) | NH4-N- RPD of duplicates | NH4-N- DQO (<=20% if conc>LOQ) | NH4-N- PR of spikes | NH4-N- DQO (80-120%) if conc>LOQ | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BN PO4-P NOTES: | PO4-P concentration | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BN NO3-N NOTES: | NO3-N concentration | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2018-08-23 | OM_8-23-2018_12-13-41PM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0105 | 0.242 | 5.27 | PO4-P | ug P/L | 0.0213 | 0.342 | 5.07 | NO3-N+NO2-N | ug N/L | 0.0108 | 0.145 | 3.59 |  | 1 |  | 5.27 | [8 empty] | 5.07 | [8 empty] | 3.59 | [7 empty] | 1
2018-08-23 | OM_8-23-2018_12-13-41PM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00995 | 0.224 | 4.86 | PO4-P | ug P/L | 0.0187 | 0.295 | 3.99 | NO3-N+NO2-N | ug N/L | 0.0109 | 0.15 | 3.83 |  | 2 |  | 4.86 | [8 empty] | 3.99 | [8 empty] | 3.83 | [7 empty] | 2
2018-08-23 | OM_8-23-2018_12-13-41PM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | -0.00953 | -0.216 |  | PO4-P | ug P/L | 0.0188 | 0.286 | 3.77 | NO3-N+NO2-N | ug N/L | 0.0106 | 0.142 | 3.41 |  | 3 |  | 0 | [8 empty] | 3.77 | [8 empty] | 3.41 | [7 empty] | 3
2018-08-23 | OM_8-23-2018_12-13-41PM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0112 | 0.24 | 5.24 | PO4-P | ug P/L | 0.0199 | 0.314 | 4.41 | NO3-N+NO2-N | ug N/L | 0.0102 | 0.139 | 3.28 |  | 4 |  | 5.24 | [8 empty] | 4.41 | [8 empty] | 3.28 | [7 empty] | 4
2018-08-23 | OM_8-23-2018_12-13-41PM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0104 | 0.25 | 5.45 | PO4-P | ug P/L | 0.019 | 0.309 | 4.3 | NO3-N+NO2-N | ug N/L | 0.0104 | 0.146 | 3.63 |  | 5 |  | 5.45 | [8 empty] | 4.3 | [8 empty] | 3.63 | [7 empty] | 5
2018-08-23 | OM_8-23-2018_12-13-41PM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00996 | 0.225 | 4.88 | PO4-P | ug P/L | 0.0193 | 0.319 | 4.54 | NO3-N+NO2-N | ug N/L | 0.0126 | 0.288 |  |  | 6 |  | 4.88 | [8 empty] | 4.54 | [8 empty] | 0 | [7 empty] | 6
2018-08-23 | OM_8-23-2018_12-13-41PM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0106 | 0.215 | 4.66 | PO4-P | ug P/L | 0.0195 | 0.319 | 4.53 | NO3-N+NO2-N | ug N/L | 0.00988 | 0.126 | 2.64 |  | 7 |  | 4.66 | [8 empty] | 4.53 | [8 empty] | 2.64 | [7 empty] | 7
2018-08-23 | OM_8-23-2018_12-13-41PM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0105 | 0.26 | 5.68 | PO4-P | ug P/L | 0.0196 | 0.326 | 4.69 | NO3-N+NO2-N | ug N/L | 0.00916 | 0.127 | 2.72 |  | 8 |  | 5.68 | [8 empty] | 4.69 | [8 empty] | 2.72 | [7 empty] | 8
2018-09-11 | OM_9-11-2018_02-40-57PM.OMN | MDL at 5 ppb | S9 | 1 | 1 | NH4-N | ug N/L | 0.0239 | 0.431 | 5.99 | PO4-P | ug P/L | 0.0168 | 0.249 | 4.92 | NO3-N+NO2-N | ug N/L | 0.00993 | 0.129 | 5.1 |  | 9 |  | 5.99 | [8 empty] | 4.92 | [8 empty] | 5.1 | [7 empty] | 9
2018-09-11 | OM_9-11-2018_02-40-57PM.OMN | MDL at 5 ppb | S9 | 1 | 1 | NH4-N | ug N/L | 0.0168 | 0.334 | 3.93 | PO4-P | ug P/L | 0.0172 | 0.273 | 5.48 | NO3-N+NO2-N | ug N/L | 0.0105 | 0.176 | 7.2 |  | 10 |  | 3.93 | [8 empty] | 5.48 | [8 empty] | 7.2 | [7 empty] | 10
2018-09-11 | OM_9-11-2018_02-40-57PM.OMN | MDL at 5 ppb | S9 | 1 | 1 | NH4-N | ug N/L | 0.0239 | 0.348 | 4.23 | PO4-P | ug P/L | 0.0174 | 0.259 | 5.14 | NO3-N+NO2-N | ug N/L | 0.0113 | 0.137 | 5.47 |  | 11 |  | 4.23 | [8 empty] | 5.14 | [8 empty] | 5.47 | [7 empty] | 11
2018-09-11 | OM_9-11-2018_02-40-57PM.OMN | MDL at 5 ppb | S9 | 1 | 1 | NH4-N | ug N/L | 0.0296 | 0.445 | 6.28 | PO4-P | ug P/L | 0.0176 | 0.264 | 5.26 | NO3-N+NO2-N | ug N/L | 0.01 | 0.121 | 4.71 |  | 12 |  | 6.28 | [8 empty] | 5.26 | [8 empty] | 4.71 | [7 empty] | 12
2018-09-11 | OM_9-11-2018_02-40-57PM.OMN | MDL at 5 ppb | S9 | 1 | 1 | NH4-N | ug N/L | 0.0333 | 0.454 | 6.48 | PO4-P | ug P/L | 0.0176 | 0.271 | 5.42 | NO3-N+NO2-N | ug N/L | 0.0101 | 0.132 | 5.2 |  | 13 |  | 6.48 | [8 empty] | 5.42 | [8 empty] | 5.2 | [7 empty] | 13
2018-09-11 | OM_9-11-2018_02-40-57PM.OMN | MDL at 5 ppb | S9 | 1 | 1 | NH4-N | ug N/L | 0.0243 | 0.369 | 4.66 | PO4-P | ug P/L | 0.0179 | 0.256 | 5.07 | NO3-N+NO2-N | ug N/L | 0.0109 | 0.148 | 5.93 |  | 14 |  | 4.66 | [8 empty] | 5.07 | [8 empty] | 5.93 | [7 empty] | 14
2018-09-11 | OM_9-11-2018_02-40-57PM.OMN | MDL at 5 ppb | S9 | 1 | 1 | NH4-N | ug N/L | 0.0403 | 0.335 | 3.94 | PO4-P | ug P/L | 0.0176 | 0.255 | 5.06 | NO3-N+NO2-N | ug N/L | 0.0104 | 0.135 | 5.38 |  | 15 |  | 3.94 | [8 empty] | 5.06 | [8 empty] | 5.38 | [7 empty] | 15
2018-09-11 | OM_9-11-2018_02-40-57PM.OMN | MDL at 5 ppb | S9 | 1 | 1 | NH4-N | ug N/L | 0.024 | 0.317 | 3.57 | PO4-P | ug P/L | 0.0171 | 0.25 | 4.95 | NO3-N+NO2-N | ug N/L | 0.0108 | 0.157 | 6.36 |  | 16 |  | 3.57 | [8 empty] | 4.95 | [8 empty] | 6.36 | [7 empty] | 16
2018-09-11 | OM_9-11-2018_12-21-43PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.00609 | 0.18 | 6.1 | Phosphate | ug/L | 0.00832 | 0.115 | 3.36 | Nitrate-Nitrite | ug/L | 0.00359 | 0.0708 | 5.55 |  | 17 |  | 6.1 | [8 empty] | 3.36 | [8 empty] | 5.55 | [7 empty] | 17
2018-09-11 | OM_9-11-2018_12-21-43PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.00593 | 0.193 | 6.49 | Phosphate | ug/L | 0.00859 | 0.131 | 4 | Nitrate-Nitrite | ug/L | 0.00378 | 0.0782 | 6.08 |  | 18 |  | 6.49 | [8 empty] | 4 | [8 empty] | 6.08 | [7 empty] | 18
2018-09-11 | OM_9-11-2018_12-21-43PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.00635 | 0.173 | 5.91 | Phosphate | ug/L | 0.00775 | 0.0971 | 2.66 | Nitrate-Nitrite | ug/L | 0.00404 | 0.0856 | 6.6 |  | 19 |  | 5.91 | [8 empty] | 2.66 | [8 empty] | 6.6 | [7 empty] | 19
2018-09-11 | OM_9-11-2018_12-21-43PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.0053 | 0.18 | 6.12 | Phosphate | ug/L | 0.00714 | 0.151 | 4.78 | Nitrate-Nitrite | ug/L | 0.00331 | 0.0722 | 5.65 |  | 20 |  | 6.12 | [8 empty] | 4.78 | [8 empty] | 5.65 | [7 empty] | 20
2018-09-11 | OM_9-11-2018_12-21-43PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.00588 | 0.199 | 6.66 | Phosphate | ug/L | 0.00816 | 0.117 | 3.47 | Nitrate-Nitrite | ug/L | 0.00363 | 0.0784 | 6.09 |  | 21 |  | 6.66 | [8 empty] | 3.47 | [8 empty] | 6.09 | [7 empty] | 21
2018-09-11 | OM_9-11-2018_12-21-43PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.00506 | 0.172 | 5.87 | Phosphate | ug/L | 0.00842 | 0.098 | 2.7 | Nitrate-Nitrite | ug/L | 0.00332 | 0.0722 | 5.65 |  | 22 |  | 5.87 | [8 empty] | 2.7 | [8 empty] | 5.65 | [7 empty] | 22
2018-09-11 | OM_9-11-2018_12-21-43PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.0436 | 0.698 |  | Phosphate | ug/L | 0.00739 | 0.108 | 3.1 | Nitrate-Nitrite | ug/L | 0.00343 | 0.0697 | 5.48 |  | 23 |  | 0 | [8 empty] | 3.1 | [8 empty] | 5.48 | [7 empty] | 23
2018-09-11 | OM_9-11-2018_12-21-43PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 1.93 | 12.1 |  | Phosphate | ug/L | 3.49 | 19.9 |  | Nitrate-Nitrite | ug/L | 2.05 | 3.97 |  |  | 24 |  | 0 | [8 empty] | 0 | [8 empty] | 0 | [7 empty] | 24
2018-10-23 | OM_10-23-2018_11-16-33AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.0038 | 0.0789 | 2.57 | Phosphate | ug/L | 0.00983 | 0.176 | 5.44 | Nitrate-Nitrite | ug/L | 0.00621 | 0.0961 | 5.69 |  | 25 |  | 2.57 | [8 empty] | 5.44 | [8 empty] | 5.69 | [7 empty] | 25
2018-10-23 | OM_10-23-2018_11-16-33AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | -0.00461 | 0.063 | 2.08 | Phosphate | ug/L | 0.00842 | 0.103 | 2.8 | Nitrate-Nitrite | ug/L | 0.00604 | 0.105 | 6.09 |  | 26 |  | 2.08 | [8 empty] | 2.8 | [8 empty] | 6.09 | [7 empty] | 26
2018-10-23 | OM_10-23-2018_11-16-33AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.00578 | 0.243 | 7.65 | Phosphate | ug/L | 0.00966 | 0.175 | 5.39 | Nitrate-Nitrite | ug/L | 0.00586 | 0.105 | 6.08 |  | 27 |  | 7.65 | [8 empty] | 5.39 | [8 empty] | 6.08 | [7 empty] | 27
2018-10-23 | OM_10-23-2018_11-16-33AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.00689 | 0.185 | 5.86 | Phosphate | ug/L | 0.0088 | 0.128 | 3.7 | Nitrate-Nitrite | ug/L | 0.00586 | 0.092 | 5.49 |  | 28 |  | 5.86 | [8 empty] | 3.7 | [8 empty] | 5.49 | [7 empty] | 28
2018-10-23 | OM_10-23-2018_11-16-33AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.00598 | 0.134 | 4.27 | Phosphate | ug/L | 0.00873 | 0.12 | 3.39 | Nitrate-Nitrite | ug/L | 0.00585 | 0.091 | 5.45 |  | 29 |  | 4.27 | [8 empty] | 3.39 | [8 empty] | 5.45 | [7 empty] | 29
2018-10-23 | OM_10-23-2018_11-16-33AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.00677 | 0.198 | 6.27 | Phosphate | ug/L | 0.00826 | 0.168 | 5.13 | Nitrate-Nitrite | ug/L | 0.00572 | 0.0892 | 5.37 |  | 30 |  | 6.27 | [8 empty] | 5.13 | [8 empty] | 5.37 | [7 empty] | 30
2018-10-23 | OM_10-23-2018_11-16-33AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | 0.00611 | 0.0681 | 2.24 | Phosphate | ug/L | 0.0091 | 0.168 | 5.12 | Nitrate-Nitrite | ug/L | 0.00612 | 0.0963 | 5.7 |  | 31 |  | 2.24 | [8 empty] | 5.12 | [8 empty] | 5.7 | [7 empty] | 31
2018-10-23 | OM_10-23-2018_11-16-33AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | Ammonium | ug/L | -0.00628 | -0.0668 |  | Phosphate | ug/L | 0.00989 | 0.164 | 5 | Nitrate-Nitrite | ug/L | 0.00616 | 0.102 | 5.97 |  | 32 |  | 0 | [8 empty] | 5 | [8 empty] | 5.97 | [7 empty] | 32
2018-10-23 | OM_10-23-2018_12-15-43PM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0164 | 0.356 | 5.47 | PO4-P | ug P/L | 0.017 | 0.279 | 6.09 | NO3-N+NO2-N | ug N/L | 0.0102 | 0.14 | 6.46 |  | 33 |  | 5.47 | [8 empty] | 6.09 | [8 empty] | 6.46 | [7 empty] | 33
2018-10-23 | OM_10-23-2018_12-15-43PM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0428 | 0.344 | 5.19 | PO4-P | ug P/L | 0.0169 | 0.266 | 5.78 | NO3-N+NO2-N | ug N/L | 0.00835 | 0.117 | 5.45 |  | 34 |  | 5.19 | [8 empty] | 5.78 | [8 empty] | 5.45 | [7 empty] | 34
2018-10-23 | OM_10-23-2018_12-15-43PM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0122 | 0.235 | 2.53 | PO4-P | ug P/L | 0.0172 | 0.272 | 5.93 | NO3-N+NO2-N | ug N/L | 0.00935 | 0.117 | 5.44 |  | 35 |  | 2.53 | [8 empty] | 5.93 | [8 empty] | 5.44 | [7 empty] | 35
2018-10-23 | OM_10-23-2018_12-15-43PM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0117 | 0.178 |  | PO4-P | ug P/L | 0.0183 | 0.273 | 5.95 | NO3-N+NO2-N | ug N/L | 0.00945 | 0.123 | 5.72 |  | 36 |  | 0 | [8 empty] | 5.95 | [8 empty] | 5.72 | [7 empty] | 36
2018-10-23 | OM_10-23-2018_12-15-43PM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0205 | 0.401 | 6.58 | PO4-P | ug P/L | 0.0164 | 0.254 | 5.51 | NO3-N+NO2-N | ug N/L | 0.00948 | 0.147 | 6.76 |  | 37 |  | 6.58 | [8 empty] | 5.51 | [8 empty] | 6.76 | [7 empty] | 37
2018-10-23 | OM_10-23-2018_12-15-43PM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0171 | 0.397 | 6.48 | PO4-P | ug P/L | 0.0175 | 0.274 | 5.97 | NO3-N+NO2-N | ug N/L | 0.00985 | 0.134 | 6.21 |  | 38 |  | 6.48 | [8 empty] | 5.97 | [8 empty] | 6.21 | [7 empty] | 38
2018-10-23 | OM_10-23-2018_12-15-43PM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0153 | 0.357 | 5.52 | PO4-P | ug P/L | 0.0171 | 0.257 | 5.58 | NO3-N+NO2-N | ug N/L | 0.0101 | 0.119 | 5.52 |  | 39 |  | 5.52 | [8 empty] | 5.58 | [8 empty] | 5.52 | [7 empty] | 39
2018-10-23 | OM_10-23-2018_12-15-43PM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0122 | 0.229 | 2.41 | PO4-P | ug P/L | 0.0169 | 0.258 | 5.6 | NO3-N+NO2-N | ug N/L | 0.00934 | 0.113 | 5.27 |  | 40 |  | 2.41 | [8 empty] | 5.6 | [8 empty] | 5.27 | [7 empty] | 40
2019-01-15 | OM_1-15-2019_11-17-46AMeditedv3.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0116 | 0.269 | 3.51 | PO4-P | ug P/L | 0.0695 | 1.72 |  | NO3-N+NO2-N | ug N/L | 0.00887 | 0.141 | 4.93 |  | 41 |  | 3.51 | [8 empty] | 0 | [8 empty] | 4.93 | [7 empty] | 41
2019-01-15 | OM_1-15-2019_11-17-46AMeditedv3.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0162 | 0.351 | 5.29 | PO4-P | ug P/L | 0.0188 | 0.278 | 4.87 | NO3-N+NO2-N | ug N/L | 0.0153 | 0.202 | 7.51 |  | 42 |  | 5.29 | [8 empty] | 4.87 | [8 empty] | 7.51 | [7 empty] | 42
2019-01-15 | OM_1-15-2019_11-17-46AMeditedv3.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0139 | 0.271 | 3.56 | PO4-P | ug P/L | 0.0194 | 0.299 | 5.38 | NO3-N+NO2-N | ug N/L | 0.0114 | 0.155 | 5.51 |  | 43 |  | 3.56 | [8 empty] | 5.38 | [8 empty] | 5.51 | [7 empty] | 43
2019-01-15 | OM_1-15-2019_11-17-46AMeditedv3.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0156 | 0.334 | 4.92 | PO4-P | ug P/L | 0.02 | 0.316 | 5.77 | NO3-N+NO2-N | ug N/L | 0.0121 | 0.166 | 6.01 |  | 44 |  | 4.92 | [8 empty] | 5.77 | [8 empty] | 6.01 | [7 empty] | 44
2019-01-15 | OM_1-15-2019_11-17-46AMeditedv3.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.011 | 0.259 | 3.31 | PO4-P | ug P/L | 0.0185 | 0.308 | 5.58 | NO3-N+NO2-N | ug N/L | 0.0243 | 0.162 | 5.82 |  | 45 |  | 3.31 | [8 empty] | 5.58 | [8 empty] | 5.82 | [7 empty] | 45
2019-01-15 | OM_1-15-2019_11-17-46AMeditedv3.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0151 | 0.386 | 6.03 | PO4-P | ug P/L | 0.0191 | 0.293 | 5.21 | NO3-N+NO2-N | ug N/L | -0.0151 | 0.22 | 8.25 |  | 46 |  | 6.03 | [8 empty] | 5.21 | [8 empty] | 8.25 | [7 empty] | 46
2019-01-15 | OM_1-15-2019_11-17-46AMeditedv3.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0151 | 0.347 | 5.19 | PO4-P | ug P/L | 0.0183 | 0.285 | 5.03 | NO3-N+NO2-N | ug N/L | 0.0103 | 0.183 | 6.7 |  | 47 |  | 5.19 | [8 empty] | 5.03 | [8 empty] | 6.7 | [7 empty] | 47
2019-01-15 | OM_1-15-2019_11-17-46AMeditedv3.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0138 | 0.292 | 4 | PO4-P | ug P/L | 0.0349 | 0.316 | 5.79 | NO3-N+NO2-N | ug N/L | 0.00976 | 0.126 | 4.33 |  | 48 |  | 4 | [8 empty] | 5.79 | [8 empty] | 4.33 | [7 empty] | 48
2019-04-17 | OM_4-17-2019_12-32-12PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | [5 empty] | Phosphate | ug/L | 0.0111 | 0.229 | 5.51 | Nitrate-Nitrite | ug/L | 0.00575 | 0.102 | 4.69 |  | 49 |  | 0 | [8 empty] | 5.51 | [8 empty] | 4.69 | [7 empty] | 49
2019-04-17 | OM_4-17-2019_12-32-12PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | [5 empty] | Phosphate | ug/L | 0.0118 | 0.205 | 4.66 | Nitrate-Nitrite | ug/L | 0.00667 | 0.109 | 5.06 |  | 50 |  | 0 | [8 empty] | 4.66 | [8 empty] | 5.06 | [7 empty] | 50
2019-04-17 | OM_4-17-2019_12-32-12PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | [5 empty] | Phosphate | ug/L | 0.0127 | 0.221 | 5.22 | Nitrate-Nitrite | ug/L | 0.00629 | 0.108 | 5.02 |  | 51 |  | 0 | [8 empty] | 5.22 | [8 empty] | 5.02 | [7 empty] | 51
2019-04-17 | OM_4-17-2019_12-32-12PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | [5 empty] | Phosphate | ug/L | 0.0123 | 0.204 | 4.62 | Nitrate-Nitrite | ug/L | 0.00579 | 0.112 | 5.22 |  | 52 |  | 0 | [8 empty] | 4.62 | [8 empty] | 5.22 | [7 empty] | 52
2019-04-17 | OM_4-17-2019_12-32-12PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | [5 empty] | Phosphate | ug/L | 0.0099 | 0.157 | 2.95 | Nitrate-Nitrite | ug/L | 0.00563 | 0.104 | 4.81 |  | 53 |  | 0 | [8 empty] | 2.95 | [8 empty] | 4.81 | [7 empty] | 53
2019-04-17 | OM_4-17-2019_12-32-12PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | [5 empty] | Phosphate | ug/L | 0.0119 | 0.205 | 4.68 | Nitrate-Nitrite | ug/L | 0.00568 | 0.112 | 5.22 |  | 54 |  | 0 | [8 empty] | 4.68 | [8 empty] | 5.22 | [7 empty] | 54
2019-04-17 | OM_4-17-2019_12-32-12PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | [5 empty] | Phosphate | ug/L | 0.0113 | 0.207 | 4.74 | Nitrate-Nitrite | ug/L | 0.00593 | 0.107 | 4.97 |  | 55 |  | 0 | [8 empty] | 4.74 | [8 empty] | 4.97 | [7 empty] | 55
2019-04-17 | OM_4-17-2019_12-32-12PMBRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | [5 empty] | Phosphate | ug/L | 0.0116 | 0.183 | 3.88 | Nitrate-Nitrite | ug/L | 0.00718 | 0.109 | 5.04 |  | 56 |  | 0 | [8 empty] | 3.88 | [8 empty] | 5.04 | [7 empty] | 56
2019-08-22 | OM_8-22-2019_02-16-33PMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0148 | 0.344 | 4.48 | PO4-P | ug P/L | 0.02 | 0.317 | 4.2 | NO3-N+NO2-N | ug N/L | 0.00886 | 0.121 | 3.52 |  | 57 |  | 4.48 | [8 empty] | 4.2 | [8 empty] | 3.52 | [7 empty] | 57
2019-08-22 | OM_8-22-2019_02-16-33PMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0623 | 0.479 | 6.87 | PO4-P | ug P/L | 0.0212 | 0.347 | 4.87 | NO3-N+NO2-N | ug N/L | 0.0101 | 0.181 | 6.41 |  | 58 |  | 6.87 | [8 empty] | 4.87 | [8 empty] | 6.41 | [7 empty] | 58
2019-08-22 | OM_8-22-2019_02-16-33PMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.201 | 1.5 |  | PO4-P | ug P/L | 0.0217 | 0.355 | 5.05 | NO3-N+NO2-N | ug N/L | 0.00957 | 0.151 | 4.96 |  | 59 |  | 0 | [8 empty] | 5.05 | [8 empty] | 4.96 | [7 empty] | 59
2019-08-22 | OM_8-22-2019_02-16-33PMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | -0.206 | -6.56 |  | PO4-P | ug P/L | 0.022 | 0.381 | 5.67 | NO3-N+NO2-N | ug N/L | 0.0105 | 0.144 | 4.64 |  | 60 |  | 0 | [8 empty] | 5.67 | [8 empty] | 4.64 | [7 empty] | 60
2019-08-22 | OM_8-22-2019_02-16-33PMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0118 | 0.297 | 3.63 | PO4-P | ug P/L | 0.0227 | 0.373 | 5.48 | NO3-N+NO2-N | ug N/L | 0.0105 | 0.15 | 4.95 |  | 61 |  | 3.63 | [8 empty] | 5.48 | [8 empty] | 4.95 | [7 empty] | 61
2019-08-22 | OM_8-22-2019_02-16-33PMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0128 | 0.117 | 0.444 | PO4-P | ug P/L | 0.0221 | 0.385 | 5.74 | NO3-N+NO2-N | ug N/L | 0.0136 | 0.227 | 8.65 |  | 62 |  | 0.444 | [8 empty] | 5.74 | [8 empty] | 8.65 | [7 empty] | 62
2019-08-22 | OM_8-22-2019_02-16-33PMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.057 | 1.37 |  | PO4-P | ug P/L | 0.0197 | 0.328 | 4.43 | NO3-N+NO2-N | ug N/L | 0.00896 | 0.136 | 4.24 |  | 63 |  | 0 | [8 empty] | 4.43 | [8 empty] | 4.24 | [7 empty] | 63
2019-08-22 | OM_8-22-2019_02-16-33PMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.244 | 1.51 |  | PO4-P | ug P/L | 0.0194 | 0.338 | 4.66 | NO3-N+NO2-N | ug N/L | 0.00785 | 0.131 | 4 |  | 64 |  | 0 | [8 empty] | 4.66 | [8 empty] | 4 | [7 empty] | 64
2019-08-22 | OM_8-22-2019_02-16-33PMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00866 | 0.219 | 2.25 | PO4-P | ug P/L | 0.0175 | 0.319 | 4.24 | NO3-N+NO2-N | ug N/L | 0.00671 | 0.11 | 3 |  | 65 |  | 2.25 | [8 empty] | 4.24 | [8 empty] | 3 | [7 empty] | 65
2019-09-23 | OM_9-23-2019_12-09-25PMedited.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0108 | 0.248 | 6.57 | PO4-P | ug P/L | 0.0204 | 0.316 | 4.46 | NO3-N+NO2-N | ug N/L | 0.0081 | 0.11 | 2.97 |  | 66 |  | 6.57 | [8 empty] | 4.46 | [8 empty] | 2.97 | [7 empty] | 66
2019-09-23 | OM_9-23-2019_12-09-25PMedited.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.016 | 0.316 | 8.15 | PO4-P | ug P/L | 0.0222 | 0.375 | 6.07 | NO3-N+NO2-N | ug N/L | 0.00825 | 0.0982 | 2.39 |  | 67 |  | 8.15 | [8 empty] | 6.07 | [8 empty] | 2.39 | [7 empty] | 67
2019-09-23 | OM_9-23-2019_12-09-25PMedited.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0163 | 0.335 | 8.58 | PO4-P | ug P/L | 0.0212 | 0.344 | 5.22 | NO3-N+NO2-N | ug N/L | 0.00883 | 0.141 | 4.51 |  | 68 |  | 8.58 | [8 empty] | 5.22 | [8 empty] | 4.51 | [7 empty] | 68
2019-09-23 | OM_9-23-2019_12-09-25PMedited.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0117 | 0.284 | 7.39 | PO4-P | ug P/L | 0.0215 | 0.368 | 5.89 | NO3-N+NO2-N | ug N/L | 0.0106 | 0.144 | 4.66 |  | 69 |  | 7.39 | [8 empty] | 5.89 | [8 empty] | 4.66 | [7 empty] | 69
2019-09-23 | OM_9-23-2019_12-09-25PMedited.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0265 | 0.0893 |  | PO4-P | ug P/L | 0.022 | 0.36 | 5.68 | NO3-N+NO2-N | ug N/L | 0.00998 | 0.148 | 4.84 |  | 70 |  | 0 | [8 empty] | 5.68 | [8 empty] | 4.84 | [7 empty] | 70
2019-09-23 | OM_9-23-2019_12-09-25PMedited.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0116 | 0.321 | 8.25 | PO4-P | ug P/L | 0.0214 | 0.343 | 5.21 | NO3-N+NO2-N | ug N/L | 0.00925 | 0.147 | 4.8 |  | 71 |  | 8.25 | [8 empty] | 5.21 | [8 empty] | 4.8 | [7 empty] | 71
2019-09-23 | OM_9-23-2019_12-09-25PMedited.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0108 | 0.214 | 5.78 | PO4-P | ug P/L | 0.0221 | 0.34 | 5.13 | NO3-N+NO2-N | ug N/L | 0.00892 | 0.0882 | 1.9 |  | 72 |  | 5.78 | [8 empty] | 5.13 | [8 empty] | 1.9 | [7 empty] | 72
2019-09-23 | OM_9-23-2019_12-09-25PMedited.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0105 | 0.241 | 6.4 | PO4-P | ug P/L | 0.0224 | 0.333 | 4.93 | NO3-N+NO2-N | ug N/L | 0.00778 | 0.107 | 2.82 |  | 73 |  | 6.4 | [8 empty] | 4.93 | [8 empty] | 2.82 | [7 empty] | 73
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00954 | 0.265 | 4.84 | PO4-P | ug P/L | 0.0209 | 0.321 | 4.45 | NO3-N+NO2-N | ug N/L | 0.00867 | 0.126 | 4.04 |  | 74 |  | 4.84 | [8 empty] | 4.45 | [8 empty] | 4.04 | [7 empty] | 74
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0128 | 0.273 | 5.03 | PO4-P | ug P/L | 0.023 | 0.372 | 5.85 | NO3-N+NO2-N | ug N/L | 0.0101 | 0.135 | 4.47 |  | 75 |  | 5.03 | [8 empty] | 5.85 | [8 empty] | 4.47 | [7 empty] | 75
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00979 | 0.253 | 4.56 | PO4-P | ug P/L | 0.0225 | 0.406 | 6.77 | NO3-N+NO2-N | ug N/L | 0.00711 | 0.0718 | 1.39 |  | 76 |  | 4.56 | [8 empty] | 6.77 | [8 empty] | 1.39 | [7 empty] | 76
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00958 | 0.236 | 4.16 | PO4-P | ug P/L | 0.0219 | 0.377 | 5.96 | NO3-N+NO2-N | ug N/L | 0.0122 | 0.205 | 7.92 |  | 77 |  | 4.16 | [8 empty] | 5.96 | [8 empty] | 7.92 | [7 empty] | 77
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0244 | 0.347 | 6.72 | PO4-P | ug P/L | 0.0212 | 0.368 | 5.73 | NO3-N+NO2-N | ug N/L | 0.0107 | 0.122 | 3.84 |  | 78 |  | 6.72 | [8 empty] | 5.73 | [8 empty] | 3.84 | [7 empty] | 78
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0145 | 0.397 | 7.86 | PO4-P | ug P/L | 0.0211 | 0.326 | 4.59 | NO3-N+NO2-N | ug N/L | 0.0104 | 0.0974 | 2.64 |  | 79 |  | 7.86 | [8 empty] | 4.59 | [8 empty] | 2.64 | [7 empty] | 79
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0102 | 0.242 | 4.3 | PO4-P | ug P/L | 0.0194 | 0.324 | 4.54 | NO3-N+NO2-N | ug N/L | 0.00694 | 0.0877 | 2.17 |  | 80 |  | 4.3 | [8 empty] | 4.54 | [8 empty] | 2.17 | [7 empty] | 80
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0103 | 0.246 | 4.39 | PO4-P | ug P/L | 0.0209 | 0.325 | 4.55 | NO3-N+NO2-N | ug N/L | 0.0084 | 0.158 | 5.6 |  | 81 |  | 4.39 | [8 empty] | 4.55 | [8 empty] | 5.6 | [7 empty] | 81
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0108 | 0.345 | 6.68 | PO4-P | ug P/L | 0.0205 | 0.324 | 4.53 | NO3-N+NO2-N | ug N/L | 0.0139 | 0.16 | 5.71 |  | 82 |  | 6.68 | [8 empty] | 4.53 | [8 empty] | 5.71 | [7 empty] | 82
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00875 | 0.164 | 5.11 | PO4-P | ug P/L | 0.019 | 0.295 | 4.69 | NO3-N+NO2-N | ug N/L | 0.00885 | 0.13 | 7.84 |  | 83 |  | 5.11 | [8 empty] | 4.69 | [8 empty] | 7.84 | [7 empty] | 83
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00884 | 0.186 | 5.89 | PO4-P | ug P/L | 0.0189 | 0.271 | 3.85 | NO3-N+NO2-N | ug N/L | 0.00824 | 0.0975 | 6.03 |  | 84 |  | 5.89 | [8 empty] | 3.85 | [8 empty] | 6.03 | [7 empty] | 84
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.01 | 0.215 | 6.9 | PO4-P | ug P/L | 0.0186 | 0.285 | 4.33 | NO3-N+NO2-N | ug N/L | 0.00827 | 0.133 | 8.03 |  | 85 |  | 6.9 | [8 empty] | 4.33 | [8 empty] | 8.03 | [7 empty] | 85
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00898 | 0.163 | 5.08 | PO4-P | ug P/L | 0.0194 | 0.293 | 4.62 | NO3-N+NO2-N | ug N/L | 0.00901 | 0.114 | 6.99 |  | 86 |  | 5.08 | [8 empty] | 4.62 | [8 empty] | 6.99 | [7 empty] | 86
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0068 | 0.122 | 3.64 | PO4-P | ug P/L | 0.0202 | 0.308 | 5.15 | NO3-N+NO2-N | ug N/L | 0.0107 | 0.132 | 7.98 |  | 87 |  | 3.64 | [8 empty] | 5.15 | [8 empty] | 7.98 | [7 empty] | 87
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00884 | 0.187 | 5.9 | PO4-P | ug P/L | 0.0203 | 0.304 | 5.01 | NO3-N+NO2-N | ug N/L | 0.00898 | 0.117 | 7.12 |  | 88 |  | 5.9 | [8 empty] | 5.01 | [8 empty] | 7.12 | [7 empty] | 88
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00661 | 0.0969 | 2.74 | PO4-P | ug P/L | 0.0196 | 0.292 | 4.56 | NO3-N+NO2-N | ug N/L | 0.00992 | 0.077 | 4.87 |  | 89 |  | 2.74 | [8 empty] | 4.56 | [8 empty] | 4.87 | [7 empty] | 89
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00719 | 0.109 | 3.17 | PO4-P | ug P/L | 0.0192 | 0.279 | 4.11 | NO3-N+NO2-N | ug N/L | 0.00989 | 0.122 | 7.42 |  | 90 |  | 3.17 | [8 empty] | 4.11 | [8 empty] | 7.42 | [7 empty] | 90
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00901 | 0.16 | 4.98 | PO4-P | ug P/L | 0.0188 | 0.276 | 4.01 | NO3-N+NO2-N | ug N/L | 0.0146 | 0.124 | 7.54 |  | 91 |  | 4.98 | [8 empty] | 4.01 | [8 empty] | 7.54 | [7 empty] | 91
2019-10-21 | OM_10-21-2019_10-10-52AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.012 | 0.317 | 6.1 | PO4-P | ug P/L | 0.0167 | 0.282 | 4.62 | NO3-N+NO2-N | ug N/L | 0.00928 | 0.143 | 6.2 |  | 92 |  | 6.1 | [8 empty] | 4.62 | [8 empty] | 6.2 | [7 empty] | 92
2019-10-21 | OM_10-21-2019_10-10-52AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0159 | 0.415 | 7.91 | PO4-P | ug P/L | 0.0153 | 0.262 | 3.93 | NO3-N+NO2-N | ug N/L | 0.0108 | 0.162 | 7.07 |  | 93 |  | 7.91 | [8 empty] | 3.93 | [8 empty] | 7.07 | [7 empty] | 93
2019-10-21 | OM_10-21-2019_10-10-52AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0122 | 0.432 | 8.23 | PO4-P | ug P/L | 0.017 | 0.273 | 4.32 | NO3-N+NO2-N | ug N/L | 0.00867 | 0.13 | 5.58 |  | 94 |  | 8.23 | [8 empty] | 4.32 | [8 empty] | 5.58 | [7 empty] | 94
2019-10-21 | OM_10-21-2019_10-10-52AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0136 | 0.209 | 4.09 | PO4-P | ug P/L | 0.0157 | 0.262 | 3.96 | NO3-N+NO2-N | ug N/L | 0.00837 | 0.144 | 6.24 |  | 95 |  | 4.09 | [8 empty] | 3.96 | [8 empty] | 6.24 | [7 empty] | 95
2019-10-21 | OM_10-21-2019_10-10-52AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00983 | 0.209 | 4.09 | PO4-P | ug P/L | 0.0152 | 0.252 | 3.57 | NO3-N+NO2-N | ug N/L | 0.00886 | 0.139 | 6 |  | 96 |  | 4.09 | [8 empty] | 3.57 | [8 empty] | 6 | [7 empty] | 96
2019-10-21 | OM_10-21-2019_10-10-52AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0102 | 0.256 | 4.96 | PO4-P | ug P/L | 0.0162 | 0.275 | 4.41 | NO3-N+NO2-N | ug N/L | 0.00856 | 0.126 | 5.43 |  | 97 |  | 4.96 | [8 empty] | 4.41 | [8 empty] | 5.43 | [7 empty] | 97
2019-10-21 | OM_10-21-2019_10-10-52AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0468 | 0.364 | 6.97 | PO4-P | ug P/L | 0.0145 | 0.235 | 3 | NO3-N+NO2-N | ug N/L | 0.00945 | 0.141 | 6.11 |  | 98 |  | 6.97 | [8 empty] | 3 | [8 empty] | 6.11 | [7 empty] | 98
2019-10-21 | OM_10-21-2019_10-10-52AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0093 | 0.24 | 4.66 | PO4-P | ug P/L | 0.0163 | 0.276 | 4.42 | NO3-N+NO2-N | ug N/L | 0.0103 | 0.127 | 5.48 |  | 99 |  | 4.66 | [8 empty] | 4.42 | [8 empty] | 5.48 | [7 empty] | 99
2019-10-21 | OM_10-21-2019_10-10-52AM_BRNedit.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0245 | 0.348 | 6.67 | PO4-P | ug P/L | 0.0168 | 0.257 | 3.76 | NO3-N+NO2-N | ug N/L | 0.00997 | 0.148 | 6.4 |  | 100 |  | 6.67 | [8 empty] | 3.76 | [8 empty] | 6.4 | [7 empty] | 100
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00968 | 0.239 | 5.82 | PO4-P | ug P/L | 0.0183 | 0.318 | 5.4 | NO3-N+NO2-N | ug N/L | 0.00736 | 0.121 | 4.72 |  | 101 |  | 5.82 | [8 empty] | 5.4 | [8 empty] | 4.72 | [7 empty] | 101
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00951 | 0.108 |  | PO4-P | ug P/L | 0.0175 | 0.291 | 4.59 | NO3-N+NO2-N | ug N/L | 0.00725 | 0.117 | 4.5 |  | 102 |  | 0 | [8 empty] | 4.59 | [8 empty] | 4.5 | [7 empty] | 102
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.01 | 0.195 | 4.78 | PO4-P | ug P/L | 0.0175 | 0.29 | 4.56 | NO3-N+NO2-N | ug N/L | 0.00684 | 0.11 | 4.12 |  | 103 |  | 4.78 | [8 empty] | 4.56 | [8 empty] | 4.12 | [7 empty] | 103
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00711 | 0.172 | 4.24 | PO4-P | ug P/L | 0.0176 | 0.279 | 4.24 | NO3-N+NO2-N | ug N/L | 0.00785 | 0.117 | 4.47 |  | 104 |  | 4.24 | [8 empty] | 4.24 | [8 empty] | 4.47 | [7 empty] | 104
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | -0.00946 | 0.123 | 3.07 | PO4-P | ug P/L | 0.0191 | 0.307 | 5.07 | NO3-N+NO2-N | ug N/L | 0.00765 | 0.112 | 4.25 |  | 105 |  | 3.07 | [8 empty] | 5.07 | [8 empty] | 4.25 | [7 empty] | 105
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.104 | 0.799 |  | PO4-P | ug P/L | 0.0192 | 0.326 | 5.61 | NO3-N+NO2-N | ug N/L | 0.00779 | 0.133 | 5.32 |  | 106 |  | 0 | [8 empty] | 5.61 | [8 empty] | 5.32 | [7 empty] | 106
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0135 | 0.185 | 4.54 | PO4-P | ug P/L | 0.0196 | 0.326 | 5.63 | NO3-N+NO2-N | ug N/L | 0.00884 | 0.131 | 5.2 |  | 107 |  | 4.54 | [8 empty] | 5.63 | [8 empty] | 5.2 | [7 empty] | 107
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00786 | 0.122 | 3.04 | PO4-P | ug P/L | 0.0188 | 0.333 | 5.83 | NO3-N+NO2-N | ug N/L | 0.00793 | 0.129 | 5.12 |  | 108 |  | 3.04 | [8 empty] | 5.83 | [8 empty] | 5.12 | [7 empty] | 108
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0119 | 0.247 | 5.01 | PO4-P | ug P/L | 0.0201 | 0.302 | 3.8 | NO3-N+NO2-N | ug N/L | 0.00949 | 0.123 | 5.24 |  | 109 |  | 5.01 | [8 empty] | 3.8 | [8 empty] | 5.24 | [7 empty] | 109
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0103 | 0.271 | 5.56 | PO4-P | ug P/L | 0.0235 | 0.425 | 7.26 | NO3-N+NO2-N | ug N/L | 0.0089 | 0.127 | 5.46 |  | 110 |  | 5.56 | [8 empty] | 7.26 | [8 empty] | 5.46 | [7 empty] | 110
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0102 | 0.237 | 4.81 | PO4-P | ug P/L | 0.0217 | 0.367 | 5.63 | NO3-N+NO2-N | ug N/L | 0.00914 | 0.141 | 6.14 |  | 111 |  | 4.81 | [8 empty] | 5.63 | [8 empty] | 6.14 | [7 empty] | 111
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0108 | 0.208 | 4.15 | PO4-P | ug P/L | 0.0161 | 0.272 | 2.97 | NO3-N+NO2-N | ug N/L | 0.00856 | 0.137 | 5.94 |  | 112 |  | 4.15 | [8 empty] | 2.97 | [8 empty] | 5.94 | [7 empty] | 112
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0139 | 0.287 | 5.92 | PO4-P | ug P/L | 0.021 | 0.34 | 4.88 | NO3-N+NO2-N | ug N/L | 0.00984 | 0.142 | 6.22 |  | 113 |  | 5.92 | [8 empty] | 4.88 | [8 empty] | 6.22 | [7 empty] | 113
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00776 | 0.185 | 3.63 | PO4-P | ug P/L | 0.0199 | 0.303 | 3.83 | NO3-N+NO2-N | ug N/L | 0.00925 | 0.133 | 5.73 |  | 114 |  | 3.63 | [8 empty] | 3.83 | [8 empty] | 5.73 | [7 empty] | 114
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0126 | 0.296 | 6.12 | PO4-P | ug P/L | 0.02 | 0.332 | 4.65 | NO3-N+NO2-N | ug N/L | 0.00759 | 0.106 | 4.42 |  | 115 |  | 6.12 | [8 empty] | 4.65 | [8 empty] | 4.42 | [7 empty] | 115
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | MDL at 5 ppb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0101 | 0.237 | 4.8 | PO4-P | ug P/L | 0.0206 | 0.38 | 6.01 | NO3-N+NO2-N | ug N/L | 0.00842 | 0.116 | 4.89 |  | 116 |  | 4.8 | [8 empty] | 6.01 | [8 empty] | 4.89 | [7 empty] | 116
2020-03-18 | OM_3-18-2020_09-56-27AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.012 | 0.262 | 4.11 | PO4-P | ug P/L | 0.0228 | 0.373 | 6.83 | NO3-N+NO2-N | ug N/L | 0.00622 | 0.0816 | 3.08 |  | 117 |  | 4.11 | [8 empty] | 6.83 | [8 empty] | 3.08 | [7 empty] | 117
2020-03-18 | OM_3-18-2020_09-56-27AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0137 | 0.316 | 5.18 | PO4-P | ug P/L | 0.0256 | 0.431 | 8.45 | NO3-N+NO2-N | ug N/L | 0.00627 | 0.127 | 5.86 |  | 118 |  | 5.18 | [8 empty] | 8.45 | [8 empty] | 5.86 | [7 empty] | 118
2020-03-18 | OM_3-18-2020_09-56-27AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0152 | 0.353 | 5.91 | PO4-P | ug P/L | 0.027 | 0.461 | 9.29 | NO3-N+NO2-N | ug N/L | 0.00563 | 0.103 | 4.4 |  | 119 |  | 5.91 | [8 empty] | 9.29 | [8 empty] | 4.4 | [7 empty] | 119
2020-03-18 | OM_3-18-2020_09-56-27AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0109 | 0.264 | 4.15 | PO4-P | ug P/L | 0.0254 | 0.498 | 10.3 | NO3-N+NO2-N | ug N/L | 0.00608 | 0.109 | 4.79 |  | 120 |  | 4.15 | [8 empty] | 10.3 | [8 empty] | 4.79 | [7 empty] | 120
2020-03-18 | OM_3-18-2020_09-56-27AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0191 | 0.326 | 5.36 | PO4-P | ug P/L | 0.0222 | 0.415 | 8.01 | NO3-N+NO2-N | ug N/L | 0.00677 | 0.102 | 4.31 |  | 121 |  | 5.36 | [8 empty] | 8.01 | [8 empty] | 4.31 | [7 empty] | 121
2020-03-18 | OM_3-18-2020_09-56-27AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00928 | 0.241 | 3.7 | PO4-P | ug P/L | 0.0198 | 0.373 | 6.84 | NO3-N+NO2-N | ug N/L | 0.00713 | 0.142 | 6.79 |  | 122 |  | 3.7 | [8 empty] | 6.84 | [8 empty] | 6.79 | [7 empty] | 122
2020-03-18 | OM_3-18-2020_09-56-27AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0202 | 0.156 | 2.01 | PO4-P | ug P/L | 0.0202 | 0.398 | 7.53 | NO3-N+NO2-N | ug N/L | 0.0111 | 0.128 | 5.95 |  | 123 |  | 2.01 | [8 empty] | 7.53 | [8 empty] | 5.95 | [7 empty] | 123
2020-03-18 | OM_3-18-2020_09-56-27AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0103 | 0.243 | 3.74 | PO4-P | ug P/L | 0.0192 | 0.365 | 6.63 | NO3-N+NO2-N | ug N/L | 0.00558 | 0.0957 | 3.95 |  | 124 |  | 3.74 | [8 empty] | 6.63 | [8 empty] | 3.95 | [7 empty] | 124
2020-09-01 | OM_9-1-2020_10-10-28AMedit2.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00954 | 0.197 | 4.18 | PO4-P | ug P/L | 0.0209 | 0.376 | 4.59 | NO3-N+NO2-N | ug N/L | 0.00576 | 0.0964 | 4.16 |  | 125 |  | 4.18 | [8 empty] | 4.59 | [8 empty] | 4.16 | [7 empty] | 125
2020-09-01 | OM_9-1-2020_10-10-28AMedit2.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00878 | 0.194 | 4.1 | PO4-P | ug P/L | 0.0192 | 0.333 | 3.46 | NO3-N+NO2-N | ug N/L | 0.00676 | 0.134 | 6.42 |  | 126 |  | 4.1 | [8 empty] | 3.46 | [8 empty] | 6.42 | [7 empty] | 126
2020-09-01 | OM_9-1-2020_10-10-28AMedit2.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.01 | 0.195 | 4.12 | PO4-P | ug P/L | 0.0208 | 0.371 | 4.45 | NO3-N+NO2-N | ug N/L | 0.00586 | 0.103 | 4.56 |  | 127 |  | 4.12 | [8 empty] | 4.45 | [8 empty] | 4.56 | [7 empty] | 127
2020-09-01 | OM_9-1-2020_10-10-28AMedit2.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.01 | 0.251 | 5.37 | PO4-P | ug P/L | 0.0218 | 0.37 | 4.43 | NO3-N+NO2-N | ug N/L | 0.00526 | 0.093 | 3.96 |  | 128 |  | 5.37 | [8 empty] | 4.43 | [8 empty] | 3.96 | [7 empty] | 128
2020-09-01 | OM_9-1-2020_10-10-28AMedit2.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.00587 | 0.0775 | 1.52 | PO4-P | ug P/L | 0.0209 | 0.375 | 4.57 | NO3-N+NO2-N | ug N/L | 0.00554 | 0.123 | 5.73 |  | 129 |  | 1.52 | [8 empty] | 4.57 | [8 empty] | 5.73 | [7 empty] | 129
2020-09-01 | OM_9-1-2020_10-10-28AMedit2.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0103 | 0.199 | 4.23 | PO4-P | ug P/L | 0.0234 | 0.381 | 4.71 | NO3-N+NO2-N | ug N/L | 0.00573 | 0.0955 | 4.11 |  | 130 |  | 4.23 | [8 empty] | 4.71 | [8 empty] | 4.11 | [7 empty] | 130
2020-09-01 | OM_9-1-2020_10-10-28AMedit2.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0185 | 0.244 | 5.23 | PO4-P | ug P/L | 0.0217 | 0.404 | 5.3 | NO3-N+NO2-N | ug N/L | 0.00616 | 0.117 | 5.4 |  | 131 |  | 5.23 | [8 empty] | 5.3 | [8 empty] | 5.4 | [7 empty] | 131
2020-09-01 | OM_9-1-2020_10-10-28AMedit2.omn | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0269 | 0.283 | 6.08 | PO4-P | ug P/L | 0.0201 | 0.356 | 4.06 | NO3-N+NO2-N | ug N/L | 0.00581 | 0.12 | 5.57 |  | 132 |  | 6.08 | [8 empty] | 4.06 | [8 empty] | 5.57 | [7 empty] | 132
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0102 | 0.232 | 4.44 | PO4-P | ug P/L | 0.021 | 0.344 | 4.43 | NO3-N+NO2-N | ug N/L | 0.00514 | 0.0831 | 3.81 |  | 133 |  | 4.44 | [8 empty] | 4.43 | [8 empty] | 3.81 | [7 empty] | 133
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0103 | 0.256 | 4.9 | PO4-P | ug P/L | 0.0224 | 0.377 | 5.3 | NO3-N+NO2-N | ug N/L | 0.00612 | 0.11 | 5.39 |  | 134 |  | 4.9 | [8 empty] | 5.3 | [8 empty] | 5.39 | [7 empty] | 134
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0105 | 0.165 | 3.16 | PO4-P | ug P/L | 0.0208 | 0.381 | 5.39 | NO3-N+NO2-N | ug N/L | 0.0064 | 0.101 | 4.84 |  | 135 |  | 3.16 | [8 empty] | 5.39 | [8 empty] | 4.84 | [7 empty] | 135
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.04 | 0.477 | 9.2 | PO4-P | ug P/L | 0.0217 | 0.382 | 5.42 | NO3-N+NO2-N | ug N/L | 0.00671 | 0.105 | 5.09 |  | 136 |  | 9.2 | [8 empty] | 5.42 | [8 empty] | 5.09 | [7 empty] | 136
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.01 | 0.234 | 4.49 | PO4-P | ug P/L | 0.0206 | 0.34 | 4.32 | NO3-N+NO2-N | ug N/L | 0.00654 | 0.122 | 6.07 |  | 137 |  | 4.49 | [8 empty] | 4.32 | [8 empty] | 6.07 | [7 empty] | 137
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0111 | 0.26 | 4.99 | PO4-P | ug P/L | 0.0191 | 0.335 | 4.2 | NO3-N+NO2-N | ug N/L | 0.00532 | 0.0895 | 4.18 |  | 138 |  | 4.99 | [8 empty] | 4.2 | [8 empty] | 4.18 | [7 empty] | 138
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | 0.0177 | 0.59 | 11.4 | PO4-P | ug P/L | 0.0195 | 0.331 | 4.08 | NO3-N+NO2-N | ug N/L | 0.00594 | 0.104 | 5 |  | 139 |  | 11.4 | [8 empty] | 4.08 | [8 empty] | 5 | [7 empty] | 139
2020-09-15 | OM_9-15-2020_09-24-32AM.OMN | MDL at 5 ppb | S4 | 1 | 1 | NH4-N | ug N/L | -0.0224 | -0.183 |  | PO4-P | ug P/L | 0.0205 | 0.338 | 4.26 | NO3-N+NO2-N | ug N/L | 0.00584 | 0.0902 | 4.22 |  | 140 |  | 0 | [8 empty] | 4.26 | [8 empty] | 4.22 | [7 empty] | 140
2020-11-12 | OM_11-12-2020_09-49-28AM.OMN | chk 5 | S4 | 1 | 1 | NH4-N | ug N/L | -0.0156 | 0.104 | 3.68 | PO4-P | ug P/L | 0.021 | 0.381 | 6.41 | NO3-N | ug N/L | 0.0188 | 0.226 | 6.35 |  | 141 |  | 3.68 | [8 empty] | 6.41 | [8 empty] | 6.35 | [7 empty] | 141
2020-11-12 | OM_11-12-2020_09-49-28AM.OMN | chk 5 | S4 | 1 | 1 | NH4-N | ug N/L | 0.00958 | 0.197 | 5.69 | PO4-P | ug P/L | 0.0217 | 0.391 | 6.62 | NO3-N | ug N/L | 0.0203 | 0.242 | 6.81 |  | 142 |  | 5.69 | [8 empty] | 6.62 | [8 empty] | 6.81 | [7 empty] | 142
2020-11-12 | OM_11-12-2020_09-49-28AM.OMN | chk 5 | S4 | 1 | 1 | NH4-N | ug N/L | 0.00833 | 0.274 | 7.36 | PO4-P | ug P/L | 0.0208 | 0.387 | 6.53 | NO3-N | ug N/L | 0.0197 | 0.251 | 7.07 |  | 143 |  | 7.36 | [8 empty] | 6.53 | [8 empty] | 7.07 | [7 empty] | 143
2020-11-12 | OM_11-12-2020_09-49-28AM.OMN | chk 5 | S4 | 1 | 1 | NH4-N | ug N/L | 0.0138 | 0.23 | 6.4 | PO4-P | ug P/L | 0.0197 | 0.382 | 6.42 | NO3-N | ug N/L | 0.019 | 0.281 | 7.92 |  | 144 |  | 6.4 | [8 empty] | 6.42 | [8 empty] | 7.92 | [7 empty] | 144
2020-11-12 | OM_11-12-2020_09-49-28AM.OMN | chk 5 | S4 | 1 | 1 | NH4-N | ug N/L | 0.01 | 0.23 | 6.41 | PO4-P | ug P/L | 0.0205 | 0.403 | 6.86 | NO3-N | ug N/L | 0.0175 | 0.24 | 6.76 |  | 145 |  | 6.41 | [8 empty] | 6.86 | [8 empty] | 6.76 | [7 empty] | 145
2020-11-12 | OM_11-12-2020_09-49-28AM.OMN | chk 5 | S4 | 1 | 1 | NH4-N | ug N/L | 0.00942 | 0.178 | 5.28 | PO4-P | ug P/L | 0.0147 | 0.3 | 4.65 | NO3-N | ug N/L | 0.0152 | 0.229 | 6.43 |  | 146 |  | 5.28 | [8 empty] | 4.65 | [8 empty] | 6.43 | [7 empty] | 146
2020-11-12 | OM_11-12-2020_09-49-28AM.OMN | chk 5 | S4 | 1 | 1 | NH4-N | ug N/L | 0.0134 | 0.353 | 9.07 | PO4-P | ug P/L | 0.0203 | 0.422 | 7.29 | NO3-N | ug N/L | 0.0148 | 0.243 | 6.83 |  | 147 |  | 9.07 | [8 empty] | 7.29 | [8 empty] | 6.83 | [7 empty] | 147
2020-11-12 | OM_11-12-2020_09-49-28AM.OMN | chk 5 | S4 | 1 | 1 | NH4-N | ug N/L | 0.00898 | 0.23 | 6.4 | PO4-P | ug P/L | 0.0214 | 0.23 | 3.12 | NO3-N | ug N/L | 0.0161 | 0.239 | 6.74 |  | 148 |  | 6.4 | [8 empty] | 3.12 | [8 empty] | 6.74 | [7 empty] | 148
2021-06-16 | OM_6-16-2021_09-50-13AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0083 | 0.204 | 3.96 | PO4-P | ug P/L | 0.02 | 0.292 | 2.07 | NO3-N | ug N/L | 0.00938 | 0.116 | 3.89 |  | 149 |  | 3.96 | [8 empty] | 2.07 | [8 empty] | 3.89 | [7 empty] | 149
2021-06-16 | OM_6-16-2021_09-50-13AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | -0.025 | -0.169 |  | PO4-P | ug P/L | 0.0507 | 0.639 | 9.77 | NO3-N | ug N/L | 0.0079 | 0.112 | 3.68 |  | 150 |  | 0 | [8 empty] | 9.77 | [8 empty] | 3.68 | [7 empty] | 150
2021-06-16 | OM_6-16-2021_09-50-13AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00968 | 0.21 | 4.08 | PO4-P | ug P/L | 0.0549 | 0.676 | 10.6 | NO3-N | ug N/L | 0.00846 | 0.125 | 4.36 |  | 151 |  | 4.08 | [8 empty] | 10.6 | [8 empty] | 4.36 | [7 empty] | 151
2021-06-16 | OM_6-16-2021_09-50-13AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00871 | 0.223 | 4.35 | PO4-P | ug P/L | 0.0201 | 0.402 | 4.52 | NO3-N | ug N/L | 0.00812 | 0.12 | 4.12 |  | 152 |  | 4.35 | [8 empty] | 4.52 | [8 empty] | 4.12 | [7 empty] | 152
2021-06-16 | OM_6-16-2021_09-50-13AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00789 | 0.242 | 4.73 | PO4-P | ug P/L | 0.0207 | 0.406 | 4.59 | NO3-N | ug N/L | 0.00831 | 0.109 | 3.5 |  | 153 |  | 4.73 | [8 empty] | 4.59 | [8 empty] | 3.5 | [7 empty] | 153
2021-06-16 | OM_6-16-2021_09-50-13AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00835 | 0.257 | 5.04 | PO4-P | ug P/L | 0.0513 | 0.608 | 9.08 | NO3-N | ug N/L | 0.00819 | 0.119 | 4.07 |  | 154 |  | 5.04 | [8 empty] | 9.08 | [8 empty] | 4.07 | [7 empty] | 154
2021-06-16 | OM_6-16-2021_09-50-13AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0117 | 0.151 | 2.89 | PO4-P | ug P/L | 0.0191 | 0.311 | 2.49 | NO3-N | ug N/L | 0.00854 | 0.117 | 3.94 |  | 155 |  | 2.89 | [8 empty] | 2.49 | [8 empty] | 3.94 | [7 empty] | 155
2021-06-16 | OM_6-16-2021_09-50-13AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0168 | 0.389 | 7.71 | PO4-P | ug P/L | 0.0451 | 0.504 | 6.78 | NO3-N | ug N/L | 0.00832 | 0.118 | 4 |  | 156 |  | 7.71 | [8 empty] | 6.78 | [8 empty] | 4 | [7 empty] | 156
2021-07-20 | OM_7-20-2021_10-09-35AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00597 | 0.124 | 3.13 | PO4-P | ug P/L | 0.022 | 0.47 | 8.86 | NO3-N | ug N/L | 0.0083 | 0.105 | 5.31 |  | 157 |  | 3.13 | [8 empty] | 8.86 | [8 empty] | 5.31 | [7 empty] | 157
2021-07-20 | OM_7-20-2021_10-09-35AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00672 | 0.106 | 2.76 | PO4-P | ug P/L | 0.0188 | 0.36 | 6.47 | NO3-N | ug N/L | 0.00789 | 0.113 | 5.71 |  | 158 |  | 2.76 | [8 empty] | 6.47 | [8 empty] | 5.71 | [7 empty] | 158
2021-07-20 | OM_7-20-2021_10-09-35AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00969 | 0.203 | 4.76 | PO4-P | ug P/L | 0.0181 | 0.333 | 5.88 | NO3-N | ug N/L | 0.00668 | 0.116 | 5.9 |  | 159 |  | 4.76 | [8 empty] | 5.88 | [8 empty] | 5.9 | [7 empty] | 159
2021-07-20 | OM_7-20-2021_10-09-35AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00642 | 0.0963 | 2.57 | PO4-P | ug P/L | 0.0186 | 0.357 | 6.41 | NO3-N | ug N/L | 0.00807 | 0.105 | 5.29 |  | 160 |  | 2.57 | [8 empty] | 6.41 | [8 empty] | 5.29 | [7 empty] | 160
2021-07-20 | OM_7-20-2021_10-09-35AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00911 | 0.19 | 4.5 | PO4-P | ug P/L | 0.0189 | 0.353 | 6.33 | NO3-N | ug N/L | 0.00779 | 0.101 | 5.08 |  | 161 |  | 4.5 | [8 empty] | 6.33 | [8 empty] | 5.08 | [7 empty] | 161
2021-07-20 | OM_7-20-2021_10-09-35AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00892 | 0.254 | 5.81 | PO4-P | ug P/L | 0.0195 | 0.345 | 6.16 | NO3-N | ug N/L | 0.00827 | 0.118 | 5.98 |  | 162 |  | 5.81 | [8 empty] | 6.16 | [8 empty] | 5.98 | [7 empty] | 162
2021-07-20 | OM_7-20-2021_10-09-35AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | -0.00864 | 0.0663 | 1.95 | PO4-P | ug P/L | 0.02 | 0.371 | 6.71 | NO3-N | ug N/L | 0.0083 | 0.109 | 5.52 |  | 163 |  | 1.95 | [8 empty] | 6.71 | [8 empty] | 5.52 | [7 empty] | 163
2021-07-20 | OM_7-20-2021_10-09-35AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00915 | 0.149 | 3.66 | PO4-P | ug P/L | 0.0187 | 0.355 | 6.37 | NO3-N | ug N/L | 0.00843 | 0.0978 | 4.9 |  | 164 |  | 3.66 | [8 empty] | 6.37 | [8 empty] | 4.9 | [7 empty] | 164
2021-07-30 | OM_7-30-2021_09-40-31AM edit1.omn | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0066 | 0.124 | 3.03 | PO4-P | ug P/L | 0.0204 | 0.375 | 5.68 | NO3-N | ug N/L | 0.00847 | 0.129 | 6.74 |  | 165 |  | 3.03 | [8 empty] | 5.68 | [8 empty] | 6.74 | [7 empty] | 165
2021-07-30 | OM_7-30-2021_09-40-31AM edit1.omn | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0113 | 0.194 | 4.6 | PO4-P | ug P/L | 0.0203 | 0.367 | 5.49 | NO3-N | ug N/L | 0.00888 | 0.127 | 6.63 |  | 166 |  | 4.6 | [8 empty] | 5.49 | [8 empty] | 6.63 | [7 empty] | 166
2021-07-30 | OM_7-30-2021_09-40-31AM edit1.omn | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00654 | 0.115 | 2.83 | PO4-P | ug P/L | 0.0187 | 0.311 | 4.22 | NO3-N | ug N/L | 0.00919 | 0.13 | 6.82 |  | 167 |  | 2.83 | [8 empty] | 4.22 | [8 empty] | 6.82 | [7 empty] | 167
2021-07-30 | OM_7-30-2021_09-40-31AM edit1.omn | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00616 | 0.146 | 3.51 | PO4-P | ug P/L | 0.0209 | 0.39 | 6.01 | NO3-N | ug N/L | 0.00934 | 0.121 | 6.35 |  | 168 |  | 3.51 | [8 empty] | 6.01 | [8 empty] | 6.35 | [7 empty] | 168
2021-07-30 | OM_7-30-2021_09-40-31AM edit1.omn | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00931 | 0.123 | 3.01 | PO4-P | ug P/L | 0.0206 | 0.387 | 5.95 | NO3-N | ug N/L | 0.00899 | 0.112 | 5.85 |  | 169 |  | 3.01 | [8 empty] | 5.95 | [8 empty] | 5.85 | [7 empty] | 169
2021-07-30 | OM_7-30-2021_09-40-31AM edit1.omn | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00516 | 0.106 | 2.62 | PO4-P | ug P/L | 0.0207 | 0.363 | 5.4 | NO3-N | ug N/L | 0.00794 | 0.0844 | 4.35 |  | 170 |  | 2.62 | [8 empty] | 5.4 | [8 empty] | 4.35 | [7 empty] | 170
2021-07-30 | OM_7-30-2021_09-40-31AM edit1.omn | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0101 | 0.217 | 5.1 | PO4-P | ug P/L | 0.0213 | 0.36 | 5.35 | NO3-N | ug N/L | 0.00855 | 0.117 | 6.1 |  | 171 |  | 5.1 | [8 empty] | 5.35 | [8 empty] | 6.1 | [7 empty] | 171
2021-07-30 | OM_7-30-2021_09-40-31AM edit1.omn | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00679 | 0.232 | 5.45 | PO4-P | ug P/L | 0.0213 | 0.374 | 5.66 | NO3-N | ug N/L | 0.00788 | 0.119 | 6.19 |  | 172 |  | 5.45 | [8 empty] | 5.66 | [8 empty] | 6.19 | [7 empty] | 172
2021-08-03 | OM_8-3-2021_09-30-01AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00662 | 0.157 | 3.74 | PO4-P | ug P/L | 0.0208 | 0.373 | 4.4 | NO3-N | ug N/L | 0.00903 | 0.134 | 6.96 |  | 173 |  | 3.74 | [8 empty] | 4.4 | [8 empty] | 6.96 | [7 empty] | 173
2021-08-03 | OM_8-3-2021_09-30-01AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00476 | 0.0567 | 1.65 | PO4-P | ug P/L | 0.021 | 0.392 | 4.84 | NO3-N | ug N/L | 0.00988 | 0.135 | 7.03 |  | 174 |  | 1.65 | [8 empty] | 4.84 | [8 empty] | 7.03 | [7 empty] | 174
2021-08-03 | OM_8-3-2021_09-30-01AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00823 | 0.18 | 4.23 | PO4-P | ug P/L | 0.0208 | 0.371 | 4.36 | NO3-N | ug N/L | 0.00939 | 0.136 | 7.06 |  | 175 |  | 4.23 | [8 empty] | 4.36 | [8 empty] | 7.06 | [7 empty] | 175
2021-08-03 | OM_8-3-2021_09-30-01AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00888 | 0.26 | 5.88 | PO4-P | ug P/L | 0.0216 | 0.403 | 5.1 | NO3-N | ug N/L | 0.0098 | 0.138 | 7.19 |  | 176 |  | 5.88 | [8 empty] | 5.1 | [8 empty] | 7.19 | [7 empty] | 176
2021-08-03 | OM_8-3-2021_09-30-01AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0106 | 0.255 | 5.78 | PO4-P | ug P/L | 0.0213 | 0.382 | 4.6 | NO3-N | ug N/L | 0.00932 | 0.124 | 6.51 |  | 177 |  | 5.78 | [8 empty] | 4.6 | [8 empty] | 6.51 | [7 empty] | 177
2021-08-03 | OM_8-3-2021_09-30-01AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0106 | 0.2 | 4.64 | PO4-P | ug P/L | 0.021 | 0.387 | 4.72 | NO3-N | ug N/L | 0.00991 | 0.159 | 8.17 |  | 178 |  | 4.64 | [8 empty] | 4.72 | [8 empty] | 8.17 | [7 empty] | 178
2021-08-03 | OM_8-3-2021_09-30-01AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00983 | 0.274 | 6.18 | PO4-P | ug P/L | 0.0212 | 0.393 | 4.86 | NO3-N | ug N/L | 0.00961 | 0.137 | 7.12 |  | 179 |  | 6.18 | [8 empty] | 4.86 | [8 empty] | 7.12 | [7 empty] | 179
2021-08-03 | OM_8-3-2021_09-30-01AM.OMN | chk 5 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00949 | 0.229 | 5.26 | PO4-P | ug P/L | 0.0209 | 0.385 | 4.68 | NO3-N | ug N/L | 0.00958 | 0.152 | 7.83 |  | 180 |  | 5.26 | [8 empty] | 4.68 | [8 empty] | 7.83 | [7 empty] | 180
2021-08-24 | OM_8-24-2021_08-50-59AM edit.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00988 | 0.304 | 7.22 | PO4-P | ug P/L | 0.0209 | 0.376 | 5.02 | NO3-N | ug N/L | 0.0122 | 0.168 | 6.26 |  | 181 |  | 7.22 | [8 empty] | 5.02 | [8 empty] | 6.26 | [7 empty] | 181
2021-08-24 | OM_8-24-2021_08-50-59AM edit.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | -0.00853 | 0.194 | 5.11 | PO4-P | ug P/L | 0.0216 | 0.395 | 5.45 | NO3-N | ug N/L | 0.0148 | 0.291 | 12 |  | 182 |  | 5.11 | [8 empty] | 5.45 | [8 empty] | 12 | [7 empty] | 182
2021-08-24 | OM_8-24-2021_08-50-59AM edit.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00907 | 0.206 | 5.34 | PO4-P | ug P/L | 0.0209 | 0.383 | 5.17 | NO3-N | ug N/L | 0.0132 | 0.225 | 8.9 |  | 183 |  | 5.34 | [8 empty] | 5.17 | [8 empty] | 8.9 | [7 empty] | 183
2021-08-24 | OM_8-24-2021_08-50-59AM edit.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00637 | 0.145 | 4.16 | PO4-P | ug P/L | 0.0206 | 0.388 | 5.3 | NO3-N | ug N/L | 0.00997 | 0.144 | 5.11 |  | 184 |  | 4.16 | [8 empty] | 5.3 | [8 empty] | 5.11 | [7 empty] | 184
2021-08-24 | OM_8-24-2021_08-50-59AM edit.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0112 | 0.176 | 4.76 | PO4-P | ug P/L | 0.0235 | 0.42 | 6.03 | NO3-N | ug N/L | 0.0117 | 0.126 | 4.28 |  | 185 |  | 4.76 | [8 empty] | 6.03 | [8 empty] | 4.28 | [7 empty] | 185
2021-08-24 | OM_8-24-2021_08-50-59AM edit.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00949 | 0.232 | 5.83 | PO4-P | ug P/L | 0.0208 | 0.376 | 5 | NO3-N | ug N/L | 0.0101 | 0.15 | 5.38 |  | 186 |  | 5.83 | [8 empty] | 5 | [8 empty] | 5.38 | [7 empty] | 186
2021-08-24 | OM_8-24-2021_08-50-59AM edit.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0079 | 0.272 | 6.6 | PO4-P | ug P/L | 0.0207 | 0.381 | 5.14 | NO3-N | ug N/L | 0.0107 | 0.142 | 5 |  | 187 |  | 6.6 | [8 empty] | 5.14 | [8 empty] | 5 | [7 empty] | 187
2021-08-24 | OM_8-24-2021_08-50-59AM edit.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00795 | 0.173 | 4.69 | PO4-P | ug P/L | 0.0213 | 0.403 | 5.63 | NO3-N | ug N/L | 0.0111 | 0.134 | 4.64 |  | 188 |  | 4.69 | [8 empty] | 5.63 | [8 empty] | 4.64 | [7 empty] | 188
2021-09-28 | OM_9-28-2021_09-15-07AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | -0.0205 | -0.0728 | 0 | PO4-P | ug P/L | 0.018 | 0.337 | 4.03 | NO3-N | ug N/L | 0.00963 | 0.144 | 5.82 |  | 189 |  | 0 | [8 empty] | 4.03 | [8 empty] | 5.82 | [7 empty] | 189
2021-09-28 | OM_9-28-2021_09-15-07AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0164 | 0.514 | 9.79 | PO4-P | ug P/L | 0.0192 | 0.335 | 3.99 | NO3-N | ug N/L | 0.0088 | 0.125 | 4.9 |  | 190 |  | 9.79 | [8 empty] | 3.99 | [8 empty] | 4.9 | [7 empty] | 190
2021-09-28 | OM_9-28-2021_09-15-07AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0146 | 0.358 | 6.8 | PO4-P | ug P/L | 0.0191 | 0.349 | 4.3 | NO3-N | ug N/L | 0.0102 | 0.154 | 6.31 |  | 191 |  | 6.8 | [8 empty] | 4.3 | [8 empty] | 6.31 | [7 empty] | 191
2021-09-28 | OM_9-28-2021_09-15-07AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.016 | 0.524 | 9.98 | PO4-P | ug P/L | 0.0191 | 0.338 | 4.05 | NO3-N | ug N/L | 0.00863 | 0.125 | 4.88 |  | 192 |  | 9.98 | [8 empty] | 4.05 | [8 empty] | 4.88 | [7 empty] | 192
2021-09-28 | OM_9-28-2021_09-15-07AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0167 | 0.365 | 6.93 | PO4-P | ug P/L | 0.0194 | 0.4 | 5.48 | NO3-N | ug N/L | 0.00926 | 0.125 | 4.92 |  | 193 |  | 6.93 | [8 empty] | 5.48 | [8 empty] | 4.92 | [7 empty] | 193
2021-09-28 | OM_9-28-2021_09-15-07AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0131 | 0.344 | 6.52 | PO4-P | ug P/L | 0.0197 | 0.374 | 4.88 | NO3-N | ug N/L | 0.00756 | 0.105 | 3.95 |  | 194 |  | 6.52 | [8 empty] | 4.88 | [8 empty] | 3.95 | [7 empty] | 194
2021-09-28 | OM_9-28-2021_09-15-07AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | -0.0138 | 0.19 | 3.59 | PO4-P | ug P/L | 0.0194 | 0.387 | 5.18 | NO3-N | ug N/L | 0.00794 | 0.12 | 4.65 |  | 195 |  | 3.59 | [8 empty] | 5.18 | [8 empty] | 4.65 | [7 empty] | 195
2021-09-28 | OM_9-28-2021_09-15-07AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0121 | 0.329 | 6.25 | PO4-P | ug P/L | 0.0185 | 0.328 | 3.82 | NO3-N | ug N/L | 0.00863 | 0.134 | 5.33 |  | 196 |  | 6.25 | [8 empty] | 3.82 | [8 empty] | 5.33 | [7 empty] | 196
2021-10-05 | OM_10-5-2021_09-07-11AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.017 | 0.38 | 6.75 | PO4-P | ug P/L | 0.0224 | 0.435 | 7.69 | NO3-N | ug N/L | 0.0137 | 0.164 | 7.17 |  | 197 |  | 6.75 | [6 empty] | 1 |  | 7.69 | [6 empty] | 1 |  | 7.17 | [7 empty] | 197
2021-10-05 | OM_10-5-2021_09-07-11AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0157 | 0.368 | 6.49 | PO4-P | ug P/L | 0.0235 | 0.421 | 7.38 | NO3-N | ug N/L | 0.0128 | 0.177 | 7.77 |  | 198 |  | 6.49 | [6 empty] | 1 |  | 7.38 | [6 empty] | 1 |  | 7.77 | [7 empty] | 198
2021-10-05 | OM_10-5-2021_09-07-11AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.015 | 0.321 | 5.52 | PO4-P | ug P/L | 0.0274 | 0.468 | 8.43 | NO3-N | ug N/L | 0.0129 | 0.171 | 7.48 |  | 199 |  | 5.52 | [6 empty] | 1 |  | 8.43 | [6 empty] | 1 |  | 7.48 | [7 empty] | 199
2021-10-05 | OM_10-5-2021_09-07-11AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0169 | 0.364 | 6.41 | PO4-P | ug P/L | 0.0238 | 0.421 | 7.38 | NO3-N | ug N/L | 0.013 | 0.17 | 7.43 |  | 200 |  | 6.41 | [6 empty] | 1 |  | 7.38 | [6 empty] | 1 |  | 7.43 | [7 empty] | 200
2021-10-05 | OM_10-5-2021_09-07-11AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0158 | 0.416 | 7.49 | PO4-P | ug P/L | 0.0245 | 0.471 | 8.52 | NO3-N | ug N/L | 0.0126 | 0.176 | 7.75 |  | 201 |  | 7.49 | [6 empty] | 1 |  | 8.52 | [6 empty] | 1 |  | 7.75 | [7 empty] | 201
2021-10-05 | OM_10-5-2021_09-07-11AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0178 | 0.405 | 7.26 | PO4-P | ug P/L | 0.0238 | 0.431 | 7.6 | NO3-N | ug N/L | 0.0116 | 0.176 | 7.77 |  | 202 |  | 7.26 | [6 empty] | 1 |  | 7.6 | [6 empty] | 1 |  | 7.77 | [7 empty] | 202
2021-10-05 | OM_10-5-2021_09-07-11AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0197 | 0.293 | 4.93 | PO4-P | ug P/L | 0.0262 | 0.542 | 10.1 | NO3-N | ug N/L | 0.0101 | 0.184 | 8.13 |  | 203 |  | 4.93 | [6 empty] | 1 |  | 10.1 | [6 empty] | 1 |  | 8.13 | [7 empty] | 203
2021-10-05 | OM_10-5-2021_09-07-11AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0182 | 0.416 | 7.5 | PO4-P | ug P/L | 0.0233 | 0.435 | 7.68 | NO3-N | ug N/L | 0.0129 | 0.181 | 8 |  | 204 |  | 7.5 | [6 empty] | 1 |  | 7.68 | [6 empty] | 1 |  | 8 | [7 empty] | 204
2021-10-21 | OM_10-21-2021_09-14-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0103 | 0.226 | 6.04 | PO4-P | ug P/L | 0.0205 | 0.377 | 5.13 | NO3-N | ug N/L | 0.00878 | 0.129 | 6.11 |  | 205 |  | 6.04 | [6 empty] | 1 |  | 5.13 | [6 empty] | 1 |  | 6.11 | [7 empty] | 205
2021-10-21 | OM_10-21-2021_09-14-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.012 | 0.224 | 6 | PO4-P | ug P/L | 0.0202 | 0.362 | 4.79 | NO3-N | ug N/L | 0.00882 | 0.129 | 6.08 |  | 206 |  | 6 | [6 empty] | 1 |  | 4.79 | [6 empty] | 1 |  | 6.08 | [7 empty] | 206
2021-10-21 | OM_10-21-2021_09-14-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0142 | 0.3 | 7.58 | PO4-P | ug P/L | 0.0192 | 0.369 | 4.94 | NO3-N | ug N/L | 0.00876 | 0.127 | 6 |  | 207 |  | 7.58 | [6 empty] | 1 |  | 4.94 | [6 empty] | 1 |  | 6 | [7 empty] | 207
2021-10-21 | OM_10-21-2021_09-14-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0115 | 0.277 | 7.1 | PO4-P | ug P/L | 0.0184 | 0.341 | 4.31 | NO3-N | ug N/L | 0.00852 | 0.125 | 5.9 |  | 208 |  | 7.1 | [6 empty] | 1 |  | 4.31 | [6 empty] | 1 |  | 5.9 | [7 empty] | 208
2021-10-21 | OM_10-21-2021_09-14-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0147 | 0.314 | 7.88 | PO4-P | ug P/L | 0.0179 | 0.33 | 4.07 | NO3-N | ug N/L | 0.00786 | 0.115 | 5.4 |  | 209 |  | 7.88 | [6 empty] | 1 |  | 4.07 | [6 empty] | 1 |  | 5.4 | [7 empty] | 209
2021-10-21 | OM_10-21-2021_09-14-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0147 | 0.311 | 7.81 | PO4-P | ug P/L | 0.0186 | 0.351 | 4.53 | NO3-N | ug N/L | 0.00826 | 0.139 | 6.58 |  | 210 |  | 7.81 | [6 empty] | 1 |  | 4.53 | [6 empty] | 1 |  | 6.58 | [7 empty] | 210
2021-10-21 | OM_10-21-2021_09-14-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0118 | 0.306 | 7.71 | PO4-P | ug P/L | 0.0183 | 0.343 | 4.35 | NO3-N | ug N/L | 0.00798 | 0.114 | 5.38 |  | 211 |  | 7.71 | [6 empty] | 1 |  | 4.35 | [6 empty] | 1 |  | 5.38 | [7 empty] | 211
2021-10-21 | OM_10-21-2021_09-14-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0131 | 0.314 | 7.89 | PO4-P | ug P/L | 0.0194 | 0.36 | 4.75 | NO3-N | ug N/L | 0.00829 | 0.11 | 5.19 |  | 212 |  | 7.89 | [6 empty] | 1 |  | 4.75 | [6 empty] | 1 |  | 5.19 | [7 empty] | 212
2021-10-28 | OM_10-28-2021_10-23-54AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0161 | 0.36 | 4.52 | PO4-P | ug P/L | 0.0183 | 0.342 | 4.99 | NO3-N | ug N/L | 0.0109 | 0.12 | 5.25 |  | 213 | recalibrated using middle and end to 250ppb | 9.4658 | [6 empty] | 1 |  | 4.99 | [6 empty] | 1 |  | 5.25 | [7 empty] | 213
2021-10-28 | OM_10-28-2021_10-23-54AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0126 | 0.274 | 3.04 | PO4-P | ug P/L | 0.0193 | 0.333 | 4.79 | NO3-N | ug N/L | 0.00991 | 0.159 | 7.24 |  | 214 | recalibrated using middle and end to 250ppb | 7.2187 | [6 empty] | 1 |  | 4.79 | [6 empty] | 1 |  | 7.24 | [7 empty] | 214
2021-10-28 | OM_10-28-2021_10-23-54AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0431 | 0.424 | 5.62 | PO4-P | ug P/L | 0.0181 | 0.331 | 4.73 | NO3-N | ug N/L | 0.00942 | 0.142 | 6.39 |  | 215 | recalibrated using middle and end to 250ppb | 11.1358 | [6 empty] | 1 |  | 4.73 | [6 empty] | 1 |  | 6.39 | [7 empty] | 215
2021-10-28 | OM_10-28-2021_10-23-54AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0128 | 0.231 | 2.3 | PO4-P | ug P/L | 0.0186 | 0.354 | 5.25 | NO3-N | ug N/L | 0.00804 | 0.134 | 5.95 |  | 216 | recalibrated using middle and end to 250ppb | 6.0939 | [6 empty] | 1 |  | 5.25 | [6 empty] | 1 |  | 5.95 | [7 empty] | 216
2021-10-28 | OM_10-28-2021_10-23-54AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0172 | 0.512 | 7.12 | PO4-P | ug P/L | 0.0181 | 0.33 | 4.72 | NO3-N | ug N/L | 0.00747 | 0.118 | 5.15 |  | 217 | recalibrated using middle and end to 250ppb | 13.429 | [6 empty] | 1 |  | 4.72 | [6 empty] | 1 |  | 5.15 | [7 empty] | 217
2021-10-28 | OM_10-28-2021_10-23-54AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0219 | 0.337 | 4.12 | PO4-P | ug P/L | 0.018 | 0.333 | 4.77 | NO3-N | ug N/L | 0.00765 | 0.116 | 5 |  | 218 | recalibrated using middle and end to 250ppb | 8.8652 | [6 empty] | 1 |  | 4.77 | [6 empty] | 1 |  | 5 | [7 empty] | 218
2021-10-28 | OM_10-28-2021_10-23-54AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0138 | 0.375 | 4.78 | PO4-P | ug P/L | 0.0201 | 0.346 | 5.08 | NO3-N | ug N/L | 0.0108 | 0.177 | 8.15 |  | 219 | recalibrated using middle and end to 250ppb | 9.8574 | [6 empty] | 1 |  | 5.08 | [6 empty] | 1 |  | 8.15 | [7 empty] | 219
2021-10-28 | OM_10-28-2021_10-23-54AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0229 | 0.3 | 3.48 | PO4-P | ug P/L | 0.0192 | 0.354 | 5.26 | NO3-N | ug N/L | 0.01 | 0.156 | 7.08 |  | 220 | recalibrated using middle and end to 250ppb | 7.8984 | [6 empty] | 1 |  | 5.26 | [6 empty] | 1 |  | 7.08 | [7 empty] | 220
2021-11-02 | OM_11-2-2021_09-37-39AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0127 | 0.341 | 6.31 | PO4-P | ug P/L | 0.0189 | 0.342 | 4.36 | NO3-N | ug N/L | 0.00707 | 0.107 | 4.39 |  | 221 |  | 5.9712 | [6 empty] | 1 |  | 4.36 | [6 empty] | 1 |  | 4.39 | [7 empty] | 221
2021-11-02 | OM_11-2-2021_09-37-39AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0123 | 0.245 | 4.22 | PO4-P | ug P/L | 0.0183 | 0.335 | 4.21 | NO3-N | ug N/L | 0.00735 | 0.133 | 5.67 |  | 222 |  | 3.5271 | [6 empty] | 1 |  | 4.21 | [6 empty] | 1 |  | 5.67 | [7 empty] | 222
2021-11-02 | OM_11-2-2021_09-37-39AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0128 | 0.302 | 5.46 | PO4-P | ug P/L | 0.0182 | 0.318 | 3.84 | NO3-N | ug N/L | 0.00715 | 0.115 | 4.8 |  | 223 |  | 4.9792 | [6 empty] | 1 |  | 3.84 | [6 empty] | 1 |  | 4.8 | [7 empty] | 223
2021-11-02 | OM_11-2-2021_09-37-39AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0109 | 0.335 | 6.17 | PO4-P | ug P/L | 0.0178 | 0.296 | 3.37 | NO3-N | ug N/L | 0.00705 | 0.122 | 5.14 |  | 224 |  | 5.8186 | [6 empty] | 1 |  | 3.37 | [6 empty] | 1 |  | 5.14 | [7 empty] | 224
2021-11-02 | OM_11-2-2021_09-37-39AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0127 | 0.242 | 4.15 | PO4-P | ug P/L | 0.0165 | 0.277 | 2.98 | NO3-N | ug N/L | 0.0068 | 0.122 | 5.13 |  | 225 |  | 3.4506 | [6 empty] | 1 |  | 2.98 | [6 empty] | 1 |  | 5.13 | [7 empty] | 225
2021-11-02 | OM_11-2-2021_09-37-39AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0107 | 0.261 | 4.57 | PO4-P | ug P/L | 0.0182 | 0.345 | 4.43 | NO3-N | ug N/L | 0.00699 | 0.13 | 5.53 |  | 226 |  | 3.935 | [6 empty] | 1 |  | 4.43 | [6 empty] | 1 |  | 5.53 | [7 empty] | 226
2021-11-02 | OM_11-2-2021_09-37-39AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0111 | 0.266 | 4.67 | PO4-P | ug P/L | 0.0171 | 0.299 | 3.44 | NO3-N | ug N/L | 0.00834 | 0.109 | 4.49 |  | 227 |  | 4.0625 | [6 empty] | 1 |  | 3.44 | [6 empty] | 1 |  | 4.49 | [7 empty] | 227
2021-11-11 | OM_11-11-2021_09-27-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0126 | 0.236 | 5.98 | PO4-P | ug P/L | 0.0186 | 0.319 | 4.08 | NO3-N | ug N/L | 0.00752 | 0.111 | 5.96 |  | 228 |  | 5.98 | [6 empty] | 1 |  | 4.08 | [6 empty] | 1 |  | 5.96 | [7 empty] | 228
2021-11-11 | OM_11-11-2021_09-27-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00983 | 0.302 | 7.24 | PO4-P | ug P/L | 0.0184 | 0.32 | 4.09 | NO3-N | ug N/L | 0.00731 | 0.0979 | 5.23 |  | 229 |  | 7.24 | [6 empty] | 1 |  | 4.09 | [6 empty] | 1 |  | 5.23 | [7 empty] | 229
2021-11-11 | OM_11-11-2021_09-27-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0104 | 0.259 | 6.41 | PO4-P | ug P/L | 0.0182 | 0.317 | 4.03 | NO3-N | ug N/L | 0.00669 | 0.1 | 5.36 |  | 230 |  | 6.41 | [6 empty] | 1 |  | 4.03 | [6 empty] | 1 |  | 5.36 | [7 empty] | 230
2021-11-11 | OM_11-11-2021_09-27-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0131 | 0.38 | 8.77 | PO4-P | ug P/L | 0.0179 | 0.316 | 4 | NO3-N | ug N/L | 0.00691 | 0.0999 | 5.34 |  | 231 |  | 8.77 | [6 empty] | 1 |  | 4 | [6 empty] | 1 |  | 5.34 | [7 empty] | 231
2021-11-11 | OM_11-11-2021_09-27-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00864 | 0.213 | 5.52 | PO4-P | ug P/L | 0.0191 | 0.334 | 4.42 | NO3-N | ug N/L | 0.00686 | 0.109 | 5.87 |  | 232 |  | 5.52 | [6 empty] | 1 |  | 4.42 | [6 empty] | 1 |  | 5.87 | [7 empty] | 232
2021-11-11 | OM_11-11-2021_09-27-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | -0.0108 | 0.253 | 6.29 | PO4-P | ug P/L | 0.0193 | 0.359 | 4.99 | NO3-N | ug N/L | 0.00702 | 0.104 | 5.59 |  | 233 |  | 6.29 | [6 empty] | 1 |  | 4.99 | [6 empty] | 1 |  | 5.59 | [7 empty] | 233
2021-11-11 | OM_11-11-2021_09-27-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0119 | 0.257 | 6.38 | PO4-P | ug P/L | 0.0185 | 0.314 | 3.96 | NO3-N | ug N/L | 0.00711 | 0.109 | 5.82 |  | 234 |  | 6.38 | [6 empty] | 1 |  | 3.96 | [6 empty] | 1 |  | 5.82 | [7 empty] | 234
2021-11-11 | OM_11-11-2021_09-27-58AM edit1.omn | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0102 | 0.274 | 6.7 | PO4-P | ug P/L | 0.0189 | 0.335 | 4.43 | NO3-N | ug N/L | 0.00721 | 0.107 | 5.75 |  | 235 |  | 6.7 | [6 empty] | 1 |  | 4.43 | [6 empty] | 1 |  | 5.75 | [7 empty] | 235
2021-12-09 | OM_12-9-2021_09-48-28AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00887 | 0.201 | 3.58 | PO4-P | ug P/L | 0.0194 | 0.332 | 4.42 | NO3-N | ug N/L | 0.00761 | 0.105 | 3.66 |  | 1 |  | 3.58 | [6 empty] | 1 |  | 4.42 | [6 empty] | 1 |  | 3.66
2021-12-09 | OM_12-9-2021_09-48-28AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00874 | 0.185 | 3.23 | PO4-P | ug P/L | 0.0191 | 0.323 | 4.21 | NO3-N | ug N/L | 0.00839 | 0.106 | 3.71 |  | 1 |  | 3.23 | [6 empty] | 1 |  | 4.21 | [6 empty] | 1 |  | 3.71
2021-12-09 | OM_12-9-2021_09-48-28AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0159 | 0.18 | 3.12 | PO4-P | ug P/L | 0.0188 | 0.337 | 4.53 | NO3-N | ug N/L | 0.00856 | 0.125 | 4.75 |  | 1 |  | 3.12 | [6 empty] | 1 |  | 4.53 | [6 empty] | 1 |  | 4.75
2021-12-09 | OM_12-9-2021_09-48-28AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.012 | 0.274 | 5.13 | PO4-P | ug P/L | 0.0184 | 0.326 | 4.27 | NO3-N | ug N/L | 0.00766 | 0.103 | 3.58 |  | 1 |  | 5.13 | [6 empty] | 1 |  | 4.27 | [6 empty] | 1 |  | 3.58
2021-12-09 | OM_12-9-2021_09-48-28AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0231 | 0.334 | 6.42 | PO4-P | ug P/L | 0.0193 | 0.328 | 4.31 | NO3-N | ug N/L | 0.00817 | 0.119 | 4.4 |  | 1 |  | 6.42 | [6 empty] | 1 |  | 4.31 | [6 empty] | 1 |  | 4.4
2021-12-09 | OM_12-9-2021_09-48-28AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00493 | 0.0942 | 1.29 | PO4-P | ug P/L | 0.0202 | 0.343 | 4.65 | NO3-N | ug N/L | 0.00807 | 0.127 | 4.85 |  | 1 |  | 1.29 | [6 empty] | 1 |  | 4.65 | [6 empty] | 1 |  | 4.85
2021-12-09 | OM_12-9-2021_09-48-28AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00942 | 0.193 | 3.4 | PO4-P | ug P/L | 0.0184 | 0.318 | 4.08 | NO3-N | ug N/L | 0.00775 | 0.108 | 3.82 |  | 1 |  | 3.4 | [6 empty] | 1 |  | 4.08 | [6 empty] | 1 |  | 3.82
2021-12-09 | OM_12-9-2021_09-48-28AM.OMN | chk 5 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0143 | 0.222 | 4.02 | PO4-P | ug P/L | 0.0191 | 0.322 | 4.17 | NO3-N | ug N/L | 0.00827 | 0.111 | 4.02 |  | 1 |  | 4.02 | [6 empty] | 1 |  | 4.17 | [6 empty] | 1 |  | 4.02
[22 empty] | Known |  | 5 | [8 empty] | 5 | [8 empty] | 5
| Mean | [20 empty] | Mean |  | 4.8457 | [8 empty] | 5.0669 | [8 empty] | 5.4348 | [6 empty] | Min | 1
| Std | [20 empty] | Std |  | 2.3956 | [8 empty] | 1.4192 | [8 empty] | 1.4649 | [6 empty] | Max | 235
| CV | [20 empty] | CV |  | 49.4385 | [8 empty] | 28.0099 | [8 empty] | 26.9543
Warning Limits | Xbar + 2SD | [19 empty] | Warning Limits | Xbar + 2SD |  | 9.637 | [8 empty] | 7.9054 | [8 empty] | 8.3646
| Xbar - 2SD | [20 empty] | Xbar - 2SD |  | 0.054421 | [8 empty] | 2.2284 | [8 empty] | 2.505
Control Limits | Xbar + 3SD | [19 empty] | Control Limits | Xbar + 3SD |  | 12.0326 | [8 empty] | 9.3246 | [8 empty] | 9.8296
| Xbar - 3SD | [20 empty] | Xbar - 3SD |  | -2.3412 | [8 empty] | 0.8092 | [8 empty] | 1.0401
Warning Limits as % | Xbar + 2SD | [19 empty] | Warning Limits as % | Xbar + 2SD |  | 198.8769 | [8 empty] | 156.0199 | [8 empty] | 153.9086
| Xbar - 2SD | [20 empty] | Xbar - 2SD |  | 1.1231 | [8 empty] | 43.9801 | [8 empty] | 46.0914
Control Limits as % | Xbar + 3SD | [19 empty] | Control Limits as % | Xbar + 3SD |  | 248.3154 | [8 empty] | 184.0298 | [8 empty] | 180.8628
| Xbar - 3SD | [20 empty] | Xbar - 3SD |  | -48.3154 | [8 empty] | 15.9702 | [8 empty] | 19.1372
Count | [20 empty] | Count |  |  | 243 | [8 empty] | 243 | [8 empty] | 243
99th percentile.inc | [20 empty] | 99th percentile.inc |  |  | 10.6504 | [8 empty] | 9.9614 | [8 empty] | 8.482
Maximum | [20 empty] | Maximum |  |  | 13.429 | [8 empty] | 10.6 | [8 empty] | 12
[21 empty] | Known |  |  | 5 | [8 empty] | 5 | [8 empty] | 5
[21 empty] | MDL |  |  | 5.6102 | [8 empty] | 3.3237 | [8 empty] | 3.4306
[21 empty] | LOQ |  |  | 23.9563 | [8 empty] | 14.1924 | [8 empty] | 14.6492
[21 empty] | ratio mean/mdl |  |  | 0.86372 | [8 empty] | 1.5245 | [8 empty] | 1.5842

## rolling blanks
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration |  | Observation #
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | cal std 0 | S10 | 1 | 1 | NH4-N | ug N/L | -0.0429 | -0.137 | 0 | PO4-P | ug P/L | 0.00485 | 0.0771 | 0 | NO3-N+NO2-N | ug N/L | 0.00176 | -0.00302 | 0 |  | 1
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | cal std 0 | S10 | 1 | 1 | NH4-N | ug N/L | 0.0311 | 0.237 | 0 | PO4-P | ug P/L | 0.00465 | 0.0656 | 0 | NO3-N+NO2-N | ug N/L | 0.0016 | 0.032 | 0 |  | 2
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | cal std 0 | S10 | 1 | 1 | NH4-N | ug N/L | 0.019 | 0.158 | 0 | PO4-P | ug P/L | 0.00502 | 0.0619 | 0 | NO3-N+NO2-N | ug N/L | -0.000907 | -0.0016 | 0 |  | 3
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | rb | S10 | 1 | 1 | NH4-N | ug N/L | -0.00446 | -0.137 | -6.06 | PO4-P | ug P/L | 0.00369 | 0.0694 | 0.214 | NO3-N+NO2-N | ug N/L | 0.00242 | 0.0248 | 0.378 |  | 4
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | rb | S10 | 1 | 1 | NH4-N | ug N/L | 0.00924 | 0.194 | 0.946 | PO4-P | ug P/L | 0.00509 | 0.0805 | 0.474 | NO3-N+NO2-N | ug N/L | 0.00182 | 0.0167 | 0.013 |  | 5
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | rb | S10 | 1 | 1 | NH4-N | ug N/L | 0.00917 | 0.207 | 1.23 | PO4-P | ug P/L | 0.00215 | 0.00923 | -1.21 | NO3-N+NO2-N | ug N/L | 0.00294 | 0.0753 | 2.66 |  | 6
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | rb | S10 | 1 | 1 | NH4-N | ug N/L | -0.00409 | 0.00101 | -3.14 | PO4-P | ug P/L | 0.00556 | 0.0888 | 0.671 | NO3-N+NO2-N | ug N/L | 0.00265 | 0.0411 | 1.12 |  | 7
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | rb | S10 | 1 | 1 | NH4-N | ug N/L | 0.00701 | 0.17 | 0.444 | PO4-P | ug P/L | 0.00451 | 0.0703 | 0.235 | NO3-N+NO2-N | ug N/L | 0.0025 | 0.036 | 0.882 |  | 8
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | chk std 0 | S10 | 1 | 1 | NH4-N | ug N/L | 0.0113 | 0.154 | 0.0955 | PO4-P | ug P/L | 0.00389 | 0.0619 | 0.0347 | NO3-N+NO2-N | ug N/L | 0.00143 | 0.0149 | -0.07 |  | 9
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | rb | S10 | 1 | 1 | NH4-N | ug N/L | 0.00484 | 0.122 | -0.565 | PO4-P | ug P/L | 0.00562 | 0.086 | 0.605 | NO3-N+NO2-N | ug N/L | -0.00238 | -0.025 | -1.87 |  | 10
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | rb | S10 | 1 | 1 | NH4-N | ug N/L | 0.358 | 0.628 | 10.2 | PO4-P | ug P/L | 0.00718 | 0.119 | 1.4 | NO3-N+NO2-N | ug N/L | 0.00199 | 0.0362 | 0.891 |  | 11
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | rb | S10 | 1 | 1 | NH4-N | ug N/L | -0.0266 | -0.58 | -15.5 | PO4-P | ug P/L | 0.00547 | 0.0923 | 0.755 | NO3-N+NO2-N | ug N/L | 0.00231 | 0.0271 | 0.48 |  | 12
2018-09-11 | OM_9-11-2018_02-40-57PM_BRNedit.omn | chk std 0 | S10 | 1 | 1 | NH4-N | ug N/L | 0.00388 | 0.0707 | -1.66 | PO4-P | ug P/L | 0.00529 | 0.086 | 0.606 | NO3-N+NO2-N | ug N/L | 0.00106 | -0.00195 | -0.831 |  | 13
2018-10-23 | OM_10-23-2018_01-37-18PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0152 | 0.433 | 8.91 | PO4-P | ug P/L | 0.00437 | 0.0869 | 1.37 | NO3-N+NO2-N | ug N/L | 0.000942 | 0.0148 | 0.873 |  | 14
2018-10-23 | OM_10-23-2018_01-37-18PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.011 | 0.251 | 4.58 | PO4-P | ug P/L | 0.00523 | 0.0958 | 1.56 | NO3-N+NO2-N | ug N/L | 0.00128 | 0.0152 | 0.89 |  | 15
2018-10-23 | OM_10-23-2018_01-37-18PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00862 | 0.246 | 4.45 | PO4-P | ug P/L | 0.00217 | 0.0236 | -0.0377 | NO3-N+NO2-N | ug N/L | 0.00512 | 0.0436 | 2.16 |  | 16
2018-10-23 | OM_10-23-2018_01-37-18PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00561 | 0.0472 | -0.285 | PO4-P | ug P/L | 0.00189 | 0.0252 | -0.00188 | NO3-N+NO2-N | ug N/L | 0.00035 | 0.000684 | 0.242 |  | 17
2018-10-23 | OM_10-23-2018_01-37-18PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.0145 | -0.0617 | -2.88 | PO4-P | ug P/L | 0.00416 | 0.0828 | 1.28 | NO3-N+NO2-N | ug N/L | 0.000865 | -0.00338 | 0.061 |  | 18
2018-10-23 | OM_10-23-2018_01-37-18PM_BRNedit.omn | chk std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00518 | 0.167 | 2.57 | PO4-P | ug P/L | 0.00383 | 0.0708 | 1.01 | NO3-N+NO2-N | ug N/L | 0.00135 | 0.0113 | 0.718 |  | 19
2018-10-23 | OM_10-23-2018_01-37-18PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.00461 | -0.0665 | -3 | PO4-P | ug P/L | 0.00363 | 0.0568 | 0.7 | NO3-N+NO2-N | ug N/L | -0.00109 | -0.0121 | -0.328 |  | 20
2018-10-23 | OM_10-23-2018_01-37-18PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.0129 | -0.203 | -6.25 | PO4-P | ug P/L | -0.00292 | -0.0186 | -0.971 | NO3-N+NO2-N | ug N/L | 0.00124 | 0.0205 | 1.13 |  | 21
2018-10-23 | OM_10-23-2018_01-37-18PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0106 | 0.298 | 5.69 | PO4-P | ug P/L | 0.004 | 0.0635 | 0.848 | NO3-N+NO2-N | ug N/L | 0.00131 | 0.0146 | 0.864 |  | 22
2018-10-23 | OM_10-23-2018_01-37-18PM_BRNedit.omn | chk std 0 | S5 | 1 | 1 | NH4-N | ug N/L | -0.0332 | -0.224 | -6.74 | PO4-P | ug P/L | 0.00385 | 0.0792 | 1.19 | NO3-N+NO2-N | ug N/L | 0.000787 | 9.5e-05 | 0.216 |  | 23
2019-01-15 | OM_1-15-2019_03-14-20PMeditedv3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0232 | 0.224 | 1.89 | PO4-P | ug P/L | 0.00516 | 0.0828 | 1.65 | NO3-N+NO2-N | ug N/L | -0.00332 | -0.061 | -3.57 |  | 24
2019-01-15 | OM_1-15-2019_03-14-20PMeditedv3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0109 | 0.307 | 3.68 | PO4-P | ug P/L | -0.00339 | -0.0632 | -1.77 | NO3-N+NO2-N | ug N/L | -0.0077 | -0.0397 | -2.68 |  | 25
2019-01-15 | OM_1-15-2019_03-14-20PMeditedv3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0117 | 0.21 | 1.59 | PO4-P | ug P/L | 0.00588 | 0.0925 | 1.88 | NO3-N+NO2-N | ug N/L | -0.0822 | -0.551 |  |  | 26
2019-01-15 | OM_1-15-2019_03-14-20PMeditedv3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0126 | 0.0141 | -2.63 | PO4-P | ug P/L | 0.00623 | 0.104 | 2.16 | NO3-N+NO2-N | ug N/L | -0.0142 | -0.0901 | -4.79 |  | 27
2019-01-15 | OM_1-15-2019_03-14-20PMeditedv3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00863 | 0.158 | 0.488 | PO4-P | ug P/L | 0.00652 | 0.125 | 2.63 | NO3-N+NO2-N | ug N/L | 0.00163 | -0.00329 | -1.16 |  | 28
2019-01-15 | OM_1-15-2019_03-14-20PMeditedv3.omn | chk std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00773 | 0.129 | -0.152 | PO4-P | ug P/L | -0.0041 | -0.0226 | -0.82 | NO3-N+NO2-N | ug N/L | 0.00144 | 0.0106 | -0.577 |  | 29
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | rb | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | -0.00735 | -0.0459 | -4.82 | NO3-N+NO2-N | ug N/L | -0.00286 | -0.0199 | -1.97 |  | 30
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | rb | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | 0.00821 | 0.174 | 0.225 | NO3-N+NO2-N | ug N/L | 0.00273 | 0.0264 | 0.138 |  | 31
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | rb | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | 0.00686 | 0.136 | -0.633 | NO3-N+NO2-N | ug N/L | -0.00224 | -0.0263 | -2.27 |  | 32
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | rb | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | -0.00797 | -0.046 | -4.82 | NO3-N+NO2-N | ug N/L | -0.00151 | -0.0142 | -1.72 |  | 33
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | chk std 0 | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | -0.00394 | -0.0197 | -4.22 | NO3-N+NO2-N | ug N/L | 0.00254 | 0.0448 | 0.979 |  | 34
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | rb | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | -0.00663 | -0.0391 | -4.66 | NO3-N+NO2-N | ug N/L | 0.00283 | 0.033 | 0.438 |  | 35
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | rb | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | 0.00833 | 0.166 | 0.0407 | NO3-N+NO2-N | ug N/L | 0.00274 | 0.0435 | 0.918 |  | 36
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | rb | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | -0.00719 | -0.0421 | -4.73 | NO3-N+NO2-N | ug N/L | 0.00254 | 0.00932 | -0.641 |  | 37
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | rb | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | -0.00731 | -0.0443 | -4.78 | NO3-N+NO2-N | ug N/L | 0.00237 | -0.0133 | -1.67 |  | 38
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | rb | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | 0.0027 | -0.000718 | -3.78 | NO3-N+NO2-N | ug N/L | 0.00296 | 0.00797 | -0.702 |  | 39
2019-04-17 | OM_4-17-2019_01-18-39PMbrnedit.omn | chk std 0 | S5 | 1 | 1 | [5 empty] | PO4-P | ug P/L | -0.00692 | -0.0403 | -4.69 | NO3-N+NO2-N | ug N/L | -0.00195 | -0.0263 | -2.27 |  | 40
2019-09-23 | OM_9-23-2019_12-09-25PMedited.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00501 | 0.129 | 0 | PO4-P | ug P/L | 0.00844 | 0.114 | 0 | NO3-N+NO2-N | ug N/L | -0.00159 | 0.00625 | 0 |  | 41
2019-09-23 | OM_9-23-2019_12-09-25PMedited.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00591 | 0.112 | 0 | PO4-P | ug P/L | 0.00781 | 0.1 | 0 | NO3-N+NO2-N | ug N/L | -0.0044 | -0.0252 | 0 |  | 42
2019-09-23 | OM_9-23-2019_02-47-08PMedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.00296 | -0.034 | 0.0273 | PO4-P | ug P/L | -0.00891 | -0.0537 | -5.57 | NO3-N+NO2-N | ug N/L | -0.00266 | -0.0411 | -4.44 |  | 43
2019-09-23 | OM_9-23-2019_02-47-08PMedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0149 | 0.221 | 5.95 | PO4-P | ug P/L | -0.00774 | -0.0469 | -5.38 | NO3-N+NO2-N | ug N/L | 0.00341 | 0.0684 | 0.931 |  | 44
2019-09-23 | OM_9-23-2019_02-47-08PMedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.00488 | 0.0765 | 2.59 | PO4-P | ug P/L | -0.00895 | -0.0568 | -5.65 | NO3-N+NO2-N | ug N/L | -0.00274 | -0.0467 | -4.72 |  | 45
2019-09-23 | OM_9-23-2019_02-47-08PMedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0036 | 0.0583 | 2.17 | PO4-P | ug P/L | 0.00297 | -0.00722 | -4.31 | NO3-N+NO2-N | ug N/L | 0.0029 | -0.0122 | -3.02 |  | 46
2019-09-23 | OM_9-23-2019_02-47-08PMedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.00395 | -0.0417 | -0.152 | PO4-P | ug P/L | -0.00972 | -0.0605 | -5.76 | NO3-N+NO2-N | ug N/L | -0.00286 | -0.0223 | -3.52 |  | 47
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0101 | 0.0713 | 0 | PO4-P | ug P/L | 0.00842 | 0.153 | 0 | NO3-N+NO2-N | ug N/L | 0.00383 | 0.0162 | 0 |  | 48
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | -0.00281 | -0.0637 | 0 | PO4-P | ug P/L | 0.00922 | 0.175 | 0 | NO3-N+NO2-N | ug N/L | 0.000942 | 0.00645 | 0 |  | 49
2019-10-07 | OM_10-7-2019_11-37-43AM_BRNedit.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00396 | 0.106 | 0 | PO4-P | ug P/L | 0.00908 | 0.166 | 0 | NO3-N+NO2-N | ug N/L | -0.00755 | -0.0709 | 0 |  | 50
2019-10-07 | OM_10-7-2019_02-15-47PM_BRNedit.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | -0.02 | -0.646 | 0 | PO4-P | ug P/L | -0.00781 | -0.0506 | 0 | NO3-N+NO2-N | ug N/L | -0.00548 | 0.0211 | 0 |  | 51
2019-10-07 | OM_10-7-2019_02-49-17PM_BRNedit2.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0137 | 0.22 | 4.3 | PO4-P | ug P/L | -0.00796 | -0.0493 | -5.96 | NO3-N+NO2-N | ug N/L | -0.00225 | -0.0185 | -7.19 |  | 52
2019-10-07 | OM_10-7-2019_02-49-17PM_BRNedit2.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0261 | 0.398 | 8.12 | PO4-P | ug P/L | -0.00738 | -0.0461 | -5.87 | NO3-N+NO2-N | ug N/L | 0.0164 | 0.0651 | -2.83 |  | 53
2019-10-07 | OM_10-7-2019_02-49-17PM_BRNedit2.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0267 | 0.392 | 7.98 | PO4-P | ug P/L | -0.00736 | -0.0442 | -5.82 | NO3-N+NO2-N | ug N/L | 0.00352 | 0.0302 | -4.65 |  | 54
2019-10-07 | OM_10-7-2019_02-49-17PM_BRNedit2.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00593 | 0.0503 | 0.649 | PO4-P | ug P/L | 0.0103 | 0.203 | 0.97 | NO3-N+NO2-N | ug N/L | 0.00521 | 0.0935 | -1.35 |  | 55
2019-10-07 | OM_10-7-2019_02-49-17PM_BRNedit2.omn | blank | S5 | 1 | 1 | NH4-N | ug N/L | -0.00866 | -0.0163 | -0.782 | PO4-P | ug P/L | -0.00753 | -0.0457 | -5.86 | NO3-N+NO2-N | ug N/L | 0.00716 | 0.152 | 1.71 |  | 56
2019-10-07 | OM_10-7-2019_02-49-17PM_BRNedit2.omn | blank | S5 | 1 | 1 | NH4-N | ug N/L | -0.00659 | 0.00556 | -0.313 | PO4-P | ug P/L | -0.00893 | -0.0563 | -6.15 | NO3-N+NO2-N | ug N/L | -0.0039 | -0.0506 | -8.86 |  | 57
2019-10-07 | OM_10-7-2019_02-49-17PM_BRNedit2.omn | blank | S5 | 1 | 1 | NH4-N | ug N/L | 0.0194 | 0.349 | 7.06 | PO4-P | ug P/L | -0.00713 | -0.0459 | -5.87 | NO3-N+NO2-N | ug N/L | -0.0072 | 0.0314 | -4.59 |  | 58
2019-10-07 | OM_10-7-2019_02-49-17PM_BRNedit2.omn | chk 0 | S5 | 1 | 1 | NH4-N | ug N/L | -0.00363 | -0.1 | -2.59 | PO4-P | ug P/L | 0.00766 | 0.103 | -1.8 | NO3-N+NO2-N | ug N/L | -0.00206 | -0.0155 | -7.03 |  | 59
2019-10-07 | OM_10-7-2019_02-49-17PM_BRNedit2.omn | blank | S5 | 1 | 1 | NH4-N | ug N/L | 0.0232 | 0.221 | 4.32 | PO4-P | ug P/L | -0.00572 | -0.0337 | -5.53 | NO3-N+NO2-N | ug N/L | -0.00187 | -0.0156 | -7.03 |  | 60
2019-10-07 | OM_10-7-2019_02-49-17PM_BRNedit2.omn | blank | S5 | 1 | 1 | NH4-N | ug N/L | -0.00602 | -0.0901 | -2.37 | PO4-P | ug P/L | -0.00651 | -0.04 | -5.7 | NO3-N+NO2-N | ug N/L | -0.0124 | -0.0604 | -9.37 |  | 61
2019-10-07 | OM_10-7-2019_06-14-10PM_BRNedit2.omn | blank | S5 | 1 | 1 | NH4-N | ug N/L | 0.0281 | 0.128 | 2.32 | PO4-P | ug P/L | 0.00978 | 0.169 | 0.0713 | NO3-N+NO2-N | ug N/L | 0.00169 | 0.0121 | -5.59 |  | 62
2019-10-07 | OM_10-7-2019_06-14-10PM_BRNedit2.omn | chk 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0164 | -0.0225 | -0.916 | PO4-P | ug P/L | 0.00954 | 0.184 | 0.47 | NO3-N+NO2-N | ug N/L | -0.0017 | 0.0158 | -5.4 |  | 63
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00408 | 0.0555 | 0 | PO4-P | ug P/L | 0.00931 | 0.177 | 0 | NO3-N+NO2-N | ug N/L | -0.00362 | -0.0409 | 0 |  | 64
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00219 | 0.0156 | 0 | PO4-P | ug P/L | 0.0091 | 0.172 | 0 | NO3-N+NO2-N | ug N/L | -0.00553 | -0.0746 | 0 |  | 65
2019-10-14 | OM_10-14-2019_10-58-57AMbrnedit.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0123 | 0.0661 | 0 | PO4-P | ug P/L | 0.0101 | 0.168 | 0 | NO3-N+NO2-N | ug N/L | 0.00507 | 0.0174 | 0 |  | 66
2019-10-14 | OM_10-14-2019_12-51-46PMbrnedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.0216 | -0.067 | -2.39 | PO4-P | ug P/L | 0.01 | 0.198 | 1.63 | NO3-N+NO2-N | ug N/L | 0.00263 | 0.00255 | -0.773 |  | 67
2019-10-14 | OM_10-14-2019_12-51-46PMbrnedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00389 | 0.0842 | 2.9 | PO4-P | ug P/L | 0.00818 | 0.141 | -0.311 | NO3-N+NO2-N | ug N/L | 0.00331 | 0.0313 | 0.848 |  | 68
2019-10-14 | OM_10-14-2019_12-51-46PMbrnedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00708 | 0.0161 | 0.519 | PO4-P | ug P/L | 0.0103 | 0.167 | 0.588 | NO3-N+NO2-N | ug N/L | 0.00318 | 0.0514 | 1.98 |  | 69
2019-10-14 | OM_10-14-2019_12-51-46PMbrnedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00256 | 0.0309 | 1.04 | PO4-P | ug P/L | 0.00944 | 0.163 | 0.449 | NO3-N+NO2-N | ug N/L | -0.00354 | -0.0409 | -3.22 |  | 70
2019-10-14 | OM_10-14-2019_12-51-46PMbrnedit.omn | chk std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00405 | 0.0694 | 2.38 | PO4-P | ug P/L | 0.00937 | 0.164 | 0.474 | NO3-N+NO2-N | ug N/L | 0.00377 | 0.0153 | -0.0545 |  | 71
2019-10-14 | OM_10-14-2019_12-51-46PMbrnedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00855 | 0.0943 | 3.26 | PO4-P | ug P/L | 0.00945 | 0.159 | 0.301 | NO3-N+NO2-N | ug N/L | -0.00223 | -0.0186 | -1.96 |  | 72
2019-10-14 | OM_10-14-2019_12-51-46PMbrnedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00851 | 0.0394 | 1.34 | PO4-P | ug P/L | 0.00759 | 0.136 | -0.495 | NO3-N+NO2-N | ug N/L | -0.00246 | 0.00748 | -0.495 |  | 73
2019-10-14 | OM_10-14-2019_12-51-46PMbrnedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00579 | 0.148 | 5.12 | PO4-P | ug P/L | 0.00925 | 0.17 | 0.702 | NO3-N+NO2-N | ug N/L | -0.00246 | -0.0226 | -2.19 |  | 74
2019-10-21 | OM_10-21-2019_02-07-06PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00419 | 0.0902 | 1.63 | PO4-P | ug P/L | 0.0109 | 0.201 | 2.89 | NO3-N+NO2-N | ug N/L | -0.00292 | 0.00532 | -1.79 |  | 75
2019-10-21 | OM_10-21-2019_02-07-06PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0106 | 0.093 | 1.7 | PO4-P | ug P/L | 0.0102 | 0.174 | 1.74 | NO3-N+NO2-N | ug N/L | -0.00459 | -0.0131 | -2.86 |  | 76
2019-10-21 | OM_10-21-2019_02-07-06PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00442 | 0.0691 | 1.08 | PO4-P | ug P/L | 0.00853 | 0.128 | -0.141 | NO3-N+NO2-N | ug N/L | 0.00323 | 0.029 | -0.421 |  | 77
2019-10-21 | OM_10-21-2019_02-07-06PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.0152 | -0.0912 | -3.07 | PO4-P | ug P/L | 0.00915 | 0.168 | 1.51 | NO3-N+NO2-N | ug N/L | 0.00235 | 0.0214 | -0.864 |  | 78
2019-10-21 | OM_10-21-2019_02-07-06PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.0085 | -0.0277 | -1.42 | PO4-P | ug P/L | 0.00804 | 0.14 | 0.357 | NO3-N+NO2-N | ug N/L | -0.00204 | -0.00971 | -2.66 |  | 79
2019-10-21 | OM_10-21-2019_02-07-06PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.00264 | 0.021 | -0.163 | PO4-P | ug P/L | 0.0101 | 0.185 | 2.21 | NO3-N+NO2-N | ug N/L | 0.00348 | 0.042 | 0.33 |  | 80
2019-10-21 | OM_10-21-2019_02-07-06PM_BRNedit.omn | chk 0 | S5 | 1 | 1 | NH4-N | ug N/L | -0.00796 | -0.127 | -4 | PO4-P | ug P/L | 0.00938 | 0.175 | 1.8 | NO3-N+NO2-N | ug N/L | 0.0031 | 0.0268 | -0.551 |  | 81
2019-10-21 | OM_10-21-2019_02-07-06PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.016 | 0.068 | 1.06 | PO4-P | ug P/L | 0.00831 | 0.135 | 0.178 | NO3-N+NO2-N | ug N/L | 0.0029 | 0.0229 | -0.776 |  | 82
2019-10-21 | OM_10-21-2019_02-07-06PM_BRNedit.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00631 | 0.0645 | 0.963 | PO4-P | ug P/L | 0.0109 | 0.197 | 2.69 | NO3-N+NO2-N | ug N/L | 0.00284 | 0.0356 | -0.0416 |  | 83
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | cal std 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00442 | 0.11 | 0 | PO4-P | ug P/L | 0.00956 | 0.163 | 0 | NO3-N+NO2-N | ug N/L | 0.00191 | 0.0228 | 0 |  | 84
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | cal std 0 | S6 | 1 | 1 | NH4-N | ug N/L | -0.0131 | -0.0218 | 0 | PO4-P | ug P/L | 0.00831 | 0.103 | 0 | NO3-N+NO2-N | ug N/L | 0.00198 | 0.0336 | 0 |  | 85
2019-10-24 | OM_10-24-2019_09-28-22AM.OMN | cal std 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00607 | 0.101 | 0 | PO4-P | ug P/L | 0.00766 | 0.101 | 0 | NO3-N+NO2-N | ug N/L | 0.00207 | 0.0341 | 0 |  | 86
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00392 | 0.101 | 1.77 | PO4-P | ug P/L | 0.00953 | 0.16 | 0.513 | NO3-N+NO2-N | ug N/L | 0.00202 | 0.0211 | -0.39 |  | 87
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.00644 | 0.0126 | -0.372 | PO4-P | ug P/L | 0.00917 | 0.13 | -0.41 | NO3-N+NO2-N | ug N/L | -0.00124 | -0.00292 | -1.62 |  | 88
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00514 | 0.0756 | 1.17 | PO4-P | ug P/L | 0.00831 | 0.119 | -0.744 | NO3-N+NO2-N | ug N/L | 0.00179 | 0.0305 | 0.0862 |  | 89
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.0188 | -0.138 | -4.04 | PO4-P | ug P/L | 0.00827 | 0.121 | -0.696 | NO3-N+NO2-N | ug N/L | -0.00132 | -0.0214 | -2.56 |  | 90
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.0218 | -0.213 | -5.88 | PO4-P | ug P/L | 0.00884 | 0.136 | -0.237 | NO3-N+NO2-N | ug N/L | 0.00244 | 0.0429 | 0.719 |  | 91
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | chk std 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00252 | 0.0415 | 0.332 | PO4-P | ug P/L | 0.00879 | 0.148 | 0.158 | NO3-N+NO2-N | ug N/L | 0.00203 | 0.033 | 0.217 |  | 92
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00481 | 0.0948 | 1.63 | PO4-P | ug P/L | 0.00842 | 0.154 | 0.343 | NO3-N+NO2-N | ug N/L | 0.00154 | 0.0205 | -0.422 |  | 93
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.0106 | -0.027 | -1.34 | PO4-P | ug P/L | 0.00896 | 0.169 | 0.801 | NO3-N+NO2-N | ug N/L | 0.00181 | 0.0244 | -0.221 |  | 94
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.00756 | 0.126 | 2.39 | PO4-P | ug P/L | 0.00956 | 0.168 | 0.773 | NO3-N+NO2-N | ug N/L | -0.0019 | -0.0244 | -2.71 |  | 95
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.0139 | -0.117 | -3.53 | PO4-P | ug P/L | 0.00814 | 0.138 | -0.149 | NO3-N+NO2-N | ug N/L | 0.00249 | 0.0582 | 1.5 |  | 96
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00829 | 0.0456 | 0.434 | PO4-P | ug P/L | 0.00725 | 0.0958 | -1.47 | NO3-N+NO2-N | ug N/L | -0.00187 | 0.00898 | -1.01 |  | 97
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | chk std 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00555 | 0.0898 | 1.51 | PO4-P | ug P/L | 0.00653 | 0.101 | -1.3 | NO3-N+NO2-N | ug N/L | -0.00149 | -0.0176 | -2.37 |  | 98
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00593 | 0.119 | 2.22 | PO4-P | ug P/L | -0.00693 | -0.0426 | -5.74 | NO3-N+NO2-N | ug N/L | -0.00118 | -0.00861 | -1.91 |  | 99
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.00277 | 0.00313 | -0.604 | PO4-P | ug P/L | -0.0068 | -0.0384 | -5.61 | NO3-N+NO2-N | ug N/L | 0.00146 | 0.0285 | -0.0124 |  | 100
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.00478 | -0.0248 | -1.28 | PO4-P | ug P/L | 0.00815 | 0.176 | 1 | NO3-N+NO2-N | ug N/L | 0.00129 | 0.031 | 0.114 |  | 101
2019-10-24 | OM_10-24-2019_11-20-21AM.OMN | chk std 0 | S6 | 1 | 1 | NH4-N | ug N/L | -0.00857 | -0.0388 | -1.63 | PO4-P | ug P/L | 0.0137 | 0.177 | 1.03 | NO3-N+NO2-N | ug N/L | -0.00124 | -0.0102 | -1.98 |  | 102
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | cal std 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0128 | 0.0994 | 0 | PO4-P | ug P/L | 0.00764 | 0.156 | 0 | NO3-N+NO2-N | ug N/L | 0.00204 | 0.00745 | 0 |  | 103
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | cal std 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00545 | 0.0989 | 0 | PO4-P | ug P/L | 0.00782 | 0.154 | 0 | NO3-N+NO2-N | ug N/L | 0.00214 | 0.0164 | 0 |  | 104
2019-10-31 | OM_10-31-2019_11-11-08AM.OMN | cal std 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00595 | 0.0852 | 0 | PO4-P | ug P/L | 0.00806 | 0.139 | 0 | NO3-N+NO2-N | ug N/L | -0.00229 | -0.0134 | 0 |  | 105
2019-10-31 | OM_10-31-2019_01-44-45PM_edit.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.136 | 1.59 |  | PO4-P | ug P/L | 0.0126 | 0.263 | 2.71 | NO3-N+NO2-N | ug N/L | -0.00181 | -0.0126 | -1.59 |  | 106
2020-03-18 | OM_3-18-2020_09-56-27AM.OMN | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.00335 | 0.0656 | 0 | PO4-P | ug P/L | -0.0154 | 0.0607 | 0 | NO3-N+NO2-N | ug N/L | 0.00421 | 0.0348 | 0 |  | 107
2020-03-18 | OM_3-18-2020_09-56-27AM.OMN | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0106 | 0.21 | 0 | PO4-P | ug P/L | -0.00916 | 0.0159 | 0 | NO3-N+NO2-N | ug N/L | 0.00228 | 0.0435 | 0 |  | 108
2020-03-18 | OM_3-18-2020_12-16-14PM.OMN | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0176 | 0.173 | 2.68 | PO4-P | ug P/L | -0.00782 | -0.0511 | -6.84 | NO3-N+NO2-N | ug N/L | 0.00144 | 0.00539 | -1.62 |  | 109
2020-03-18 | OM_3-18-2020_12-16-14PM.OMN | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0188 | 0.123 | 1.69 | PO4-P | ug P/L | 0.0114 | 0.247 | 1.79 | NO3-N+NO2-N | ug N/L | 0.00153 | 0.0362 | 0.282 |  | 110
2020-03-18 | OM_3-18-2020_12-16-14PM.OMN | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00533 | 0.0789 | 0.828 | PO4-P | ug P/L | -0.00848 | -0.0586 | -7.06 | NO3-N+NO2-N | ug N/L | -0.00292 | -0.0323 | -3.94 |  | 111
2020-03-18 | OM_3-18-2020_12-16-14PM.OMN | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0325 | 0.613 |  | PO4-P | ug P/L | 0.0121 | 0.186 | 0.0267 | NO3-N+NO2-N | ug N/L | 0.00469 | 0.0247 | -0.424 |  | 112
2020-03-18 | OM_3-18-2020_12-16-14PM.OMN | chk 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.0034 | 0.0838 | 0.926 | PO4-P | ug P/L | 0.00974 | 0.206 | 0.61 | NO3-N+NO2-N | ug N/L | 0.00303 | 0.0563 | 1.52 |  | 113
2020-09-01 | OM_9-1-2020_10-10-28AMedit2.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.006 | 0.171 | 0 | PO4-P | ug P/L | 0.0102 | 0.19 | 0 | NO3-N+NO2-N | ug N/L | 0.00303 | 0.0356 | 0 |  | 114
2020-09-01 | OM_9-1-2020_10-10-28AMedit2.omn | cal std 0 | S5 | 1 | 1 | NH4-N | ug N/L | 0.024 | 0.618 | 0 | PO4-P | ug P/L | -0.0072 | -0.0446 | 0 | NO3-N+NO2-N | ug N/L | 0.00225 | 0.0059 | 0 |  | 115
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00573 | 0.0951 | -1.01 | PO4-P | ug P/L | -0.00659 | -0.0405 | -6.22 | NO3-N+NO2-N | ug N/L | -0.0024 | -0.00859 | -2.08 |  | 116
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.00735 | -0.114 | -5.75 | PO4-P | ug P/L | -0.00581 | 0.0642 | -3.51 | NO3-N+NO2-N | ug N/L | 0.00438 | 0.0128 | -0.811 |  | 117
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | -0.0194 | -0.109 | -5.62 | PO4-P | ug P/L | 0.00455 | 0.0164 | -4.75 | NO3-N+NO2-N | ug N/L | 0.00272 | 0.0319 | 0.326 |  | 118
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00239 | 0.0229 | -2.64 | PO4-P | ug P/L | -0.00592 | 0.0421 | -4.08 | NO3-N+NO2-N | ug N/L | 0.00331 | 0.0286 | 0.131 |  | 119
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00267 | 0.0552 | -1.91 | PO4-P | ug P/L | 0.0114 | 0.24 | 1.04 | NO3-N+NO2-N | ug N/L | 0.00198 | 0.000447 | -1.55 |  | 120
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00988 | 0.147 | 0.159 | PO4-P | ug P/L | 0.0106 | 0.2 | 0.00273 | NO3-N+NO2-N | ug N/L | 0.00122 | 0.021 | -0.324 |  | 121
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00526 | 0.0787 | -1.38 | PO4-P | ug P/L | 0.0101 | 0.208 | 0.216 | NO3-N+NO2-N | ug N/L | 0.00199 | 0.0335 | 0.418 |  | 122
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.0183 | 0.222 | 1.87 | PO4-P | ug P/L | 0.012 | 0.241 | 1.07 | NO3-N+NO2-N | ug N/L | 0.00253 | 0.0155 | -0.65 |  | 123
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00502 | 0.103 | -0.834 | PO4-P | ug P/L | -0.00763 | -0.0489 | -6.44 | NO3-N+NO2-N | ug N/L | 0.00256 | 0.0223 | -0.248 |  | 124
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | rb | S5 | 1 | 1 | NH4-N | ug N/L | 0.00704 | 0.14 | 0.0015 | PO4-P | ug P/L | -0.0083 | -0.0549 | -6.6 | NO3-N+NO2-N | ug N/L | 0.00191 | -0.0111 | -2.23 |  | 125
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00592 | 0.083 | 1.52 | PO4-P | ug P/L | 0.0073 | 0.0908 | -2.4 | NO3-N | ug N/L | -0.00181 | -0.0122 | -2.89 |  | 126
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.012 | 0.12 | 2.27 | PO4-P | ug P/L | -0.0162 | -0.171 | -8.21 | NO3-N | ug N/L | 0.00201 | 0.00758 | -1.84 |  | 127
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.00304 | -0.00463 | -0.253 | PO4-P | ug P/L | 0.00822 | 0.1 | -2.19 | NO3-N | ug N/L | 0.0023 | 0.0407 | -0.0915 |  | 128
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00763 | 0.205 | 3.98 | PO4-P | ug P/L | 0.00873 | 0.0992 | -2.21 | NO3-N | ug N/L | 0.00408 | 0.0665 | 1.28 |  | 129
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.00313 | -0.0506 | -1.18 | PO4-P | ug P/L | 0.00854 | 0.0919 | -2.38 | NO3-N | ug N/L | 0.00267 | 0.054 | 0.615 |  | 130
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | chk 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0241 | 0.176 | 3.39 | PO4-P | ug P/L | -0.00817 | -0.0305 | -5.09 | NO3-N | ug N/L | -0.00126 | -0.00538 | -2.53 |  | 131
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.00961 | -0.165 | -3.5 | PO4-P | ug P/L | 0.00781 | 0.146 | -1.18 | NO3-N | ug N/L | -0.00249 | -0.0223 | -3.42 |  | 132
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00535 | 0.141 | 2.69 | PO4-P | ug P/L | 0.0088 | 0.12 | -1.75 | NO3-N | ug N/L | -0.00232 | -0.0196 | -3.28 |  | 133
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.0106 | 0.076 | 1.38 | PO4-P | ug P/L | 0.00807 | 0.0981 | -2.24 | NO3-N | ug N/L | -0.00155 | -0.0126 | -2.91 |  | 134
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | chk 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.0103 | 0.279 | 5.48 | PO4-P | ug P/L | 0.00848 | 0.192 | -0.163 | NO3-N | ug N/L | 0.00118 | 0.0105 | -1.69 |  | 135
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.0132 | 0.0932 | 2.5 | PO4-P | ug P/L | 0.00984 | 0.117 | 1.19 | NO3-N | ug N/L | -0.00205 | -0.0202 | -1.43 |  | 136
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00335 | 0.0431 | 1.47 | PO4-P | ug P/L | -0.00593 | -0.0345 | -2.11 | NO3-N | ug N/L | 0.00113 | 0.0108 | 0.234 |  | 137
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.0137 | 0.209 | 4.89 | PO4-P | ug P/L | 0.00369 | -0.000945 | -1.38 | NO3-N | ug N/L | 0.00291 | 0.0353 | 1.55 |  | 138
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.0033 | 0.0816 | 2.26 | PO4-P | ug P/L | -0.00705 | -0.0428 | -2.29 | NO3-N | ug N/L | 0.00164 | 0.0264 | 1.07 |  | 139
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.0029 | 0.00107 | 0.602 | PO4-P | ug P/L | -0.00811 | -0.047 | -2.38 | NO3-N | ug N/L | -0.00151 | -0.0109 | -0.931 |  | 140
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | chk 0 | S6 | 1 | 1 | NH4-N | ug N/L | -0.00917 | -0.0314 | -0.0662 | PO4-P | ug P/L | 0.00913 | 0.19 | 2.79 | NO3-N | ug N/L | -0.00203 | -0.0124 | -1.01 |  | 141
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.00922 | -0.0353 | -0.147 | PO4-P | ug P/L | -0.00668 | -0.0392 | -2.21 | NO3-N | ug N/L | 0.000934 | -0.00538 | -0.632 |  | 142
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00815 | 0.307 | 6.91 | PO4-P | ug P/L | -0.00851 | -0.0519 | -2.49 | NO3-N | ug N/L | 0.00257 | 0.0463 | 2.14 |  | 143
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | chk 0 | S6 | 1 | 1 | NH4-N | ug N/L | -0.0081 | -0.0265 | 0.0339 | PO4-P | ug P/L | 0.00943 | 0.19 | 2.77 | NO3-N | ug N/L | 0.00105 | 0.000918 | -0.294 |  | 144
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00764 | 0.109 | 2.72 | PO4-P | ug P/L | 0.00972 | 0.19 | 1.47 | NO3-N | ug N/L | -0.00249 | -0.0183 | -1.22 |  | 145
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00512 | 0.122 | 3.01 | PO4-P | ug P/L | 0.00913 | 0.18 | 1.25 | NO3-N | ug N/L | 0.00101 | -0.00105 | -0.28 |  | 146
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00764 | 0.0755 | 1.96 | PO4-P | ug P/L | 0.00783 | 0.148 | 0.515 | NO3-N | ug N/L | -0.00242 | -0.0173 | -1.16 |  | 147
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00962 | 0.193 | 4.59 | PO4-P | ug P/L | 0.00859 | 0.158 | 0.738 | NO3-N | ug N/L | -0.00144 | -0.0169 | -1.14 |  | 148
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | chk 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00318 | 0.057 | 1.55 | PO4-P | ug P/L | 0.00817 | 0.166 | 0.939 | NO3-N | ug N/L | 0.00125 | 0.0232 | 1.04 |  | 149
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.00773 | -0.0464 | -0.761 | PO4-P | ug P/L | -0.00803 | -0.0507 | -3.99 | NO3-N | ug N/L | -0.00289 | -0.0111 | -0.824 |  | 150
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00719 | 0.14 | 3.42 | PO4-P | ug P/L | 0.00916 | 0.178 | 1.21 | NO3-N | ug N/L | 0.00122 | 0.0142 | 0.547 |  | 151
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.0183 | -0.0755 | -1.41 | PO4-P | ug P/L | -0.00983 | -0.0706 | -4.45 | NO3-N | ug N/L | 0.00258 | 0.0285 | 1.32 |  | 152
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00327 | 0.0496 | 1.38 | PO4-P | ug P/L | 0.00965 | 0.194 | 1.58 | NO3-N | ug N/L | -0.00223 | -0.014 | -0.983 |  | 153
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00219 | 0.0141 | 0.591 | PO4-P | ug P/L | 0.00848 | 0.16 | 0.791 | NO3-N | ug N/L | 0.00225 | 0.0362 | 1.74 |  | 154
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | chk 0 | S6 | 1 | 1 | NH4-N | ug N/L | -0.00378 | -0.0983 | -1.92 | PO4-P | ug P/L | 0.0101 | 0.253 | 2.91 | NO3-N | ug N/L | 0.00244 | 0.0461 | 2.27 |  | 155
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00754 | 0.14 | 3.39 | PO4-P | ug P/L | 0.00893 | 0.187 | 0.218 | NO3-N | ug N/L | -0.00286 | -0.0158 | -0.133 |  | 156
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.0136 | -0.0912 | -1.43 | PO4-P | ug P/L | -0.00524 | 0.0341 | -3.24 | NO3-N | ug N/L | 0.00243 | 0.00489 | 0.848 |  | 157
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.0144 | 0.209 | 4.84 | PO4-P | ug P/L | 0.00793 | 0.154 | -0.544 | NO3-N | ug N/L | -0.00176 | 0.0125 | 1.21 |  | 158
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00499 | 0.232 | 5.3 | PO4-P | ug P/L | 0.00852 | 0.175 | -0.0676 | NO3-N | ug N/L | 0.00101 | 0.0101 | 1.1 |  | 159
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00468 | 0.139 | 3.37 | PO4-P | ug P/L | 0.00905 | 0.188 | 0.238 | NO3-N | ug N/L | -0.00395 | -0.0168 | -0.183 |  | 160
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | chk 0 | S6 | 1 | 1 | NH4-N | ug N/L | 0.00556 | 0.131 | 3.21 | PO4-P | ug P/L | 0.0082 | 0.16 | -0.396 | NO3-N | ug N/L | 0.00155 | 0.0359 | 2.32 |  | 161
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | -0.0123 | -0.0932 | -1.47 | PO4-P | ug P/L | -0.0076 | -0.0437 | -5 | NO3-N | ug N/L | -0.00208 | -0.00631 | 0.316 |  | 162
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.0127 | 0.0602 | 1.73 | PO4-P | ug P/L | -0.00731 | -0.0433 | -4.99 | NO3-N | ug N/L | 0.00833 | 0.04 | 2.51 |  | 163
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.0474 | 0.293 | 6.58 | PO4-P | ug P/L | 0.0095 | 0.17 | -0.183 | NO3-N | ug N/L | 0.0042 | 0.0708 | 3.98 |  | 164
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00558 | 0.0386 | 1.28 | PO4-P | ug P/L | 0.00869 | 0.171 | -0.159 | NO3-N | ug N/L | 0.00323 | 0.0559 | 3.27 |  | 165
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00586 | 0.0787 | 2.11 | PO4-P | ug P/L | 0.00895 | 0.182 | 0.1 | NO3-N | ug N/L | 0.00291 | 0.0561 | 3.28 |  | 166
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | rb | S6 | 1 | 1 | NH4-N | ug N/L | 0.00693 | 0.202 | 4.69 | PO4-P | ug P/L | 0.00802 | 0.165 | -0.3 | NO3-N | ug N/L | -0.00172 | 0.01 | 1.09 |  | 167
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | chk 0 | S6 | 1 | 1 | NH4-N | ug N/L | -0.00532 | -0.149 | -2.63 | PO4-P | ug P/L | 0.00811 | 0.176 | -0.0401 | NO3-N | ug N/L | -0.0028 | -0.0202 | -0.343 |  | 168
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.0145 | 0.119 | 11.2 | PO4-P | ug P/L | -0.00758 | -0.0462 | -4.82 | NO3-N | ug N/L | 0.00313 | 0.007 | -1.76 |  | 169
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.0184 | 0.11 | 11 | PO4-P | ug P/L | 0.00796 | 0.166 | 0.112 | NO3-N | ug N/L | 0.00611 | 0.166 | 5.67 |  | 170
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.021 | 0.0738 | 10.4 | PO4-P | ug P/L | -0.00737 | -0.0425 | -4.73 | NO3-N | ug N/L | 0.000969 | -0.00598 | -2.37 |  | 171
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.00401 | 0.0625 | 10.2 | PO4-P | ug P/L | 0.0104 | 0.228 | 1.57 | NO3-N | ug N/L | 0.00278 | 0.0468 | 0.104 |  | 172
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.00831 | -0.0901 | 7.59 | PO4-P | ug P/L | 0.00854 | 0.188 | 0.626 | NO3-N | ug N/L | 0.00343 | 0.0866 | 1.97 |  | 173
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | chk 0 | S7 | 1 | 1 | NH4-N | ug N/L | 0.0233 | 0.137 | 11.5 | PO4-P | ug P/L | 0.00755 | 0.153 | -0.176 | NO3-N | ug N/L | 0.0017 | 0.013 | -1.48 |  | 174
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.0131 | 0.038 | 9.78 | PO4-P | ug P/L | 0.00989 | 0.181 | 0.464 | NO3-N | ug N/L | -0.00205 | -0.0182 | -2.95 |  | 175
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.00704 | -0.15 | 6.56 | PO4-P | ug P/L | 0.0099 | 0.198 | 0.868 | NO3-N | ug N/L | 0.00377 | 0.0749 | 1.42 |  | 176
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.0126 | 0.212 | 12.8 | PO4-P | ug P/L | 0.00922 | 0.172 | 0.257 | NO3-N | ug N/L | -0.00319 | -0.0251 | -3.27 |  | 177
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | chk 0 | S7 | 1 | 1 | NH4-N | ug N/L | 0.0303 | 0.164 | 11.9 | PO4-P | ug P/L | 0.00727 | 0.143 | -0.412 | NO3-N | ug N/L | -0.00284 | -0.0217 | -3.11 |  | 178
2021-09-28 | OM_9-28-2021_10-45-41AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.0232 | 0.353 | 6.7 | PO4-P | ug P/L | 0.00861 | 0.164 | 0.0358 | NO3-N | ug N/L | 0.00114 | 0.0145 | -0.398 |  | 179
2021-09-28 | OM_9-28-2021_10-45-41AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | -0.0111 | -0.0818 | -1.61 | PO4-P | ug P/L | 0.00905 | 0.174 | 0.284 | NO3-N | ug N/L | 0.00175 | 0.0148 | -0.38 |  | 180
2021-09-28 | OM_9-28-2021_10-45-41AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | -0.0172 | -0.295 | -5.69 | PO4-P | ug P/L | 0.00981 | 0.188 | 0.607 | NO3-N | ug N/L | -0.00191 | -0.0262 | -2.35 |  | 181
2021-09-28 | OM_9-28-2021_10-45-41AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | -0.0178 | -0.137 | -2.68 | PO4-P | ug P/L | -0.00833 | -0.0486 | -4.85 | NO3-N | ug N/L | 0.00196 | 0.0238 | 0.0521 |  | 182
2021-09-28 | OM_9-28-2021_10-45-41AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.0112 | 0.167 | 3.14 | PO4-P | ug P/L | 0.00774 | 0.126 | -0.832 | NO3-N | ug N/L | 0.00177 | 0.0393 | 0.793 |  | 183
2021-09-28 | OM_9-28-2021_10-45-41AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.00556 | 0.134 | 2.51 | PO4-P | ug P/L | 0.00943 | 0.205 | 0.983 | NO3-N | ug N/L | 0.00277 | 0.029 | 0.301 |  | 184
2021-09-28 | OM_9-28-2021_10-45-41AM edit1.omn | rb | S7 | 1 | 1 | NH4-N | ug N/L | 0.0209 | 0.286 | 5.42 | PO4-P | ug P/L | 0.00961 | 0.211 | 1.13 | NO3-N | ug N/L | 0.00332 | 0.0659 | 2.07 |  | 185
2021-09-28 | OM_9-28-2021_01-13-20PM edit1.omn | chk 0 | S7 | 1 | 1 | NH4-N | ug N/L | -0.00488 | -0.0809 | -1.6 | PO4-P | ug P/L | 0.00877 | 0.177 | 0.345 | NO3-N | ug N/L | 0.00157 | 0.0283 | 0.268 |  | 186
| Mean | [8 empty] | 1.2685 |  |  |  |  | -0.97412 |  |  |  |  | -0.70151 | Min | 1
| Std | [8 empty] | 3.8914 |  |  |  |  | 2.5509 |  |  |  |  | 2.1349 | Max | 186
| CV | [8 empty] | 306.7684 |  |  |  |  | -261.8665 |  |  |  |  | -304.3326
Warning Limits | Xbar + 2SD | [8 empty] | 9.0514 |  |  |  |  | 4.1277 |  |  |  |  | 3.5683
| Xbar - 2SD | [8 empty] | -6.5143 |  |  |  |  | -6.0759 |  |  |  |  | -4.9713
Control Limits | Xbar + 3SD | [8 empty] | 12.9428 |  |  |  |  | 6.6786 |  |  |  |  | 5.7032
| Xbar - 3SD | [8 empty] | -10.4057 |  |  |  |  | -8.6268 |  |  |  |  | -7.1062
Count | [9 empty] | 173 |  |  |  |  | 186 |  |  |  |  | 185
99th percentile.inc | [9 empty] | 11.612 |  |  |  |  | 2.805 |  |  |  |  | 3.392
Maximum | [9 empty] | 12.8 |  |  |  |  | 2.91 |  |  |  |  | 5.67

## rolling spikes
NEED TO SCREEN FOR SPIKE SIZE - large concentration and small spike a problem - eg stroubles creek nitrate samples and lake bottom ammonia
spike is usually ~20 ppb
flagged when the spiked sample concentration is more than 20 + 200 = 220
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Reduction Efficiency | Analyst Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | NH4-N NOTES: | NH4-N concentration | NH4-N- % Error in check stds | NH4-N- DQO (<=20%) | NH4-N- RPD of duplicates | NH4-N- DQO (<=20% if conc>LOQ) | NH4-N- PR of spikes | NH4-N- DQO (80-120%) if conc>LOQ | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BN PO4-P NOTES: | PO4-P concentration | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BN NO3-N NOTES: | NO3-N concentration | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spk | 15 | 1 | 1 | NH4-N | ug N/L | 0.0493 | 1.04 | 22.1 | PO4-P | ug P/L | 0.0944 | 1.53 | 34.5 | NO3-N+NO2-N | ug N/L | 0.204 | 3.32 | 200 |  | 1 |  | 22.1 |  |  |  |  | 96.96 | PASS | 1 |  | 34.5 |  |  |  |  | 93.95 | PASS | 1 |  | 200 |  |  |  |  | 130 | FAIL |  | 1
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 30 | 1 | 1 | NH4-N | ug N/L | 0.0617 | 1.37 | 29.4 | PO4-P | ug P/L | 0.0672 | 1.02 | 21.4 | NO3-N+NO2-N | ug N/L | 0.0247 | 0.416 | 23.2 |  | 1 |  | 29.4 |  |  |  |  | 96.94 | PASS | 1 |  | 21.4 |  |  |  |  | 83.14 | PASS | 1 |  | 23.2 |  |  |  |  | 100.57 | PASS |  | 2
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spk | 45 | 1 | 1 | NH4-N | ug N/L | 0.0591 | 1.24 | 26.5 | PO4-P | ug P/L | 0.0759 | 1.21 | 26.2 | NO3-N+NO2-N | ug N/L | 0.141 | 2.27 | 135 |  | 1 |  | 26.5 |  |  |  |  | 87.1 | PASS | 1 |  | 26.2 |  |  |  |  | 84.22 | PASS | 1 |  | 135 |  |  |  |  | 38.5 | FAIL |  | 3
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spk | 60 | 1 | 1 | NH4-N | ug N/L | 0.0663 | 1.44 | 30.9 | PO4-P | ug P/L | 0.0697 | 1.08 | 22.9 | NO3-N+NO2-N | ug N/L | 0.0202 | 0.339 | 18.6 |  | 1 |  | 30.9 |  |  |  |  | 103.09 | PASS | 1 |  | 22.9 |  |  |  |  | 87.39 | PASS | 1 |  | 18.6 |  |  |  |  | 92.53 | PASS |  | 4
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spk | 75 | 1 | 1 | NH4-N | ug N/L | 0.0628 | 1.37 | 29.4 | PO4-P | ug P/L | 0.0579 | 0.93 | 19 | NO3-N+NO2-N | ug N/L | 0.0208 | 0.366 | 20.2 |  | 1 |  | 29.4 |  |  |  |  | 123.19 | FAIL | 1 |  | 19 |  |  |  |  | 78.05 | FAIL | 1 |  | 20.2 |  |  |  |  | 90.02 | PASS |  | 5
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spk | 90 | 1 | 1 | NH4-N | ug N/L | 0.053 | 1.15 | 24.5 | PO4-P | ug P/L | 0.075 | 1.21 | 26.1 | NO3-N+NO2-N | ug N/L | 0.0189 | 0.324 | 17.7 |  | 1 |  | 24.5 | [6 empty] | 1 |  | 26.1 | [6 empty] | 1 |  | 17.7 | [7 empty] | 6
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spk | 105 | 1 | 1 | NH4-N | ug N/L | 0.0559 | 1.23 | 26.4 | PO4-P | ug P/L | 0.0675 | 1.06 | 22.5 | NO3-N+NO2-N | ug N/L | 0.0374 | 0.618 | 35.3 |  | 1 |  | 26.4 |  |  |  |  | 101.34 | PASS | 1 |  | 22.5 |  |  |  |  | 84.75 | PASS | 1 |  | 35.3 |  |  |  |  | 76.03 | FAIL |  | 7
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spk | 120 | 1 | 1 | NH4-N | ug N/L | 0.0542 | 1.07 | 22.8 | PO4-P | ug P/L | 0.0626 | 1.01 | 21.1 | NO3-N+NO2-N | ug N/L | 0.019 | 0.324 | 17.7 |  | 1 |  | 22.8 |  |  |  |  | 99.98 | PASS | 1 |  | 21.1 |  |  |  |  | 90.91 | PASS | 1 |  | 17.7 |  |  |  |  | 89.485 | PASS |  | 8
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spk | 135 | 1 | 1 | NH4-N | ug N/L | 0.056 | 1.24 | 26.5 | PO4-P | ug P/L | 0.0701 | 1.14 | 24.4 | NO3-N+NO2-N | ug N/L | 0.0386 | 0.648 | 37.1 |  | 1 |  | 26.5 |  |  |  |  | 109.25 | PASS | 1 |  | 24.4 |  |  |  |  | 98.14 | PASS | 1 |  | 37.1 |  |  |  |  | 86.21 | PASS |  | 9
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | spk | 150 | 1 | 1 | NH4-N | ug N/L | 0.0589 | 1.29 | 27.5 | PO4-P | ug P/L | 0.0699 | 1.13 | 24.2 | NO3-N+NO2-N | ug N/L | 0.0313 | 0.535 | 30.3 |  | 1 |  | 27.5 |  |  |  |  | 131 | FAIL | 1 |  | 24.2 |  |  |  |  | 89.92 | PASS | 1 |  | 30.3 |  |  |  |  | 79.53 | FAIL |  | 10
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | F 6.2m 06jul20 SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.0985 | 1.63 | 31.4 | PO4-P | ug P/L | 0.0622 | 1.01 | 21.8 | NO3-N+NO2-N | ug N/L | 0.0223 | 0.391 | 21.5 |  | 1 |  | 31.4 |  |  |  |  | 54.14 | FAIL | 1 |  | 21.8 |  |  |  |  | 82.18 | PASS | 1 |  | 21.5 |  |  |  |  | 94.5 | PASS |  | 11
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | F 3.8m 20jul20 SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.0636 | 1.31 | 25.2 | PO4-P | ug P/L | 0.0687 | 1.11 | 24.6 | NO3-N+NO2-N | ug N/L | 0.0195 | 0.312 | 17 |  | 1 |  | 25.2 |  |  |  |  | 97.32 | PASS | 1 |  | 24.6 |  |  |  |  | 92.01 | PASS | 1 |  | 17 |  |  |  |  | 77.45 | PASS |  | 12
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 6m 30jul20 SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.0738 | 1.8 | 34.7 | PO4-P | ug P/L | 0.0689 | 1.06 | 23.3 | NO3-N+NO2-N | ug N/L | 0.0211 | 0.342 | 18.7 |  | 1 |  | 34.7 |  |  |  |  | 88.47 | PASS | 1 |  | 23.3 |  |  |  |  | 91.28 | PASS | 1 |  | 18.7 |  |  |  |  | 88.82 | PASS |  | 13
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 0.1m 14may20 SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.0911 | 1.99 | 38.5 | PO4-P | ug P/L | 0.0567 | 0.893 | 18.9 | NO3-N+NO2-N | ug N/L | 0.0174 | 0.297 | 16.1 |  | 1 |  | 38.5 |  |  |  |  | 149.35 | PASS | 1 |  | 18.9 |  |  |  |  | 76.64 | PASS | 1 |  | 16.1 |  |  |  |  | 78.865 | PASS |  | 14
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 3m30jul20 SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0647 | 1.43 | 27.6 | PO4-P | ug P/L | 0.0677 | 1.08 | 23.8 | NO3-N+NO2-N | ug N/L | 0.0195 | 0.331 | 18.1 |  | 1 |  | 27.6 |  |  |  |  | 95.86 | PASS | 1 |  | 23.8 |  |  |  |  | 83.58 | PASS | 1 |  | 18.1 |  |  |  |  | 75.41 | PASS |  | 15
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | F 3.8m 31jan20 SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.0696 | 1.33 | 25.7 | PO4-P | ug P/L | 0.0679 | 1.09 | 24 | NO3-N+NO2-N | ug N/L | 0.0171 | 0.287 | 15.5 |  | 1 |  | 25.7 |  |  |  |  | 105.52 | PASS | 1 |  | 24 |  |  |  |  | 92 | PASS | 1 |  | 15.5 |  |  |  |  | 71.3 | PASS |  | 16
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | F 6.2m 22jun20 SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.103 | 2.09 | 40.4 | PO4-P | ug P/L | 0.047 | 0.721 | 14.3 | NO3-N+NO2-N | ug N/L | 0.0185 | 0.307 | 16.7 |  | 1 |  | 40.4 |  |  |  |  | 116.04 | PASS | 1 |  | 14.3 |  |  |  |  | 43.58 | PASS | 1 |  | 16.7 |  |  |  |  | 58.82 | PASS |  | 17
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 6m 12aug20 SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.147 | 3.33 | 64.3 | PO4-P | ug P/L | 0.0614 | 0.983 | 21.2 | NO3-N+NO2-N | ug N/L | 0.018 | 0.312 | 17 |  | 1 |  | 64.3 |  |  |  |  | 83.93 | PASS | 1 |  | 21.2 |  |  |  |  | 80.72 | PASS | 1 |  | 17 |  |  |  |  | 70.7 | PASS |  | 18
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | F 50 4.5m 20apr20 SPK | 150 | 1 | 1 | NH4-N | ug N/L | 0.102 | 2.34 | 45.1 | PO4-P | ug P/L | 0.0599 | 0.941 | 20.1 | NO3-N+NO2-N | ug N/L | 0.0294 | 0.49 | 27.2 |  | 1 |  | 45.1 |  |  |  |  | 123.51 | FAIL | 1 |  | 20.1 |  |  |  |  | 81.96 | PASS | 1 |  | 27.2 |  |  |  |  | 75.72 | FAIL |  | 19
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spk | 15 | 1 | 1 | NH4-N | ug N/L | 0.0489 | 1.06 | 24.4 | PO4-P | ug P/L | 0.0995 | 1.51 | 31.1 | NO3-N | ug N/L | 0.131 | 1.6 | 45.4 |  | 1 |  | 24.4 |  |  |  |  | 99.99 | PASS | 1 |  | 31.1 |  |  |  |  | 92.61 | PASS | 1 |  | 45.4 |  |  |  |  | 73.54 | FAIL |  | 20
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spk | 30 | 1 | 1 | NH4-N | ug N/L | 1.73 | 36.2 | 767 | PO4-P | ug P/L | 0.0436 | 0.756 | 14.5 | NO3-N | ug N/L | 0.0599 | 0.766 | 21.7 |  | 1 |  | 767 |  |  |  |  | 46.7 | FAIL | 1 |  | 14.5 |  |  |  |  | 59.75 | PASS | 1 |  | 21.7 |  |  |  |  | 80.37 | PASS |  | 21
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spk | 45 | 1 | 1 | NH4-N | ug N/L | 0.0682 | 1.75 | 39.3 | PO4-P | ug P/L | 0.0758 | 1.27 | 25.7 | NO3-N | ug N/L | 0.069 | 0.848 | 24 |  | 1 |  | 39.3 |  |  |  |  | 97.93 | PASS | 1 |  | 25.7 |  |  |  |  | 90.42 | PASS | 1 |  | 24 |  |  |  |  | 96.95 | PASS |  | 22
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spk | 60 | 1 | 1 | NH4-N | ug N/L | 0.102 | 2.3 | 51.2 | PO4-P | ug P/L | 0.0693 | 1.14 | 23 | NO3-N | ug N/L | 0.0612 | 0.796 | 22.5 |  | 1 |  | 51.2 | [6 empty] | 1 |  | 23 | [6 empty] | 1 |  | 22.5 | [7 empty] | 23
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spk | 75 | 1 | 1 | NH4-N | ug N/L | 0.269 | 6.02 | 131 | PO4-P | ug P/L | 0.0573 | 0.999 | 19.8 | NO3-N | ug N/L | 0.0652 | 0.844 | 23.9 |  | 1 |  | 131 |  |  |  |  | 93.1 | PASS | 1 |  | 19.8 |  |  |  |  | 79.83 | PASS | 1 |  | 23.9 |  |  |  |  | 96.79 | PASS |  | 24
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spk | 90 | 1 | 1 | NH4-N | ug N/L | 2.34 | 47.7 | 1000 | PO4-P | ug P/L | 0.0634 | 1.06 | 21.2 | NO3-N | ug N/L | 0.0638 | 0.851 | 24.1 |  | 1 |  | 1000 |  |  |  |  | 100 | PASS | 1 |  | 21.2 |  |  |  |  | 93.67 | PASS | 1 |  | 24.1 |  |  |  |  | 84.81 | PASS |  | 25
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spk | 105 | 1 | 1 | NH4-N | ug N/L | 0.154 | 3.34 | 73.6 | PO4-P | ug P/L | 0.0551 | 0.938 | 18.5 | NO3-N | ug N/L | 0.0827 | 1.07 | 30.3 |  | 1 |  | 73.6 |  |  |  |  | 95.86 | PASS | 1 |  | 18.5 |  |  |  |  | 71.1 | PASS | 1 |  | 30.3 |  |  |  |  | 101.03 | PASS |  | 26
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spk | 120 | 1 | 1 | NH4-N | ug N/L | 1.58 | 34.6 | 734 | PO4-P | ug P/L | 0.0525 | 0.895 | 17.6 | NO3-N | ug N/L | 0.0629 | 0.827 | 23.4 |  | 1 |  | 734 |  |  |  |  | 78.4 | FAIL | 1 |  | 17.6 |  |  |  |  | 70.56 | PASS | 1 |  | 23.4 |  |  |  |  | 93.39 | PASS |  | 27
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spk | 135 | 1 | 1 | NH4-N | ug N/L | 0.0889 | 2.19 | 48.8 | PO4-P | ug P/L | 0.0783 | 1.31 | 26.5 | NO3-N | ug N/L | 0.0826 | 1 | 28.5 |  | 1 |  | 48.8 |  |  |  |  | 116.88 | PASS | 1 |  | 26.5 |  |  |  |  | 97.7 | PASS | 1 |  | 28.5 |  |  |  |  | 98.9 | PASS |  | 28
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | spk | 150 | 1 | 1 | NH4-N | ug N/L | 0.156 | 3.49 | 77 | PO4-P | ug P/L | 0.0567 | 0.945 | 18.7 | NO3-N | ug N/L | 0.0618 | 0.783 | 22.2 |  | 1 |  | 77 |  |  |  |  | 97.2 | PASS | 1 |  | 18.7 |  |  |  |  | 87.02 | PASS | 1 |  | 22.2 |  |  |  |  | 90.82 | PASS |  | 29
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.146 | 3.13 | 63.1 | PO4-P | ug P/L | 0.074 | 1.21 | 24.7 | NO3-N | ug N/L | 0.136 | 1.65 | 42.2 |  | 1 |  | 63.1 |  |  |  |  | 81.31 | PASS | 1 |  | 24.7 |  |  |  |  | 84.42 | PASS | 1 |  | 42.2 |  |  |  |  | 106.22 | PASS |  | 30
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.0538 | 1.22 | 23.9 | PO4-P | ug P/L | 0.064 | 1.05 | 21.5 | NO3-N | ug N/L | 0.0675 | 0.832 | 20.8 |  | 1 |  | 23.9 |  |  |  |  | 106.39 | PASS | 1 |  | 21.5 |  |  |  |  | 82 | PASS | 1 |  | 20.8 |  |  |  |  | 92.63 | PASS |  | 31
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.0521 | 1.19 | 23.3 | PO4-P | ug P/L | 0.0652 | 1.09 | 22.2 | NO3-N | ug N/L | 0.162 | 1.92 | 49.3 |  | 1 |  | 23.3 |  |  |  |  | 105.13 | PASS | 1 |  | 22.2 |  |  |  |  | 81.07 | PASS | 1 |  | 49.3 |  |  |  |  | 92.93 | PASS |  | 32
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 4.66 | 94.9 | 1720 | PO4-P | ug P/L | 0.22 | 3.22 | 66.7 | NO3-N | ug N/L | 0.0811 | 1.03 | 25.8 |  | 3 | oor | 1720 | [6 empty] | 1 |  | 66.7 |  |  |  |  | 16.67 | FAIL | 1 |  | 25.8 |  |  |  |  | 91.78 | PASS |  | 33
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.313 | 6.9 | 140 | PO4-P | ug P/L | 0.0733 | 1.2 | 24.5 | NO3-N | ug N/L | 0.0707 | 0.845 | 21.1 |  | 1 |  | 140 |  |  |  |  | 129 | FAIL | 1 |  | 24.5 |  |  |  |  | 80.45 | PASS | 1 |  | 21.1 |  |  |  |  | 89.71 | PASS |  | 34
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 3.22 | 60.4 | 1150 | PO4-P | ug P/L | 0.0612 | 1.02 | 20.8 | NO3-N | ug N/L | 0.0864 | 1.04 | 26.2 |  | 1 |  | 1150 |  |  |  |  | 165 | FAIL | 1 |  | 20.8 |  |  |  |  | 76.48 | PASS | 1 |  | 26.2 |  |  |  |  | 89.32 | PASS |  | 35
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.0768 | 1.93 | 36.6 | PO4-P | ug P/L | 0.0613 | 1 | 20.4 | NO3-N | ug N/L | 0.106 | 1.29 | 32.8 |  | 1 |  | 36.6 |  |  |  |  | 137.86 | PASS | 1 |  | 20.4 |  |  |  |  | 74.64 | PASS | 1 |  | 32.8 |  |  |  |  | 97.78 | PASS |  | 36
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.219 | 4.86 | 98.3 | PO4-P | ug P/L | 0.0658 | 1.08 | 22.1 | NO3-N | ug N/L | 0.149 | 1.8 | 46.3 |  | 1 |  | 98.3 |  |  |  |  | 101.33 | PASS | 1 |  | 22.1 |  |  |  |  | 76.91 | PASS | 1 |  | 46.3 |  |  |  |  | 81.13 | PASS |  | 37
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.0596 | 1.41 | 27.7 | PO4-P | ug P/L | 0.0652 | 1.08 | 22 | NO3-N | ug N/L | 0.0709 | 0.867 | 21.7 |  | 1 |  | 27.7 |  |  |  |  | 110.87 | PASS | 1 |  | 22 |  |  |  |  | 80.1 | PASS | 1 |  | 21.7 |  |  |  |  | 87.17 | PASS |  | 38
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.0527 | 1.14 | 22.8 | PO4-P | ug P/L | 0.0709 | 1.22 | 22.8 | NO3-N | ug N/L | 0.035 | 0.451 | 21.6 |  | 1 |  | 22.8 |  |  |  |  | 85.83 | PASS | 1 |  | 22.8 |  |  |  |  | 78.98 | PASS | 1 |  | 21.6 |  |  |  |  | 92.51 | PASS |  | 39
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.0525 | 1.22 | 24.5 | PO4-P | ug P/L | 0.089 | 1.32 | 24.9 | NO3-N | ug N/L | 0.0589 | 0.734 | 36.6 |  | 1 |  | 24.5 |  |  |  |  | 93.1 | PASS | 1 |  | 24.9 |  |  |  |  | 115.29 | PASS | 1 |  | 36.6 |  |  |  |  | 92.16 | PASS |  | 40
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.0724 | 1.72 | 34.7 | PO4-P | ug P/L | 0.0667 | 1.12 | 20.6 | NO3-N | ug N/L | 0.0362 | 0.461 | 22.2 |  | 1 |  | 34.7 |  |  |  |  | 133.02 | PASS | 1 |  | 20.6 |  |  |  |  | 94.51 | PASS | 1 |  | 22.2 |  |  |  |  | 93.42 | PASS |  | 41
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.0512 | 1.16 | 23.3 | PO4-P | ug P/L | 0.0675 | 1.14 | 20.9 | NO3-N | ug N/L | 0.0404 | 0.529 | 25.7 |  | 1 |  | 23.3 |  |  |  |  | 86.73 | PASS | 1 |  | 20.9 |  |  |  |  | 97.79 | PASS | 1 |  | 25.7 |  |  |  |  | 106.22 | PASS |  | 42
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.523 | 11.3 | 228 | PO4-P | ug P/L | 0.1 | 1.7 | 33.5 | NO3-N | ug N/L | 0.0337 | 0.443 | 21.2 |  | 1 |  | 228 |  |  |  |  | 97.8 | PASS | 1 |  | 33.5 |  |  |  |  | 97.85 | PASS | 1 |  | 21.2 |  |  |  |  | 94.67 | PASS |  | 43
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0906 | 1.85 | 37.1 | PO4-P | ug P/L | 0.0802 | 1.33 | 25.2 | NO3-N | ug N/L | 0.127 | 1.64 | 84.3 |  | 1 |  | 37.1 |  |  |  |  | 70.71 | FAIL | 1 |  | 25.2 |  |  |  |  | 84.67 | PASS | 1 |  | 84.3 |  |  |  |  | 89.43 | PASS |  | 44
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.0527 | 1.24 | 24.8 | PO4-P | ug P/L | 0.0663 | 1.16 | 21.3 | NO3-N | ug N/L | 0.031 | 0.403 | 19.1 |  | 1 |  | 24.8 |  |  |  |  | 89.03 | PASS | 1 |  | 21.3 |  |  |  |  | 103.53 | PASS | 1 |  | 19.1 |  |  |  |  | 89.31 | PASS |  | 45
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.0601 | 1.22 | 24.5 | PO4-P | ug P/L | 0.0702 | 1.23 | 22.9 | NO3-N | ug N/L | 0.0312 | 0.409 | 19.4 |  | 1 |  | 24.5 |  |  |  |  | 82.6 | PASS | 1 |  | 22.9 |  |  |  |  | 107.34 | PASS | 1 |  | 19.4 |  |  |  |  | 92.54 | PASS |  | 46
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.308 | 6.62 | 134 | PO4-P | ug P/L | 0.0637 | 1.09 | 19.9 | NO3-N | ug N/L | 0.0351 | 0.48 | 23.2 |  | 1 |  | 134 |  |  |  |  | 88.4 | PASS | 1 |  | 19.9 |  |  |  |  | 106.89 | PASS | 1 |  | 23.2 |  |  |  |  | 93.12 | PASS |  | 47
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | SPK | 150 | 1 | 1 | NH4-N | ug N/L | 0.0631 | 1.23 | 24.6 | PO4-P | ug P/L | 0.0689 | 1.19 | 22.1 | NO3-N | ug N/L | 0.0287 | 0.406 | 19.2 |  | 1 |  | 24.6 |  |  |  |  | 103.21 | PASS | 1 |  | 22.1 |  |  |  |  | 91.51 | PASS | 1 |  | 19.2 |  |  |  |  | 97.22 | PASS |  | 48
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.41 | 8.85 | 183 | PO4-P | ug P/L | 0.0546 | 0.913 | 18.5 | NO3-N | ug N/L | 0.0305 | 0.389 | 20.6 |  | 1 |  | 183 |  |  |  |  | 108.3 | PASS | 1 |  | 18.5 |  |  |  |  | 73.75 | PASS | 1 |  | 20.6 |  |  |  |  | 89.96 | PASS |  | 49
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.123 | 2.68 | 55.7 | PO4-P | ug P/L | 0.0842 | 1.42 | 29.5 | NO3-N | ug N/L | 0.0314 | 0.405 | 21.4 |  | 1 |  | 55.7 |  |  |  |  | 88.57 | PASS | 1 |  | 29.5 |  |  |  |  | 90.95 | PASS | 1 |  | 21.4 |  |  |  |  | 92.64 | PASS |  | 50
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.05 | 1.11 | 23.4 | PO4-P | ug P/L | 0.0646 | 1.08 | 22.2 | NO3-N | ug N/L | 0.0315 | 0.396 | 20.9 |  | 1 |  | 23.4 |  |  |  |  | 71.99 | PASS | 1 |  | 22.2 |  |  |  |  | 87.67 | PASS | 1 |  | 20.9 |  |  |  |  | 102.82 | PASS |  | 51
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.0671 | 1.49 | 31.3 | PO4-P | ug P/L | 0.0597 | 1.01 | 20.7 | NO3-N | ug N/L | 0.0341 | 0.415 | 21.9 |  | 1 |  | 31.3 |  |  |  |  | 82.63 | PASS | 1 |  | 20.7 |  |  |  |  | 86.62 | PASS | 1 |  | 21.9 |  |  |  |  | 91.39 | PASS |  | 52
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.0726 | 1.45 | 30.4 | PO4-P | ug P/L | 0.092 | 1.5 | 31.2 | NO3-N | ug N/L | 0.0298 | 0.373 | 19.7 |  | 1 |  | 30.4 |  |  |  |  | 100.04 | PASS | 1 |  | 31.2 |  |  |  |  | 83.12 | PASS | 1 |  | 19.7 |  |  |  |  | 88.57 | PASS |  | 53
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.0604 | 1.24 | 26 | PO4-P | ug P/L | 0.0592 | 0.981 | 20 | NO3-N | ug N/L | 0.0311 | 0.383 | 20.2 |  | 1 |  | 26 |  |  |  |  | 80.1 | PASS | 1 |  | 20 |  |  |  |  | 83.6 | PASS | 1 |  | 20.2 |  |  |  |  | 83.87 | PASS |  | 54
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.0569 | 1.22 | 25.7 | PO4-P | ug P/L | 0.0659 | 1.07 | 22 | NO3-N | ug N/L | 0.0331 | 0.428 | 22.6 |  | 1 |  | 25.7 |  |  |  |  | 96.42 | PASS | 1 |  | 22 |  |  |  |  | 88.7 | PASS | 1 |  | 22.6 |  |  |  |  | 92.01 | PASS |  | 55
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.0673 | 1.45 | 32.6 | PO4-P | ug P/L | 0.0902 | 1.58 | 33 | NO3-N | ug N/L | 1.87 | 23.7 | 1280 |  | 1 |  | 32.6 |  |  |  |  | 104.26 | PASS | 1 |  | 33 |  |  |  |  | 93.8 | PASS | 1 |  | 1280 |  |  |  |  | -22 | FAIL |  | 56
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.0652 | 1.3 | 29.3 | PO4-P | ug P/L | 0.0825 | 1.44 | 29.9 | NO3-N | ug N/L | 1.86 | 23.5 | 1270 |  | 1 |  | 29.3 |  |  |  |  | 87.93 | PASS | 1 |  | 29.9 |  |  |  |  | 95.49 | PASS | 1 |  | 1270 |  |  |  |  | 77 | FAIL |  | 57
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.0601 | 1.27 | 28.6 | PO4-P | ug P/L | 0.0774 | 1.35 | 27.9 | NO3-N | ug N/L | 1.83 | 23.4 | 1260 |  | 1 |  | 28.6 |  |  |  |  | 104.51 | PASS | 1 |  | 27.9 |  |  |  |  | 92.89 | PASS | 1 |  | 1260 |  |  |  |  | -24 | FAIL |  | 58
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.0766 | 1.49 | 33.5 | PO4-P | ug P/L | 0.0839 | 1.47 | 30.5 | NO3-N | ug N/L | 1.83 | 23.3 | 1260 |  | 1 |  | 33.5 |  |  |  |  | 104.35 | PASS | 1 |  | 30.5 |  |  |  |  | 92.05 | PASS | 1 |  | 1260 |  |  |  |  | -24 | FAIL |  | 59
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.117 | 2.52 | 56.6 | PO4-P | ug P/L | 0.106 | 1.72 | 36.3 | NO3-N | ug N/L | 1.9 | 24.2 | 1310 |  | 1 |  | 56.6 | [6 empty] | 1 |  | 36.3 | [6 empty] | 1 |  | 1310 | [7 empty] | 60
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0898 | 1.7 | 38.3 | PO4-P | ug P/L | 0.0791 | 1.36 | 28.1 | NO3-N | ug N/L | 1.84 | 23.5 | 1270 |  | 1 |  | 38.3 |  |  |  |  | 76.83 | FAIL | 1 |  | 28.1 |  |  |  |  | 90.31 | PASS | 1 |  | 1270 |  |  |  |  | 127 | FAIL |  | 61
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.104 | 2.23 | 50.1 | PO4-P | ug P/L | 0.0803 | 1.35 | 27.8 | NO3-N | ug N/L | 1.84 | 23.3 | 1260 |  | 1 |  | 50.1 |  |  |  |  | 97.01 | PASS | 1 |  | 27.8 |  |  |  |  | 98.38 | PASS | 1 |  | 1260 |  |  |  |  | -24 | FAIL |  | 62
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.0636 | 1.23 | 27.8 | PO4-P | ug P/L | 0.0843 | 1.59 | 33.3 | NO3-N | ug N/L | 1.47 | 18.8 | 1010 |  | 1 |  | 27.8 |  |  |  |  | 89.28 | PASS | 1 |  | 33.3 |  |  |  |  | 116.83 | PASS | 1 |  | 1010 |  |  |  |  | -149 | FAIL |  | 63
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.0647 | 1.25 | 28.2 | PO4-P | ug P/L | 0.0792 | 1.39 | 28.8 | NO3-N | ug N/L | 1.45 | 18.6 | 1000 |  | 1 |  | 28.2 |  |  |  |  | 93.32 | PASS | 1 |  | 28.8 |  |  |  |  | 80.88 | PASS | 1 |  | 1000 |  |  |  |  | -100 | FAIL |  | 64
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | SPK | 150 | 1 | 1 | NH4-N | ug N/L | 0.125 | 2.69 | 60.3 | PO4-P | ug P/L | 0.187 | 3.12 | 68.1 | NO3-N | ug N/L | 1.42 | 18.2 | 987 |  | 1 |  | 60.3 |  |  |  |  | 96.03 | PASS | 1 |  | 68.1 |  |  |  |  | 73.31 | FAIL | 1 |  | 987 |  |  |  |  | -16.3 | FAIL |  | 65
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.103 | 2.24 | 47.1 | PO4-P | ug P/L | 0.0655 | 1.15 | 22.1 | NO3-N | ug N/L | 1.45 | 18.5 | 881 |  | 1 |  | 47.1 |  |  |  |  | 108.71 | PASS | 1 |  | 22.1 |  |  |  |  | 87.51 | PASS | 1 |  | 881 |  |  |  |  | 48.1 | FAIL |  | 66
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.343 | 7.37 | 154 | PO4-P | ug P/L | 0.196 | 3.28 | 70.1 | NO3-N | ug N/L | 4.45 | 57.2 | 2750 |  | 1 |  | 154 |  |  |  |  | 110.4 | PASS | 1 |  | 70.1 |  |  |  |  | 98.01 | PASS | 1 |  | 2750 |  |  |  |  | 75 | FAIL |  | 67
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.0662 | 1.34 | 28.4 | PO4-P | ug P/L | 0.0675 | 1.19 | 22.8 | NO3-N | ug N/L | 1.71 | 21.9 | 1050 |  | 1 |  | 28.4 |  |  |  |  | 93.34 | PASS | 1 |  | 22.8 |  |  |  |  | 82.63 | PASS | 1 |  | 1050 |  |  |  |  | -95 | FAIL |  | 68
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.123 | 2.65 | 55.7 | PO4-P | ug P/L | 0.0639 | 1.1 | 20.7 | NO3-N | ug N/L | 1.26 | 16.1 | 767 |  | 1 |  | 55.7 |  |  |  |  | 55.07 | FAIL | 1 |  | 20.7 |  |  |  |  | 78.32 | PASS | 1 |  | 767 |  |  |  |  | 236.7 | FAIL |  | 69
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.0674 | 1.41 | 29.8 | PO4-P | ug P/L | 0.0629 | 1.11 | 21.1 | NO3-N | ug N/L | 0.606 | 7.77 | 370 |  | 1 |  | 29.8 | [6 empty] | 1 |  | 21.1 | [6 empty] | 1 |  | 370 | [7 empty] | 70
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0627 | 1.34 | 28.3 | PO4-P | ug P/L | 0.0622 | 1.1 | 20.9 | NO3-N | ug N/L | 0.0626 | 0.798 | 38.5 | 2932.3308 | 1 |  | 28.3 | [6 empty] | 1 |  | 20.9 | [6 empty] | 1 |  | 38.5 | [7 empty] | 71
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.117 | 2.61 | 54.7 | PO4-P | ug P/L | 0.0618 | 1.08 | 20.3 | NO3-N | ug N/L | 1.28 | 16.3 | 780 |  | 1 |  | 54.7 |  |  |  |  | 113.47 | PASS | 1 |  | 20.3 |  |  |  |  | 84.38 | PASS | 1 |  | 780 |  |  |  |  | 148 | FAIL |  | 72
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.0824 | 1.75 | 36.8 | PO4-P | ug P/L | 0.0671 | 1.18 | 22.6 | NO3-N | ug N/L | 0.979 | 12.6 | 598 |  | 1 |  | 36.8 |  |  |  |  | 70.68 | FAIL | 1 |  | 22.6 |  |  |  |  | 88.21 | PASS | 1 |  | 598 |  |  |  |  | 19.8 | FAIL |  | 73
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.0764 | 1.59 | 33.6 | PO4-P | ug P/L | 0.0582 | 1.03 | 19.3 | NO3-N | ug N/L | 0.734 | 9.4 | 448 |  | 1 |  | 33.6 |  |  |  |  | 111.86 | PASS | 1 |  | 19.3 |  |  |  |  | 82.93 | PASS | 1 |  | 448 |  |  |  |  | 59.8 | FAIL |  | 74
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 150 | 1 | 1 | NH4-N | ug N/L | 0.0712 | 1.51 | 31.9 | PO4-P | ug P/L | 0.0636 | 1.11 | 21.1 | NO3-N | ug N/L | 0.661 | 8.51 | 405 |  | 1 |  | 31.9 |  |  |  |  | 100.19 | PASS | 1 |  | 21.1 |  |  |  |  | 87.46 | PASS | 1 |  | 405 |  |  |  |  | 25.5 | FAIL |  | 75
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 165 | 1 | 1 | NH4-N | ug N/L | 0.0913 | 1.65 | 34.8 | PO4-P | ug P/L | 0.0823 | 1.4 | 27.6 | NO3-N | ug N/L | 1.47 | 19 | 904 |  | 1 |  | 34.8 |  |  |  |  | 77.98 | FAIL | 1 |  | 27.6 |  |  |  |  | 87.76 | PASS | 1 |  | 904 |  |  |  |  | 110.4 | PASS |  | 76
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 180 | 1 | 1 | NH4-N | ug N/L | 0.0833 | 2.26 | 47.5 | PO4-P | ug P/L | 0.0953 | 1.48 | 29.5 | NO3-N | ug N/L | 1.65 | 21 | 1000 |  | 1 |  | 47.5 |  |  |  |  | 167.25 | FAIL | 1 |  | 29.5 |  |  |  |  | 86.45 | PASS | 1 |  | 1000 |  |  |  |  | -150 | FAIL |  | 77
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | SPK | 195 | 1 | 1 | NH4-N | ug N/L | 0.0852 | 1.97 | 41.4 | PO4-P | ug P/L | 0.0927 | 1.59 | 32 | NO3-N | ug N/L | 1.56 | 20 | 956 |  | 1 |  | 41.4 |  |  |  |  | 111.14 | PASS | 1 |  | 32 |  |  |  |  | 94.7 | PASS | 1 |  | 956 |  |  |  |  | 75.6 | FAIL |  | 78
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.199 | 2.6 | 53.6 | PO4-P | ug P/L | 0.0634 | 1.14 | 22.8 | NO3-N | ug N/L | 3.52 | 45.5 | 2240 |  | 1 |  | 53.6 |  |  |  |  | 88.86 | PASS | 1 |  | 22.8 |  |  |  |  | 89.28 | PASS | 1 |  | 2240 |  |  |  |  | -76 | FAIL |  | 79
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.976 | 20.3 | 360 | PO4-P | ug P/L | 0.265 | 4.38 | 98.3 | NO3-N | ug N/L | 3.2 | 41.2 | 2020 |  | 1 |  | 360 |  |  |  |  | 71 | FAIL | 1 |  | 98.3 |  |  |  |  | 108.33 | PASS | 1 |  | 2020 |  |  |  |  | -48 | FAIL |  | 80
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.119 | 2.42 | 50.5 | PO4-P | ug P/L | 0.0786 | 1.32 | 26.9 | NO3-N | ug N/L | 1.42 | 18.1 | 865 |  | 1 |  | 50.5 |  |  |  |  | 92.55 | PASS | 1 |  | 26.9 |  |  |  |  | 89.84 | PASS | 1 |  | 865 |  |  |  |  | 136.5 | FAIL |  | 81
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.121 | 2.66 | 54.7 | PO4-P | ug P/L | 0.0619 | 1.09 | 21.6 | NO3-N | ug N/L | 2.21 | 28.3 | 1370 |  | 1 |  | 54.7 |  |  |  |  | 97.97 | PASS | 1 |  | 21.6 |  |  |  |  | 87.31 | PASS | 1 |  | 1370 |  |  |  |  | -13 | FAIL |  | 82
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.215 | 4.63 | 88.5 | PO4-P | ug P/L | 0.0688 | 1.19 | 24 | NO3-N | ug N/L | 3.84 | 49.7 | 2460 |  | 1 |  | 88.5 |  |  |  |  | 85.35 | PASS | 1 |  | 24 |  |  |  |  | 92.2 | PASS | 1 |  | 2460 |  |  |  |  | -54 | FAIL |  | 83
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0588 | 1.21 | 29.9 | PO4-P | ug P/L | 0.0867 | 1.46 | 30.3 | NO3-N | ug N/L | 1.3 | 16.6 | 793 |  | 1 |  | 29.9 |  |  |  |  | 63.49 | FAIL | 1 |  | 30.3 |  |  |  |  | 79.53 | FAIL | 1 |  | 793 |  |  |  |  | 94.3 | PASS |  | 84
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.0826 | 1.7 | 38.3 | PO4-P | ug P/L | 0.0953 | 1.6 | 33.6 | NO3-N | ug N/L | 1.25 | 16 | 763 |  | 1 |  | 38.3 |  |  |  |  | 65.33 | FAIL | 1 |  | 33.6 |  |  |  |  | 87.86 | PASS | 1 |  | 763 |  |  |  |  | 16.3 | FAIL |  | 85
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.139 | 3.14 | 63 | PO4-P | ug P/L | 0.0709 | 1.26 | 25.6 | NO3-N | ug N/L | 2.19 | 28.3 | 1370 |  | 1 |  | 63 |  |  |  |  | 100.3 | PASS | 1 |  | 25.6 |  |  |  |  | 93.81 | PASS | 1 |  | 1370 |  |  |  |  | 87 | PASS |  | 86
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.215 | 4.65 | 88.8 | PO4-P | ug P/L | 0.152 | 2.52 | 54.8 | NO3-N | ug N/L | 0.983 | 12.6 | 597 |  | 1 |  | 88.8 |  |  |  |  | 110.38 | PASS | 1 |  | 54.8 |  |  |  |  | 94.98 | PASS | 1 |  | 597 |  |  |  |  | 144.7 | FAIL |  | 87
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | SPK | 150 | 1 | 1 | NH4-N | ug N/L | 9.77 | 226 | 4370 | PO4-P | ug P/L | 0.953 | 15.3 | 354 | NO3-N | ug N/L | 3.29 | 42.5 | 2090 |  | 1 |  | 4370 |  |  |  |  | 187 | FAIL | 1 |  | 354 |  |  |  |  | 95.4 | PASS | 1 |  | 2090 |  |  |  |  | 9 | FAIL |  | 88
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.115 | 2.38 | 51.2 | PO4-P | ug P/L | 0.065 | 1.18 | 23.5 | NO3-N | ug N/L | 0.942 | 12.6 | 611 |  | 1 |  | 51.2 |  |  |  |  | 74.12 | FAIL | 1 |  | 23.5 |  |  |  |  | 86.45 | PASS | 1 |  | 611 |  |  |  |  | 1.1 | FAIL |  | 89
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.117 | 2.4 | 51.6 | PO4-P | ug P/L | 0.0709 | 1.24 | 24.8 | NO3-N | ug N/L | 1.02 | 13.7 | 664 |  | 1 |  | 51.6 |  |  |  |  | 109.66 | PASS | 1 |  | 24.8 |  |  |  |  | 91.23 | PASS | 1 |  | 664 |  |  |  |  | 26.4 | FAIL |  | 90
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.0645 | 1.35 | 29.7 | PO4-P | ug P/L | 0.057 | 1.05 | 20.5 | NO3-N | ug N/L | 0.395 | 5.29 | 254 |  | 1 |  | 29.7 |  |  |  |  | 100.47 | PASS | 1 |  | 20.5 |  |  |  |  | 90.5 | PASS | 1 |  | 254 |  |  |  |  | 75.4 | FAIL |  | 91
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.0945 | 2.03 | 44.1 | PO4-P | ug P/L | 0.0658 | 1.18 | 23.4 | NO3-N | ug N/L | 0.972 | 13.1 | 635 |  | 1 |  | 44.1 |  |  |  |  | 101.91 | PASS | 1 |  | 23.4 |  |  |  |  | 97.29 | PASS | 1 |  | 635 |  |  |  |  | 158.5 | FAIL |  | 92
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.0568 | 1.23 | 27.1 | PO4-P | ug P/L | 0.0615 | 1.15 | 22.6 | NO3-N | ug N/L | 0.711 | 9.59 | 463 |  | 1 |  | 27.1 |  |  |  |  | 122.16 | PASS | 1 |  | 22.6 |  |  |  |  | 93.11 | PASS | 1 |  | 463 |  |  |  |  | 66.3 | FAIL |  | 93
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.061 | 1.2 | 26.5 | PO4-P | ug P/L | 0.0607 | 1.1 | 21.5 | NO3-N | ug N/L | 1.22 | 16.3 | 796 |  | 1 |  | 26.5 |  |  |  |  | 128.7 | PASS | 1 |  | 21.5 |  |  |  |  | 91.6 | PASS | 1 |  | 796 |  |  |  |  | 109.6 | PASS |  | 94
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.0984 | 1.82 | 39.8 | PO4-P | ug P/L | 0.0726 | 1.32 | 26.6 | NO3-N | ug N/L | 0.747 | 10.1 | 487 |  | 1 |  | 39.8 |  |  |  |  | 110.98 | PASS | 1 |  | 26.6 |  |  |  |  | 93.46 | PASS | 1 |  | 487 |  |  |  |  | 53.7 | FAIL |  | 95
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.0681 | 1.6 | 32 | PO4-P | ug P/L | 0.055 | 0.993 | 20.3 | NO3-N | ug N/L | 0.0334 | 0.46 | 21.7 |  | 1 |  | 32 |  |  |  |  | 107.7333 | PASS | 1 |  | 20.3 |  |  |  |  | 93.55 | PASS | 1 |  | 21.7 |  |  |  |  | 105.37 | PASS |  | 96
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.0778 | 1.75 | 35.3 | PO4-P | ug P/L | 0.0659 | 1.1 | 22.6 | NO3-N | ug N/L | 0.0345 | 0.52 | 24.7 |  | 1 |  | 35.3 |  |  |  |  | 104.8633 | PASS | 1 |  | 22.6 |  |  |  |  | 83.1133 | PASS | 1 |  | 24.7 |  |  |  |  | 116.55 | PASS |  | 97
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 1.33 | 28.8 | 597 | PO4-P | ug P/L | 0.0611 | 1.13 | 23.3 | NO3-N | ug N/L | 0.0314 | 0.43 | 20.2 |  | 1 |  | 597 |  |  |  |  | 91.7 | PASS | 1 |  | 23.3 |  |  |  |  | 62.5967 | FAIL | 1 |  | 20.2 |  |  |  |  | 97.54 | PASS |  | 98
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.0545 | 1.25 | 24.7 | PO4-P | ug P/L | 0.0711 | 1.2 | 25 | NO3-N | ug N/L | 0.03 | 0.414 | 19.5 |  | 1 |  | 24.7 |  |  |  |  | 92.4967 | PASS | 1 |  | 25 |  |  |  |  | 84.0467 | PASS | 1 |  | 19.5 |  |  |  |  | 98.8567 | PASS |  | 99
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.0624 | 1.42 | 28.3 | PO4-P | ug P/L | 0.0652 | 1.25 | 26.1 | NO3-N | ug N/L | 0.544 | 7.36 | 361 |  | 1 |  | 28.3 |  |  |  |  | 81.7633 | PASS | 1 |  | 26.1 |  |  |  |  | 111.89 | PASS | 1 |  | 361 |  |  |  |  | 89.4333 | PASS |  | 100
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0551 | 1.15 | 22.7 | PO4-P | ug P/L | 0.106 | 1.8 | 38.4 | NO3-N | ug N/L | 0.266 | 3.64 | 178 |  | 1 |  | 22.7 |  |  |  |  | 87.87 | PASS | 1 |  | 38.4 |  |  |  |  | 102.5067 | PASS | 1 |  | 178 |  |  |  |  | 81.8 | PASS |  | 101
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.0481 | 1.05 | 20.6 | PO4-P | ug P/L | 0.0586 | 1.03 | 21.2 | NO3-N | ug N/L | 0.0639 | 0.886 | 42.7 |  | 1 |  | 20.6 |  |  |  |  | 81.0467 | PASS | 1 |  | 21.2 |  |  |  |  | 93.8 | PASS | 1 |  | 42.7 |  |  |  |  | 93.87 | PASS |  | 102
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 1.74 | 37.8 | 783 | PO4-P | ug P/L | 0.0595 | 1.04 | 21.3 | NO3-N | ug N/L | 0.0313 | 0.439 | 20.7 |  | 1 |  | 783 |  |  |  |  | 67.6333 | FAIL | 1 |  | 21.3 |  |  |  |  | 90.3967 | PASS | 1 |  | 20.7 |  |  |  |  | 85.0033 | PASS |  | 103
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.066 | 1.33 | 26.4 | PO4-P | ug P/L | 0.0632 | 1.15 | 23.9 | NO3-N | ug N/L | 0.0296 | 0.425 | 20 |  | 1 |  | 26.4 |  |  |  |  | 90.1067 | PASS | 1 |  | 23.9 |  |  |  |  | 107.99 | PASS | 1 |  | 20 |  |  |  |  | 103.568 | PASS |  | 104
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | SPK | 150 | 1 | 1 | NH4-N | ug N/L | 0.07 | 1.49 | 29.8 | PO4-P | ug P/L | 0.0745 | 1.29 | 27 | NO3-N | ug N/L | 0.206 | 2.8 | 137 |  | 1 |  | 29.8 |  |  |  |  | 112.0467 | PASS | 1 |  | 27 |  |  |  |  | 92.8333 | PASS | 1 |  | 137 |  |  |  |  | 67.0333 | FAIL |  | 105
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.0521 | 1.21 | 26.7 | PO4-P | ug P/L | 0.0558 | 0.968 | 18.5 | NO3-N | ug N/L | 0.0421 | 0.586 | 28.3 |  | 1 |  | 26.7 |  |  |  |  | 112.92 | PASS | 1 |  | 18.5 |  |  |  |  | 83.1 | PASS | 1 |  | 28.3 |  |  |  |  | 96.78 | PASS |  | 106
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.0577 | 1.23 | 27.1 | PO4-P | ug P/L | 0.0689 | 1.25 | 24.7 | NO3-N | ug N/L | 1.42 | 19.2 | 950 |  | 1 |  | 27.1 |  |  |  |  | 99.86 | PASS | 1 |  | 24.7 |  |  |  |  | 94.52 | PASS | 1 |  | 950 |  |  |  |  | 145 | FAIL |  | 107
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.0559 | 1.36 | 29.8 | PO4-P | ug P/L | 0.0682 | 1.23 | 24.3 | NO3-N | ug N/L | 1.52 | 20.7 | 1030 |  | 1 |  | 29.8 |  |  |  |  | 116.08 | PASS | 1 |  | 24.3 |  |  |  |  | 90.33 | PASS | 1 |  | 1030 |  |  |  |  | 3 | FAIL |  | 108
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.0499 | 1.13 | 25.1 | PO4-P | ug P/L | 0.0644 | 1.15 | 22.5 | NO3-N | ug N/L | 0.0919 | 1.23 | 59.7 |  | 1 |  | 25.1 |  |  |  |  | 103.16 | PASS | 1 |  | 22.5 |  |  |  |  | 92.25 | PASS | 1 |  | 59.7 |  |  |  |  | 98.97 | PASS |  | 109
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.0982 | 1.84 | 40 | PO4-P | ug P/L | 0.067 | 1.19 | 23.4 | NO3-N | ug N/L | 1.48 | 20.1 | 997 |  | 1 |  | 40 |  |  |  |  | 101 | PASS | 1 |  | 23.4 |  |  |  |  | 77.99 | PASS | 1 |  | 997 |  |  |  |  | -165.3 | FAIL |  | 110
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0636 | 1.48 | 32.5 | PO4-P | ug P/L | 0.0616 | 1.02 | 19.7 | NO3-N | ug N/L | 1.49 | 20.4 | 1010 |  | 1 |  | 32.5 | [6 empty] | 1 |  | 19.7 | [6 empty] | 1 |  | 1010 | [7 empty] | 111
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | SPK | 105 | 1 | 1 | NH4-N | ug N/L | -0.00907 | 0.31 | 7.8 | PO4-P | ug P/L | 0.0136 | 0.261 | 2.5 | NO3-N | ug N/L | 0.0362 | 0.574 | 27.7 |  | 1 |  | 7.8 | [6 empty] | 1 |  | 2.5 | [6 empty] | 1 |  | 27.7 | [7 empty] | 112
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.013 | 0.386 | 9.4 | PO4-P | ug P/L | 0.0113 | 0.23 | 1.81 | NO3-N | ug N/L | 0.0306 | 0.433 | 20.9 |  | 1 |  | 9.4 | [6 empty] | 1 |  | 1.81 | [6 empty] | 1 |  | 20.9 | [7 empty] | 113
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.00949 | 0.212 | 5.75 | PO4-P | ug P/L | 0.0128 | 0.263 | 2.55 | NO3-N | ug N/L | 0.043 | 0.598 | 28.9 |  | 1 |  | 5.75 | [6 empty] | 1 |  | 2.55 | [6 empty] | 1 |  | 28.9 | [7 empty] | 114
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | SPK | 150 | 1 | 1 | NH4-N | ug N/L | 0.0204 | 0.373 | 9.12 | PO4-P | ug P/L | 0.0227 | 0.452 | 6.8 | NO3-N | ug N/L | 0.0972 | 1.34 | 64.8 |  | 1 |  | 9.12 | [6 empty] | 1 |  | 6.8 | [6 empty] | 1 |  | 64.8 | [7 empty] | 115
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.0676 | 1.52 | 26.9 | PO4-P | ug P/L | 0.0647 | 1.12 | 22.5 | NO3-N | ug N/L | 1.32 | 18.1 | 936 |  | 2 | restricted recal begin to middle | 35.3394 |  |  |  |  | 135.2825 | PASS | 1 |  | 22.5 |  |  |  |  | 93.15 | PASS | 1 |  | 936 |  |  |  |  | -66.4 | FAIL |  | 116
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.0832 | 1.79 | 32.1 | PO4-P | ug P/L | 0.0722 | 1.26 | 25.7 | NO3-N | ug N/L | 1.17 | 16 | 823 |  | 2 | restricted recal begin to middle | 41.8965 |  |  |  |  | 137.0891 | FAIL | 1 |  | 25.7 |  |  |  |  | 94.02 | PASS | 1 |  | 823 |  |  |  |  | 87.3 | PASS |  | 117
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.064 | 1.54 | 27.2 | PO4-P | ug P/L | 0.0698 | 1.2 | 24.3 | NO3-N | ug N/L | 1.34 | 18.3 | 946 |  | 2 | restricted recal begin to middle | 35.8273 |  |  |  |  | 126.5226 | FAIL | 1 |  | 24.3 |  |  |  |  | 91.03 | PASS | 1 |  | 946 |  |  |  |  | 64.6 | FAIL |  | 118
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.0551 | 1.28 | 22.3 | PO4-P | ug P/L | 0.0555 | 0.954 | 18.8 | NO3-N | ug N/L | 0.865 | 12 | 616 |  | 2 | restricted recal middle to end | 29.4582 |  |  |  |  | 119.1104 | PASS | 1 |  | 18.8 |  |  |  |  | 88.23 | PASS | 1 |  | 616 |  |  |  |  | 136.6 | FAIL |  | 119
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0595 | 1.36 | 23.9 | PO4-P | ug P/L | 0.0679 | 1.19 | 24.1 | NO3-N | ug N/L | 1.29 | 17.8 | 916 |  | 2 | restricted recal middle to end | 35.3438 |  |  |  |  | 151.2233 | PASS | 1 |  | 24.1 |  |  |  |  | 92.56 | PASS | 1 |  | 916 |  |  |  |  | 21.6 | FAIL |  | 120
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.0405 | 0.915 | 15.2 | PO4-P | ug P/L | 0.0618 | 1.07 | 21.4 | NO3-N | ug N/L | 0.0318 | 0.458 | 22.6 |  | 2 | restricted recal middle to end | 23.8851 |  |  |  |  | 91.2136 | PASS | 1 |  | 21.4 |  |  |  |  | 92.94 | PASS | 1 |  | 22.6 |  |  |  |  | 88.56 | PASS |  | 121
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.0718 | 1.56 | 27.6 | PO4-P | ug P/L | 0.0696 | 1.22 | 24.7 | NO3-N | ug N/L | 1.22 | 16.9 | 871 |  | 2 | restricted recal middle to end | 40.464 |  |  |  |  | 96.7717 | PASS | 1 |  | 24.7 |  |  |  |  | 90.02 | PASS | 1 |  | 871 |  |  |  |  | 22.1 | FAIL |  | 122
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | SPK | 150 | 1 | 1 | NH4-N | ug N/L | 0.055 | 1.39 | 24.5 | PO4-P | ug P/L | 0.0624 | 1.08 | 21.6 | NO3-N | ug N/L | 0.0341 | 0.475 | 23.5 |  | 2 | restricted recal middle to end | 36.113 |  |  |  |  | 140.1118 | PASS | 1 |  | 21.6 |  |  |  |  | 94.51 | PASS | 1 |  | 23.5 |  |  |  |  | 96.5 | PASS |  | 123
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | SPK | 165 | 1 | 1 | NH4-N | ug N/L | 0.116 | 2.51 | 46.1 | PO4-P | ug P/L | 0.0761 | 1.31 | 26.9 | NO3-N | ug N/L | 0.877 | 12.1 | 626 |  | 2 | restricted recal middle to end | 64.5329 |  |  |  |  | 92.3826 | PASS | 1 |  | 26.9 |  |  |  |  | 88.49 | PASS | 1 |  | 626 |  |  |  |  | 22.6 | FAIL |  | 124
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.0542 | 1.11 | 22.4 | PO4-P | ug P/L | 0.0671 | 1.11 | 20.8 | NO3-N | ug N/L | 0.0788 | 1.2 | 59.5 |  | 1 |  | 25.2623 |  |  |  |  | 103.56 | PASS | 1 |  | 20.8 |  |  |  |  | 100.08 | PASS | 1 |  | 59.5 |  |  |  |  | 100.45 | PASS |  | 125
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.0548 | 1.3 | 26.3 | PO4-P | ug P/L | 0.0655 | 1.1 | 20.7 | NO3-N | ug N/L | 0.104 | 1.52 | 76 |  | 1 |  | 29.9502 |  |  |  |  | 122.8903 | PASS | 1 |  | 20.7 |  |  |  |  | 95.62 | PASS | 1 |  | 76 |  |  |  |  | 85.6 | PASS |  | 126
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.0514 | 1.3 | 26.4 | PO4-P | ug P/L | 0.0721 | 1.16 | 22 | NO3-N | ug N/L | 0.0649 | 0.919 | 45.4 |  | 1 |  | 29.9502 |  |  |  |  | 128.9954 | PASS | 1 |  | 22 |  |  |  |  | 97.2 | PASS | 1 |  | 45.4 |  |  |  |  | 101.04 | PASS |  | 127
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.0553 | 1.3 | 26.3 | PO4-P | ug P/L | 0.0805 | 1.33 | 25.5 | NO3-N | ug N/L | 0.0363 | 0.557 | 27.1 |  | 1 |  | 29.9502 |  |  |  |  | 113.4977 | PASS | 1 |  | 25.5 |  |  |  |  | 96.45 | PASS | 1 |  | 27.1 |  |  |  |  | 96.31 | PASS |  | 128
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.0522 | 1.35 | 27.6 | PO4-P | ug P/L | 0.0746 | 1.27 | 24.4 | NO3-N | ug N/L | 0.177 | 2.53 | 127 |  | 1 |  | 31.1787 |  |  |  |  | 141.6307 | PASS | 1 |  | 24.4 |  |  |  |  | 93.34 | PASS | 1 |  | 127 |  |  |  |  | 42.7 | FAIL |  | 129
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0666 | 1.44 | 29.4 | PO4-P | ug P/L | 0.0832 | 1.35 | 26.1 | NO3-N | ug N/L | 0.103 | 1.5 | 74.9 |  | 1 |  | 37.3941 |  |  |  |  | 117.7086 | PASS | 1 |  | 26.1 |  |  |  |  | 108.21 | PASS | 1 |  | 74.9 |  |  |  |  | 94.99 | PASS |  | 130
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 0.0666 | 1.49 | 30.4 | PO4-P | ug P/L | 0.0636 | 1.03 | 20.1 | NO3-N | ug N/L | 0.756 | 10.5 | 582 |  | 1 |  | 30.4 |  |  |  |  | 105.19 | PASS | 1 |  | 20.1 |  |  |  |  | 91.36 | PASS | 1 |  | 582 |  |  |  |  | -41.8 | FAIL |  | 131
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.0724 | 1.74 | 35.3 | PO4-P | ug P/L | 0.0615 | 1.03 | 20.2 | NO3-N | ug N/L | 1.07 | 14.8 | 830 |  | 1 |  | 35.3 |  |  |  |  | 108.03 | PASS | 1 |  | 20.2 |  |  |  |  | 96.87 | PASS | 1 |  | 830 |  |  |  |  | 43 | FAIL |  | 132
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.0515 | 1.18 | 24.4 | PO4-P | ug P/L | 0.0649 | 1.03 | 20.1 | NO3-N | ug N/L | 0.027 | 0.39 | 21.3 |  | 1 |  | 24.4 |  |  |  |  | 104.34 | PASS | 1 |  | 20.1 |  |  |  |  | 93.81 | PASS | 1 |  | 21.3 |  |  |  |  | 97.53 | PASS |  | 133
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.0617 | 1.2 | 24.7 | PO4-P | ug P/L | 0.0624 | 1.01 | 19.7 | NO3-N | ug N/L | 0.0268 | 0.378 | 20.6 |  | 1 |  | 24.7 |  |  |  |  | 101.22 | PASS | 1 |  | 19.7 |  |  |  |  | 95.12 | PASS | 1 |  | 20.6 |  |  |  |  | 89.86 | PASS |  | 134
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 0.0592 | 1.39 | 28.4 | PO4-P | ug P/L | 0.0642 | 1.03 | 20.1 | NO3-N | ug N/L | 0.0276 | 0.394 | 21.5 |  | 1 |  | 28.4 |  |  |  |  | 103.69 | PASS | 1 |  | 20.1 |  |  |  |  | 92.61 | PASS | 1 |  | 21.5 |  |  |  |  | 91.85 | PASS |  | 135
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0799 | 1.6 | 32.6 | PO4-P | ug P/L | 0.0626 | 1.02 | 19.9 | NO3-N | ug N/L | 0.0355 | 0.499 | 27.2 |  | 1 |  | 32.6 |  |  |  |  | 90.76 | PASS | 1 |  | 19.9 |  |  |  |  | 91.69 | PASS | 1 |  | 27.2 |  |  |  |  | 94.47 | PASS |  | 136
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.0637 | 1.48 | 30.1 | PO4-P | ug P/L | 0.0606 | 0.965 | 18.7 | NO3-N | ug N/L | 0.0306 | 0.444 | 24.2 |  | 1 |  | 30.1 |  |  |  |  | 112.91 | PASS | 1 |  | 18.7 |  |  |  |  | 90.64 | PASS | 1 |  | 24.2 |  |  |  |  | 86.82 | PASS |  | 137
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.0504 | 1.16 | 24.1 | PO4-P | ug P/L | 0.0631 | 1.01 | 19.7 | NO3-N | ug N/L | 0.202 | 2.84 | 156 |  | 1 |  | 24.1 |  |  |  |  | 89.41 | PASS | 1 |  | 19.7 |  |  |  |  | 91.07 | PASS | 1 |  | 156 |  |  |  |  | 70.6 | FAIL |  | 138
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.0538 | 1.31 | 27 | PO4-P | ug P/L | 0.061 | 0.95 | 18.4 | NO3-N | ug N/L | 0.0273 | 0.381 | 20.8 |  | 1 |  | 27 |  |  |  |  | 126.4 | PASS | 1 |  | 18.4 |  |  |  |  | 86.54 | PASS | 1 |  | 20.8 |  |  |  |  | 86.58 | PASS |  | 139
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 150 | 1 | 1 | NH4-N | ug N/L | 0.0505 | 1.13 | 23.4 | PO4-P | ug P/L | 0.0625 | 1 | 19.6 | NO3-N | ug N/L | 0.193 | 2.71 | 149 |  | 1 |  | 23.4 |  |  |  |  | 89.34 | PASS | 1 |  | 19.6 |  |  |  |  | 92.16 | PASS | 1 |  | 149 |  |  |  |  | 84.9 | PASS |  | 140
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 165 | 1 | 1 | NH4-N | ug N/L | 0.0538 | 1.15 | 23.8 | PO4-P | ug P/L | 0.0597 | 0.938 | 18.1 | NO3-N | ug N/L | 0.185 | 2.58 | 142 |  | 1 |  | 23.8 |  |  |  |  | 89.93 | PASS | 1 |  | 18.1 |  |  |  |  | 86.11 | PASS | 1 |  | 142 |  |  |  |  | 69.2 | FAIL |  | 141
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | SPK | 180 | 1 | 1 | NH4-N | ug N/L | 0.0871 | 1.92 | 38.8 | PO4-P | ug P/L | 0.0618 | 1.02 | 20 | NO3-N | ug N/L | 0.834 | 11.6 | 649 |  | 1 |  | 38.8 |  |  |  |  | 107.88 | PASS | 1 |  | 20 |  |  |  |  | 90.6 | PASS | 1 |  | 649 |  |  |  |  | 24.9 | FAIL |  | 142
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | SPK | 15 | 1 | 1 | NH4-N | ug N/L | 1.7 | 31.8 | 668 | PO4-P | ug P/L | 0.071 | 1.14 | 23.4 | NO3-N | ug N/L | 0.0341 | 0.436 | 19 |  | 1 |  | 668 |  |  |  |  | 81.8 | PASS | 1 |  | 23.4 |  |  |  |  | 100.19 | PASS | 1 |  | 19 |  |  |  |  | 90.15 | PASS |  | 143
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | SPK | 30 | 1 | 1 | NH4-N | ug N/L | 0.0929 | 1.75 | 36.6 | PO4-P | ug P/L | 0.066 | 1.05 | 21.1 | NO3-N | ug N/L | 0.112 | 1.44 | 67.5 |  | 1 |  | 36.6 |  |  |  |  | 91.16 | PASS | 1 |  | 21.1 |  |  |  |  | 100.61 | PASS | 1 |  | 67.5 |  |  |  |  | 92.25 | PASS |  | 144
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | SPK | 45 | 1 | 1 | NH4-N | ug N/L | 0.0595 | 1.44 | 29.9 | PO4-P | ug P/L | 0.0716 | 1.16 | 23.7 | NO3-N | ug N/L | 0.0346 | 0.444 | 19.3 |  | 1 |  | 29.9 |  |  |  |  | 120.04 | PASS | 1 |  | 23.7 |  |  |  |  | 102.17 | PASS | 1 |  | 19.3 |  |  |  |  | 90.23 | PASS |  | 145
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | SPK | 60 | 1 | 1 | NH4-N | ug N/L | 0.225 | 4.74 | 100 | PO4-P | ug P/L | 0.0685 | 1.1 | 22.3 | NO3-N | ug N/L | 0.0401 | 0.5 | 22 |  | 1 |  | 100 |  |  |  |  | 77.5 | FAIL | 1 |  | 22.3 |  |  |  |  | 92.08 | PASS | 1 |  | 22 |  |  |  |  | 97.1 | PASS |  | 146
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | SPK | 75 | 1 | 1 | NH4-N | ug N/L | 1.93 | 42.1 | 878 | PO4-P | ug P/L | 0.0665 | 1.09 | 22.1 | NO3-N | ug N/L | 0.0346 | 0.448 | 19.5 |  | 1 |  | 878 |  |  |  |  | 77.8 | FAIL | 1 |  | 22.1 |  |  |  |  | 93.61 | PASS | 1 |  | 19.5 |  |  |  |  | 95.07 | PASS |  | 147
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | SPK | 90 | 1 | 1 | NH4-N | ug N/L | 0.0553 | 1.18 | 24.4 | PO4-P | ug P/L | 0.129 | 1.99 | 43 | NO3-N | ug N/L | 0.254 | 3.23 | 162 |  | 1 |  | 24.4 |  |  |  |  | 96.24 | PASS | 1 |  | 43 |  |  |  |  | 94.8 | PASS | 1 |  | 162 |  |  |  |  | 101.2 | PASS |  | 148
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | SPK | 105 | 1 | 1 | NH4-N | ug N/L | 0.478 | 10.3 | 217 | PO4-P | ug P/L | 0.0412 | 0.658 | 12.1 | NO3-N | ug N/L | 0.0306 | 0.405 | 17.5 |  | 1 |  | 217 |  |  |  |  | 91.7 | PASS | 1 |  | 12.1 |  |  |  |  | 42.51 | PASS | 1 |  | 17.5 |  |  |  |  | 84.675 | PASS |  | 149
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | SPK | 120 | 1 | 1 | NH4-N | ug N/L | 0.426 | 9.28 | 196 | PO4-P | ug P/L | 0.0682 | 1.07 | 21.8 | NO3-N | ug N/L | 0.031 | 0.411 | 17.8 |  | 1 |  | 196 |  |  |  |  | 59.6 | FAIL | 1 |  | 21.8 |  |  |  |  | 94.38 | PASS | 1 |  | 17.8 |  |  |  |  | 85.08 | PASS |  | 150
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | SPK | 135 | 1 | 1 | NH4-N | ug N/L | 0.0514 | 1.08 | 22.2 | PO4-P | ug P/L | 0.069 | 1.1 | 22.4 | NO3-N | ug N/L | 0.0289 | 0.381 | 16.3 |  | 1 |  | 22.2 |  |  |  |  | 95.42 | PASS | 1 |  | 22.4 |  |  |  |  | 95.54 | PASS | 1 |  | 16.3 |  |  |  |  | 81.995 | PASS |  | 151
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | SPK | 150 | 1 | 1 | NH4-N | ug N/L | 2.31 | 50 | 1040 | PO4-P | ug P/L | 0.0653 | 1.01 | 20.2 | NO3-N | ug N/L | 0.0302 | 0.425 | 18.4 |  | 1 |  | 1040 |  |  |  |  | 54 | FAIL | 1 |  | 20.2 |  |  |  |  | 72.32 | PASS | 1 |  | 18.4 |  |  |  |  | 89.78 | PASS |  | 152
Known
| Mean | [20 empty] | Mean |  | 132.5869 |  |  |  |  | 100.7865 |  |  |  | 26.6754 |  |  |  |  | 88.7368 |  |  |  | 376.9941 |  |  |  |  | 68.3483 |  | Min | 1
| Std | [20 empty] | Std |  | 420.3562 |  |  |  |  | 21.9652 |  |  |  | 28.7498 |  |  |  |  | 12.2103 |  |  |  | 549.5104 |  |  |  |  | 58.9497 |  | Max | 152
| CV | [20 empty] | CV |  | 317.0419 |  |  |  |  | 21.7938 |  |  |  | 107.7766 |  |  |  |  | 13.7601 |  |  |  | 145.761 |  |  |  |  | 86.2489
Warning Limits | Xbar + 2SD | [19 empty] | Warning Limits | Xbar + 2SD |  | 973.2994 |  |  |  |  | 144.7168 |  |  |  | 84.1751 |  |  |  |  | 113.1573 |  |  |  | 1476.015 |  |  |  |  | 186.2476
| Xbar - 2SD | [20 empty] | Xbar - 2SD |  | -708.1255 |  |  |  |  | 56.8561 |  |  |  | -30.8243 |  |  |  |  | 64.3162 |  |  |  | -722.0268 |  |  |  |  | -49.551
Control Limits | Xbar + 3SD | [19 empty] | Control Limits | Xbar + 3SD |  | 1393.6556 |  |  |  |  | 166.682 |  |  |  | 112.9249 |  |  |  |  | 125.3676 |  |  |  | 2025.5254 |  |  |  |  | 245.1972
| Xbar - 3SD | [20 empty] | Xbar - 3SD |  | -1128.4818 |  |  |  |  | 34.8909 |  |  |  | -59.5741 |  |  |  |  | 52.106 |  |  |  | -1271.5372 |  |  |  |  | -108.5007
Warning Limits as % | Xbar + 2SD | [19 empty] | Warning Limits as % | Xbar + 2SD |  | 734.0839 |  |  |  |  | 143.5876 |  |  |  | 315.5532 |  |  |  |  | 127.5202 |  |  |  | 391.5221 |  |  |  |  | 272.4979
| Xbar - 2SD | [20 empty] | Xbar - 2SD |  | -534.0839 |  |  |  |  | 56.4124 |  |  |  | -115.5532 |  |  |  |  | 72.4798 |  |  |  | -191.5221 |  |  |  |  | -72.4979
Control Limits as % | Xbar + 3SD | [19 empty] | Control Limits as % | Xbar + 3SD |  | 1051.1258 |  |  |  |  | 165.3813 |  |  |  | 423.3297 |  |  |  |  | 141.2803 |  |  |  | 537.2831 |  |  |  |  | 358.7468
| Xbar - 3SD | [20 empty] | Xbar - 3SD |  | -851.1258 |  |  |  |  | 34.6187 |  |  |  | -223.3297 |  |  |  |  | 58.7197 |  |  |  | -337.2831 |  |  |  |  | -158.7468
Coefficient of Variation % | 100* std/mean | [19 empty] | Coefficient of Variation % | 100* std/mean |  | 317.0419 |  |  |  |  | 21.7938 |  |  |  | 107.7766 |  |  |  |  | 13.7601 |  |  |  | 145.761 |  |  |  |  | 86.2489
3 x Coefficient of Variation % | [20 empty] | 3 x Coefficient of Variation % |  |  | 951.1258 |  |  |  |  | 65.3813 |  |  |  | 323.3297 |  |  |  |  | 41.2803 |  |  |  | 437.2831 |  |  |  |  | 258.7468
Count | [20 empty] | Count |  |  | 152 |  |  |  |  | 141 |  |  |  | 152 |  |  |  |  | 142 |  |  |  | 152 |  |  |  |  | 142
99th percentile.inc | [20 empty] | 99th percentile.inc |  |  | 1429.3 |  |  |  |  | 166.35 |  |  |  | 83.918 |  |  |  |  | 113.896 |  |  |  | 2347.8 |  |  |  |  | 154.195
Maximum | [20 empty] | Maximum |  |  | 4370 |  |  |  |  | 187 |  |  |  | 354 |  |  |  |  | 116.83 |  |  |  | 2750 |  |  |  |  | 236.7
[21 empty] | Known |  |  | 0 |  |  |  |  | 0 |  |  |  | 0 |  |  |  |  | 0 |  |  |  | 0 |  |  |  |  | 0
[21 empty] | MDL |  |  | 988.3822 |  |  |  |  | 51.6902 |  |  |  | 67.5994 |  |  |  |  | 28.7318 |  |  |  | 1292.062 |  |  |  |  | 138.7136
[21 empty] | LOQ |  |  | 4203.5623 |  |  |  |  | 219.6518 |  |  |  | 287.4983 |  |  |  |  | 122.1028 |  |  |  | 5495.1044 |  |  |  |  | 589.4966
[21 empty] | ratio mean/mdl |  |  | 0.13415 |  |  |  |  | 1.9498 |  |  |  | 0.39461 |  |  |  |  | 3.0884 |  |  |  | 0.29178 |  |  |  |  | 0.49273

## rolling dups
NEED TO SCREEN FOR BD or LOQ for both determinations
Detection Date | Run File Name | Sample ID | Cup Number | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Reduction Efficiency | Analyst Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | NH4-N NOTES: | NH4-N concentration | NH4-N- % Error in check stds | NH4-N- DQO (<=20%) | NH4-N- RPD of duplicates | NH4-N- DQO (<=20% if conc>LOQ) | NH4-N- PR of spikes | NH4-N- DQO (80-120%) if conc>LOQ | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BN PO4-P NOTES: | PO4-P concentration | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BN NO3-N NOTES: | NO3-N concentration | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ |  | Observation #
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 14 | 1 | 1 | NH4-N | ug N/L | 0.0385 | 0.866 | 18.3 | PO4-P | ug P/L | 0.0237 | 0.441 | 6.26 | NO3-N+NO2-N | ug N/L | 0.00808 | 0.125 | 5.86 |  | 1 |  | 18.3 |  |  | 7.874 | PASS |  |  | 1 |  | 6.26 |  |  | 14.2002 | PASS |  |  | 1 |  | 5.86 |  |  | 20.3008 | PASS |  |  |  | 1
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 29 | 1 | 1 | NH4-N | ug N/L | 0.104 | 2.21 | 47.8 | PO4-P | ug P/L | 0.0326 | 0.533 | 8.66 | NO3-N+NO2-N | ug N/L | 0.131 | 2.11 | 126 |  | 1 |  | 47.8 |  |  | 0.41754 | PASS |  |  | 1 |  | 8.66 |  |  | 0.23121 | PASS |  |  | 1 |  | 126 |  |  | 0 | PASS |  |  |  | 2
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 44 | 1 | 1 | NH4-N | ug N/L | 0.00804 | 0.189 | 3.42 | PO4-P | ug P/L | 0.0152 | 0.268 | 1.79 | NO3-N+NO2-N | ug N/L | 0.00332 | 0.0518 | 1.51 |  | 1 |  | 3.42 |  |  | 85 | PASS |  |  | 1 |  | 1.79 |  |  | 46.6809 | PASS |  |  | 1 |  | 1.51 |  |  | 45.7282 | PASS |  |  |  | 3
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 59 | 1 | 1 | NH4-N | ug N/L | 0.00574 | 0.107 | 1.62 | PO4-P | ug P/L | 0.0235 | 0.387 | 4.87 | NO3-N+NO2-N | ug N/L | 0.00316 | 0.0614 | 2.08 |  | 1 |  | 1.62 |  |  | 32.5581 | PASS |  |  | 1 |  | 4.87 |  |  | 24.1877 | PASS |  |  | 1 |  | 2.08 |  |  | 41.1594 | PASS |  |  |  | 4
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 74 | 1 | 1 | NH4-N | ug N/L | 0.0201 | 0.427 | 8.63 | PO4-P | ug P/L | 0.0213 | 0.372 | 4.47 | NO3-N+NO2-N | ug N/L | -0.00602 | -0.0687 | -5.66 |  | 1 |  | 8.63 |  |  | 0.34702 | PASS |  |  | 1 |  | 4.47 |  |  | 4.5902 | PASS |  |  | 1 |  | -5.66 |  |  | 175.0207 | FAIL |  |  |  | 5
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 89 | 1 | 1 | NH4-N | ug N/L | -0.00959 | -0.0932 | -2.78 | PO4-P | ug P/L | 0.021 | 0.37 | 4.43 | NO3-N+NO2-N | ug N/L | 0.00321 | 0.0588 | 1.93 |  | 1 |  | -2.78 | [6 empty] | 1 |  | 4.43 | [6 empty] | 1 |  | 1.93 | [7 empty] | 6
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 104 | 1 | 1 | NH4-N | ug N/L | 0.0145 | 0.353 | 7.03 | PO4-P | ug P/L | 0.018 | 0.313 | 2.95 | NO3-N+NO2-N | ug N/L | -0.0024 | -0.00775 | -2.03 |  | 1 |  | 7.03 |  |  | 3.6206 | PASS |  |  | 1 |  | 2.95 |  |  | 17.0543 | PASS |  |  | 1 |  | -2.03 |  |  | 13.7615 | PASS |  |  |  | 7
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 119 | 1 | 1 | NH4-N | ug N/L | 0.0438 | 0.961 | 20.3 | PO4-P | ug P/L | 0.0206 | 0.358 | 4.12 | NO3-N+NO2-N | ug N/L | 0.0698 | 1.12 | 65.5 |  | 1 |  | 20.3 |  |  | 4.3373 | PASS |  |  | 1 |  | 4.12 |  |  | 6.3455 | PASS |  |  | 1 |  | 65.5 |  |  | 3.8922 | PASS |  |  |  | 8
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 134 | 1 | 1 | NH4-N | ug N/L | 0.143 | 3.23 | 70.1 | PO4-P | ug P/L | 0.0157 | 0.264 | 1.68 | NO3-N+NO2-N | ug N/L | 0.00372 | 0.0589 | 1.93 |  | 1 |  | 70.1 |  |  | 10.5105 | PASS |  |  | 1 |  | 1.68 |  |  | 73.4463 | PASS |  |  | 1 |  | 1.93 |  |  | 60.6498 | PASS |  |  |  | 9
2020-09-01 | OM_9-1-2020_11-53-48AMedit3.omn | dup | 149 | 1 | 1 | NH4-N | ug N/L | 0.791 | 16.8 | 367 | PO4-P | ug P/L | 0.0273 | 0.432 | 6.03 | NO3-N+NO2-N | ug N/L | 0.00237 | 0.0431 | 0.991 |  | 1 |  | 367 |  |  | 0.27285 | PASS |  |  | 1 |  | 6.03 |  |  | 13.1681 | PASS |  |  | 1 |  | 0.991 |  |  | 186.6563 | FAIL |  |  |  | 10
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | F 3.8m 22jun20 DUP | 14 | 1 | 1 | NH4-N | ug N/L | -0.00845 | 0.0965 | 1.83 | PO4-P | ug P/L | 0.0363 | 0.633 | 12 | NO3-N+NO2-N | ug N/L | 0.00381 | 0.0563 | 2.28 |  | 1 |  | 1.83 |  |  | 51.5464 | PASS |  |  | 1 |  | 12 |  |  | 6.0086 | PASS |  |  | 1 |  | 2.28 |  |  | 17.1429 | PASS |  |  |  | 11
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 3m 14may20 DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0425 | 0.816 | 15.7 | PO4-P | ug P/L | 0.0227 | 0.39 | 5.63 | NO3-N+NO2-N | ug N/L | 0.00217 | 0.0264 | 0.558 |  | 1 |  | 15.7 |  |  | 16.3743 | PASS |  |  | 1 |  | 5.63 |  |  | 8.6661 | PASS |  |  | 1 |  | 0.558 |  |  | 140.1288 | FAIL |  |  |  | 12
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 0.1m 23jul20 DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.0106 | 0.157 | 3.01 | PO4-P | ug P/L | 0.0198 | 0.337 | 4.23 | NO3-N+NO2-N | ug N/L | 0.00334 | 0.0109 | -0.333 |  | 1 |  | 3.01 |  |  | 72.4576 | PASS |  |  | 1 |  | 4.23 |  |  | 0.95012 | PASS |  |  | 1 |  | -0.333 |  |  | 271.7286 | FAIL |  |  |  | 13
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 6m 02jul20 DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.0175 | 0.386 | 7.42 | PO4-P | ug P/L | 0.0167 | 0.293 | 3.1 | NO3-N+NO2-N | ug N/L | 0.00615 | 0.131 | 6.59 |  | 1 |  | 7.42 |  |  | 43.129 | PASS |  |  | 1 |  | 3.1 |  |  | 2.2329 | PASS |  |  | 1 |  | 6.59 |  |  | 53.4615 | PASS |  |  |  | 14
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | F 3.8m 03aug20 DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.204 | 4.29 | 82.9 | PO4-P | ug P/L | 0.0161 | 0.308 | 3.47 | NO3-N+NO2-N | ug N/L | 0.00321 | 0.0669 | 2.88 |  | 1 |  | 82.9 |  |  | 1.556 | PASS |  |  | 1 |  | 3.47 |  |  | 35.2542 | PASS |  |  | 1 |  | 2.88 |  |  | 111.4355 | FAIL |  |  |  | 15
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 9m 30mar20 DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.284 | 5.94 | 115 | PO4-P | ug P/L | 0.0172 | 0.319 | 3.76 | NO3-N+NO2-N | ug N/L | 0.00543 | 0.107 | 5.2 |  | 1 |  | 115 |  |  | 2.6432 | PASS |  |  | 1 |  | 3.76 |  |  | 31.3846 | PASS |  |  | 1 |  | 5.2 |  |  | 18.9474 | PASS |  |  |  | 16
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 11m 02jul20 DUP | 104 | 1 | 1 | NH4-N | ug N/L | 2.8 | 56.6 | 1070 | PO4-P | ug P/L | 0.0896 | 1.28 | 28.9 | NO3-N+NO2-N | ug N/L | 0.00667 | 0.121 | 6 |  | 1 |  | 1070 |  |  | 0.93897 | PASS |  |  | 1 |  | 28.9 |  |  | 5.7143 | PASS |  |  | 1 |  | 6 |  |  | 40 | PASS |  |  |  | 17
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 11m 04jun20 DUP | 119 | 1 | 1 | NH4-N | ug N/L | 2.08 | 44.3 | 840 | PO4-P | ug P/L | 0.133 | 1.93 | 46.1 | NO3-N+NO2-N | ug N/L | 0.00925 | 0.168 | 8.7 |  | 1 |  | 840 |  |  | 2.8169 | PASS |  |  | 1 |  | 46.1 |  |  | 0.65288 | PASS |  |  | 1 |  | 8.7 |  |  | 16.1491 | PASS |  |  |  | 18
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 3m 04jun20 DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.0313 | 0.505 | 9.73 | PO4-P | ug P/L | 0.0219 | 0.397 | 5.83 | NO3-N+NO2-N | ug N/L | 0.00595 | 0.118 | 5.82 |  | 1 |  | 9.73 |  |  | 12.2528 | PASS |  |  | 1 |  | 5.83 |  |  | 25.7502 | PASS |  |  | 1 |  | 5.82 |  |  | 68.2028 | PASS |  |  |  | 19
2020-09-15 | OM_9-15-2020_11-43-18AM.OMN | B 6m 14may20 DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.0235 | 0.481 | 9.27 | PO4-P | ug P/L | 0.0171 | 0.298 | 3.22 | NO3-N+NO2-N | ug N/L | 0.00261 | 0.0408 | 1.38 |  | 1 |  | 9.27 |  |  | 18.8569 | PASS |  |  | 1 |  | 3.22 |  |  | 34.2342 | PASS |  |  | 1 |  | 1.38 |  |  | 89.3082 | PASS |  |  |  | 20
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | dup | 14 | 1 | 1 | NH4-N | ug N/L | 0.0198 | 0.0921 | 3.42 | PO4-P | ug P/L | 0.0207 | 0.409 | 7 | NO3-N | ug N/L | 0.00651 | 0.118 | 3.3 |  | 1 |  | 3.42 |  |  | 19.7628 | PASS |  |  | 1 |  | 7 |  |  | 6.4962 | PASS |  |  | 1 |  | 3.3 |  |  | 8.8608 | PASS |  |  |  | 21
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | dup | 29 | 1 | 1 | NH4-N | ug N/L | 0.0145 | 0.359 | 9.21 | PO4-P | ug P/L | 0.0238 | 0.44 | 7.68 | NO3-N | ug N/L | 0.0129 | 0.238 | 6.7 |  | 1 |  | 9.21 |  |  | 13.9373 | PASS |  |  | 1 |  | 7.68 |  |  | 5.1997 | PASS |  |  | 1 |  | 6.7 |  |  | 13.206 | PASS |  |  |  | 22
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | dup | 44 | 1 | 1 | NH4-N | ug N/L | 1.89 | 41.2 | 869 | PO4-P | ug P/L | 0.0147 | 0.297 | 4.57 | NO3-N | ug N/L | 0.0665 | 0.807 | 22.9 |  | 1 |  | 869 |  |  | 2.9189 | PASS |  |  | 1 |  | 4.57 |  |  | 10.5991 | PASS |  |  | 1 |  | 22.9 |  |  | 3.1042 | PASS |  |  |  | 23
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | dup | 59 | 1 | 1 | NH4-N | ug N/L | 1.51 | 33 | 700 | PO4-P | ug P/L | 0.0157 | 0.308 | 4.8 | NO3-N | ug N/L | 0.0173 | 0.259 | 7.29 |  | 1 |  | 700 | [6 empty] | 1 |  | 4.8 | [6 empty] | 1 |  | 7.29 | [7 empty] | 24
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | dup | 74 | 1 | 1 | NH4-N | ug N/L | 0.0203 | 0.334 | 8.67 | PO4-P | ug P/L | 0.0337 | 0.569 | 10.5 | NO3-N | ug N/L | 0.00671 | 0.128 | 3.59 |  | 1 |  | 8.67 |  |  | 28.7599 | PASS |  |  | 1 |  | 10.5 |  |  | 15.7895 | PASS |  |  | 1 |  | 3.59 |  |  | 0.83916 | PASS |  |  |  | 25
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | dup | 89 | 1 | 1 | NH4-N | ug N/L | 0.0168 | 0.388 | 9.82 | PO4-P | ug P/L | 0.0329 | 0.57 | 10.5 | NO3-N | ug N/L | -0.00585 | -0.0167 | -0.531 |  | 1 |  | 9.82 |  |  | 5.0104 | PASS |  |  | 1 |  | 10.5 |  |  | 2.8169 | PASS |  |  | 1 |  | -0.531 |  |  | 241.4925 | FAIL |  |  |  | 26
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | dup | 104 | 1 | 1 | NH4-N | ug N/L | 2.48 | 54.6 | 1140 | PO4-P | ug P/L | 0.0162 | 0.304 | 4.73 | NO3-N | ug N/L | 0.0123 | 0.166 | 4.65 |  | 1 |  | 1140 |  |  | 0.88106 | PASS |  |  | 1 |  | 4.73 |  |  | 27.2146 | PASS |  |  | 1 |  | 4.65 |  |  | 11.9312 | PASS |  |  |  | 27
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | dup | 119 | 1 | 1 | NH4-N | ug N/L | 0.022 | 0.38 | 9.66 | PO4-P | ug P/L | 0.0207 | 0.412 | 7.07 | NO3-N | ug N/L | 0.00678 | 0.117 | 3.27 |  | 1 |  | 9.66 |  |  | 11.8421 | PASS |  |  | 1 |  | 7.07 |  |  | 14.0697 | PASS |  |  | 1 |  | 3.27 |  |  | 22.7642 | PASS |  |  |  | 28
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | dup | 134 | 1 | 1 | NH4-N | ug N/L | 0.155 | 3.47 | 76.5 | PO4-P | ug P/L | 0.018 | 0.316 | 4.98 | NO3-N | ug N/L | 0.0141 | 0.182 | 5.1 |  | 1 |  | 76.5 |  |  | 3.0524 | PASS |  |  | 1 |  | 4.98 |  |  | 10.4662 | PASS |  |  | 1 |  | 5.1 |  |  | 0.19627 | PASS |  |  |  | 29
2020-11-12 | OM_11-12-2020_11-05-25AM.OMN | dup | 149 | 1 | 1 | NH4-N | ug N/L | 0.232 | 5.11 | 112 | PO4-P | ug P/L | 0.0159 | 0.299 | 4.61 | NO3-N | ug N/L | 0.0207 | 0.295 | 8.31 |  | 1 |  | 112 |  |  | 6.8966 | PASS |  |  | 1 |  | 4.61 |  |  | 66.8596 | PASS |  |  | 1 |  | 8.31 |  |  | 33.2632 | PASS |  |  |  | 30
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 1.58 | 31.8 | 628 | PO4-P | ug P/L | 0.0201 | 0.326 | 6.39 | NO3-N | ug N/L | 0.0206 | 0.254 | 5.65 |  | 1 |  | 628 |  |  | 0.15911 | PASS |  |  | 1 |  | 6.39 |  |  | 5.9226 | PASS |  |  | 1 |  | 5.65 |  |  | 11.98 | PASS |  |  |  | 31
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0321 | 0.86 | 16.5 | PO4-P | ug P/L | 0.0172 | 0.322 | 6.32 | NO3-N | ug N/L | 0.273 | 3.29 | 85.5 |  | 1 |  | 16.5 |  |  | 34.8754 | FAIL |  |  | 1 |  | 6.32 |  |  | 0.47356 | PASS |  |  | 1 |  | 85.5 |  |  | 1.4134 | PASS |  |  |  | 32
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.0224 | 0.768 | 7.75 | PO4-P | ug P/L | 0.0213 | 0.373 | 7.35 | NO3-N | ug N/L | 0.0157 | 0.203 | 4.3 |  | 1 |  | 7.75 |  |  | 20.394 | PASS |  |  | 1 |  | 7.35 |  |  | 4.3912 | PASS |  |  | 1 |  | 4.3 |  |  | 11.0429 | PASS |  |  |  | 33
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.1 | 2.16 | 43.1 | PO4-P | ug P/L | 0.0218 | 0.308 | 6.02 | NO3-N | ug N/L | 0.0211 | 0.267 | 5.98 |  | 1 |  | 43.1 |  |  | 0 | PASS |  |  | 1 |  | 6.02 |  |  | 14.4838 | PASS |  |  | 1 |  | 5.98 |  |  | 14.4298 | PASS |  |  |  | 34
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.994 | 20.5 | 411 | PO4-P | ug P/L | 0.0173 | 0.217 | 4.15 | NO3-N | ug N/L | 0.00992 | 0.166 | 3.36 |  | 1 |  | 411 |  |  | 0.72727 | PASS |  |  | 1 |  | 4.15 |  |  | 34.4965 | PASS |  |  | 1 |  | 3.36 |  |  | 28.9608 | PASS |  |  |  | 35
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.223 | 4.88 | 98.7 | PO4-P | ug P/L | 0.026 | 0.405 | 8.02 | NO3-N | ug N/L | 0.0055 | 0.0839 | 1.2 |  | 1 |  | 98.7 |  |  | 2.3035 | PASS |  |  | 1 |  | 8.02 |  |  | 12.8355 | PASS |  |  | 1 |  | 1.2 |  |  | 22.8782 | PASS |  |  |  | 36
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0215 | 0.516 | 9.44 | PO4-P | ug P/L | 0.0533 | 0.899 | 18.3 | NO3-N | ug N/L | 0.427 | 5.11 | 133 |  | 1 |  | 9.44 |  |  | 44.8644 | PASS |  |  | 1 |  | 18.3 |  |  | 1.6529 | PASS |  |  | 1 |  | 133 |  |  | 0 | PASS |  |  |  | 37
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.28 | 6.02 | 122 | PO4-P | ug P/L | 0.0182 | 0.287 | 5.59 | NO3-N | ug N/L | 0.0106 | 0.136 | 2.57 |  | 1 |  | 122 |  |  | 3.2258 | PASS |  |  | 1 |  | 5.59 |  |  | 9.3777 | PASS |  |  | 1 |  | 2.57 |  |  | 9.2764 | PASS |  |  |  | 38
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.642 | 13.8 | 278 | PO4-P | ug P/L | 0.0192 | 0.314 | 6.15 | NO3-N | ug N/L | 0.00981 | 0.104 | 1.73 |  | 1 |  | 278 |  |  | 0 | PASS |  |  | 1 |  | 6.15 |  |  | 0.6483 | PASS |  |  | 1 |  | 1.73 |  |  | 25.6927 | PASS |  |  |  | 39
2021-01-27 | OM_1-27-2021_10-59-09AM.OMN | DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.145 | 0.27 | 9.84 | PO4-P | ug P/L | 0.016 | 0.297 | 5.79 | NO3-N | ug N/L | 0.0438 | 0.52 | 12.6 |  | 1 |  | 9.84 |  |  | 38.6901 | PASS |  |  | 1 |  | 5.79 |  |  | 0 | PASS |  |  | 1 |  | 12.6 |  |  | 10.5263 | PASS |  |  |  | 40
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.0231 | 0.518 | 10.3 | PO4-P | ug P/L | 0.0166 | 0.231 | 0.707 | NO3-N | ug N/L | 0.00374 | 0.0567 | 0.758 |  | 1 |  | 10.3 |  |  | 28.3333 | FAIL |  |  | 1 |  | 0.707 |  |  | 160.5963 | FAIL |  |  | 1 |  | 0.758 |  |  | 87.1212 | PASS |  |  |  | 41
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.206 | 4.41 | 88.9 | PO4-P | ug P/L | 0.0223 | 0.312 | 2.52 | NO3-N | ug N/L | 0.00507 | 0.0348 | -0.405 |  | 1 |  | 88.9 |  |  | 2.5541 | PASS |  |  | 1 |  | 2.52 |  |  | 21.2766 | PASS |  |  | 1 |  | -0.405 |  |  | 236.3636 | FAIL |  |  |  | 42
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.0184 | 0.558 | 11.1 | PO4-P | ug P/L | 0.0223 | 0.408 | 4.64 | NO3-N | ug N/L | 0.00585 | 0.106 | 3.36 |  | 1 |  | 11.1 |  |  | 36.9264 | FAIL |  |  | 1 |  | 4.64 |  |  | 7.3743 | PASS |  |  | 1 |  | 3.36 |  |  | 65.6126 | PASS |  |  |  | 43
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.0151 | 0.382 | 7.57 | PO4-P | ug P/L | 0.0163 | 0.273 | 1.65 | NO3-N | ug N/L | 0.13 | 1.67 | 86.3 |  | 1 |  | 7.57 |  |  | 43.2129 | PASS |  |  | 1 |  | 1.65 |  |  | 7.5802 | PASS |  |  | 1 |  | 86.3 |  |  | 2.8202 | PASS |  |  |  | 44
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 4.05 | 85.3 | 1710 | PO4-P | ug P/L | 0.0842 | 1.25 | 23.4 | NO3-N | ug N/L | 0.00511 | 0.073 | 1.62 |  | 1 |  | 1710 |  |  | 0 | PASS |  |  | 1 |  | 23.4 |  |  | 7.0103 | PASS |  |  | 1 |  | 1.62 |  |  | 62.987 | PASS |  |  |  | 45
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.306 | 6.75 | 136 | PO4-P | ug P/L | 0.0131 | 0.157 | -0.937 | NO3-N | ug N/L | 0.00699 | 0.102 | 3.14 |  | 1 |  | 136 |  |  | 3.7453 | PASS |  |  | 1 |  | -0.937 |  |  | 67.9056 | PASS |  |  | 1 |  | 3.14 |  |  | 0.95087 | PASS |  |  |  | 46
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0172 | 0.449 | 8.91 | PO4-P | ug P/L | 0.0133 | 0.197 | -0.0347 | NO3-N | ug N/L | 0.0222 | 0.317 | 14.5 |  | 1 |  | 8.91 |  |  | 7.6632 | PASS |  |  | 1 |  | -0.0347 |  |  | 224.1686 | FAIL |  |  | 1 |  | 14.5 |  |  | 29.912 | FAIL |  |  |  | 47
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.304 | 6.62 | 133 | PO4-P | ug P/L | 0.0147 | 0.207 | 0.191 | NO3-N | ug N/L | 0.0151 | 0.199 | 8.27 |  | 1 |  | 133 |  |  | 2.963 | PASS |  |  | 1 |  | 0.191 |  |  | 367.1772 | FAIL |  |  | 1 |  | 8.27 |  |  | 16.1201 | PASS |  |  |  | 48
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.0203 | 0.478 | 9.5 | PO4-P | ug P/L | 0.014 | 0.219 | 0.442 | NO3-N | ug N/L | 0.0705 | 0.922 | 46.5 |  | 1 |  | 9.5 |  |  | 0.9429 | PASS |  |  | 1 |  | 0.442 |  |  | 566.805 | FAIL |  |  | 1 |  | 46.5 |  |  | 102.9316 | FAIL |  |  |  | 49
2021-06-16 | OM_6-16-2021_11-10-05AM.OMN | DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.00647 | 0.189 | 3.67 | PO4-P | ug P/L | 0.0171 | 0.248 | 1.09 | NO3-N | ug N/L | 0.00246 | 0.0539 | 0.607 |  | 1 |  | 3.67 |  |  | 27.9503 | PASS |  |  | 1 |  | 1.09 |  |  | 35.4717 | PASS |  |  | 1 |  | 0.607 |  |  | 172.0503 | FAIL |  |  |  | 50
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.836 | 17.8 | 367 | PO4-P | ug P/L | 0.0222 | 0.36 | 6.47 | NO3-N | ug N/L | 0.00207 | 0.032 | 1.37 |  | 1 |  | 367 |  |  | 0.81411 | PASS |  |  | 1 |  | 6.47 |  |  | 45.6768 | PASS |  |  | 1 |  | 1.37 |  |  | 27.1293 | PASS |  |  |  | 51
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0108 | 0.308 | 6.93 | PO4-P | ug P/L | 0.033 | 0.594 | 11.6 | NO3-N | ug N/L | 0.00425 | 0.0584 | 2.79 |  | 1 |  | 6.93 |  |  | 1.4327 | PASS |  |  | 1 |  | 11.6 |  |  | 2.6201 | PASS |  |  | 1 |  | 2.79 |  |  | 2.1739 | PASS |  |  |  | 52
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.716 | 14.8 | 305 | PO4-P | ug P/L | 0.0142 | 0.253 | 4.15 | NO3-N | ug N/L | 0.00398 | 0.0933 | 4.66 |  | 1 |  | 305 |  |  | 5.7336 | PASS |  |  | 1 |  | 4.15 |  |  | 2.934 | PASS |  |  | 1 |  | 4.66 |  |  | 35.3535 | PASS |  |  |  | 53
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.0672 | 1.42 | 29.9 | PO4-P | ug P/L | 0.0141 | 0.255 | 4.2 | NO3-N | ug N/L | 0.00265 | 0.0444 | 2.04 |  | 1 |  | 29.9 |  |  | 4.894 | PASS |  |  | 1 |  | 4.2 |  |  | 13.7472 | PASS |  |  | 1 |  | 2.04 |  |  | 14.5455 | PASS |  |  |  | 54
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.013 | 0.256 | 5.86 | PO4-P | ug P/L | 0.0392 | 0.695 | 13.8 | NO3-N | ug N/L | 0.00373 | 0.0465 | 2.15 |  | 1 |  | 5.86 |  |  | 59.3037 | PASS |  |  | 1 |  | 13.8 |  |  | 4.4444 | PASS |  |  | 1 |  | 2.15 |  |  | 25.9109 | PASS |  |  |  | 55
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.0428 | 1.02 | 21.7 | PO4-P | ug P/L | 0.0149 | 0.298 | 5.12 | NO3-N | ug N/L | 0.00395 | 0.0531 | 2.5 |  | 1 |  | 21.7 |  |  | 21.4286 | FAIL |  |  | 1 |  | 5.12 |  |  | 53.4653 | PASS |  |  | 1 |  | 2.5 |  |  | 38.7097 | PASS |  |  |  | 56
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.626 | 11.9 | 246 | PO4-P | ug P/L | 0.0131 | 0.209 | 3.18 | NO3-N | ug N/L | 0.00352 | 0.0432 | 1.98 |  | 1 |  | 246 |  |  | 0.81633 | PASS |  |  | 1 |  | 3.18 |  |  | 59.4475 | PASS |  |  | 1 |  | 1.98 |  |  | 73.8854 | PASS |  |  |  | 57
2021-07-20 | OM_7-20-2021_12-13-40PM edit 1.omn | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.593 | 12.7 | 262 | PO4-P | ug P/L | 0.0118 | 0.22 | 3.44 | NO3-N | ug N/L | 0.00728 | 0.1 | 5.05 |  | 1 |  | 262 |  |  | 1.1516 | PASS |  |  | 1 |  | 3.44 |  |  | 3.9886 | PASS |  |  | 1 |  | 5.05 |  |  | 17.8539 | PASS |  |  |  | 58
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.0263 | 0.521 | 11.9 | PO4-P | ug P/L | 0.0491 | 0.887 | 17.3 | NO3-N | ug N/L | 1.85 | 23.4 | 1260 |  | 1 |  | 11.9 |  |  | 19.0114 | PASS |  |  | 1 |  | 17.3 |  |  | 1.1494 | PASS |  |  | 1 |  | 1260 |  |  | 1.5748 | PASS |  |  |  | 59
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0296 | 0.709 | 16.1 | PO4-P | ug P/L | 0.0683 | 1.18 | 24 | NO3-N | ug N/L | 1.24 | 15.8 | 852 |  | 1 |  | 16.1 |  |  | 29.1815 | FAIL |  |  | 1 |  | 24 |  |  | 5.2738 | PASS |  |  | 1 |  | 852 |  |  | 32.2835 | FAIL |  |  |  | 60
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.0169 | 0.318 | 7.39 | PO4-P | ug P/L | 0.0335 | 0.626 | 11.4 | NO3-N | ug N/L | 1.86 | 23.6 | 1280 |  | 1 |  | 7.39 |  |  | 35.6865 | PASS |  |  | 1 |  | 11.4 |  |  | 7.2727 | PASS |  |  | 1 |  | 1280 |  |  | 1.5748 | PASS |  |  |  | 61
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.0232 | 0.318 | 7.39 | PO4-P | ug P/L | 0.0285 | 0.55 | 9.67 | NO3-N | ug N/L | 1.89 | 24 | 1300 |  | 1 |  | 7.39 |  |  | 70.294 | PASS |  |  | 1 |  | 9.67 |  |  | 1.6684 | PASS |  |  | 1 |  | 1300 |  |  | 3.9216 | PASS |  |  |  | 62
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.0647 | 1.35 | 30.4 | PO4-P | ug P/L | 0.0319 | 0.596 | 10.7 | NO3-N | ug N/L | 1.9 | 24.1 | 1300 |  | 1 |  | 30.4 | [6 empty] | 1 |  | 10.7 | [6 empty] | 1 |  | 1300 | [7 empty] | 63
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.0781 | 1.63 | 36.7 | PO4-P | ug P/L | 0.0491 | 0.895 | 17.5 | NO3-N | ug N/L | 1.85 | 23.6 | 1280 |  | 1 |  | 36.7 |  |  | 4.7872 | PASS |  |  | 1 |  | 17.5 |  |  | 0.5698 | PASS |  |  | 1 |  | 1280 |  |  | 0.77821 | PASS |  |  |  | 64
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0667 | 1.49 | 33.5 | PO4-P | ug P/L | 0.038 | 0.711 | 13.3 | NO3-N | ug N/L | 1.91 | 24.4 | 1320 |  | 1 |  | 33.5 |  |  | 3.5191 | PASS |  |  | 1 |  | 13.3 |  |  | 0 | PASS |  |  | 1 |  | 1320 |  |  | 3.861 | PASS |  |  |  | 65
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.0723 | 1.46 | 32.8 | PO4-P | ug P/L | 0.0358 | 0.655 | 12.1 | NO3-N | ug N/L | 1.84 | 23.6 | 1270 |  | 1 |  | 32.8 |  |  | 1.8462 | PASS |  |  | 1 |  | 12.1 |  |  | 5.6225 | PASS |  |  | 1 |  | 1270 |  |  | 0 | PASS |  |  |  | 66
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.0232 | 0.509 | 11.6 | PO4-P | ug P/L | 0.0423 | 0.781 | 14.9 | NO3-N | ug N/L | 1.48 | 19 | 1030 |  | 1 |  | 11.6 |  |  | 4.2194 | PASS |  |  | 1 |  | 14.9 |  |  | 9.5847 | PASS |  |  | 1 |  | 1030 |  |  | 4.3651 | PASS |  |  |  | 67
2021-07-30 | OM_7-30-2021_11-01-19AM edit1.omn | DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.15 | 3.22 | 72 | PO4-P | ug P/L | 0.0247 | 0.468 | 7.79 | NO3-N | ug N/L | 0.89 | 11.4 | 617 |  | 1 |  | 72 |  |  | 3.8217 | PASS |  |  | 1 |  | 7.79 |  |  | 1.9072 | PASS |  |  | 1 |  | 617 |  |  | 0 | PASS |  |  |  | 68
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.0908 | 1.96 | 41.2 | PO4-P | ug P/L | 0.0163 | 0.34 | 3.67 | NO3-N | ug N/L | 1.3 | 16.5 | 786 |  | 1 |  | 41.2 |  |  | 3.1063 | PASS |  |  | 1 |  | 3.67 |  |  | 31.6514 | PASS |  |  | 1 |  | 786 |  |  | 5.6246 | PASS |  |  |  | 69
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0752 | 1.63 | 34.4 | PO4-P | ug P/L | 0.0126 | 0.219 | 0.934 | NO3-N | ug N/L | 0.114 | 1.49 | 71.4 |  | 1 |  | 34.4 |  |  | 6.1972 | PASS |  |  | 1 |  | 0.934 |  |  | 58.6989 | PASS |  |  | 1 |  | 71.4 |  |  | 7.5581 | PASS |  |  |  | 70
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.0308 | 0.595 | 12.9 | PO4-P | ug P/L | 0.0119 | 0.268 | 2.03 | NO3-N | ug N/L | 0.0688 | 0.909 | 43.8 |  | 1 |  | 12.9 |  |  | 13.2231 | PASS |  |  | 1 |  | 2.03 |  |  | 45.3172 | PASS |  |  | 1 |  | 43.8 |  |  | 4.8998 | PASS |  |  |  | 71
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.023 | 0.496 | 10.8 | PO4-P | ug P/L | 0.0267 | 0.52 | 7.73 | NO3-N | ug N/L | 1.85 | 23.6 | 1130 |  | 1 |  | 10.8 |  |  | 9.6916 | PASS |  |  | 1 |  | 7.73 |  |  | 1.8277 | PASS |  |  | 1 |  | 1130 |  |  | 0.88889 | PASS |  |  |  | 72
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.0275 | 0.623 | 13.5 | PO4-P | ug P/L | 0.0275 | 0.518 | 7.68 | NO3-N | ug N/L | 1.76 | 22.5 | 1080 |  | 1 |  | 13.5 | [6 empty] | 1 |  | 7.68 | [6 empty] | 1 |  | 1080 | [7 empty] | 73
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.02 | 0.413 | 9.08 | PO4-P | ug P/L | 0.0461 | 0.849 | 15.2 | NO3-N | ug N/L | 1.83 | 23.4 | 1120 |  | 1 |  | 9.08 | [6 empty] | 1 |  | 15.2 | [6 empty] | 1 |  | 1120 | [7 empty] | 74
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0273 | 0.594 | 12.8 | PO4-P | ug P/L | 0.0171 | 0.348 | 3.85 | NO3-N | ug N/L | 0.666 | 8.49 | 404 |  | 1 |  | 12.8 |  |  | 10.6996 | PASS |  |  | 1 |  | 3.85 |  |  | 14.4847 | PASS |  |  | 1 |  | 404 |  |  | 1.4742 | PASS |  |  |  | 75
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.0225 | 0.599 | 13 | PO4-P | ug P/L | 0.0219 | 0.418 | 5.41 | NO3-N | ug N/L | 0.289 | 3.7 | 177 |  | 1 |  | 13 |  |  | 8 | PASS |  |  | 1 |  | 5.41 |  |  | 3.4514 | PASS |  |  | 1 |  | 177 |  |  | 1.6807 | PASS |  |  |  | 76
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.0288 | 0.636 | 13.7 | PO4-P | ug P/L | 0.0137 | 0.301 | 2.77 | NO3-N | ug N/L | 0.746 | 9.65 | 460 |  | 1 |  | 13.7 |  |  | 31.9018 | FAIL |  |  | 1 |  | 2.77 |  |  | 27.5154 | PASS |  |  | 1 |  | 460 |  |  | 1.7241 | PASS |  |  |  | 77
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.0295 | 0.583 | 12.6 | PO4-P | ug P/L | 0.014 | 0.293 | 2.61 | NO3-N | ug N/L | 0.659 | 8.48 | 404 |  | 1 |  | 12.6 |  |  | 16.0584 | PASS |  |  | 1 |  | 2.61 |  |  | 3.3898 | PASS |  |  | 1 |  | 404 |  |  | 1.9608 | PASS |  |  |  | 78
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 164 | 1 | 1 | NH4-N | ug N/L | 0.0406 | 0.941 | 20.1 | PO4-P | ug P/L | 0.0285 | 0.533 | 8.01 | NO3-N | ug N/L | 1.65 | 21.3 | 1020 |  | 1 |  | 20.1 |  |  | 0 | PASS |  |  | 1 |  | 8.01 |  |  | 10.0771 | PASS |  |  | 1 |  | 1020 |  |  | 2.381 | PASS |  |  |  | 79
2021-08-03 | OM_8-3-2021_10-46-23AM edit1.omn | DUP | 194 | 1 | 1 | NH4-N | ug N/L | 0.0455 | 0.715 | 15.4 | PO4-P | ug P/L | 0.0448 | 0.809 | 14.2 | NO3-N | ug N/L | 1.65 | 21.1 | 1010 |  | 1 |  | 15.4 |  |  | 18.8235 | PASS |  |  | 1 |  | 14.2 |  |  | 6.8027 | PASS |  |  | 1 |  | 1010 |  |  | 2.9268 | PASS |  |  |  | 80
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.126 | 2.69 | 55.3 | PO4-P | ug P/L | 0.0192 | 0.4 | 5.57 | NO3-N | ug N/L | 0.948 | 12.1 | 574 |  | 1 |  | 55.3 |  |  | 0.54397 | PASS |  |  | 1 |  | 5.57 |  |  | 10.9848 | PASS |  |  | 1 |  | 574 |  |  | 3.5461 | PASS |  |  |  | 81
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0381 | 0.656 | 20.3 | PO4-P | ug P/L | 0.0263 | 0.528 | 8.54 | NO3-N | ug N/L | 1.25 | 16 | 761 |  | 1 |  | 20.3 |  |  | 15.9574 | PASS |  |  | 1 |  | 8.54 |  |  | 17.1647 | PASS |  |  | 1 |  | 761 |  |  | 1.723 | PASS |  |  |  | 82
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.0318 | 0.682 | 20.8 | PO4-P | ug P/L | 0.059 | 1.03 | 20.2 | NO3-N | ug N/L | 3.36 | 43.2 | 2120 |  | 1 |  | 20.8 |  |  | 1.9417 | PASS |  |  | 1 |  | 20.2 |  |  | 8.5308 | PASS |  |  | 1 |  | 2120 |  |  | 6.8337 | PASS |  |  |  | 83
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.231 | 4.88 | 92.9 | PO4-P | ug P/L | 0.0228 | 0.452 | 6.78 | NO3-N | ug N/L | 2.25 | 28.9 | 1400 |  | 1 |  | 92.9 |  |  | 8.3548 | PASS |  |  | 1 |  | 6.78 |  |  | 7.7959 | PASS |  |  | 1 |  | 1400 |  |  | 1.4388 | PASS |  |  |  | 84
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.11 | 2.02 | 43.7 | PO4-P | ug P/L | 0.0798 | 1.34 | 27.4 | NO3-N | ug N/L | 0.944 | 12 | 566 |  | 1 |  | 43.7 |  |  | 16.3866 | PASS |  |  | 1 |  | 27.4 |  |  | 15.7983 | PASS |  |  | 1 |  | 566 |  |  | 2.8674 | PASS |  |  |  | 85
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.023 | 0.543 | 18.4 | PO4-P | ug P/L | 0.0258 | 0.506 | 8.03 | NO3-N | ug N/L | 2.52 | 32.3 | 1570 |  | 1 |  | 18.4 |  |  | 5.5866 | PASS |  |  | 1 |  | 8.03 |  |  | 2.2167 | PASS |  |  | 1 |  | 1570 |  |  | 1.2821 | PASS |  |  |  | 86
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0944 | 2.02 | 43.7 | PO4-P | ug P/L | 0.0517 | 0.914 | 17.5 | NO3-N | ug N/L | 1.33 | 17.1 | 815 |  | 1 |  | 43.7 |  |  | 7.109 | PASS |  |  | 1 |  | 17.5 |  |  | 4.0816 | PASS |  |  | 1 |  | 815 |  |  | 1.6079 | PASS |  |  |  | 87
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.03 | 0.556 | 18.6 | PO4-P | ug P/L | 0.0145 | 0.324 | 3.79 | NO3-N | ug N/L | 4.21 | 54.9 | 2740 |  | 1 |  | 18.6 |  |  | 8.9888 | PASS |  |  | 1 |  | 3.79 |  |  | 1.328 | PASS |  |  | 1 |  | 2740 |  |  | 1.1009 | PASS |  |  |  | 88
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.348 | 7.13 | 131 | PO4-P | ug P/L | 0.0241 | 0.47 | 7.2 | NO3-N | ug N/L | 2.28 | 29.5 | 1430 |  | 1 |  | 131 |  |  | 0 | PASS |  |  | 1 |  | 7.2 |  |  | 8.7649 | PASS |  |  | 1 |  | 1430 |  |  | 2.7586 | PASS |  |  |  | 89
2021-08-24 | OM_8-24-2021_10-19-33AM edit1.omn | DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.146 | 3.15 | 63 | PO4-P | ug P/L | 0.0658 | 1.16 | 23.2 | NO3-N | ug N/L | 1.27 | 16.4 | 782 |  | 1 |  | 63 |  |  | 14.4681 | PASS |  |  | 1 |  | 23.2 |  |  | 3.3898 | PASS |  |  | 1 |  | 782 |  |  | 1.1576 | PASS |  |  |  | 90
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.0523 | 1.13 | 25.1 | PO4-P | ug P/L | 0.0232 | 0.424 | 6.03 | NO3-N | ug N/L | 0.976 | 13 | 631 |  | 1 |  | 25.1 |  |  | 9.1667 | PASS |  |  | 1 |  | 6.03 |  |  | 1.4815 | PASS |  |  | 1 |  | 631 |  |  | 1.5974 | PASS |  |  |  | 91
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0233 | 0.633 | 14 | PO4-P | ug P/L | 0.0147 | 0.303 | 3.24 | NO3-N | ug N/L | 0.898 | 12 | 582 |  | 1 |  | 14 |  |  | 22.2222 | FAIL |  |  | 1 |  | 3.24 |  |  | 39.0062 | PASS |  |  | 1 |  | 582 |  |  | 1.21 | PASS |  |  |  | 92
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.0556 | 1.14 | 25.3 | PO4-P | ug P/L | 0.0185 | 0.393 | 5.31 | NO3-N | ug N/L | 0.964 | 13 | 628 |  | 1 |  | 25.3 |  |  | 3.6217 | PASS |  |  | 1 |  | 5.31 |  |  | 3.8781 | PASS |  |  | 1 |  | 628 |  |  | 0.63898 | PASS |  |  |  | 93
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.0527 | 1.2 | 26.6 | PO4-P | ug P/L | 0.0179 | 0.325 | 3.76 | NO3-N | ug N/L | 0.916 | 12.3 | 598 |  | 1 |  | 26.6 |  |  | 7.4074 | PASS |  |  | 1 |  | 3.76 |  |  | 45.5852 | PASS |  |  | 1 |  | 598 |  |  | 3.4014 | PASS |  |  |  | 94
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.0392 | 0.817 | 18.1 | PO4-P | ug P/L | 0.0218 | 0.395 | 5.37 | NO3-N | ug N/L | 0.875 | 11.8 | 572 |  | 1 |  | 18.1 |  |  | 4.3243 | PASS |  |  | 1 |  | 5.37 |  |  | 22.6259 | PASS |  |  | 1 |  | 572 |  |  | 0.87796 | PASS |  |  |  | 95
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.00679 | 0.0995 | 1.8 | PO4-P | ug P/L | 0.0136 | 0.215 | 1.23 | NO3-N | ug N/L | 1.2 | 16.2 | 787 |  | 1 |  | 1.8 |  |  | 34.1014 | PASS |  |  | 1 |  | 1.23 |  |  | 52.2523 | PASS |  |  | 1 |  | 787 |  |  | 1.6656 | PASS |  |  |  | 96
2021-09-28 | OM_9-28-2021_10-45-41AM edit2.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0191 | 0.452 | 9.9 | PO4-P | ug P/L | 0.019 | 0.38 | 5.03 | NO3-N | ug N/L | 1.24 | 16.7 | 815 |  | 1 |  | 9.9 |  |  | 0 | PASS |  |  | 1 |  | 5.03 |  |  | 8.0153 | PASS |  |  | 1 |  | 815 |  |  | 1.9827 | PASS |  |  |  | 97
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.0179 | 0.45 | 8.2 | PO4-P | ug P/L | 0.025 | 0.465 | 8.36 | NO3-N | ug N/L | 0.00424 | 0.0719 | 2.62 |  | 1 |  | 8.2 |  |  | 5.2257 | PASS |  |  | 1 |  | 8.36 |  |  | 19.4226 | PASS |  |  | 1 |  | 2.62 |  |  | 37.1041 | PASS |  |  |  | 98
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0143 | 0.33 | 5.7 | PO4-P | ug P/L | 0.0214 | 0.383 | 6.52 | NO3-N | ug N/L | 0.003 | 0.0434 | 1.22 |  | 1 |  | 5.7 |  |  | 53.4704 | PASS |  |  | 1 |  | 6.52 |  |  | 18.6085 | PASS |  |  | 1 |  | 1.22 |  |  | 180.3384 | FAIL |  |  |  | 99
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 1.12 | 23.3 | 483 | PO4-P | ug P/L | 0.0575 | 0.961 | 19.5 | NO3-N | ug N/L | 0.00365 | 0.0691 | 2.49 |  | 1 |  | 483 |  |  | 0.20725 | PASS |  |  | 1 |  | 19.5 |  |  | 20.9632 | FAIL |  |  | 1 |  | 2.49 |  |  | 7.9332 | PASS |  |  |  | 100
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 1.92 | 41.6 | 862 | PO4-P | ug P/L | 0.0186 | 0.333 | 5.4 | NO3-N | ug N/L | 0.00287 | 0.0533 | 1.71 |  | 1 |  | 862 |  |  | 0.34863 | PASS |  |  | 1 |  | 5.4 |  |  | 24.5322 | PASS |  |  | 1 |  | 1.71 |  |  | 25 | PASS |  |  |  | 101
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.0174 | 0.424 | 7.66 | PO4-P | ug P/L | 0.0159 | 0.302 | 4.71 | NO3-N | ug N/L | 0.00449 | 0.0479 | 1.44 |  | 1 |  | 7.66 |  |  | 30.7229 | PASS |  |  | 1 |  | 4.71 |  |  | 7.489 | PASS |  |  | 1 |  | 1.44 |  |  | 85.259 | PASS |  |  |  | 102
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.0153 | 0.382 | 6.79 | PO4-P | ug P/L | 0.022 | 0.42 | 7.37 | NO3-N | ug N/L | 0.003 | 0.0504 | 1.57 |  | 1 |  | 6.79 |  |  | 4.7448 | PASS |  |  | 1 |  | 7.37 |  |  | 0.13578 | PASS |  |  | 1 |  | 1.57 |  |  | 28.4153 | PASS |  |  |  | 103
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0176 | 0.397 | 7.11 | PO4-P | ug P/L | 0.0218 | 0.414 | 7.22 | NO3-N | ug N/L | 0.00737 | 0.139 | 5.94 |  | 1 |  | 7.11 |  |  | 33.2943 | PASS |  |  | 1 |  | 7.22 |  |  | 3.1357 | PASS |  |  | 1 |  | 5.94 |  |  | 42.449 | PASS |  |  |  | 104
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.3 | 6.57 | 135 | PO4-P | ug P/L | 0.0175 | 0.322 | 5.15 | NO3-N | ug N/L | 0.0162 | 0.224 | 10.1 |  | 1 |  | 135 |  |  | 0.73801 | PASS |  |  | 1 |  | 5.15 |  |  | 32.1108 | PASS |  |  | 1 |  | 10.1 |  |  | 13.8249 | PASS |  |  |  | 105
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.0112 | 0.334 | 5.79 | PO4-P | ug P/L | 0.0367 | 0.301 | 4.68 | NO3-N | ug N/L | 0.00199 | 0.0316 | 0.641 |  | 1 |  | 5.79 |  |  | 16.0448 | PASS |  |  | 1 |  | 4.68 |  |  | 8 | PASS |  |  | 1 |  | 0.641 |  |  | 147.9441 | FAIL |  |  |  | 106
2021-10-05 | OM_10-5-2021_10-30-18AM edit2.omn | DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.0185 | 0.531 | 9.87 | PO4-P | ug P/L | 0.0109 | 0.206 | 2.54 | NO3-N | ug N/L | 0.0056 | 0.0786 | 2.95 |  | 1 |  | 9.87 |  |  | 97.0655 | PASS |  |  | 1 |  | 2.54 |  |  | 3.6072 | PASS |  |  | 1 |  | 2.95 |  |  | 23.0624 | PASS |  |  |  | 107
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.0179 | 0.325 | 8.11 | PO4-P | ug P/L | 0.0272 | 0.548 | 8.98 | NO3-N | ug N/L | 1.49 | 20.2 | 998 |  | 1 |  | 8.11 |  |  | 17.8643 | PASS |  |  | 1 |  | 8.98 |  |  | 1.2325 | PASS |  |  | 1 |  | 998 |  |  | 5.0781 | PASS |  |  |  | 108
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | -0.0135 | 0.168 | 4.82 | PO4-P | ug P/L | 0.0119 | 0.239 | 2.02 | NO3-N | ug N/L | 0.0487 | 0.691 | 33.4 |  | 1 |  | 4.82 |  |  | 35.122 | PASS |  |  | 1 |  | 2.02 |  |  | 8.9835 | PASS |  |  | 1 |  | 33.4 |  |  | 17.5896 | PASS |  |  |  | 109
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.0123 | 0.391 | 9.5 | PO4-P | ug P/L | 0.023 | 0.475 | 7.34 | NO3-N | ug N/L | 1.61 | 21.8 | 1080 |  | 1 |  | 9.5 |  |  | 6.1224 | PASS |  |  | 1 |  | 7.34 |  |  | 3.0429 | PASS |  |  | 1 |  | 1080 |  |  | 2.8169 | PASS |  |  |  | 110
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.0145 | 0.391 | 9.49 | PO4-P | ug P/L | 0.0225 | 0.453 | 6.85 | NO3-N | ug N/L | 1.58 | 21.3 | 1060 |  | 1 |  | 9.49 |  |  | 10.3047 | PASS |  |  | 1 |  | 6.85 |  |  | 0.72727 | PASS |  |  | 1 |  | 1060 |  |  | 0.94787 | PASS |  |  |  | 111
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.00763 | 0.199 | 5.46 | PO4-P | ug P/L | 0.0175 | 0.336 | 4.19 | NO3-N | ug N/L | 0.0698 | 0.962 | 46.6 |  | 1 |  | 5.46 |  |  | 3.2432 | PASS |  |  | 1 |  | 4.19 |  |  | 10.4072 | PASS |  |  | 1 |  | 46.6 |  |  | 6.8812 | PASS |  |  |  | 112
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.0161 | 0.388 | 9.44 | PO4-P | ug P/L | 0.0224 | 0.404 | 5.74 | NO3-N | ug N/L | 1.55 | 21.2 | 1050 |  | 1 |  | 9.44 | [6 empty] | 1 |  | 5.74 | [6 empty] | 1 |  | 1050 | [7 empty] | 113
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0123 | 0.458 | 10.9 | PO4-P | ug P/L | 0.0208 | 0.439 | 6.53 | NO3-N | ug N/L | 1.14 | 15.5 | 767 |  | 1 |  | 10.9 |  |  | 24.8996 | FAIL |  |  | 1 |  | 6.53 |  |  | 6.1563 | PASS |  |  | 1 |  | 767 |  |  | 0 | PASS |  |  |  | 114
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.0194 | 0.528 | 12.4 | PO4-P | ug P/L | 0.0114 | 0.23 | 1.81 | NO3-N | ug N/L | 0.0353 | 0.5 | 24.1 |  | 1 |  | 12.4 |  |  | 12.8302 | PASS |  |  | 1 |  | 1.81 |  |  | 13.8817 | PASS |  |  | 1 |  | 24.1 |  |  | 14.2582 | PASS |  |  |  | 115
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.011 | 0.365 | 8.95 | PO4-P | ug P/L | 0.0104 | 0.213 | 1.43 | NO3-N | ug N/L | 0.0357 | 0.512 | 24.7 |  | 1 |  | 8.95 |  |  | 46.0481 | PASS |  |  | 1 |  | 1.43 |  |  | 15.0376 | PASS |  |  | 1 |  | 24.7 |  |  | 6.2745 | PASS |  |  |  | 116
2021-10-21 | OM_10-21-2021_10-40-55AM edit2.omn | DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.00918 | 0.219 | 5.87 | PO4-P | ug P/L | 0.0142 | 0.294 | 3.25 | NO3-N | ug N/L | 0.0329 | 0.45 | 21.7 |  | 1 |  | 5.87 |  |  | 48.2224 | PASS |  |  | 1 |  | 3.25 |  |  | 17.0284 | PASS |  |  | 1 |  | 21.7 |  |  | 4.2353 | PASS |  |  |  | 117
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.00946 | 0.151 | 0.333 | PO4-P | ug P/L | 0.0167 | 0.337 | 4.88 | NO3-N | ug N/L | 1.37 | 18.8 | 969 |  | 2 | restricted recal begin to middle | 1.1263 |  |  | 10.7448 | PASS |  |  | 1 |  | 4.88 |  |  | 19.9262 | PASS |  |  | 1 |  | 969 |  |  | 6.3898 | PASS |  |  |  | 118
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0138 | 0.387 | 4.93 | PO4-P | ug P/L | 0.0149 | 0.301 | 4.06 | NO3-N | ug N/L | 1.36 | 18.6 | 962 |  | 2 | restricted recal begin to middle | 7.1395 |  |  | 45.9865 | PASS |  |  | 1 |  | 4.06 |  |  | 4.5728 | PASS |  |  | 1 |  | 962 |  |  | 2.207 | PASS |  |  |  | 119
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.0251 | 0.595 | 8.96 | PO4-P | ug P/L | 0.0312 | 0.579 | 10.3 | NO3-N | ug N/L | 1.4 | 19.2 | 991 |  | 2 | restricted recal begin to middle | 12.3994 |  |  | 20.1652 | FAIL |  |  | 1 |  | 10.3 |  |  | 1.9231 | PASS |  |  | 1 |  | 991 |  |  | 5.3886 | PASS |  |  |  | 120
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.014 | 0.419 | 5.56 | PO4-P | ug P/L | 0.00806 | 0.2 | 1.77 | NO3-N | ug N/L | 0.943 | 13 | 669 |  | 2 | restricted recal begin to middle | 7.9511 |  |  | 7.9562 | PASS |  |  | 1 |  | 1.77 |  |  | 21.2121 | PASS |  |  | 1 |  | 669 |  |  | 6.8006 | PASS |  |  |  | 121
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.0139 | 0.252 | 2.3 | PO4-P | ug P/L | 0.00991 | 0.218 | 2.19 | NO3-N | ug N/L | 0.895 | 12.4 | 638 |  | 2 | restricted recal begin to middle | 3.7056 |  |  | 21.3278 | PASS |  |  | 1 |  | 2.19 |  |  | 2.3095 | PASS |  |  | 1 |  | 638 |  |  | 2.8617 | PASS |  |  |  | 122
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.0385 | 0.965 | 16.2 | PO4-P | ug P/L | 0.0236 | 0.449 | 7.4 | NO3-N | ug N/L | 1.38 | 19 | 983 |  | 2 | restricted recal middle to end | 25.1771 |  |  | 12.9382 | PASS |  |  | 1 |  | 7.4 |  |  | 9.7803 | PASS |  |  | 1 |  | 983 |  |  | 5.4336 | PASS |  |  |  | 123
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0261 | 0.569 | 8.47 | PO4-P | ug P/L | 0.0225 | 0.43 | 6.97 | NO3-N | ug N/L | 1.31 | 18.1 | 936 |  | 2 | restricted recal middle to end | 14.9125 |  |  | 2.7521 | PASS |  |  | 1 |  | 6.97 |  |  | 3.1073 | PASS |  |  | 1 |  | 936 |  |  | 2.7071 | PASS |  |  |  | 124
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.0177 | 0.435 | 5.86 | PO4-P | ug P/L | 0.0104 | 0.224 | 2.33 | NO3-N | ug N/L | 0.925 | 12.9 | 664 |  | 2 | restricted recal middle to end | 11.4227 |  |  | 27.2603 | FAIL |  |  | 1 |  | 2.33 |  |  | 5.8333 | PASS |  |  | 1 |  | 664 |  |  | 1.8237 | PASS |  |  |  | 125
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.015 | 0.374 | 4.67 | PO4-P | ug P/L | 0.0125 | 0.266 | 3.27 | NO3-N | ug N/L | 0.6 | 8.26 | 425 |  | 2 | restricted recal middle to end | 9.8313 |  |  | 12.9091 | PASS |  |  | 1 |  | 3.27 |  |  | 8.2111 | PASS |  |  | 1 |  | 425 |  |  | 0.23557 | PASS |  |  |  | 126
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.0241 | 0.547 | 8.04 | PO4-P | ug P/L | 0.0167 | 0.326 | 4.62 | NO3-N | ug N/L | 1.25 | 17.3 | 895 |  | 2 | restricted recal middle to end | 14.3401 |  |  | 6.154 | PASS |  |  | 1 |  | 4.62 |  |  | 0 | PASS |  |  | 1 |  | 895 |  |  | 3.1886 | PASS |  |  |  | 127
2021-10-28 | OM_10-28-2021_11-37-08AM edit1.omn | DUP | 164 | 1 | 1 | NH4-N | ug N/L | 0.0533 | 1.09 | 18.7 | PO4-P | ug P/L | 0.0233 | 0.429 | 6.95 | NO3-N | ug N/L | 1.26 | 17.4 | 896 |  | 2 | restricted recal middle to end | 28.4022 |  |  | 16.6107 | PASS |  |  | 1 |  | 6.95 |  |  | 1.1577 | PASS |  |  | 1 |  | 896 |  |  | 2.5344 | PASS |  |  |  | 128
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.0125 | 0.277 | 4.12 | PO4-P | ug P/L | 0.012 | 0.202 | 1.4 | NO3-N | ug N/L | 0.0318 | 0.498 | 24.2 |  | 1 |  | 4.3427 |  |  | 59.283 | PASS |  |  | 1 |  | 1.4 |  |  | 124.5283 | FAIL |  |  | 1 |  | 24.2 |  |  | 39.0123 | FAIL |  |  |  | 129
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0111 | 0.276 | 4.11 | PO4-P | ug P/L | 0.0178 | 0.295 | 3.4 | NO3-N | ug N/L | 0.0686 | 0.971 | 48.1 |  | 1 |  | 4.3172 |  |  | 54.9815 | PASS |  |  | 1 |  | 3.4 |  |  | 38.0952 | PASS |  |  | 1 |  | 48.1 |  |  | 0.83507 | PASS |  |  |  | 130
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.019 | 0.329 | 5.25 | PO4-P | ug P/L | 0.0153 | 0.285 | 3.17 | NO3-N | ug N/L | 0.0908 | 1.31 | 65.1 |  | 1 |  | 5.6661 |  |  | 23.3612 | PASS |  |  | 1 |  | 3.17 |  |  | 42.0922 | PASS |  |  | 1 |  | 65.1 |  |  | 1.373 | PASS |  |  |  | 131
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.0131 | 0.163 | 1.65 | PO4-P | ug P/L | 0.0237 | 0.427 | 6.24 | NO3-N | ug N/L | 0.0597 | 0.859 | 42.4 |  | 1 |  | 1.4332 |  |  | 123.6219 | FAIL |  |  | 1 |  | 6.24 |  |  | 19.5251 | PASS |  |  | 1 |  | 42.4 |  |  | 1.1862 | PASS |  |  |  | 132
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.0107 | 0.358 | 5.88 | PO4-P | ug P/L | 0.0141 | 0.26 | 2.65 | NO3-N | ug N/L | 0.0116 | 0.174 | 7.79 |  | 1 |  | 6.4031 |  |  | 69.7007 | PASS |  |  | 1 |  | 2.65 |  |  | 14.3608 | PASS |  |  | 1 |  | 7.79 |  |  | 7.1809 | PASS |  |  |  | 133
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.00864 | 0.178 | 1.97 | PO4-P | ug P/L | 0.0181 | 0.339 | 4.33 | NO3-N | ug N/L | 0.0121 | 0.189 | 8.54 |  | 1 |  | 4.7063 |  |  | 32.1994 | PASS |  |  | 1 |  | 4.33 |  |  | 8.6747 | PASS |  |  | 1 |  | 8.54 |  |  | 1.1641 | PASS |  |  |  | 134
2021-11-02 | OM_11-2-2021_11-41-13AM edit2 kb.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0122 | 0.276 | 4.11 | PO4-P | ug P/L | 0.02 | 0.366 | 4.93 | NO3-N | ug N/L | 0.044 | 0.621 | 30.4 |  | 1 |  | 7.271 |  |  | 46.0542 | PASS |  |  | 1 |  | 4.93 |  |  | 31.1841 | PASS |  |  | 1 |  | 30.4 |  |  | 3.2362 | PASS |  |  |  | 135
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.0244 | 0.466 | 10.5 | PO4-P | ug P/L | 0.0118 | 0.234 | 2.15 | NO3-N | ug N/L | 0.781 | 10.8 | 603 |  | 1 |  | 10.5 |  |  | 27.1605 | FAIL |  |  | 1 |  | 2.15 |  |  | 25.7218 | PASS |  |  | 1 |  | 603 |  |  | 1.3356 | PASS |  |  |  | 136
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 0.0208 | 0.465 | 10.4 | PO4-P | ug P/L | 0.0126 | 0.275 | 3.09 | NO3-N | ug N/L | 0.741 | 10.2 | 566 |  | 1 |  | 10.4 |  |  | 10.2072 | PASS |  |  | 1 |  | 3.09 |  |  | 9.8472 | PASS |  |  | 1 |  | 566 |  |  | 0.35398 | PASS |  |  |  | 137
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 0.0263 | 0.509 | 11.3 | PO4-P | ug P/L | 0.0096 | 0.22 | 1.84 | NO3-N | ug N/L | 1.11 | 15.3 | 859 |  | 1 |  | 11.3 |  |  | 35.0365 | FAIL |  |  | 1 |  | 1.84 |  |  | 29.6296 | PASS |  |  | 1 |  | 859 |  |  | 1.5249 | PASS |  |  |  | 138
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.0275 | 0.696 | 14.9 | PO4-P | ug P/L | 0.011 | 0.234 | 2.14 | NO3-N | ug N/L | 0.776 | 10.7 | 597 |  | 1 |  | 14.9 |  |  | 16.6154 | PASS |  |  | 1 |  | 2.14 |  |  | 28.1124 | PASS |  |  | 1 |  | 597 |  |  | 2.0305 | PASS |  |  |  | 139
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.0124 | 0.0866 | 3.06 | PO4-P | ug P/L | 0.0106 | 0.223 | 1.9 | NO3-N | ug N/L | 0.0022 | 0.0408 | 2.08 |  | 1 |  | 3.06 |  |  | 91.0062 | PASS |  |  | 1 |  | 1.9 |  |  | 36.0248 | PASS |  |  | 1 |  | 2.08 |  |  | 3.9216 | PASS |  |  |  | 140
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 0.0113 | 0.252 | 6.29 | PO4-P | ug P/L | 0.0122 | 0.268 | 2.92 | NO3-N | ug N/L | 0.608 | 8.45 | 469 |  | 1 |  | 6.29 |  |  | 23.6861 | PASS |  |  | 1 |  | 2.92 |  |  | 18.7266 | PASS |  |  | 1 |  | 469 |  |  | 2.5918 | PASS |  |  |  | 141
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 0.0113 | 0.281 | 6.85 | PO4-P | ug P/L | 0.0157 | 0.276 | 3.11 | NO3-N | ug N/L | 0.00272 | 0.0427 | 2.19 |  | 1 |  | 6.85 |  |  | 40.2332 | PASS |  |  | 1 |  | 3.11 |  |  | 55.4415 | PASS |  |  | 1 |  | 2.19 |  |  | 30.2326 | PASS |  |  |  | 142
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.0167 | 0.397 | 9.09 | PO4-P | ug P/L | 0.0118 | 0.23 | 2.07 | NO3-N | ug N/L | 0.497 | 6.94 | 385 |  | 1 |  | 9.09 |  |  | 3.242 | PASS |  |  | 1 |  | 2.07 |  |  | 2.445 | PASS |  |  | 1 |  | 385 |  |  | 0 | PASS |  |  |  | 143
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.0207 | 0.475 | 10.6 | PO4-P | ug P/L | 0.0113 | 0.226 | 1.97 | NO3-N | ug N/L | 0.56 | 7.82 | 434 |  | 1 |  | 10.6 |  |  | 31.746 | FAIL |  |  | 1 |  | 1.97 |  |  | 3.9801 | PASS |  |  | 1 |  | 434 |  |  | 0.68886 | PASS |  |  |  | 144
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.0321 | 0.367 | 8.52 | PO4-P | ug P/L | 0.0251 | 0.422 | 6.4 | NO3-N | ug N/L | 0.172 | 2.42 | 133 |  | 1 |  | 8.52 |  |  | 22.8908 | PASS |  |  | 1 |  | 6.4 |  |  | 163.5847 | FAIL |  |  | 1 |  | 133 |  |  | 0 | PASS |  |  |  | 145
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 164 | 1 | 1 | NH4-N | ug N/L | -0.0139 | 0.139 | 4.07 | PO4-P | ug P/L | 0.00732 | 0.181 | 0.956 | NO3-N | ug N/L | 0.0151 | 0.226 | 12.3 |  | 1 |  | 4.07 |  |  | 47.2795 | PASS |  |  | 1 |  | 0.956 |  |  | 62.7423 | PASS |  |  | 1 |  | 12.3 |  |  | 8.5603 | PASS |  |  |  | 146
2021-11-11 | OM_11-11-2021_10-48-14AM edit1.omn | DUP | 179 | 1 | 1 | NH4-N | ug N/L | 0.0238 | 0.384 | 8.85 | PO4-P | ug P/L | 0.0113 | 0.227 | 1.99 | NO3-N | ug N/L | 0.503 | 7.04 | 390 |  | 1 |  | 8.85 |  |  | 34.965 | PASS |  |  | 1 |  | 1.99 |  |  | 23.1111 | PASS |  |  | 1 |  | 390 |  |  | 32.2581 | FAIL |  |  |  | 147
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | DUP | 14 | 1 | 1 | NH4-N | ug N/L | 0.0146 | 0.349 | 6.73 | PO4-P | ug P/L | 0.0107 | 0.214 | 1.68 | NO3-N | ug N/L | 0.00347 | 0.0574 | 1.33 |  | 1 |  | 6.73 |  |  | 61.6132 | PASS |  |  | 1 |  | 1.68 |  |  | 7.4074 | PASS |  |  | 1 |  | 1.33 |  |  | 60.733 | PASS |  |  |  | 148
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | DUP | 29 | 1 | 1 | NH4-N | ug N/L | 1.38 | 28.6 | 601 | PO4-P | ug P/L | 0.0149 | 0.239 | 2.28 | NO3-N | ug N/L | 0.00249 | 0.0206 | -0.364 |  | 1 |  | 601 |  |  | 0.8285 | PASS |  |  | 1 |  | 2.28 |  |  | 38.0107 | PASS |  |  | 3 | both BD | -0.364 |  |  | 484.375 | FAIL |  |  |  | 149
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | DUP | 44 | 1 | 1 | NH4-N | ug N/L | 4.9 | 107 | 2160 | PO4-P | ug P/L | 0.0566 | 0.801 | 15.4 | NO3-N | ug N/L | 0.00682 | 0.11 | 3.73 |  | 1 |  | 2160 |  |  | 0 | PASS |  |  | 1 |  | 15.4 |  |  | 6.7114 | PASS |  |  | 1 |  | 3.73 |  |  | 31.677 | PASS |  |  |  | 150
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | DUP | 59 | 1 | 1 | NH4-N | ug N/L | 0.0203 | 0.386 | 7.52 | PO4-P | ug P/L | 0.065 | 1.05 | 21.2 | NO3-N | ug N/L | 0.26 | 3.29 | 164 |  | 1 |  | 7.52 |  |  | 17.0274 | PASS |  |  | 1 |  | 21.2 |  |  | 1.8692 | PASS |  |  | 1 |  | 164 |  |  | 3.096 | PASS |  |  |  | 151
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | DUP | 74 | 1 | 1 | NH4-N | ug N/L | 0.015 | 0.381 | 7.4 | PO4-P | ug P/L | 0.0155 | 0.272 | 3.03 | NO3-N | ug N/L | 0.00308 | 0.046 | 0.802 |  | 1 |  | 7.4 |  |  | 193.8789 | FAIL |  |  | 1 |  | 3.03 |  |  | 3.8835 | PASS |  |  | 1 |  | 0.802 |  |  | 19.674 | PASS |  |  |  | 152
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | DUP | 89 | 1 | 1 | NH4-N | ug N/L | 1.64 | 34.8 | 730 | PO4-P | ug P/L | 0.0192 | 0.29 | 3.46 | NO3-N | ug N/L | 0.00331 | 0.0669 | 1.76 |  | 1 |  | 730 |  |  | 1.3605 | PASS |  |  | 1 |  | 3.46 |  |  | 8.7481 | PASS |  |  | 1 |  | 1.76 |  |  | 135.7177 | FAIL |  |  |  | 153
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | DUP | 104 | 1 | 1 | NH4-N | ug N/L | 2.4 | 52.3 | 1090 | PO4-P | ug P/L | 0.0546 | 0.766 | 14.6 | NO3-N | ug N/L | 0.00209 | 0.0403 | 0.539 |  | 1 |  | 1090 |  |  | 0.92166 | PASS |  |  | 1 |  | 14.6 |  |  | 12.8205 | PASS |  |  | 1 |  | 0.539 |  |  | 18.9757 | PASS |  |  |  | 154
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | DUP | 119 | 1 | 1 | NH4-N | ug N/L | 0.0469 | 1.02 | 21.2 | PO4-P | ug P/L | 0.0102 | 0.206 | 1.49 | NO3-N | ug N/L | 0.0825 | 1.07 | 49.3 |  | 1 |  | 21.2 |  |  | 6.8293 | PASS |  |  | 1 |  | 1.49 |  |  | 7.6655 | PASS |  |  | 1 |  | 49.3 |  |  | 1.0194 | PASS |  |  |  | 155
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | DUP | 134 | 1 | 1 | NH4-N | ug N/L | 0.873 | 18.8 | 396 | PO4-P | ug P/L | 0.0156 | 0.287 | 3.4 | NO3-N | ug N/L | 0.00314 | 0.0501 | 0.992 |  | 1 |  | 396 |  |  | 1.5038 | PASS |  |  | 1 |  | 3.4 |  |  | 5.7489 | PASS |  |  | 1 |  | 0.992 |  |  | 0.40241 | PASS |  |  |  | 156
2021-12-09 | OM_12-9-2021_11-05-35AM edit1.omn | DUP | 149 | 1 | 1 | NH4-N | ug N/L | 0.284 | 6.07 | 129 | PO4-P | ug P/L | 0.0182 | 0.33 | 4.39 | NO3-N | ug N/L | 0.00221 | 0.03 | 0.0647 |  | 1 |  | 129 |  |  | 2.2989 | PASS |  |  | 1 |  | 4.39 |  |  | 17.4636 | PASS |  |  | 1 |  | 0.0647 |  |  | 149.9129 | FAIL |  |  |  | 157
Known
| Mean | [20 empty] | Mean |  | 119.4897 |  |  | 19.9787 | [5 empty] | 6.7548 |  |  | 25.6695 | [5 empty] | 359.2132 |  |  | 31.1602 |  |  |  | Min | 1
| Std | [20 empty] | Std |  | 302.8428 |  |  | 26.2037 | [5 empty] | 6.3502 |  |  | 60.5655 | [5 empty] | 505.6647 |  |  | 62.7685 |  |  |  | Max | 157
| CV | [20 empty] | CV |  | 253.4468 |  |  | 131.1583 | [5 empty] | 94.0105 |  |  | 235.9437 | [5 empty] | 140.7701 |  |  | 201.4384
Warning Limits | Xbar + 2SD | [19 empty] | Warning Limits | Xbar + 2SD |  | 725.1752 |  |  | 72.3861 | [5 empty] | 19.4551 |  |  | 146.8006 | [5 empty] | 1370.5427 |  |  | 156.6972
| Xbar - 2SD | [20 empty] | Xbar - 2SD |  | -486.1958 |  |  | -32.4287 | [5 empty] | -5.9456 |  |  | -95.4616 | [5 empty] | -652.1162 |  |  | -94.3769
Control Limits | Xbar + 3SD | [19 empty] | Control Limits | Xbar + 3SD |  | 1028.0179 |  |  | 98.5899 | [5 empty] | 25.8053 |  |  | 207.3661 | [5 empty] | 1876.2075 |  |  | 219.4657
| Xbar - 3SD | [20 empty] | Xbar - 3SD |  | -789.0386 |  |  | -58.6325 | [5 empty] | -12.2958 |  |  | -156.0271 | [5 empty] | -1157.781 |  |  | -157.1454
Warning Limits as % | Xbar + 2SD | [19 empty] | Warning Limits as % | Xbar + 2SD |  | 606.8936 |  |  | 362.3166 | [5 empty] | 288.021 |  |  | 571.8874 | [5 empty] | 381.5401 |  |  | 502.8768
| Xbar - 2SD | [20 empty] | Xbar - 2SD |  | -406.8936 |  |  | -162.3166 | [5 empty] | -88.021 |  |  | -371.8874 | [5 empty] | -181.5401 |  |  | -302.8768
Control Limits as % | Xbar + 3SD | [19 empty] | Control Limits as % | Xbar + 3SD |  | 860.3404 |  |  | 493.475 | [5 empty] | 382.0315 |  |  | 807.8311 | [5 empty] | 522.3102 |  |  | 704.3151
| Xbar - 3SD | [20 empty] | Xbar - 3SD |  | -660.3404 |  |  | -293.475 | [5 empty] | -182.0315 |  |  | -607.8311 | [5 empty] | -322.3102 |  |  | -504.3151
Count | [20 empty] | Count |  |  | 157 |  |  | 151 | [5 empty] | 157 |  |  | 151 | [5 empty] | 157 |  |  | 151