## data for export
Name | TIC concentration in mg/L | TC concentration in mg/L | TOC concentration in mg/L | TNb concentration in mg/L
rolling MDL | 0.56888 | 0.88705 | 0.76186 | 0.051691
rolling LOQ | 2.3304 | 3.6337 | 3.1209 | 0.21175

## 23oct23
RUN NOTES:
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area |  | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0.037222 | -0.013215 | -0.050437 | -0.0043643
TIC | TC | TNb |  | 0.0006 | 1131.4 | 0.0012 | 2139.3 | 6e-05 | 1201.5 |  | 0.2 | 2.9132 | 5.6491 | 2.736 | 0.33269
66.5 | 380.5 | 257.5 |  | 0.0018 | 3750.7 | 0.0036 | 7099.9 | 0.00018 | 2904.5 |  | 0.6 | 3.1473 | 6.275 | 3.1276 | 0.27324
|  |  |  | 0.002997 | 5723.211 | 0.005994 | 11054.087 | 0.0002997 | 5500.005 |  | 0.333 | 8.6238 | 17.6141 | 8.9904 | 0.93812
|  |  |  | 0.004203 | 8379.889 | 0.008406 | 16022.613 | 0.0004203 | 7626.113 |  | 0.467 | 8.9852 | 18.2117 | 9.2265 | 0.92933
|  |  |  | 0.0054 | 10880.2 | 0.0108 | 20232.9 | 0.00054 | 9222.5 |  | 0.6 | 9.0709 | 17.9024 | 8.8315 | 0.87551
|  |  | Slope |  | 4.9851e-07 |  | 5.3122e-07 |  | 5.7196e-08
|  |  | Intercept |  | 1.8611e-05 |  | -6.6074e-06 |  | -2.1822e-06
|  |  | RSQ |  | 0.99881 |  | 0.99929 |  | 0.99603
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7407 | 13351 |  | 5299 | 10.163 | 19.315 | 9.152 |  | 0.73 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 13:16 |  | 1 |  | 12.3704 | 23.6189 | 11.2485 | 1.003
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7250 | 13284 |  | 4855 | 9.962 | 19.221 | 9.259 |  | 0.653 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 13:25 |  | 1 |  | 12.1095 | 23.5002 | 11.3908 | 0.91834 |  |  | 0.21932 |  |  |  |  | 0.83289 |  |  |  |  | 1.4811 |  |  |  |  | 10.8792 |  |  | 12.1228 | 23.5985 | 11.4757 | 0.87097
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7266 | 13395 |  | 4358 | 9.982 | 19.378 | 9.396 |  | 0.566 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 13:36 |  | 1 |  | 12.1361 | 23.6968 | 11.5607 | 0.82359
4 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3109 | 1350 |  | 721 | 2.8 | 1.422 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 13:53 |  | 1 |  | 3.137 | 1.4211 | -1.7159 | 0.078112
5 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 529 | 1376 |  | 451 | 0.82 | 1.444 | 0.624 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 14:02 |  | 1 |  | 0.56465 | 1.4487 | 0.88404 | 0.047226 |  |  | 12.5758 |  |  |  |  | 5.6305 |  |  |  |  | 0.91857 |  |  |  |  | 105.4198 |  |  | 0.60254 | 1.4907 | 0.88812 | 0.030926
6 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 605 | 1455 |  | 166 | 0.879 | 1.512 | 0.632 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 14:12 |  | 1 |  | 0.64042 | 1.5326 | 0.8922 | 0.014625
7 | 3 |  | Flush | [20 empty] | 2023-10-23 | 14:18 |  | 1
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 58 | 522 |  | 569 | 0.459 | 0.721 | 0.262 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 14:34 |  | 1 |  | 0.09505 | 0.54137 | 0.44633 | 0.060725
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 50 | 439 |  | 334 | 0.453 | 0.651 | 0.197 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 14:43 |  | 1 |  | 0.087074 | 0.45319 | 0.36612 | 0.033842 |  |  | 31.7817 |  |  |  |  | 31.7881 |  |  |  |  | 54.6775 |  |  |  |  | 69.7523 |  |  | 0.10352 | 0.39104 | 0.28752 | 0.025092
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 83 | 322 |  | 181 | 0.478 | 0.551 | 0.073 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 14:51 |  | 1 |  | 0.11998 | 0.32889 | 0.20891 | 0.016341
11 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 579 | 2262 |  | 1521 | 2.147 | 5.487 | 3.34 |  | 0.108 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 15:07 |  | 1 |  | 1.5363 | 5.975 | 4.4388 | 0.42406
12 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1143 | 2286 |  | 1374 | 3.229 | 5.538 | 2.309 |  | 0.069 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 15:16 |  | 1 |  | 2.9421 | 6.0388 | 3.0967 | 0.38202 |  | 0.68486 | 1.2628 |  |  |  | 0.88956 | 0.48265 |  |  |  | 2.9509 | 0.26423 |  |  |  | 20.7163 | 10.9764 |  |  | 2.9608 | 6.0534 | 3.0926 | 0.36215
13 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1158 | 2297 |  | 1235 | 3.258 | 5.561 | 2.303 |  | 0.033 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 15:26 |  | 1 |  | 2.9795 | 6.068 | 3.0885 | 0.34227
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3725 | 7143 |  | 3471 | 2.727 | 5.275 | 2.548 |  | 0.206 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 15:45 |  | 1 |  | 3.126 | 6.3131 | 3.1872 | 0.32724
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3748 | 7643 |  | 3249 | 2.742 | 5.628 | 2.886 |  | 0.186 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 15:55 |  | 1 |  | 3.1451 | 6.7558 | 3.6107 | 0.30608 |  | 7.1206 | 4.266 |  |  |  | 11.3202 | 2.2932 |  |  |  | 15.5199 | 8.3754 |  |  |  | 0.89808 | 2.236 |  |  | 3.2136 | 6.6792 | 3.4656 | 0.30269
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3913 | 7470 |  | 3178 | 2.847 | 5.506 | 2.659 |  | 0.18 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 16:06 |  | 1 |  | 3.2822 | 6.6026 | 3.3205 | 0.29931
17 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 4610 | 11231 |  | 5816 | 5.934 | 14.705 | 8.771 |  | 0.739 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 16:25 |  | 1 |  | 6.9572 | 17.8963 | 10.9391 | 0.9924
18 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 5768 | 11340 |  | 5697 | 7.268 | 14.843 | 7.576 |  | 0.72 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 16:35 |  | 1 |  | 8.6908 | 18.0702 | 9.3794 | 0.97196 |  | 3.4437 | 0.017227 |  |  |  | 0.10212 | 0.57547 |  |  |  | 3.6479 | 1.0955 |  |  |  | 7.5088 | 0.90532 |  |  | 8.6901 | 18.0184 | 9.3283 | 0.96758
19 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 5767 | 11275 |  | 5646 | 7.266 | 14.76 | 7.494 |  | 0.712 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 16:46 |  | 1 |  | 8.6893 | 17.9665 | 9.2772 | 0.9632
20 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8185 | 16344 |  | 8026 | 7.167 | 15.124 | 7.956 |  | 0.774 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 17:06 |  | 1 |  | 8.7772 | 18.5773 | 9.8001 | 0.97831
21 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8385 | 16297 |  | 7954 | 7.332 | 15.081 | 7.749 |  | 0.766 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 17:17 |  | 1 |  | 8.9907 | 18.5238 | 9.5331 | 0.96949 |  | 0.57269 | 1.3444 |  |  |  | 3.422 | 0.98988 |  |  |  | 6.2713 | 0.65433 |  |  |  | 8.0957 | 0.6924 |  |  | 9.0515 | 18.616 | 9.5644 | 0.97286
22 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8499 | 16459 |  | 8009 | 7.426 | 15.227 | 7.801 |  | 0.773 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 17:28 |  | 1 |  | 9.1124 | 18.7081 | 9.5957 | 0.97623
23 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 10590 | 20494 |  | 9602 | 7.116 | 14.701 | 7.585 |  | 0.74 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 17:49 |  | 1 |  | 8.8298 | 18.1336 | 9.3038 | 0.91169
24 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 10913 | 20655 |  | 9586 | 7.322 | 14.814 | 7.492 |  | 0.739 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 18:01 |  | 1 |  | 9.0982 | 18.2761 | 9.1779 | 0.91016 |  | 1.5246 | 0.85475 |  |  |  | 1.7036 | 0.3337 |  |  |  | 1.8826 | 0.18552 |  |  |  | 0.55177 | 1.1482 |  |  | 9.1372 | 18.3067 | 9.1694 | 0.90497
25 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11007 | 20724 |  | 9477 | 7.383 | 14.863 | 7.481 |  | 0.729 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 18:13 |  | 1 |  | 9.1763 | 18.3372 | 9.1609 | 0.89977
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5656 | 12373 |  | 5441 | 7.924 | 17.935 | 10.011 |  | 0.755 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 18:32 |  | 1 |  | 9.4607 | 21.8871 | 12.4264 | 1.0301
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6019 | 12095 |  | 5292 | 8.387 | 17.542 | 9.154 |  | 0.729 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 18:42 |  | 1 |  | 10.0639 | 21.3948 | 11.3309 | 1.0017 | 100 |  | 1.6649 |  |  | 100 |  | 0.40472 |  |  | 100 |  | 2.2076 |  |  | 100 |  | 0.8602 |  |  | 9.9808 | 21.4382 | 11.4574 | 0.99737
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5919 | 12144 |  | 5247 | 8.26 | 17.611 | 9.351 |  | 0.721 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 18:52 |  | 1 |  | 9.8977 | 21.4816 | 11.5839 | 0.99308
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4909 | 7133 |  | 3384 | 4.181 | 6.322 | 2.141 |  | 0.238 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 19:11 |  | 1 |  | 4.9316 | 7.5651 | 2.6335 | 0.38274
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3455 | 7105 |  | 3285 | 3.065 | 6.298 | 3.233 |  | 0.228 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 19:21 |  | 1 |  | 3.482 | 7.5354 | 4.0534 | 0.37141 |  |  | 2.8786 |  |  |  |  | 0.90644 |  |  |  |  | 4.2759 |  |  |  |  | 1.5206 |  |  | 3.5328 | 7.5014 | 3.9686 | 0.36861
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3557 | 7041 |  | 3236 | 3.144 | 6.243 | 3.1 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 19:32 |  | 1 |  | 3.5837 | 7.4674 | 3.8837 | 0.36581
32 | 3 |  | Flush | [20 empty] | 2023-10-23 | 19:38 |  | 1
33 | 9 | B50 12sep23 0.1m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2801 | 7019 |  | 2522 | 2.564 | 6.225 | 3.661 |  | 0.148 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 19:57 |  | 1 |  | 2.8299 | 7.444 | 4.6141 | 0.28413
34 | 9 | B50 12sep23 0.1m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3541 | 6893 |  | 2430 | 3.131 | 6.118 | 2.987 |  | 0.138 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 20:07 |  | 1 |  | 3.5677 | 7.3101 | 3.7424 | 0.27361 |  |  | 1.1807 |  |  |  |  | 0.072642 |  |  |  |  | 1.253 |  |  |  |  | 1.093 |  |  | 3.5468 | 7.3128 | 3.766 | 0.27212
35 | 9 | B50 12sep23 0.1m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3499 | 6898 |  | 2404 | 3.1 | 6.122 | 3.022 |  | 0.135 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 20:18 |  | 1 |  | 3.5258 | 7.3154 | 3.7896 | 0.27063
36 | 10 | F50 12sep23 9m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5716 | 8246 |  | 11290 | 4.8 | 7.265 | 2.465 |  | 1.065 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 20:36 |  | 1 |  | 5.7362 | 8.7476 | 3.0114 | 1.2871
37 | 10 | F50 12sep23 9m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6482 | 8164 |  | 11471 | 5.388 | 7.195 | 1.807 |  | 1.084 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 20:47 |  | 1 |  | 6.5 | 8.6605 | 2.1605 | 1.3078 |  |  | 1.2968 |  |  |  |  | 0.073578 |  |  |  |  | 4.0862 |  |  |  |  | 1.0095 |  |  | 6.4581 | 8.6637 | 2.2056 | 1.3145
38 | 10 | F50 12sep23 9m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6398 | 8170 |  | 11587 | 5.323 | 7.2 | 1.877 |  | 1.096 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 20:58 |  | 1 |  | 6.4162 | 8.6669 | 2.2506 | 1.3211
39 | 11 | B50 12sep23 9m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5786 | 7835 |  | 9237 | 4.854 | 6.916 | 2.063 |  | 0.85 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 21:16 |  | 1 |  | 5.806 | 8.3109 | 2.5049 | 1.0523
40 | 11 | B50 12sep23 9m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5817 | 7916 |  | 9340 | 4.877 | 6.985 | 2.107 |  | 0.861 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 21:27 |  | 1 |  | 5.8369 | 8.397 | 2.5601 | 1.0641 |  |  | 1.0819 |  |  |  |  | 0.16435 |  |  |  |  | 2.9489 |  |  |  |  | 2.6729 |  |  | 5.8055 | 8.4039 | 2.5984 | 1.0785
41 | 11 | B50 12sep23 9m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5754 | 7929 |  | 9592 | 4.829 | 6.996 | 2.166 |  | 0.887 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 21:38 |  | 1 |  | 5.7741 | 8.4108 | 2.6367 | 1.0929
42 | 12 | F50 29may23 9m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4615 | 6632 |  | 3604 | 3.955 | 5.897 | 1.942 |  | 0.261 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 21:56 |  | 1 |  | 4.6385 | 7.0328 | 2.3943 | 0.4079
43 | 12 | F50 29may23 9m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4398 | 6639 |  | 3577 | 3.789 | 5.903 | 2.114 |  | 0.258 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 22:07 |  | 1 |  | 4.4222 | 7.0403 | 2.6181 | 0.40481 |  |  | 0.49479 |  |  |  |  | 0.28714 |  |  |  |  | 1.6219 |  |  |  |  | 0.11297 |  |  | 4.4331 | 7.0302 | 2.5971 | 0.40504
44 | 12 | F50 29may23 9m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4420 | 6620 |  | 3581 | 3.805 | 5.887 | 2.081 |  | 0.258 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 22:18 |  | 1 |  | 4.4441 | 7.0201 | 2.576 | 0.40527
45 | 13 | F50 22aug23 0.1m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4291 | 10250 |  | 4896 | 3.707 | 8.962 | 5.256 |  | 0.396 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 22:36 |  | 1 |  | 4.3155 | 10.8767 | 6.5612 | 0.5557
46 | 13 | F50 22aug23 0.1m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4597 | 10349 |  | 4851 | 3.942 | 9.046 | 5.105 |  | 0.391 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 22:47 |  | 1 |  | 4.6206 | 10.9819 | 6.3613 | 0.55055 |  |  | 1.3469 |  |  |  |  | 0.33918 |  |  |  |  | 0.38644 |  |  |  |  | 1.1914 |  |  | 4.5897 | 10.9633 | 6.3737 | 0.54729
47 | 13 | F50 22aug23 0.1m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4535 | 10314 |  | 4794 | 3.894 | 9.016 | 5.122 |  | 0.385 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 22:58 |  | 1 |  | 4.5587 | 10.9447 | 6.386 | 0.54403
48 | 14 | F50 14aug23 3.8m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4988 | 8376 |  | 3047 | 4.242 | 7.375 | 3.133 |  | 0.203 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 23:16 |  | 1 |  | 5.0104 | 8.8857 | 3.8753 | 0.34419
49 | 14 | F50 14aug23 3.8m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5229 | 8328 |  | 2997 | 4.426 | 7.334 | 2.908 |  | 0.197 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 23:27 |  | 1 |  | 5.2507 | 8.8347 | 3.584 | 0.33847 |  |  | 1.4515 |  |  |  |  | 0.59948 |  |  |  |  | 0.66204 |  |  |  |  | 0.30464 |  |  | 5.2891 | 8.8613 | 3.5722 | 0.33795
50 | 14 | F50 14aug23 3.8m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5306 | 8378 |  | 2988 | 4.486 | 7.376 | 2.89 |  | 0.197 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 23:38 |  | 1 |  | 5.3275 | 8.8878 | 3.5604 | 0.33744
51 | 15 | F100 16may23 0.1m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4412 | 5623 |  | 1130 | 3.8 | 5.042 | 1.242 |  | 0.002 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 23:56 |  | 1 |  | 4.4361 | 5.9608 | 1.5247 | 0.1249
52 | 15 | F100 16may23 0.1m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4431 | 5648 |  | 1042 | 3.814 | 5.064 | 1.25 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 0:07 |  | 1 |  | 4.4551 | 5.9874 | 1.5323 | 0.11483 |  |  | 0.94439 |  |  |  |  | 0.053219 |  |  |  |  | 2.8981 |  |  |  |  | 2.5218 |  |  | 4.4341 | 5.989 | 1.5549 | 0.1134
53 | 15 | F100 16may23 0.1m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4389 | 5651 |  | 1017 | 3.782 | 5.066 | 1.284 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 0:18 |  | 1 |  | 4.4132 | 5.9906 | 1.5774 | 0.11197
54 | 16 | F50 12sep23 8m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6072 | 7949 |  | 11606 | 5.074 | 7.013 | 1.939 |  | 1.098 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 0:36 |  | 1 |  | 6.0912 | 8.4321 | 2.3409 | 1.3233
55 | 16 | F50 12sep23 8m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6522 | 7953 |  | 11361 | 5.418 | 7.016 | 1.598 |  | 1.072 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 0:47 |  | 1 |  | 6.5398 | 8.4363 | 1.8965 | 1.2952 |  |  | 3.6226 |  |  |  |  | 0.352 |  |  |  |  | 11.8129 |  |  |  |  | 2.0887 |  |  | 6.6605 | 8.4512 | 1.7907 | 1.3089
56 | 16 | F50 12sep23 8m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6764 | 7981 |  | 11600 | 5.604 | 7.04 | 1.436 |  | 1.097 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 0:58 |  | 1 |  | 6.7811 | 8.4661 | 1.6849 | 1.3226
57 | 17 | F50 12sep23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5064 | 12674 |  | 10455 | 4.3 | 11.016 | 6.716 |  | 0.977 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 1:17 |  | 1 |  | 5.0862 | 13.4521 | 8.3659 | 1.1916
58 | 17 | F50 12sep23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4871 | 12683 |  | 10459 | 4.152 | 11.024 | 6.872 |  | 0.978 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 1:28 |  | 1 |  | 4.8937 | 13.4616 | 8.5679 | 1.1921 |  |  | 1.0651 |  |  |  |  | 0.81623 |  |  |  |  | 0.67437 |  |  |  |  | 0.50989 |  |  | 4.8678 | 13.4069 | 8.5391 | 1.189
59 | 17 | F50 12sep23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4819 | 12580 |  | 10406 | 4.112 | 10.936 | 6.824 |  | 0.972 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 1:39 |  | 1 |  | 4.8419 | 13.3522 | 8.5103 | 1.186
60 | 18 | F50 16may23 3.8m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4164 | 6700 |  | 2164 | 3.61 | 5.954 | 2.345 |  | 0.11 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 1:58 |  | 1 |  | 4.1889 | 7.1051 | 2.9162 | 0.24318
61 | 18 | F50 16may23 3.8m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4157 | 6746 |  | 2231 | 3.604 | 5.993 | 2.389 |  | 0.117 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 2:08 |  | 1 |  | 4.1819 | 7.154 | 2.9721 | 0.25084 |  |  | 1.1992 |  |  |  |  | 0.26768 |  |  |  |  | 1.0286 |  |  |  |  | 0.8175 |  |  | 4.1569 | 7.1444 | 2.9874 | 0.25187
62 | 18 | F50 16may23 3.8m | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4107 | 6728 |  | 2249 | 3.566 | 5.978 | 2.413 |  | 0.119 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 2:19 |  | 1 |  | 4.132 | 7.1348 | 3.0028 | 0.2529
63 | 19 | SPK 13 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6503 | 15457 |  | 7288 | 5.404 | 13.373 | 7.97 |  | 0.646 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 2:38 |  | 1 |  | 6.5209 | 16.4088 | 9.8879 | 0.82932
64 | 19 | SPK 13 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7427 | 15475 |  | 7143 | 6.113 | 13.389 | 7.276 |  | 0.631 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 2:49 |  | 1 |  | 7.4422 | 16.4279 | 8.9858 | 0.81273 |  |  | 0.91517 |  | 96.1758 |  |  | 0.93978 |  | 94.8455 |  |  | 2.4503 |  | 93.5151 |  |  | 0.84094 |  | 92.0743 | 7.4083 | 16.5055 | 9.0972 | 0.81617
65 | 19 | SPK 13 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7359 | 15621 |  | 7203 | 6.061 | 13.512 | 7.451 |  | 0.637 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 3:01 |  | 1 |  | 7.3744 | 16.583 | 9.2087 | 0.8196
66 | 20 | DUP 18 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4701 | 6937 |  | 2163 | 4.021 | 6.156 | 2.135 |  | 0.11 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 3:20 |  | 1 |  | 4.7243 | 7.3569 | 2.6326 | 0.24306
67 | 20 | DUP 18 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4282 | 6891 |  | 2141 | 3.7 | 6.117 | 2.417 |  | 0.108 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 3:30 |  | 1 |  | 4.3065 | 7.308 | 3.0015 | 0.24055 |  |  | 0.34667 | 3.7076 |  |  |  | 0.42071 | 2.4747 |  |  |  | 0.52685 | 0.73331 |  |  |  | 0.95564 | 5.076 |  | 4.314 | 7.3234 | 3.0094 | 0.2394
68 | 20 | DUP 18 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4297 | 6920 |  | 2121 | 3.712 | 6.141 | 2.43 |  | 0.106 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 3:41 |  | 1 |  | 4.3215 | 7.3388 | 3.0174 | 0.23826
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1050 | 636 |  | 523 | 1.221 | 0.817 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 3:58 |  | 1 |  | 1.0841 | 0.66249 | -0.42161 | 0.055462
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 267 | 624 |  | 506 | 0.62 | 0.807 | 0.187 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 4:08 |  | 1 |  | 0.30343 | 0.64974 | 0.34631 | 0.053518 |  |  | 15.4624 |  |  |  |  | 8.7022 |  |  |  |  | 35.7499 |  |  |  |  | 3.1556 |  |  | 0.32885 | 0.62265 | 0.2938 | 0.054376
71 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 318 | 573 |  | 521 | 0.659 | 0.764 | 0.106 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 4:17 |  | 1 |  | 0.35428 | 0.59556 | 0.24128 | 0.055234
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3857 | 11253 |  | 5079 | 5.623 | 16.353 | 10.73 |  | 0.692 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 4:35 |  | 1 |  | 6.4713 | 19.9039 | 13.4326 | 0.96105
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5606 | 11347 |  | 5086 | 7.86 | 16.486 | 8.626 |  | 0.693 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 4:45 |  | 1 |  | 9.3776 | 20.0703 | 10.6927 | 0.96239 | 96.1049 |  | 4.5204 |  |  | 93.168 |  | 0.98412 |  |  | 90.3179 |  | 6.0732 |  |  | 94.9331 |  | 3.364 |  |  | 9.5945 | 19.9721 | 10.3776 | 0.94647
74 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5867 | 11236 |  | 4919 | 8.193 | 16.33 | 8.137 |  | 0.664 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 4:57 |  | 1 |  | 9.8113 | 19.8738 | 10.0625 | 0.93055
75 | 3 |  | Flush | [20 empty] | 2023-10-24 | 5:03 |  | 1
76 | 21 | F50 22aug23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5036 | 10233 |  | 5982 | 4.279 | 8.948 | 4.669 |  | 0.51 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 5:22 |  | 1 |  | 5.0583 | 10.8587 | 5.8004 | 0.67993
77 | 21 | F50 22aug23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5991 | 10124 |  | 5895 | 5.011 | 8.856 | 3.845 |  | 0.501 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 5:33 |  | 1 |  | 6.0104 | 10.7429 | 4.7324 | 0.66997 |  |  | 4.8715 |  |  |  |  | 0.8371 |  |  |  |  | 4.5337 |  |  |  |  | 1.2724 |  |  | 6.1605 | 10.788 | 4.6275 | 0.67426
78 | 21 | F50 22aug23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6292 | 10209 |  | 5970 | 5.242 | 8.928 | 3.685 |  | 0.508 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 5:44 |  | 1 |  | 6.3105 | 10.8332 | 4.5226 | 0.67855
79 | 22 | F50 31jul23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5615 | 11528 |  | 6526 | 4.722 | 10.045 | 5.323 |  | 0.567 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 6:03 |  | 1 |  | 5.6355 | 12.2345 | 6.599 | 0.74216
80 | 22 | F50 31jul23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5659 | 11613 |  | 6591 | 4.756 | 10.117 | 5.361 |  | 0.573 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 6:14 |  | 1 |  | 5.6794 | 12.3248 | 6.6454 | 0.74959 |  |  | 2.3085 |  |  |  |  | 0.08624 |  |  |  |  | 1.7747 |  |  |  |  | 0.030526 |  |  | 5.6146 | 12.3195 | 6.7049 | 0.74948
81 | 22 | F50 31jul23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5529 | 11603 |  | 6589 | 4.657 | 10.108 | 5.452 |  | 0.573 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 6:25 |  | 1 |  | 5.5498 | 12.3142 | 6.7644 | 0.74936
82 | 23 | B50 24jul23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5434 | 8411 |  | 12949 | 4.584 | 7.404 | 2.82 |  | 1.238 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 6:44 |  | 1 |  | 5.4551 | 8.9229 | 3.4678 | 1.4769
83 | 23 | B50 24jul23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5646 | 8435 |  | 12977 | 4.746 | 7.425 | 2.678 |  | 1.241 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 6:54 |  | 1 |  | 5.6664 | 8.9484 | 3.282 | 1.4801 |  |  | 0.35129 |  |  |  |  | 0.26154 |  |  |  |  | 1.3285 |  |  |  |  | 1.7088 |  |  | 5.6764 | 8.9367 | 3.2603 | 1.4929
84 | 23 | B50 24jul23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5666 | 8413 |  | 13200 | 4.761 | 7.406 | 2.645 |  | 1.265 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 7:06 |  | 1 |  | 5.6864 | 8.925 | 3.2386 | 1.5056
85 | 24 | F50 22aug23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5794 | 9997 |  | 6782 | 4.86 | 8.747 | 3.887 |  | 0.593 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 7:24 |  | 1 |  | 5.814 | 10.6079 | 4.7939 | 0.77144
86 | 24 | F50 22aug23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6069 | 10034 |  | 6874 | 5.071 | 8.779 | 3.708 |  | 0.603 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 7:35 |  | 1 |  | 6.0882 | 10.6472 | 4.559 | 0.78196 |  |  | 0.85522 |  |  |  |  | 0.29891 |  |  |  |  | 1.8196 |  |  |  |  | 0.30673 |  |  | 6.0623 | 10.6632 | 4.6009 | 0.78316
87 | 24 | F50 22aug23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6017 | 10064 |  | 6895 | 5.031 | 8.805 | 3.773 |  | 0.605 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 7:46 |  | 1 |  | 6.0363 | 10.6791 | 4.6428 | 0.78437
88 | 25 | B50 22may23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3633 | 5525 |  | 2191 | 3.202 | 4.959 | 1.757 |  | 0.113 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 8:05 |  | 1 |  | 3.6594 | 5.8567 | 2.1973 | 0.24627
89 | 25 | B50 22may23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3139 | 5463 |  | 2075 | 2.823 | 4.907 | 2.083 |  | 0.101 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 8:15 |  | 1 |  | 3.1669 | 5.7909 | 2.624 | 0.233 |  |  | 0.40844 |  |  |  |  | 1.1129 |  |  |  |  | 1.9566 |  |  |  |  | 2.1336 |  |  | 3.1734 | 5.8233 | 2.6499 | 0.23054
90 | 25 | B50 22may23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3152 | 5524 |  | 2032 | 2.833 | 4.959 | 2.125 |  | 0.097 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 8:26 |  | 1 |  | 3.1799 | 5.8557 | 2.6758 | 0.22808
91 | 26 | F50 24jul23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4682 | 6762 |  | 6241 | 4.007 | 6.007 | 2 |  | 0.537 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 8:44 |  | 1 |  | 4.7053 | 7.171 | 2.4656 | 0.70955
92 | 26 | F50 24jul23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5219 | 6761 |  | 6240 | 4.419 | 6.006 | 1.587 |  | 0.537 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 8:55 |  | 1 |  | 5.2407 | 7.1699 | 1.9292 | 0.70944 |  |  | 0.07607 |  |  |  |  | 0.25159 |  |  |  |  | 0.72684 |  |  |  |  | 0.9396 |  |  | 5.2427 | 7.1789 | 1.9362 | 0.70612
93 | 26 | F50 24jul23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5223 | 6778 |  | 6182 | 4.422 | 6.021 | 1.599 |  | 0.531 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 9:06 |  | 1 |  | 5.2447 | 7.188 | 1.9432 | 0.7028
94 | 27 | F50 22aug23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5863 | 9108 |  | 8564 | 4.913 | 7.994 | 3.081 |  | 0.78 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 9:25 |  | 1 |  | 5.8828 | 9.6634 | 3.7806 | 0.97529
95 | 27 | F50 22aug23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6070 | 9076 |  | 8404 | 5.071 | 7.967 | 2.896 |  | 0.763 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 9:35 |  | 1 |  | 6.0892 | 9.6294 | 3.5402 | 0.95698 |  |  | 3.412 |  |  |  |  | 0.57208 |  |  |  |  | 4.5094 |  |  |  |  | 0.4532 |  |  | 6.1949 | 9.657 | 3.4622 | 0.95916
96 | 27 | F50 22aug23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6282 | 9128 |  | 8442 | 5.234 | 8.012 | 2.778 |  | 0.767 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 9:46 |  | 1 |  | 6.3006 | 9.6847 | 3.3841 | 0.96133
97 | 28 | B50 22aug23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5848 | 8194 |  | 13184 | 4.901 | 7.221 | 2.32 |  | 1.263 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 10:05 |  | 1 |  | 5.8678 | 8.6924 | 2.8245 | 1.5038
98 | 28 | B50 22aug23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6018 | 8244 |  | 13373 | 5.032 | 7.263 | 2.231 |  | 1.283 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 10:16 |  | 1 |  | 6.0373 | 8.7455 | 2.7081 | 1.5254 |  |  | 0.26388 |  |  |  |  | 1.3823 |  |  |  |  | 5.1515 |  |  |  |  | 0.63947 |  |  | 6.0453 | 8.6855 | 2.6401 | 1.5205
99 | 28 | B50 22aug23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6034 | 8131 |  | 13288 | 5.044 | 7.167 | 2.122 |  | 1.274 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 10:29 |  | 1 |  | 6.0533 | 8.6254 | 2.5721 | 1.5157
100 | 29 | F50 24jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4656 | 10566 |  | 6624 | 3.987 | 9.23 | 5.244 |  | 0.577 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 10:48 |  | 1 |  | 4.6794 | 11.2124 | 6.5331 | 0.75337
101 | 29 | F50 24jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4495 | 10671 |  | 6673 | 3.864 | 9.319 | 5.456 |  | 0.582 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 10:58 |  | 1 |  | 4.5189 | 11.324 | 6.8051 | 0.75897 |  |  | 0.28724 |  |  |  |  | 0.14083 |  |  |  |  | 0.043728 |  |  |  |  | 0.96928 |  |  | 4.5124 | 11.316 | 6.8036 | 0.75531
102 | 29 | F50 24jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4482 | 10656 |  | 6609 | 3.854 | 9.306 | 5.453 |  | 0.575 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 11:09 |  | 1 |  | 4.5059 | 11.3081 | 6.8022 | 0.75165
103 | 30 | B50 14aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3470 | 7179 |  | 2649 | 3.077 | 6.36 | 3.283 |  | 0.161 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 11:29 |  | 1 |  | 3.4969 | 7.614 | 4.1171 | 0.29866
104 | 30 | B50 14aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3178 | 7250 |  | 2739 | 2.853 | 6.421 | 3.567 |  | 0.17 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 11:39 |  | 1 |  | 3.2058 | 7.6894 | 4.4836 | 0.30895 |  |  | 3.2432 |  |  |  |  | 1.2653 |  |  |  |  | 4.6176 |  |  |  |  | 2.7399 |  |  | 3.2586 | 7.6411 | 4.3825 | 0.30478
105 | 30 | B50 14aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3284 | 7159 |  | 2666 | 2.935 | 6.343 | 3.409 |  | 0.163 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 11:50 |  | 1 |  | 3.3115 | 7.5927 | 4.2813 | 0.3006
106 | 31 | SPK 25 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5870 | 12384 |  | 4721 | 4.918 | 10.77 | 5.852 |  | 0.378 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 12:09 |  | 1 |  | 5.8898 | 13.144 | 7.2542 | 0.53568
107 | 31 | SPK 25 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6579 | 12836 |  | 4760 | 5.462 | 11.153 | 5.69 |  | 0.382 | 1 | 0 | 0 |  | 0 |  | TCpe | 2023-10-24 | 12:25 |  | 1 |  | 6.5967 | 13.6242 | 7.0275 | 0.54014 |  |  | 5.168 |  | 121.9741 |  |  | 4.7174 |  | 126.7792 |  |  | 14.9552 |  | 131.5844 |  |  | 1.0431 |  | 103.8781 | 6.7717 | 13.3102 | 6.5386 | 0.53734
108 | 31 | SPK 25 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6930 | 12245 |  | 4711 | 5.731 | 10.652 | 4.921 |  | 0.377 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 12:36 |  | 1 |  | 6.9466 | 12.9963 | 6.0496 | 0.53453
109 | 32 | DUP 30 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3997 | 7259 |  | 2821 | 3.481 | 6.428 | 2.947 |  | 0.179 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 12:55 |  | 1 |  | 4.0223 | 7.699 | 3.6766 | 0.31833
110 | 32 | DUP 30 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3377 | 7215 |  | 2804 | 3.005 | 6.391 | 3.386 |  | 0.177 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 13:05 |  | 1 |  | 3.4042 | 7.6522 | 4.248 | 0.31639 |  |  | 2.1312 | 3.3099 |  |  |  | 0.76071 | 0.52697 |  |  |  | 3.0191 | 1.5939 |  |  |  | 3.0267 | 5.2625 |  | 3.3683 | 7.6815 | 4.3132 | 0.32125
111 | 32 | DUP 30 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3305 | 7270 |  | 2889 | 2.951 | 6.437 | 3.487 |  | 0.186 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 13:16 |  | 1 |  | 3.3324 | 7.7107 | 4.3783 | 0.32611
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 905 | 783 |  | 614 | 1.109 | 0.942 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 13:33 |  | 1 |  | 0.93953 | 0.81867 | -0.12086 | 0.065872
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | [9 empty] | 1
114 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5
115 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3
116 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3
117 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3

## 24oct23
RUN NOTES:
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area |  | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0.035525 | 0.16791 | 0.13238 | 0.00058105
TIC | TC | TNb |  | 0.0006 | 1166 | 0.0012 | 2027.3 | 6e-05 | 1057.2 |  | 0.2 | 2.9697 | 5.6835 | 2.7139 | 0.31106
65 | 218 | 377 |  | 0.0018 | 3716 | 0.0036 | 6932.4 | 0.00018 | 2840.6 |  | 0.6 | 3.09 | 6.1398 | 3.0498 | 0.27778
|  |  |  | 0.002997 | 5804.71 | 0.005994 | 10969.312 | 0.0002997 | 5305.418 |  | 0.333 | 8.667 | 17.3579 | 8.6909 | 0.93404
|  |  |  | 0.004203 | 8506.79 | 0.008406 | 16183.388 | 0.0004203 | 7383.888 |  | 0.467 | 9.0392 | 18.1752 | 9.1359 | 0.92671
|  |  |  | 0.0054 | 10946.5 | 0.0108 | 20688.9 | 0.00054 | 8993.1 |  | 0.6 | 9.0448 | 18.0458 | 9.001 | 0.87837
|  |  | Slope |  | 4.9414e-07 |  | 5.1929e-07 |  | 5.8571e-08
|  |  | Intercept |  | 1.7762e-05 |  | 8.3953e-05 |  | 2.9053e-07
|  |  | RSQ |  | 0.99924 |  | 0.99919 |  | 0.9971
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4699 | 12246 |  | 5242 | 6.7 | 17.755 | 11.055 |  | 0.72 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 14:16 |  | 1 |  | 7.7991 | 21.4772 | 13.6781 | 1.0244
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6758 | 12343 |  | 5241 | 9.332 | 17.892 | 8.56 |  | 0.72 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 14:26 |  | 1 |  | 11.1906 | 21.6451 | 10.4545 | 1.0242 |  |  | 4.1793 |  |  |  |  | 1.0105 |  |  |  |  | 2.4971 |  |  |  |  | 1.2879 |  |  | 11.4294 | 21.755 | 10.3256 | 1.0308
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7048 | 12470 |  | 5309 | 9.703 | 18.071 | 8.368 |  | 0.732 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 14:37 |  | 1 |  | 11.6682 | 21.865 | 10.1967 | 1.0375
4 | 3 | Rinse | TIC |  |  | 0.5 |  | 3141 |  |  |  | 2.825 |  |  |  |  | 1 |  | 0 |  |  |  |  | 2023-10-24 | 14:48 |  | 1 |  | 3.1397 | #DIV/0! | #DIV/0! | #DIV/0!
5 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 435 | 459 |  | 541 | 0.749 | 0.668 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 15:04 |  | 1 |  | 0.46543 | 0.64461 | 0.17919 | 0.063955 |  |  | 69.1536 |  |  |  |  | 8.3939 |  |  |  |  | 68.634 |  |  |  |  | 4.1125 |  |  | 0.34585 | 0.61865 | 0.2728 | 0.062666
6 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 193 | 409 |  | 519 | 0.563 | 0.625 | 0.063 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 15:14 |  | 1 |  | 0.22626 | 0.59268 | 0.36642 | 0.061378
7 | 3 |  | Flush | [20 empty] | 2023-10-24 | 15:19 |  | 1
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 47 | 252 |  | 403 | 0.451 | 0.492 | 0.041 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 15:34 |  | 1 |  | 0.081974 | 0.42963 | 0.34765 | 0.047789
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 63 | 235 |  | 368 | 0.463 | 0.477 | 0.014 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 15:42 |  | 1 |  | 0.097787 | 0.41197 | 0.31419 | 0.043689 |  |  | 3.9625 |  |  |  |  | 8.9552 |  |  |  |  | 13.3303 |  |  |  |  | 4.7125 |  |  | 0.099763 | 0.39432 | 0.29455 | 0.044744
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 67 | 201 |  | 386 | 0.466 | 0.449 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 15:51 |  | 1 |  | 0.10174 | 0.37666 | 0.27492 | 0.045798
11 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 587 | 2031 |  | 1198 | 2.162 | 4.998 | 2.836 |  | 0.023 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 16:07 |  | 1 |  | 1.5391 | 5.6931 | 4.154 | 0.35229
12 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1204 | 2103 |  | 1202 | 3.347 | 5.15 | 1.803 |  | 0.024 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 16:16 |  | 1 |  | 3.0635 | 5.8801 | 2.8165 | 0.35346 |  | 1.1298 | 0.97249 |  |  |  | 1.5009 | 1.0105 |  |  |  | 3.1363 | 3.1234 |  |  |  | 18.407 | 0.98932 |  |  | 3.0487 | 5.9099 | 2.8612 | 0.35522
13 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1192 | 2126 |  | 1214 | 3.323 | 5.198 | 1.876 |  | 0.027 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 16:26 |  | 1 |  | 3.0339 | 5.9398 | 2.9059 | 0.35698
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3745 | 7024 |  | 3305 | 2.74 | 5.191 | 2.451 |  | 0.191 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 16:45 |  | 1 |  | 3.1139 | 6.2191 | 3.1052 | 0.32311
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3719 | 7124 |  | 3262 | 2.723 | 5.261 | 2.538 |  | 0.188 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 16:55 |  | 1 |  | 3.0925 | 6.3056 | 3.2132 | 0.31891 |  | 5.1409 | 3.9165 |  |  |  | 6.1033 | 1.9033 |  |  |  | 7.0656 | 0.073724 |  |  |  | 7.3137 | 1.8799 |  |  | 3.1542 | 6.3662 | 3.212 | 0.32194
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3869 | 7264 |  | 3324 | 2.82 | 5.36 | 2.541 |  | 0.193 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 17:07 |  | 1 |  | 3.216 | 6.4268 | 3.2108 | 0.32497
17 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 4756 | 11031 |  | 5496 | 6.101 | 14.45 | 8.349 |  | 0.689 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 17:25 |  | 1 |  | 7.1108 | 17.4541 | 10.3433 | 0.96756
18 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 5871 | 11163 |  | 5580 | 7.386 | 14.618 | 7.232 |  | 0.702 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 17:35 |  | 1 |  | 8.7654 | 17.66 | 8.8946 | 0.98233 |  | 2.9863 | 0.78179 |  |  |  | 2.3092 | 0.86022 |  |  |  | 1.6321 | 0.93757 |  |  |  | 8.6887 | 0.8451 |  |  | 8.7312 | 17.5844 | 8.8531 | 0.9782
19 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 5825 | 11066 |  | 5533 | 7.333 | 14.495 | 7.162 |  | 0.695 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 17:46 |  | 1 |  | 8.6971 | 17.5087 | 8.8116 | 0.97406
20 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8420 | 16257 |  | 7663 | 7.36 | 15.044 | 7.684 |  | 0.734 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 18:06 |  | 1 |  | 8.9474 | 18.257 | 9.3096 | 0.96171
21 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8514 | 16363 |  | 7555 | 7.437 | 15.14 | 7.703 |  | 0.722 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 18:17 |  | 1 |  | 9.0469 | 18.3749 | 9.328 | 0.94817 |  | 1.1497 | 1.2437 |  |  |  | 2.231 | 0.29005 |  |  |  | 3.3124 | 0.64362 |  |  |  | 5.8048 | 0.85611 |  |  | 9.1035 | 18.4016 | 9.2981 | 0.95224
22 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8621 | 16411 |  | 7620 | 7.525 | 15.184 | 7.658 |  | 0.729 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 18:29 |  | 1 |  | 9.1601 | 18.4283 | 9.2682 | 0.95632
23 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 10649 | 20821 |  | 9526 | 7.154 | 14.932 | 7.778 |  | 0.734 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 18:50 |  | 1 |  | 8.7998 | 18.1601 | 9.3603 | 0.9304
24 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11096 | 20972 |  | 9439 | 7.439 | 15.038 | 7.599 |  | 0.726 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 19:02 |  | 1 |  | 9.1679 | 18.2908 | 9.1229 | 0.9219 |  | 1.2116 | 1.2929 |  |  |  | 1.5122 | 0.20367 |  |  |  | 1.8127 | 0.87912 |  |  |  | 2.5041 | 0.13756 |  |  | 9.109 | 18.2722 | 9.1631 | 0.92254
25 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 10953 | 20929 |  | 9452 | 7.348 | 15.008 | 7.66 |  | 0.727 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 19:15 |  | 1 |  | 9.0502 | 18.2536 | 9.2034 | 0.92317
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5819 | 11940 |  | 5036 | 8.132 | 17.324 | 9.191 |  | 0.685 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 19:34 |  | 1 |  | 9.6439 | 20.9475 | 11.3036 | 0.98418
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6101 | 11991 |  | 5024 | 8.492 | 17.395 | 8.903 |  | 0.682 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 19:44 |  | 1 |  | 10.1084 | 21.0358 | 10.9274 | 0.98184 | 100 |  | 0.24412 |  |  | 100 |  | 0.25476 |  |  | 100 |  | 0.2646 |  |  | 100 |  | 0.98932 |  |  | 10.1208 | 21.0626 | 10.9419 | 0.98672
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6116 | 12022 |  | 5074 | 8.512 | 17.439 | 8.927 |  | 0.691 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 19:55 |  | 1 |  | 10.1331 | 21.0895 | 10.9564 | 0.9916
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5301 | 7311 |  | 2869 | 4.482 | 6.472 | 1.99 |  | 0.184 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 20:14 |  | 1 |  | 5.2744 | 7.7609 | 2.4865 | 0.33666
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3980 | 7254 |  | 2934 | 3.468 | 6.424 | 2.956 |  | 0.191 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 20:25 |  | 1 |  | 3.9689 | 7.7017 | 3.7329 | 0.34428 |  |  | 1.8014 |  |  |  |  | 0.18897 |  |  |  |  | 2.3495 |  |  |  |  | 1.1292 |  |  | 4.005 | 7.6945 | 3.6895 | 0.34234
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4053 | 7240 |  | 2901 | 3.524 | 6.412 | 2.888 |  | 0.187 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 20:36 |  | 1 |  | 4.041 | 7.6872 | 3.6462 | 0.34041
32 | 3 |  | Flush | [20 empty] | 2023-10-24 | 20:42 |  | 1
33 | 9 | B50 21MAR23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2240 | 5172 |  | 2726 | 2.133 | 4.661 | 2.527 |  | 0.169 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 21:00 |  | 1 |  | 2.2493 | 5.5394 | 3.2902 | 0.31991
34 | 9 | B50 21MAR23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2692 | 5183 |  | 2748 | 2.48 | 4.669 | 2.189 |  | 0.171 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 21:10 |  | 1 |  | 2.696 | 5.5509 | 2.8549 | 0.32249 |  |  | 4.5146 |  |  |  |  | 1.129 |  |  |  |  | 6.7659 |  |  |  |  | 1.5005 |  |  | 2.7582 | 5.5197 | 2.7615 | 0.32009
35 | 9 | B50 21MAR23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2818 | 5123 |  | 2707 | 2.576 | 4.618 | 2.042 |  | 0.167 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 21:20 |  | 1 |  | 2.8205 | 5.4885 | 2.668 | 0.31768
36 | 10 | B50 22AUG22 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3056 | 7318 |  | 2985 | 2.76 | 6.479 | 3.719 |  | 0.196 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 21:38 |  | 1 |  | 3.0557 | 7.7682 | 4.7125 | 0.35025
37 | 10 | B50 22AUG22 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3258 | 7277 |  | 3042 | 2.914 | 6.443 | 3.529 |  | 0.202 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 21:49 |  | 1 |  | 3.2554 | 7.7256 | 4.4703 | 0.35693 |  |  | 1.5962 |  |  |  |  | 0.22827 |  |  |  |  | 0.77978 |  |  |  |  | 2.7279 |  |  | 3.2815 | 7.7345 | 4.4529 | 0.35212
38 | 10 | B50 22AUG22 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3311 | 7294 |  | 2960 | 2.955 | 6.458 | 3.502 |  | 0.194 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 21:59 |  | 1 |  | 3.3077 | 7.7433 | 4.4356 | 0.34732
39 | 11 | F50 24JUL23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4689 | 6751 |  | 6189 | 4.012 | 5.998 | 1.986 |  | 0.531 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 22:18 |  | 1 |  | 4.6696 | 7.1793 | 2.5098 | 0.72557
40 | 11 | F50 24JUL23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5271 | 6827 |  | 6192 | 4.459 | 6.062 | 1.604 |  | 0.532 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 22:28 |  | 1 |  | 5.2448 | 7.2583 | 2.0135 | 0.72592 |  |  | 0.22586 |  |  |  |  | 0.48769 |  |  |  |  | 2.3705 |  |  |  |  | 1.4793 |  |  | 5.2507 | 7.2406 | 1.9899 | 0.72059
41 | 11 | F50 24JUL23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5283 | 6793 |  | 6101 | 4.468 | 6.033 | 1.565 |  | 0.522 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 22:39 |  | 1 |  | 5.2566 | 7.223 | 1.9663 | 0.71526
42 | 12 | B50 11OCT22 7 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5429 | 7929 |  | 15387 | 4.58 | 6.996 | 2.417 |  | 1.493 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 22:58 |  | 1 |  | 5.4009 | 8.4028 | 3.0019 | 1.803
43 | 12 | B50 11OCT22 7 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5535 | 8031 |  | 16093 | 4.662 | 7.082 | 2.421 |  | 1.567 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 23:08 |  | 1 |  | 5.5057 | 8.5087 | 3.0031 | 1.8857 |  |  | 1.9576 |  |  |  |  | 0.52624 |  |  |  |  | 2.0459 |  |  |  |  | 1.9128 |  |  | 5.4523 | 8.4864 | 3.0341 | 1.8679
44 | 12 | B50 11OCT22 7 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5427 | 7988 |  | 15788 | 4.578 | 7.046 | 2.468 |  | 1.535 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 23:20 |  | 1 |  | 5.3989 | 8.4641 | 3.0651 | 1.85
45 | 13 | F50 31JUL23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5032 | 8487 |  | 3765 | 4.275 | 7.469 | 3.194 |  | 0.278 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 23:38 |  | 1 |  | 5.0086 | 8.9823 | 3.9737 | 0.44162
46 | 13 | F50 31JUL23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5115 | 8517 |  | 3691 | 4.339 | 7.494 | 3.155 |  | 0.27 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 23:48 |  | 1 |  | 5.0906 | 9.0135 | 3.9229 | 0.43295 |  |  | 0.2527 |  |  |  |  | 0.48278 |  |  |  |  | 1.4292 |  |  |  |  | 0.70101 |  |  | 5.0842 | 9.0353 | 3.9511 | 0.43447
47 | 13 | F50 31JUL23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5102 | 8559 |  | 3717 | 4.329 | 7.529 | 3.201 |  | 0.273 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 23:59 |  | 1 |  | 5.0777 | 9.0571 | 3.9793 | 0.436
48 | 14 | F200 14AUG23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6339 | 9292 |  | 2726 | 5.278 | 8.151 | 2.873 |  | 0.169 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 0:18 |  | 1 |  | 6.3003 | 9.8184 | 3.5181 | 0.31991
49 | 14 | F200 14AUG23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6849 | 9322 |  | 2699 | 5.669 | 8.176 | 2.507 |  | 0.166 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 0:29 |  | 1 |  | 6.8043 | 9.8495 | 3.0453 | 0.31675 |  |  | 0.84598 |  |  |  |  | 0.81523 |  |  |  |  | 0.74656 |  |  |  |  | 1.603 |  |  | 6.7756 | 9.8095 | 3.0339 | 0.31423
50 | 14 | F200 14AUG23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6791 | 9245 |  | 2656 | 5.625 | 8.111 | 2.485 |  | 0.162 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 0:40 |  | 1 |  | 6.747 | 9.7696 | 3.0226 | 0.31171
51 | 15 | B50 12SEP23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5064 | 6462 |  | 1701 | 4.3 | 5.753 | 1.453 |  | 0.062 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 1:10 |  | 1 |  | 5.0402 | 6.8792 | 1.839 | 0.19984
52 | 15 | B50 12SEP23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4824 | 6469 |  | 1664 | 4.116 | 5.759 | 1.643 |  | 0.058 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 1:20 |  | 1 |  | 4.803 | 6.8865 | 2.0835 | 0.19551 |  |  | 1.2422 |  |  |  |  | 0.84101 |  |  |  |  | 5.483 |  |  |  |  | 0.77591 |  |  | 4.7734 | 6.9155 | 2.1422 | 0.19627
53 | 15 | B50 12SEP23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4764 | 6525 |  | 1677 | 4.07 | 5.806 | 1.737 |  | 0.059 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 1:32 |  | 1 |  | 4.7437 | 6.9446 | 2.2009 | 0.19703
54 | 16 | F50 24JUL23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4971 | 8702 |  | 2923 | 4.229 | 7.651 | 3.422 |  | 0.19 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 1:50 |  | 1 |  | 4.9483 | 9.2056 | 4.2573 | 0.34299
55 | 16 | F50 24JUL23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5195 | 8827 |  | 2899 | 4.401 | 7.756 | 3.356 |  | 0.187 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 2:01 |  | 1 |  | 5.1697 | 9.3354 | 4.1658 | 0.34018 |  |  | 1.8328 |  |  |  |  | 0.83788 |  |  |  |  | 0.38319 |  |  |  |  | 2.4752 |  |  | 5.1227 | 9.2965 | 4.1738 | 0.33602
56 | 16 | F50 24JUL23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5100 | 8752 |  | 2828 | 4.328 | 7.693 | 3.365 |  | 0.18 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 2:12 |  | 1 |  | 5.0758 | 9.2575 | 4.1818 | 0.33186
57 | 17 | F50 14AUG23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5502 | 7251 |  | 5735 | 4.636 | 6.422 | 1.786 |  | 0.484 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 2:30 |  | 1 |  | 5.4731 | 7.6986 | 2.2256 | 0.67239
58 | 17 | F50 14AUG23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5598 | 7254 |  | 5643 | 4.709 | 6.424 | 1.714 |  | 0.474 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 2:41 |  | 1 |  | 5.5679 | 7.7017 | 2.1338 | 0.66161 |  |  | 3.8984 |  |  |  |  | 0.75802 |  |  |  |  | 14.0185 |  |  |  |  | 1.7028 |  |  | 5.6786 | 7.6727 | 1.994 | 0.66729
59 | 17 | F50 14AUG23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5822 | 7198 |  | 5740 | 4.881 | 6.377 | 1.495 |  | 0.484 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 2:52 |  | 1 |  | 5.7893 | 7.6436 | 1.8543 | 0.67298
60 | 18 | F50 14AUG23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5868 | 7438 |  | 7422 | 4.917 | 6.58 | 1.663 |  | 0.66 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 3:10 |  | 1 |  | 5.8348 | 7.8928 | 2.0581 | 0.87001
61 | 18 | F50 14AUG23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6057 | 7372 |  | 7427 | 5.061 | 6.524 | 1.463 |  | 0.661 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 3:21 |  | 1 |  | 6.0216 | 7.8243 | 1.8027 | 0.87059 |  |  | 1.4016 |  |  |  |  | 0.34571 |  |  |  |  | 6.4117 |  |  |  |  | 0.75067 |  |  | 6.0641 | 7.8108 | 1.7467 | 0.87387
62 | 18 | F50 14AUG23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6143 | 7346 |  | 7483 | 5.127 | 6.502 | 1.374 |  | 0.667 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 3:32 |  | 1 |  | 6.1065 | 7.7973 | 1.6908 | 0.87715
63 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8348 | 15028 |  | 5825 | 6.819 | 13.011 | 6.192 |  | 0.493 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 3:51 |  | 1 |  | 8.2857 | 15.7757 | 7.4899 | 0.68293
64 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9469 | 15280 |  | 5778 | 7.679 | 13.224 | 5.545 |  | 0.488 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 4:02 |  | 1 |  | 9.3936 | 16.0374 | 6.6438 | 0.67743 |  |  | 0.60149 |  | 150.7293 |  |  | 0.98268 |  | 117.7885 |  |  | 1.5241 |  | 90.0593 |  |  | 1.393 |  | 81.4404 | 9.3654 | 15.959 | 6.5935 | 0.67274
65 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9412 | 15129 |  | 5698 | 7.636 | 13.096 | 5.46 |  | 0.48 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 4:13 |  | 1 |  | 9.3372 | 15.8805 | 6.5433 | 0.66806
66 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6624 | 7626 |  | 7463 | 5.496 | 6.739 | 1.243 |  | 0.664 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 4:32 |  | 1 |  | 6.5819 | 8.0881 | 1.5062 | 0.87481
67 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5909 | 7574 |  | 7398 | 4.948 | 6.695 | 1.747 |  | 0.658 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 4:43 |  | 1 |  | 5.8753 | 8.0341 | 2.1588 | 0.8672 |  |  | 2.95 | 1.6763 |  |  |  | 0.11628 | 2.8766 |  |  |  | 8.0253 | 17.203 |  |  |  | 1.5793 | 1.5535 |  | 5.9632 | 8.0388 | 2.0755 | 0.8604
68 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6087 | 7583 |  | 7282 | 5.084 | 6.703 | 1.618 |  | 0.646 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 4:54 |  | 1 |  | 6.0512 | 8.0434 | 1.9922 | 0.85361
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1519 | 517 |  | 500 | 1.581 | 0.716 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 5:11 |  | 1 |  | 1.5367 | 0.70485 | -0.83188 | 0.059152
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 305 | 588 |  | 507 | 0.649 | 0.776 | 0.127 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 5:20 |  | 1 |  | 0.33695 | 0.77859 | 0.44164 | 0.059972 |  |  | 0.29287 |  |  |  |  | 5.7639 |  |  |  |  | 10.6379 |  |  |  |  | 4.1878 |  |  | 0.33745 | 0.75678 | 0.41933 | 0.058742
71 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 306 | 546 |  | 486 | 0.65 | 0.741 | 0.091 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 5:30 |  | 1 |  | 0.33794 | 0.73497 | 0.39703 | 0.057512
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3877 | 11212 |  | 4867 | 5.649 | 16.295 | 10.646 |  | 0.655 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 5:48 |  | 1 |  | 6.4452 | 19.6874 | 13.2422 | 0.95118
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5626 | 11341 |  | 4941 | 7.885 | 16.478 | 8.593 |  | 0.668 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 5:58 |  | 1 |  | 9.326 | 19.9107 | 10.5847 | 0.96563 | 93.5827 |  | 3.1465 |  |  | 94.2573 |  | 0.41817 |  |  | 94.9559 |  | 3.6676 |  |  | 97.2668 |  | 1.2205 |  |  | 9.4751 | 19.8691 | 10.3941 | 0.95977
74 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5807 | 11293 |  | 4881 | 8.116 | 16.41 | 8.293 |  | 0.657 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 6:09 |  | 1 |  | 9.6241 | 19.8276 | 10.2035 | 0.95392
75 | 3 |  | Flush | [20 empty] | 2023-10-25 | 6:15 |  | 1
76 | 21 | B50 24JUL23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2810 | 6069 |  | 2164 | 2.571 | 5.42 | 2.849 |  | 0.11 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 6:33 |  | 1 |  | 2.8126 | 6.471 | 3.6584 | 0.25408
77 | 21 | B50 24JUL23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3234 | 6076 |  | 2118 | 2.896 | 5.426 | 2.529 |  | 0.105 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 6:43 |  | 1 |  | 3.2316 | 6.4783 | 3.2467 | 0.24869 |  |  | 1.9982 |  |  |  |  | 1.0525 |  |  |  |  | 0.10219 |  |  |  |  | 0.47215 |  |  | 3.2642 | 6.5126 | 3.2483 | 0.2481
78 | 21 | B50 24JUL23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3300 | 6142 |  | 2108 | 2.946 | 5.482 | 2.536 |  | 0.104 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 6:54 |  | 1 |  | 3.2969 | 6.5469 | 3.25 | 0.24752
79 | 22 | B50 21MAR23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2624 | 4928 |  | 2487 | 2.428 | 4.453 | 2.026 |  | 0.144 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 7:13 |  | 1 |  | 2.6288 | 5.286 | 2.6572 | 0.29191
80 | 22 | B50 21MAR23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4283 | 4903 |  | 2536 | 3.701 | 4.433 | 0.732 |  | 0.149 | 1 | 0 | 0 |  | 0 |  | TICpe | 2023-10-25 | 7:28 |  | 1 |  | 4.2683 | 5.2601 | 0.99171 | 0.29765 |  |  | 64.2813 |  |  |  |  | 0.098674 |  |  |  |  | 102.4146 |  |  |  |  | 3.3614 |  |  | 3.2301 | 5.2626 | 2.0325 | 0.29273
81 | 22 | B50 21MAR23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2182 | 4908 |  | 2452 | 2.089 | 4.436 | 2.347 |  | 0.14 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 7:39 |  | 1 |  | 2.192 | 5.2652 | 3.0733 | 0.28781
82 | 23 | C50 19AUG22 20 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10270 | 14383 |  | 7250 | 8.294 | 12.464 | 4.17 |  | 0.642 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 7:58 |  | 1 |  | 10.1852 | 15.1058 | 4.9206 | 0.84986
83 | 23 | C50 19AUG22 20 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12366 | 14339 |  | 7157 | 9.902 | 12.426 | 2.524 |  | 0.633 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 8:09 |  | 1 |  | 12.2566 | 15.0601 | 2.8034 | 0.83897 |  |  | 3.5949 |  |  |  |  | 0.068986 |  |  |  |  | 17.8354 |  |  |  |  | 0.44581 |  |  | 12.481 | 15.0549 | 2.5739 | 0.84084
84 | 23 | C50 19AUG22 20 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 12820 | 14329 |  | 7189 | 10.25 | 12.418 | 2.168 |  | 0.636 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 8:20 |  | 1 |  | 12.7053 | 15.0497 | 2.3444 | 0.84271
85 | 24 | C50 30JUN22 21 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11408 | 13731 |  | 5019 | 9.167 | 11.911 | 2.744 |  | 0.409 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 8:39 |  | 1 |  | 11.3099 | 14.4286 | 3.1188 | 0.58852
86 | 24 | C50 30JUN22 21 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11299 | 13790 |  | 4991 | 9.084 | 11.961 | 2.878 |  | 0.406 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 8:50 |  | 1 |  | 11.2021 | 14.4899 | 3.2878 | 0.58524 |  |  | 0.83275 |  |  |  |  | 0.57506 |  |  |  |  | 0.29801 |  |  |  |  | 2.6077 |  |  | 11.1557 | 14.4483 | 3.2927 | 0.59297
87 | 24 | C50 30JUN22 21 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11205 | 13710 |  | 5123 | 9.011 | 11.893 | 2.882 |  | 0.42 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 9:02 |  | 1 |  | 11.1092 | 14.4068 | 3.2976 | 0.6007
88 | 25 | B50 31JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5062 | 6512 |  | 2260 | 4.299 | 5.795 | 1.497 |  | 0.12 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 9:20 |  | 1 |  | 5.0382 | 6.9311 | 1.8929 | 0.26532
89 | 25 | B50 31JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3488 | 6578 |  | 2336 | 3.091 | 5.851 | 2.76 |  | 0.128 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 9:31 |  | 1 |  | 3.4827 | 6.9997 | 3.517 | 0.27422 |  |  | 0.11357 |  |  |  |  | 0.88631 |  |  |  |  | 1.8666 |  |  |  |  | 0.59984 |  |  | 3.4807 | 7.0308 | 3.5501 | 0.2734
90 | 25 | B50 31JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3484 | 6638 |  | 2322 | 3.088 | 5.902 | 2.814 |  | 0.127 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 9:42 |  | 1 |  | 3.4787 | 7.062 | 3.5833 | 0.27258
91 | 26 | B50 24JUL23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3554 | 6034 |  | 2141 | 3.141 | 5.39 | 2.249 |  | 0.108 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 10:00 |  | 1 |  | 3.5479 | 6.4347 | 2.8868 | 0.25138
92 | 26 | B50 24JUL23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3740 | 6028 |  | 2143 | 3.284 | 5.386 | 2.101 |  | 0.108 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 10:10 |  | 1 |  | 3.7317 | 6.4285 | 2.6967 | 0.25162 |  |  | 2.0602 |  |  |  |  | 0.48351 |  |  |  |  | 3.8997 |  |  |  |  | 0.74213 |  |  | 3.6937 | 6.444 | 2.7504 | 0.25255
93 | 26 | B50 24JUL23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3663 | 6058 |  | 2159 | 3.225 | 5.411 | 2.186 |  | 0.11 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 10:21 |  | 1 |  | 3.6556 | 6.4596 | 2.804 | 0.25349
94 | 27 | B50 31JUL23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3714 | 6343 |  | 2575 | 3.264 | 5.652 | 2.388 |  | 0.153 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 10:39 |  | 1 |  | 3.706 | 6.7556 | 3.0496 | 0.30222
95 | 27 | B50 31JUL23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3827 | 6263 |  | 2587 | 3.351 | 5.585 | 2.234 |  | 0.155 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 10:50 |  | 1 |  | 3.8177 | 6.6725 | 2.8548 | 0.30363 |  |  | 2.9082 |  |  |  |  | 0.63614 |  |  |  |  | 2.4854 |  |  |  |  | 1.3032 |  |  | 3.874 | 6.6938 | 2.8198 | 0.30562
96 | 27 | B50 31JUL23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3941 | 6304 |  | 2621 | 3.438 | 5.62 | 2.181 |  | 0.158 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 11:00 |  | 1 |  | 3.9303 | 6.7151 | 2.7848 | 0.30761
97 | 28 | B50 29MAY23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4123 | 6817 |  | 6698 | 3.578 | 6.054 | 2.476 |  | 0.584 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 11:19 |  | 1 |  | 4.1102 | 7.2479 | 3.1377 | 0.7852
98 | 28 | B50 29MAY23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4434 | 6839 |  | 6657 | 3.817 | 6.073 | 2.256 |  | 0.58 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 11:29 |  | 1 |  | 4.4176 | 7.2707 | 2.8532 | 0.78039 |  |  | 0.022374 |  |  |  |  | 0.64487 |  |  |  |  | 1.6164 |  |  |  |  | 1.0304 |  |  | 4.4171 | 7.2474 | 2.8303 | 0.78444
99 | 28 | B50 29MAY23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4433 | 6794 |  | 6726 | 3.815 | 6.034 | 2.219 |  | 0.587 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 11:40 |  | 1 |  | 4.4166 | 7.224 | 2.8074 | 0.78848
100 | 29 | B50 14AUG23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4311 | 6428 |  | 2042 | 3.722 | 5.724 | 2.003 |  | 0.098 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 11:59 | this is sample 20 F5014aug23 1.6? | 2 | misidentified | 4.296 | 6.8439 | 2.5479 | 0.23978
101 | 29 | B50 14AUG23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4432 | 6475 |  | 2068 | 3.815 | 5.764 | 1.949 |  | 0.1 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 12:09 | this is sample 20 F5014aug23 1.6? | 2 | misidentified | 4.4156 | 6.8927 | 2.4771 | 0.24283 |  |  | 1.4002 |  |  |  |  | 0.43601 |  |  |  |  | 1.3061 |  |  |  |  | 1.9484 |  |  | 4.4467 | 6.9078 | 2.461 | 0.24049
102 | 29 | B50 14AUG23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4495 | 6504 |  | 2028 | 3.864 | 5.789 | 1.925 |  | 0.096 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 12:20 | this is sample 20 F5014aug23 1.6? | 2 | misidentified | 4.4779 | 6.9228 | 2.445 | 0.23814
103 | 30 | F50 14AUG23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6503 | 12944 |  | 4413 | 5.404 | 11.244 | 5.841 |  | 0.346 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 12:39 | this is the spike? | 2 | misidentified | 6.4623 | 13.6113 | 7.1489 | 0.51753
104 | 30 | F50 14AUG23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7414 | 12845 |  | 4399 | 6.103 | 11.16 | 5.057 |  | 0.344 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 12:50 | this is the spike? | 2 | misidentified | 7.3627 | 13.5084 | 6.1458 | 0.51589 |  |  | 0.17435 |  | 136.5222 |  |  | 0.80403 |  | 110.9034 |  |  | 1.5532 |  | 89.9827 |  |  | 0.4305 |  | 82.7499 | 7.3691 | 13.563 | 6.1939 | 0.517
105 | 30 | F50 14AUG23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7427 | 12950 |  | 4418 | 6.113 | 11.25 | 5.137 |  | 0.346 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 13:01 | this is the spike? | 2 | misidentified | 7.3755 | 13.6175 | 6.242 | 0.51811
106 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5211 | 9504 |  | 3933 | 4.413 | 8.33 | 3.917 |  | 0.295 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 13:20 | this is the dup? | 2 | misidentified | 5.1855 | 10.0385 | 4.8531 | 0.4613
107 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4771 | 9562 |  | 3795 | 4.075 | 8.379 | 4.304 |  | 0.281 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 13:30 | this is the dup? | 2 | misidentified | 4.7506 | 10.0988 | 5.3482 | 0.44513 |  |  | 2.464 | 5.385 | 43.351 |  |  | 0.31932 | 37.3731 | 52.3767 |  |  | 1.548 | 74.6117 | 62.9413 |  |  | 3.871 | 61.4718 | 61.5336 | 4.6928 | 10.0827 | 5.3899 | 0.45392
108 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4654 | 9531 |  | 3945 | 3.985 | 8.353 | 4.368 |  | 0.297 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 13:41 | this is the dup? | 2 | misidentified | 4.635 | 10.0666 | 5.4316 | 0.46271
109 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4604 | 9447 |  | 4017 | 3.947 | 8.282 | 4.335 |  | 0.304 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 14:00 | this is left over from previous run? TIC =3.5 TC = 7.7 | 2 | misidentified | 4.5856 | 9.9793 | 5.3938 | 0.47114
110 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4643 | 9442 |  | 3955 | 3.977 | 8.278 | 4.3 |  | 0.298 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 14:10 | this is left over from previous run? TIC =3.5 TC = 7.7 | 2 | misidentified | 4.6241 | 9.9742 | 5.35 | 0.46388 |  |  | 1.1264 | 45.2397 |  |  |  | 1.5601 | 29.7295 |  |  |  | 1.9334 | 13.6533 |  |  |  | 0.47865 | 10.593 |  | 4.6503 | 10.0526 | 5.4023 | 0.46499
111 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4696 | 9593 |  | 3974 | 4.017 | 8.406 | 4.388 |  | 0.3 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 14:21 | this is left over from previous run? TIC =3.5 TC = 7.7 | 2 | misidentified | 4.6765 | 10.131 | 5.4545 | 0.4661
112 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 905 | 783 |  | 614 | 1.109 | 0.942 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-24 | 13:33 |  | 1 |  | 0.92992 | 0.98111 | 0.051191 | 0.072506
113 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | [9 empty] | 1
114 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5
115 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3
116 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3
117 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3

## 25_26oct23
RUN NOTES:
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area |  | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | -0.0081185 | 0.11648 | 0.12459 | -0.0077108
TIC | TC | TNb |  | 0.0006 | 1226 | 0.0012 | 2094 | 6e-05 | 1144.4 |  | 0.2 | 3.0172 | 5.7174 | 2.7002 | 0.31552
75 | 250 | 379 |  | 0.0018 | 3864 | 0.0036 | 7152.5 | 0.00018 | 3217.7 |  | 0.6 | 3.1843 | 6.2752 | 3.0909 | 0.30736
|  |  |  | 0.002997 | 5742.55 | 0.005994 | 10914 | 0.0002997 | 4997.586 |  | 0.333 | 8.5328 | 17.1608 | 8.628 | 0.86653
|  |  |  | 0.004203 | 8521.95 | 0.008406 | 16258 | 0.0004203 | 7595 |  | 0.467 | 9.0335 | 18.1673 | 9.1338 | 0.94332
|  |  |  | 0.0054 | 10966.5 | 0.0108 | 20793 | 0.00054 | 9077.2 |  | 0.6 | 9.0499 | 18.0574 | 9.0075 | 0.87876
|  |  | Slope |  | 4.9551e-07 |  | 5.1826e-07 |  | 5.851e-08
|  |  | Intercept |  | -4.0593e-06 |  | 5.8238e-05 |  | -3.8554e-06
|  |  | RSQ |  | 0.99828 |  | 0.99864 |  | 0.99662
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5859 | 13043 |  | 5065 | 8.183 | 18.881 | 10.698 |  | 0.69 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 15:03 |  | 1 |  | 9.6638 | 22.7265 | 13.0627 | 0.975
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7383 | 13085 |  | 5310 | 10.132 | 18.941 | 8.809 |  | 0.732 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 15:14 |  | 1 |  | 12.181 | 22.799 | 10.618 | 1.0228 |  |  | 2.1197 |  |  |  |  | 0.31774 |  |  |  |  | 1.7903 |  |  |  |  | 1.8327 |  |  | 12.3115 | 22.8353 | 10.5238 | 1.0322
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7541 | 13127 |  | 5407 | 10.334 | 18.999 | 8.665 |  | 0.749 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 15:25 |  | 1 |  | 12.4419 | 22.8716 | 10.4296 | 1.0417
4 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3285 | 1082 |  | 747 | 2.935 | 1.196 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 15:43 |  | 1 |  | 3.2474 | 1.238 | -2.0094 | 0.079704
5 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 590 | 1059 |  | 683 | 0.868 | 1.175 | 0.308 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 15:52 |  | 1 |  | 0.57658 | 1.2142 | 0.63757 | 0.072214 |  |  | 4.5355 |  |  |  |  | 4.1793 |  |  |  |  | 3.8561 |  |  |  |  | 5.3235 |  |  | 0.58996 | 1.2401 | 0.65011 | 0.070342
6 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 617 | 1109 |  | 651 | 0.888 | 1.218 | 0.329 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 16:02 |  | 1 |  | 0.60334 | 1.266 | 0.66264 | 0.06847
7 | 3 |  | Flush | [20 empty] | 2023-10-25 | 16:08 |  | 1
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 121 | 248 |  | 453 | 0.508 | 0.488 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 16:23 |  | 1 |  | 0.11179 | 0.37353 | 0.26174 | 0.0453
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 62 | 264 |  | 421 | 0.462 | 0.503 | 0.04 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 16:31 |  | 1 |  | 0.053325 | 0.39012 | 0.33679 | 0.041555 |  |  | 38.9175 |  |  |  |  | 7.7269 |  |  |  |  | 17.7083 |  |  |  |  | 26.8278 |  |  | 0.066208 | 0.37561 | 0.3094 | 0.03664
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 88 | 236 |  | 337 | 0.483 | 0.478 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 16:40 |  | 1 |  | 0.079091 | 0.3611 | 0.282 | 0.031725
11 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 633 | 2126 |  | 1272 | 2.251 | 5.2 | 2.949 |  | 0.043 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 16:56 |  | 1 |  | 1.548 | 5.8003 | 4.2523 | 0.35285
12 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1185 | 2189 |  | 1284 | 3.31 | 5.333 | 2.023 |  | 0.046 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 17:05 |  | 1 |  | 2.9156 | 5.9636 | 3.048 | 0.35636 |  | 3.0502 | 5.8566 |  |  |  | 0.39121 | 0.43358 |  |  |  | 3.4007 | 5.0452 |  |  |  | 19.9565 | 1.951 |  |  | 3.0036 | 5.9765 | 2.973 | 0.35987
13 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1256 | 2199 |  | 1308 | 3.446 | 5.353 | 1.906 |  | 0.052 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 17:15 |  | 1 |  | 3.0915 | 5.9895 | 2.898 | 0.36338
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3829 | 7456 |  | 3610 | 2.794 | 5.496 | 2.702 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 17:34 |  | 1 |  | 3.1554 | 6.5373 | 3.3819 | 0.34561
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3943 | 7460 |  | 3664 | 2.867 | 5.499 | 2.632 |  | 0.223 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 17:45 |  | 1 |  | 3.2496 | 6.5408 | 3.2912 | 0.35088 |  | 8.6215 | 0.55755 |  |  |  | 8.9052 | 0.19828 |  |  |  | 9.189 | 0.9502 |  |  |  | 17.2353 | 0.47136 |  |  | 3.2586 | 6.5343 | 3.2757 | 0.35171
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3965 | 7445 |  | 3681 | 2.881 | 5.488 | 2.607 |  | 0.224 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 17:56 |  | 1 |  | 3.2677 | 6.5278 | 3.2601 | 0.35253
17 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 4696 | 10978 |  | 5232 | 6.032 | 14.383 | 8.351 |  | 0.647 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 18:15 |  | 1 |  | 6.9755 | 17.2604 | 10.2849 | 0.90772
18 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 5735 | 11036 |  | 5277 | 7.229 | 14.456 | 7.227 |  | 0.654 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 18:25 |  | 1 |  | 8.5216 | 17.3507 | 8.8291 | 0.91563 |  | 4.365 | 1.9881 |  |  |  | 3.2225 | 0.79515 |  |  |  | 2.08 | 0.37 |  |  |  | 1.209 | 1.0416 |  |  | 8.6072 | 17.42 | 8.8128 | 0.91088
19 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 5850 | 11125 |  | 5223 | 7.362 | 14.569 | 7.207 |  | 0.646 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 18:36 |  | 1 |  | 8.6927 | 17.4892 | 8.7965 | 0.90614
20 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8338 | 16395 |  | 7750 | 7.293 | 15.169 | 7.876 |  | 0.744 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 18:56 |  | 1 |  | 8.8383 | 18.3194 | 9.481 | 0.96274
21 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8521 | 16479 |  | 7783 | 7.443 | 15.245 | 7.802 |  | 0.747 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 19:07 |  | 1 |  | 9.0325 | 18.4126 | 9.3801 | 0.96687 |  | 1.1982 | 1.6543 |  |  |  | 2.3693 | 0.15057 |  |  |  | 3.5404 | 1.3191 |  |  |  | 8.0638 | 1.1723 |  |  | 9.1078 | 18.4265 | 9.3186 | 0.97257
22 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8663 | 16504 |  | 7874 | 7.56 | 15.269 | 7.708 |  | 0.758 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 19:19 |  | 1 |  | 9.1832 | 18.4403 | 9.2572 | 0.97827
23 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 10969 | 20834 |  | 9321 | 7.359 | 14.941 | 7.582 |  | 0.716 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 19:46 |  | 1 |  | 9.052 | 18.0929 | 9.0409 | 0.90253
24 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11026 | 21011 |  | 9478 | 7.395 | 15.066 | 7.671 |  | 0.729 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 19:58 |  | 1 |  | 9.0991 | 18.2458 | 9.1467 | 0.91784 |  | 1.3805 | 0.55212 |  |  |  | 1.7588 | 0.77339 |  |  |  | 2.1371 | 0.99302 |  |  |  | 2.5675 | 1.1409 |  |  | 9.1242 | 18.3166 | 9.1923 | 0.92311
25 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11087 | 21175 |  | 9586 | 7.434 | 15.181 | 7.747 |  | 0.739 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 20:11 |  | 1 |  | 9.1494 | 18.3874 | 9.238 | 0.92837
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5982 | 12029 |  | 5118 | 8.341 | 17.449 | 9.108 |  | 0.699 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 20:29 |  | 1 |  | 9.8669 | 20.9747 | 11.1078 | 0.98533
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6338 | 12143 |  | 5190 | 8.795 | 17.609 | 8.814 |  | 0.711 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 20:39 |  | 1 |  | 10.4549 | 21.1717 | 10.7167 | 0.99938 | 100 |  | 0.62994 |  |  | 100 |  | 0.1875 |  |  | 100 |  | 0.24604 |  |  | 100 |  | 3.9199 |  |  | 10.488 | 21.1915 | 10.7036 | 0.98017
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6378 | 12166 |  | 4993 | 8.847 | 17.643 | 8.796 |  | 0.677 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 20:50 |  | 1 |  | 10.521 | 21.2114 | 10.6904 | 0.96095
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5194 | 6896 |  | 2933 | 4.4 | 6.121 | 1.721 |  | 0.191 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 21:09 |  | 1 |  | 5.1392 | 7.2644 | 2.1251 | 0.33551
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3650 | 6883 |  | 2981 | 3.215 | 6.11 | 2.895 |  | 0.196 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 21:20 |  | 1 |  | 3.6091 | 7.2509 | 3.6418 | 0.34113 |  |  | 0.87484 |  |  |  |  | 0.11443 |  |  |  |  | 1.1046 |  |  |  |  | 0.61939 |  |  | 3.625 | 7.2467 | 3.6218 | 0.34007
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3682 | 6875 |  | 2963 | 3.24 | 6.103 | 2.863 |  | 0.194 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 21:31 |  | 1 |  | 3.6408 | 7.2426 | 3.6018 | 0.33902
32 | 3 |  | Flush | [20 empty] | 2023-10-25 | 21:36 |  | 1
33 | 9 | B50 19SEP22 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5010 | 8740 |  | 17360 | 4.258 | 7.683 | 3.425 |  | 1.7 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 21:55 |  | 1 |  | 4.9569 | 9.1757 | 4.2188 | 2.0238
34 | 9 | B50 19SEP22 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6324 | 8884 |  | 17677 | 5.267 | 7.805 | 2.538 |  | 1.733 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 22:06 |  | 1 |  | 6.2591 | 9.325 | 3.0659 | 2.0609 |  |  | 1.6021 |  |  |  |  | 0.34518 |  |  |  |  | 4.4416 |  |  |  |  | 0.41931 |  |  | 6.3096 | 9.3089 | 2.9993 | 2.0652
35 | 9 | B50 19SEP22 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6426 | 8853 |  | 17751 | 5.345 | 7.779 | 2.434 |  | 1.741 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 22:17 |  | 1 |  | 6.3602 | 9.2928 | 2.9327 | 2.0695
36 | 10 | F50 22AUG23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6339 | 7810 |  | 8541 | 5.278 | 6.895 | 1.617 |  | 0.777 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 22:36 |  | 1 |  | 6.274 | 8.2117 | 1.9378 | 0.99176
37 | 10 | F50 22AUG23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6579 | 7978 |  | 8814 | 5.462 | 7.038 | 1.576 |  | 0.806 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 22:46 |  | 1 |  | 6.5118 | 8.3859 | 1.8741 | 1.0237 |  |  | 1.302 |  |  |  |  | 0.99374 |  |  |  |  | 0.070127 |  |  |  |  | 0.58469 |  |  | 6.4697 | 8.3444 | 1.8747 | 1.0207
38 | 10 | F50 22AUG23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6494 | 7898 |  | 8763 | 5.397 | 6.969 | 1.573 |  | 0.801 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 22:57 |  | 1 |  | 6.4276 | 8.3029 | 1.8754 | 1.0177
39 | 11 | F50 23MAY22 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4873 | 8987 |  | 3973 | 4.153 | 7.892 | 3.739 |  | 0.299 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 23:16 |  | 1 |  | 4.8211 | 9.4317 | 4.6106 | 0.45721
40 | 11 | F50 23MAY22 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4623 | 8987 |  | 4003 | 3.961 | 7.893 | 3.931 |  | 0.303 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 23:27 |  | 1 |  | 4.5734 | 9.4317 | 4.8584 | 0.46072 |  |  | 1.2648 |  |  |  |  | 0.08788 |  |  |  |  | 1.3447 |  |  |  |  | 3.2524 |  |  | 4.5446 | 9.4359 | 4.8912 | 0.45335
41 | 11 | F50 23MAY22 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4565 | 8995 |  | 3877 | 3.917 | 7.899 | 3.981 |  | 0.289 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 23:38 |  | 1 |  | 4.5159 | 9.44 | 4.9241 | 0.44598
42 | 12 | F100 29MAY23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3945 | 7856 |  | 1547 | 3.441 | 6.934 | 3.492 |  | 0.046 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-25 | 23:56 |  | 1 |  | 3.9015 | 8.2594 | 4.358 | 0.17332
43 | 12 | F100 29MAY23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3738 | 7959 |  | 1613 | 3.282 | 7.021 | 3.739 |  | 0.053 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 0:07 |  | 1 |  | 3.6963 | 8.3662 | 4.6699 | 0.18104 |  |  | 4.5076 |  |  |  |  | 0.29691 |  |  |  |  | 3.1668 |  |  |  |  | 3.5524 |  |  | 3.7815 | 8.3786 | 4.5971 | 0.17788
44 | 12 | F100 29MAY23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3910 | 7983 |  | 1559 | 3.414 | 7.042 | 3.627 |  | 0.047 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 0:18 |  | 1 |  | 3.8668 | 8.3911 | 4.5243 | 0.17472
45 | 13 | F50 31JUL23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4219 | 9829 |  | 4213 | 3.651 | 8.606 | 4.955 |  | 0.325 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 0:36 |  | 1 |  | 4.173 | 10.3045 | 6.1315 | 0.4853
46 | 13 | F50 31JUL23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4309 | 9901 |  | 4469 | 3.721 | 8.666 | 4.945 |  | 0.351 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 0:47 |  | 1 |  | 4.2622 | 10.3791 | 6.1169 | 0.51525 |  |  | 2.6839 |  |  |  |  | 0.67135 |  |  |  |  | 3.0775 |  |  |  |  | 5.3397 |  |  | 4.3202 | 10.3444 | 6.0242 | 0.50186
47 | 13 | F50 31JUL23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4426 | 9834 |  | 4240 | 3.81 | 8.61 | 4.8 |  | 0.327 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 0:58 |  | 1 |  | 4.3781 | 10.3097 | 5.9315 | 0.48846
48 | 14 | F50 14AUG23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4538 | 9799 |  | 3881 | 3.897 | 8.58 | 4.684 |  | 0.29 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 1:16 |  | 1 |  | 4.4891 | 10.2734 | 5.7843 | 0.44645
49 | 14 | F50 14AUG23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4618 | 9871 |  | 4051 | 3.958 | 8.641 | 4.683 |  | 0.308 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 1:27 |  | 1 |  | 4.5684 | 10.348 | 5.7796 | 0.46634 |  |  | 1.6992 |  |  |  |  | 0.75408 |  |  |  |  | 2.7366 |  |  |  |  | 0.075252 |  |  | 4.6076 | 10.3091 | 5.7016 | 0.46651
50 | 14 | F50 14AUG23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4697 | 9796 |  | 4054 | 4.018 | 8.578 | 4.559 |  | 0.308 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 1:38 |  | 1 |  | 4.6467 | 10.2703 | 5.6236 | 0.46669
51 | 15 | B50 22AUG23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4708 | 6494 |  | 1568 | 4.027 | 5.78 | 1.754 |  | 0.048 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 1:56 |  | 1 |  | 4.6576 | 6.8477 | 2.1901 | 0.17578
52 | 15 | B50 22AUG23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4816 | 6561 |  | 1529 | 4.11 | 5.837 | 1.727 |  | 0.044 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 2:06 |  | 1 |  | 4.7646 | 6.9171 | 2.1525 | 0.17121 |  |  | 2.0806 |  |  |  |  | 1.206 |  |  |  |  | 0.70317 |  |  |  |  | 0.96147 |  |  | 4.7156 | 6.8757 | 2.1601 | 0.17039
53 | 15 | B50 22AUG23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4717 | 6481 |  | 1515 | 4.034 | 5.769 | 1.735 |  | 0.042 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 2:17 |  | 1 |  | 4.6665 | 6.8342 | 2.1677 | 0.16958
54 | 16 | F50 29MAY23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4443 | 6495 |  | 2607 | 3.823 | 5.781 | 1.958 |  | 0.157 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 2:36 |  | 1 |  | 4.395 | 6.8487 | 2.4537 | 0.29736
55 | 16 | F50 29MAY23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4547 | 6552 |  | 2595 | 3.903 | 5.829 | 1.926 |  | 0.155 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 2:46 |  | 1 |  | 4.498 | 6.9078 | 2.4097 | 0.29596 |  |  | 1.0084 |  |  |  |  | 0 |  |  |  |  | 1.9098 |  |  |  |  | 0.59486 |  |  | 4.5208 | 6.9078 | 2.3869 | 0.29508
56 | 16 | F50 29MAY23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4593 | 6552 |  | 2580 | 3.939 | 5.829 | 1.891 |  | 0.154 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 2:57 |  | 1 |  | 4.5436 | 6.9078 | 2.3642 | 0.2942
57 | 17 | B50 29MAY23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3773 | 6941 |  | 2554 | 3.31 | 6.159 | 2.85 |  | 0.151 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 3:16 |  | 1 |  | 3.731 | 7.311 | 3.58 | 0.29116
58 | 17 | B50 29MAY23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3685 | 7053 |  | 2634 | 3.242 | 6.253 | 3.011 |  | 0.159 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 3:26 |  | 1 |  | 3.6438 | 7.4271 | 3.7833 | 0.30052 |  |  | 0.90156 |  |  |  |  | 0.72836 |  |  |  |  | 0.56182 |  |  |  |  | 0.23391 |  |  | 3.6274 | 7.4001 | 3.7727 | 0.30017
59 | 17 | B50 29MAY23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3652 | 7001 |  | 2628 | 3.217 | 6.21 | 2.993 |  | 0.159 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 3:37 |  | 1 |  | 3.6111 | 7.3732 | 3.7621 | 0.29982
60 | 18 | B50 29MAY23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2950 | 5841 |  | 2074 | 2.678 | 5.227 | 2.549 |  | 0.101 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 3:55 |  | 1 |  | 2.9154 | 6.1708 | 3.2554 | 0.23499
61 | 18 | B50 29MAY23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2952 | 5832 |  | 2019 | 2.679 | 5.219 | 2.54 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 4:05 |  | 1 |  | 2.9174 | 6.1615 | 3.2441 | 0.22855 |  |  | 0.40847 |  |  |  |  | 0.084078 |  |  |  |  | 0.52495 |  |  |  |  | 2.1738 |  |  | 2.9114 | 6.1641 | 3.2527 | 0.2261
62 | 18 | B50 29MAY23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2940 | 5837 |  | 1977 | 2.67 | 5.224 | 2.554 |  | 0.091 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 4:16 |  | 1 |  | 2.9055 | 6.1667 | 3.2612 | 0.22364
63 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6206 | 14949 |  | 6784 | 5.176 | 12.943 | 7.767 |  | 0.593 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 4:35 |  | 1 |  | 6.1422 | 15.6115 | 9.4693 | 0.78616
64 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7299 | 14989 |  | 6808 | 6.014 | 12.977 | 6.963 |  | 0.596 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 4:46 |  | 1 |  | 7.2253 | 15.6529 | 8.4276 | 0.78896 |  |  | 2.2244 |  | 101.7399 |  |  | 0.16541 |  | 91.0419 |  |  | 1.6343 |  | 80.3438 |  |  | 1.0587 |  | 96.673 | 7.3066 | 15.6659 | 8.3593 | 0.78481
65 | 19 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7463 | 15014 |  | 6737 | 6.14 | 12.998 | 6.857 |  | 0.589 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 4:57 |  | 1 |  | 7.3879 | 15.6789 | 8.291 | 0.78066
66 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3972 | 6153 |  | 2149 | 3.462 | 5.491 | 2.029 |  | 0.109 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 5:16 |  | 1 |  | 3.9282 | 6.4942 | 2.566 | 0.24377
67 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3113 | 6230 |  | 2139 | 2.803 | 5.557 | 2.753 |  | 0.108 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 5:26 |  | 1 |  | 3.0769 | 6.574 | 3.4971 | 0.2426 |  |  | 3.3068 | 3.8884 |  |  |  | 0.3633 | 6.2552 |  |  |  | 2.1569 | 8.3258 |  |  |  | 1.7517 | 6.1697 |  | 3.0269 | 6.5621 | 3.5352 | 0.24049
68 | 20 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3012 | 6207 |  | 2103 | 2.726 | 5.537 | 2.811 |  | 0.104 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 5:37 |  | 1 |  | 2.9768 | 6.5502 | 3.5734 | 0.23838
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 816 | 559 |  | 502 | 1.041 | 0.752 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 5:54 |  | 1 |  | 0.80055 | 0.69589 | -0.10466 | 0.051034
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 204 | 533 |  | 485 | 0.571 | 0.73 | 0.159 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 6:03 |  | 1 |  | 0.19405 | 0.66894 | 0.47489 | 0.049044 |  |  | 28.8452 |  |  |  |  | 4.5429 |  |  |  |  | 7.4959 |  |  |  |  | 10.8147 |  |  | 0.22675 | 0.68449 | 0.45774 | 0.046528
71 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 270 | 563 |  | 442 | 0.622 | 0.755 | 0.134 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 6:12 |  | 1 |  | 0.25946 | 0.70004 | 0.44058 | 0.044012
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 3957 | 11345 |  | 4630 | 5.751 | 16.483 | 10.732 |  | 0.614 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 6:31 |  | 1 |  | 6.5222 | 19.7931 | 13.2708 | 0.89016
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5709 | 11341 |  | 4563 | 7.992 | 16.478 | 8.486 |  | 0.602 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 6:41 |  | 1 |  | 9.416 | 19.7862 | 10.3702 | 0.87709 | 91.8056 |  | 4.392 |  |  | 93.7636 |  | 0.96447 |  |  | 95.9113 |  | 2.2534 |  |  | 91.2992 |  | 3.7315 |  |  | 9.6274 | 19.8821 | 10.2546 | 0.89377
74 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5965 | 11452 |  | 4734 | 8.318 | 16.634 | 8.316 |  | 0.632 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 6:52 |  | 1 |  | 9.8389 | 19.9779 | 10.1391 | 0.91044
75 | 3 |  | Flush | [20 empty] | 2023-10-26 | 6:58 |  | 1
76 | 21 | B50 14AUG23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5089 | 8177 |  | 8537 | 4.319 | 7.206 | 2.886 |  | 0.777 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 7:17 |  | 1 |  | 5.0352 | 8.5921 | 3.557 | 0.99129
77 | 21 | B50 14AUG23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6283 | 8203 |  | 8637 | 5.235 | 7.228 | 1.993 |  | 0.787 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 7:27 |  | 1 |  | 6.2185 | 8.6191 | 2.4006 | 1.003 |  |  | 0.75184 |  |  |  |  | 1.1003 |  |  |  |  | 5.7428 |  |  |  |  | 1.5742 |  |  | 6.1952 | 8.6668 | 2.4716 | 1.011
78 | 21 | B50 14AUG23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6236 | 8295 |  | 8773 | 5.199 | 7.306 | 2.107 |  | 0.802 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 7:39 |  | 1 |  | 6.1719 | 8.7144 | 2.5426 | 1.0189
79 | 22 | F50 12SEP23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5174 | 9964 |  | 4453 | 4.384 | 8.72 | 4.336 |  | 0.35 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 7:57 |  | 1 |  | 5.1194 | 10.4444 | 5.325 | 0.51338
80 | 22 | F50 12SEP23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4944 | 10075 |  | 4712 | 4.208 | 8.814 | 4.606 |  | 0.377 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 8:08 |  | 1 |  | 4.8915 | 10.5595 | 5.668 | 0.54369 |  |  | 3.2291 |  |  |  |  | 0.71915 |  |  |  |  | 4.2562 |  |  |  |  | 1.6493 |  |  | 4.9718 | 10.5216 | 5.5499 | 0.53924
81 | 22 | F50 12SEP23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5106 | 10002 |  | 4636 | 4.332 | 8.752 | 4.42 |  | 0.369 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 8:19 |  | 1 |  | 5.052 | 10.4838 | 5.4318 | 0.5348
82 | 23 | F50 12SEP23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6374 | 8431 |  | 11445 | 5.305 | 7.421 | 2.117 |  | 1.081 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 8:38 |  | 1 |  | 6.3086 | 8.8554 | 2.5468 | 1.3316
83 | 23 | F50 12SEP23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6952 | 8508 |  | 11622 | 5.748 | 7.486 | 1.738 |  | 1.1 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 8:48 |  | 1 |  | 6.8815 | 8.9352 | 2.0538 | 1.3523 |  |  | 4.944 |  |  |  |  | 1.3783 |  |  |  |  | 25.9129 |  |  |  |  | 2.3433 |  |  | 7.0559 | 8.8741 | 1.8182 | 1.3683
84 | 23 | F50 12SEP23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7304 | 8390 |  | 11896 | 6.019 | 7.387 | 1.368 |  | 1.128 | 1 | 0 | 0 |  | 0 |  | TICpe | 2023-10-26 | 9:17 |  | 1 |  | 7.2303 | 8.8129 | 1.5826 | 1.3844
85 | 24 | F100 5JUL22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7137 | 6292 |  | 5636 | 5.89 | 5.609 | 0 |  | 0.473 | 1 | 0 | 0 |  | 0 |  | TICpe | 2023-10-26 | 9:41 |  | 1 |  | 7.0648 | 6.6383 | -0.4265 | 0.65182
86 | 24 | F100 5JUL22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4790 | 6316 |  | 5698 | 4.089 | 5.63 | 1.54 |  | 0.48 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 9:52 |  | 1 |  | 4.7389 | 6.6632 | 1.9243 | 0.65907 |  |  | 1.8574 |  |  |  |  | 0.4209 |  |  |  |  | 3.031 |  |  |  |  | 3.2145 |  |  | 4.6953 | 6.6492 | 1.9539 | 0.66984
87 | 24 | F100 5JUL22 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4702 | 6289 |  | 5882 | 4.022 | 5.606 | 1.584 |  | 0.499 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 10:02 |  | 1 |  | 4.6517 | 6.6352 | 1.9835 | 0.6806
88 | 25 | F50 12SEP23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5176 | 8423 |  | 3318 | 4.386 | 7.414 | 3.028 |  | 0.231 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 10:21 |  | 1 |  | 5.1214 | 8.8471 | 3.7257 | 0.38056
89 | 25 | F50 12SEP23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5527 | 8470 |  | 3452 | 4.655 | 7.454 | 2.799 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 10:31 |  | 1 |  | 5.4692 | 8.8958 | 3.4266 | 0.39624 |  |  | 0.25336 |  |  |  |  | 0.023301 |  |  |  |  | 0.345 |  |  |  |  | 1.0688 |  |  | 5.4762 | 8.8969 | 3.4207 | 0.39414
90 | 25 | F50 12SEP23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5541 | 8472 |  | 3416 | 4.666 | 7.456 | 2.79 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 10:43 |  | 1 |  | 5.4831 | 8.8979 | 3.4148 | 0.39203
91 | 26 | F100 31JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5135 | 6668 |  | 1770 | 4.355 | 5.928 | 1.573 |  | 0.069 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 11:01 |  | 1 |  | 5.0808 | 7.028 | 1.9473 | 0.19942
92 | 26 | F100 31JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5268 | 6630 |  | 1839 | 4.457 | 5.896 | 1.439 |  | 0.076 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 11:12 |  | 1 |  | 5.2126 | 6.9886 | 1.7761 | 0.20749 |  |  | 1.5712 |  |  |  |  | 0.84184 |  |  |  |  | 7.6017 |  |  |  |  | 3.0345 |  |  | 5.1719 | 7.0182 | 1.8462 | 0.20439
93 | 26 | F100 31JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5186 | 6687 |  | 1786 | 4.394 | 5.943 | 1.55 |  | 0.071 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 11:23 |  | 1 |  | 5.1313 | 7.0477 | 1.9164 | 0.20129
94 | 27 | B50 31JUL23 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3545 | 6480 |  | 2288 | 3.135 | 5.768 | 2.634 |  | 0.123 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 11:41 |  | 1 |  | 3.505 | 6.8332 | 3.3281 | 0.26003
95 | 27 | B50 31JUL23 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3132 | 6545 |  | 2347 | 2.818 | 5.823 | 3.006 |  | 0.129 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 11:52 |  | 1 |  | 3.0958 | 6.9005 | 3.8048 | 0.26694 |  |  | 3.1198 |  |  |  |  | 0.51202 |  |  |  |  | 3.5674 |  |  |  |  | 0.21943 |  |  | 3.1448 | 6.8829 | 3.7381 | 0.26664
96 | 27 | B50 31JUL23 3.0 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3231 | 6511 |  | 2342 | 2.894 | 5.795 | 2.901 |  | 0.129 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 12:03 |  | 1 |  | 3.1939 | 6.8653 | 3.6714 | 0.26635
97 | 28 | F50 31JUL23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4877 | 11262 |  | 6221 | 4.156 | 9.819 | 5.663 |  | 0.535 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 12:22 |  | 1 |  | 4.8251 | 11.7898 | 6.9647 | 0.72027
98 | 28 | F50 31JUL23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5602 | 11316 |  | 6308 | 4.712 | 9.865 | 5.153 |  | 0.544 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 12:32 |  | 1 |  | 5.5436 | 11.8458 | 6.3022 | 0.73045 |  |  | 0.53487 |  |  |  |  | 0.11382 |  |  |  |  | 0.68792 |  |  |  |  | 0.73966 |  |  | 5.5584 | 11.8391 | 6.2806 | 0.72776
99 | 28 | F50 31JUL23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5632 | 11303 |  | 6262 | 4.736 | 9.854 | 5.118 |  | 0.539 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 12:44 |  | 1 |  | 5.5733 | 11.8323 | 6.259 | 0.72507
100 | 29 | F5031JUL23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5409 | 7804 |  | 5038 | 4.565 | 6.89 | 2.325 |  | 0.411 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 13:03 |  | 1 |  | 5.3523 | 8.2055 | 2.8532 | 0.58184
101 | 29 | F5031JUL23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5452 | 7863 |  | 5088 | 4.598 | 6.94 | 2.343 |  | 0.416 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 13:13 |  | 1 |  | 5.3949 | 8.2667 | 2.8717 | 0.58769 |  |  | 0.73209 |  |  |  |  | 1.814 |  |  |  |  | 3.8151 |  |  |  |  | 1.3842 |  |  | 5.4147 | 8.3423 | 2.9276 | 0.59179
102 | 29 | F5031JUL23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5492 | 8009 |  | 5158 | 4.628 | 7.063 | 2.435 |  | 0.423 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 13:24 |  | 1 |  | 5.4346 | 8.418 | 2.9834 | 0.59588
103 | 30 | F200 24JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5692 | 7789 |  | 2041 | 4.782 | 6.877 | 2.096 |  | 0.097 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 13:43 |  | 1 |  | 5.6328 | 8.19 | 2.5572 | 0.23113
104 | 30 | F200 24JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5741 | 7813 |  | 2097 | 4.819 | 6.898 | 2.079 |  | 0.103 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 13:53 |  | 1 |  | 5.6813 | 8.2148 | 2.5335 | 0.23768 |  |  | 3.294 |  |  |  |  | 0.76673 |  |  |  |  | 5.1436 |  |  |  |  | 2.0949 |  |  | 5.7765 | 8.2465 | 2.47 | 0.2402
105 | 30 | F200 24JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5933 | 7874 |  | 2140 | 4.966 | 6.95 | 1.983 |  | 0.108 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 14:04 |  | 1 |  | 5.8716 | 8.2781 | 2.4065 | 0.24271
106 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8717 | 14852 |  | 5473 | 7.102 | 12.861 | 5.759 |  | 0.456 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 14:24 |  | 1 |  | 8.6306 | 15.5109 | 6.8803 | 0.63274
107 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9926 | 14973 |  | 5576 | 8.03 | 12.963 | 4.934 |  | 0.467 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 14:35 |  | 1 |  | 9.8287 | 15.6364 | 5.8076 | 0.6448 |  |  | 1.8481 |  | 151.1171 |  |  | 0.22513 |  | 114.9665 |  |  | 2.583 |  | 78.816 |  |  | 0.090701 |  | 85.5854 | 9.9204 | 15.654 | 5.7336 | 0.64509
108 | 31 | SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 10111 | 15007 |  | 5581 | 8.172 | 12.992 | 4.82 |  | 0.468 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 14:48 |  | 1 |  | 10.0121 | 15.6716 | 5.6595 | 0.64538
109 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6879 | 8173 |  | 1976 | 5.692 | 7.203 | 1.51 |  | 0.091 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 15:07 |  | 1 |  | 6.8091 | 8.588 | 1.7789 | 0.22352
110 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6070 | 8170 |  | 2017 | 5.072 | 7.2 | 2.128 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 15:17 |  | 1 |  | 6.0074 | 8.5849 | 2.5775 | 0.22832 |  |  | 2.5247 | 5.1888 |  |  |  | 1.0569 | 4.551 |  |  |  | 2.4504 | 3.0429 |  |  |  | 5.7471 | 7.9003 |  | 6.0842 | 8.6305 | 2.5463 | 0.22194
111 | 32 | DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6225 | 8258 |  | 1908 | 5.19 | 7.274 | 2.084 |  | 0.083 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 15:28 |  | 1 |  | 6.161 | 8.6761 | 2.5151 | 0.21556
112 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5638 | 10310 |  | 4844 | 7.901 | 15.021 | 7.121 |  | 0.651 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 15:47 |  | 1 |  | 9.2988 | 18.0051 | 8.7063 | 0.93189
113 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6585 | 10305 |  | 4760 | 9.111 | 15.015 | 5.904 |  | 0.636 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 15:57 |  | 1 |  | 10.8629 | 17.9964 | 7.1335 | 0.91551 |  |  | 1.194 |  |  |  |  | 0.98387 |  |  |  |  | 0.663 |  |  |  |  | 0.74285 |  |  | 10.9282 | 18.0854 | 7.1573 | 0.91892
114 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6664 | 10408 |  | 4795 | 9.213 | 15.16 | 5.947 |  | 0.642 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 16:08 |  | 1 |  | 10.9934 | 18.1744 | 7.181 | 0.92234
115 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2975 | 880 |  | 543 | 2.697 | 1.024 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 16:26 |  | 1 |  | 2.9402 | 1.0286 | -1.9115 | 0.055831
116 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 546 | 923 |  | 601 | 0.834 | 1.06 | 0.227 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 16:35 |  | 1 |  | 0.53298 | 1.0732 | 0.54021 | 0.062619 |  |  | 6.1282 |  |  |  |  | 5.559 |  |  |  |  | 18.5581 |  |  |  |  | 10.0077 |  |  | 0.54983 | 1.0442 | 0.49434 | 0.059635
117 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 580 | 867 |  | 550 | 0.859 | 1.013 | 0.153 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 16:45 |  | 1 |  | 0.56667 | 1.0151 | 0.44847 | 0.056651
118 | 3 |  | Flush | [20 empty] | 2023-10-26 | 16:51
119 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 137 | 184 |  | 345 | 0.52 | 0.434 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:06 |  | 1 |  | 0.12765 | 0.3072 | 0.17954 | 0.032661
120 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 58 | 211 |  | 333 | 0.459 | 0.457 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:14 |  | 1 |  | 0.049361 | 0.33518 | 0.28582 | 0.031257 |  |  | 32.7042 |  |  |  |  | 8.0419 |  |  |  |  | 4.3026 |  |  |  |  | 4.3939 |  |  | 0.042424 | 0.32223 | 0.2798 | 0.031959
121 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 44 | 186 |  | 345 | 0.449 | 0.436 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:23 |  | 1 |  | 0.035486 | 0.30927 | 0.27378 | 0.032661
122 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 673 | 2138 |  | 1292 | 2.328 | 5.224 | 2.895 |  | 0.048 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:39 |  | 1 |  | 1.6471 | 5.8314 | 4.1843 | 0.3587
123 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1199 | 2208 |  | 1292 | 3.337 | 5.374 | 2.037 |  | 0.048 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:48 |  | 1 |  | 2.9503 | 6.0128 | 3.0625 | 0.3587 |  |  | 8.9829 |  |  |  |  | 0.8222 |  |  |  |  | 11.2695 |  |  |  |  | 0.48816 |  |  | 3.089 | 5.9882 | 2.8992 | 0.35958
124 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1311 | 2189 |  | 1298 | 3.552 | 5.332 | 1.78 |  | 0.049 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:57 |  | 1 |  | 3.2278 | 5.9636 | 2.7358 | 0.36045
125 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3981 | 7398 |  | 3567 | 2.891 | 5.455 | 2.564 |  | 0.214 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 18:17 |  | 1 |  | 3.2809 | 6.4872 | 3.2063 | 0.34142
126 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3959 | 7606 |  | 3690 | 2.877 | 5.602 | 2.725 |  | 0.225 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 18:27 |  | 1 |  | 3.2628 | 6.6669 | 3.4041 | 0.35341 |  |  | 2.1533 |  |  |  |  | 1.0943 |  |  |  |  | 4.3087 |  |  |  |  | 1.7536 |  |  | 3.2983 | 6.6306 | 3.3323 | 0.35034
127 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4045 | 7522 |  | 3627 | 2.932 | 5.543 | 2.611 |  | 0.219 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 18:39 |  | 1 |  | 3.3338 | 6.5943 | 3.2605 | 0.34727
128 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.33 | 0.33 | 4854 | 11070 |  | 5394 | 6.271 | 14.632 | 8.361 |  | 0.679 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 18:57 |  | 1 |  | 7.2762 | 17.5618 | 10.2856 | 0.94469
129 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.33 | 0.33 | 5943 | 11109 |  | 5321 | 7.537 | 14.682 | 7.145 |  | 0.667 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 19:08 |  | 1 |  | 8.9114 | 17.6231 | 8.7117 | 0.93175 |  |  | 1.9357 |  |  |  |  | 0.57758 |  |  |  |  | 0.83103 |  |  |  |  | 0.58817 |  |  | 8.9985 | 17.6741 | 8.6756 | 0.9345
130 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.33 | 0.33 | 6059 | 11174 |  | 5352 | 7.672 | 14.765 | 7.093 |  | 0.672 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 19:19 |  | 1 |  | 9.0856 | 17.7252 | 8.6396 | 0.93725
131 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.47 | 0.47 | 8722 | 16771 |  | 7855 | 7.56 | 15.411 | 7.851 |  | 0.751 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 19:39 |  | 1 |  | 9.1868 | 18.6171 | 9.4303 | 0.96967
132 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.47 | 0.47 | 8909 | 16740 |  | 7922 | 7.713 | 15.384 | 7.671 |  | 0.758 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 19:50 |  | 1 |  | 9.3839 | 18.5829 | 9.199 | 0.97801 |  |  | 0.57463 |  |  |  |  | 0.6565 |  |  |  |  | 1.8969 |  |  |  |  | 0.61286 |  |  | 9.357 | 18.6441 | 9.287 | 0.97502
133 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.47 | 0.47 | 8858 | 16851 |  | 7874 | 7.671 | 15.483 | 7.812 |  | 0.753 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 20:02 |  | 1 |  | 9.3301 | 18.7053 | 9.3751 | 0.97203
134 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11206 | 21419 |  | 10124 | 7.51 | 15.354 | 7.844 |  | 0.786 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 20:23 |  | 1 |  | 9.2477 | 18.5982 | 9.3505 | 0.98084
135 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11454 | 21520 |  | 10125 | 7.669 | 15.425 | 7.756 |  | 0.786 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 20:36 |  | 1 |  | 9.4525 | 18.6854 | 9.2329 | 0.98093 |  |  | 0.11351 |  |  |  |  | 0.15705 |  |  |  |  | 0.2016 |  |  |  |  | 0.66385 |  |  | 9.4579 | 18.7001 | 9.2422 | 0.9842
136 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11467 | 21554 |  | 10192 | 7.677 | 15.449 | 7.773 |  | 0.792 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 20:48 |  | 1 |  | 9.4633 | 18.7148 | 9.2515 | 0.98747
137 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6569 | 10337 |  | 5002 | 9.091 | 15.06 | 5.969 |  | 0.678 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 21:07 |  | 1 |  | 10.8365 | 18.0517 | 7.2152 | 0.96271
138 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6948 | 10294 |  | 4778 | 9.575 | 14.999 | 5.424 |  | 0.64 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 21:17 |  | 1 |  | 11.4625 | 17.9774 | 6.515 | 0.91902 | 109.8852 |  | 1.1034 |  |  | 84.5201 |  | 0.40442 |  |  | 75.5141 |  | 3.1136 |  |  | 94.0489 |  | 0.44467 |  |  | 11.5261 | 17.9412 | 6.4151 | 0.92107
139 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7025 | 10252 |  | 4799 | 9.674 | 14.94 | 5.265 |  | 0.643 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 21:28 |  | 1 |  | 11.5897 | 17.9049 | 6.3152 | 0.92312
140 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5617 | 7148 |  | 3452 | 4.724 | 6.334 | 1.61 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 21:47 |  | 1 |  | 5.5584 | 7.5256 | 1.9671 | 0.39624
141 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3927 | 7107 |  | 3543 | 3.428 | 6.299 | 2.872 |  | 0.255 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 21:57 |  | 1 |  | 3.8836 | 7.4831 | 3.5994 | 0.40689 |  |  | 1.9958 |  |  |  |  | 0.92377 |  |  |  |  | 0.24599 |  |  |  |  | 2.0336 |  |  | 3.9228 | 7.5178 | 3.595 | 0.4028
142 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4006 | 7174 |  | 3473 | 3.488 | 6.356 | 2.868 |  | 0.247 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 22:08 |  | 1 |  | 3.9619 | 7.5525 | 3.5906 | 0.3987
143 | 9 | F50 24JUL23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5092 | 7730 |  | 3844 | 4.321 | 6.827 | 2.505 |  | 0.286 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 22:27 |  | 1 |  | 5.0382 | 8.1288 | 3.0907 | 0.44212
144 | 9 | F50 24JUL23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5271 | 7754 |  | 3752 | 4.459 | 6.848 | 2.389 |  | 0.276 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 22:38 |  | 1 |  | 5.2155 | 8.1537 | 2.9381 | 0.43135 |  |  | 1.5086 |  |  |  |  | 1.2407 |  |  |  |  | 6.3136 |  |  |  |  | 1.8619 |  |  | 5.2552 | 8.1034 | 2.8482 | 0.42737
145 | 9 | F50 24JUL23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5351 | 7657 |  | 3684 | 4.52 | 6.766 | 2.246 |  | 0.269 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 22:49 |  | 1 |  | 5.2948 | 8.0531 | 2.7583 | 0.42339
146 | 10 | B50 31JUL23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5479 | 8379 |  | 9029 | 4.618 | 7.377 | 2.76 |  | 0.828 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 23:08 |  | 1 |  | 5.4217 | 8.8015 | 3.3798 | 1.0489
147 | 10 | B50 31JUL23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5541 | 8456 |  | 9198 | 4.666 | 7.443 | 2.777 |  | 0.846 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 23:18 |  | 1 |  | 5.4831 | 8.8813 | 3.3982 | 1.0686 |  |  | 0.95336 |  |  |  |  | 0.10509 |  |  |  |  | 1.8369 |  |  |  |  | 2.7418 |  |  | 5.5094 | 8.8767 | 3.3673 | 1.0542
148 | 10 | B50 31JUL23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5594 | 8447 |  | 8951 | 4.706 | 7.434 | 2.728 |  | 0.82 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 23:29 |  | 1 |  | 5.5356 | 8.872 | 3.3364 | 1.0397
149 | 11 | B50 14AUG23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3826 | 7375 |  | 3312 | 3.35 | 6.527 | 3.176 |  | 0.23 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 23:48 |  | 1 |  | 3.7835 | 7.7608 | 3.9773 | 0.37986
150 | 11 | B50 14AUG23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3443 | 7435 |  | 3342 | 3.056 | 6.577 | 3.521 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 23:58 |  | 1 |  | 3.404 | 7.823 | 4.4191 | 0.38337 |  |  | 0.64256 |  |  |  |  | 0.14585 |  |  |  |  | 0.23508 |  |  |  |  | 1.0432 |  |  | 3.3931 | 7.8173 | 4.4243 | 0.38138
151 | 11 | B50 14AUG23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3421 | 7424 |  | 3308 | 3.039 | 6.568 | 3.529 |  | 0.23 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 0:09 |  | 1 |  | 3.3822 | 7.8116 | 4.4295 | 0.37939
152 | 12 | B50 29MAY23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2990 | 5822 |  | 2190 | 2.709 | 5.211 | 2.502 |  | 0.113 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 0:27 |  | 1 |  | 2.955 | 6.1511 | 3.1961 | 0.24856
153 | 12 | B50 29MAY23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2857 | 5904 |  | 2306 | 2.606 | 5.28 | 2.674 |  | 0.125 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 0:38 |  | 1 |  | 2.8232 | 6.2361 | 3.4129 | 0.26214 |  |  | 0.14031 |  |  |  |  | 0.28296 |  |  |  |  | 0.63446 |  |  |  |  | 5.2689 |  |  | 2.8252 | 6.2273 | 3.4021 | 0.25541
154 | 12 | B50 29MAY23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2861 | 5887 |  | 2191 | 2.609 | 5.266 | 2.657 |  | 0.113 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 0:48 |  | 1 |  | 2.8272 | 6.2185 | 3.3913 | 0.24868
155 | 13 | F 2MAY22 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4043 | 15402 |  | 3193 | 3.516 | 13.327 | 9.81 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 1:08 |  | 1 |  | 3.9986 | 16.081 | 12.0825 | 0.36594
156 | 13 | F 2MAY22 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4469 | 15502 |  | 2974 | 3.843 | 13.411 | 9.568 |  | 0.195 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 1:19 |  | 1 |  | 4.4207 | 16.1847 | 11.7639 | 0.34031 |  |  | 2.313 |  |  |  |  | 0.96529 |  |  |  |  | 0.46346 |  |  |  |  | 5.061 |  |  | 4.3702 | 16.1069 | 11.7367 | 0.34914
157 | 13 | F 2MAY22 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4367 | 15352 |  | 3125 | 3.765 | 13.284 | 9.519 |  | 0.211 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 1:31 |  | 1 |  | 4.3197 | 16.0292 | 11.7095 | 0.35798
158 | 14 | C50 17NOV22 21 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11661 | 15921 |  | 2796 | 9.361 | 13.767 | 4.406 |  | 0.176 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 1:50 |  | 1 |  | 11.5482 | 16.619 | 5.0708 | 0.31948
159 | 14 | C50 17NOV22 21 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 13845 | 15777 |  | 2852 | 11.037 | 13.644 | 2.607 |  | 0.182 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 2:01 |  | 1 |  | 13.7126 | 16.4697 | 2.7572 | 0.32603 |  |  | 0.76315 |  |  |  |  | 0.11322 |  |  |  |  | 3.1832 |  |  |  |  | 0 |  |  | 13.7651 | 16.4791 | 2.714 | 0.32603
160 | 14 | C50 17NOV22 21 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 13951 | 15795 |  | 2852 | 11.118 | 13.66 | 2.543 |  | 0.182 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 2:13 |  | 1 |  | 13.8176 | 16.4884 | 2.6708 | 0.32603
161 | 15 | F50 14AUG23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7752 | 7661 |  | 6831 | 6.362 | 6.769 | 0.407 |  | 0.598 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 2:32 |  | 1 |  | 7.6743 | 8.0573 | 0.38303 | 0.79166
162 | 15 | F50 14AUG23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6266 | 7802 |  | 6949 | 5.222 | 6.889 | 1.666 |  | 0.611 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 2:42 |  | 1 |  | 6.2016 | 8.2034 | 2.0018 | 0.80546 |  |  | 0.94729 |  |  |  |  | 0.34057 |  |  |  |  | 4.2276 |  |  |  |  | 0.78146 |  |  | 6.1724 | 8.2174 | 2.0451 | 0.80862
163 | 15 | F50 14AUG23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6207 | 7829 |  | 7003 | 5.177 | 6.911 | 1.734 |  | 0.616 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 2:53 |  | 1 |  | 6.1431 | 8.2314 | 2.0883 | 0.81178
164 | 16 | F50 29MAY23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4803 | 6465 |  | 2874 | 4.099 | 5.756 | 1.656 |  | 0.185 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 3:12 |  | 1 |  | 4.7517 | 6.8176 | 2.0659 | 0.32861
165 | 16 | F50 29MAY23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4408 | 6486 |  | 2893 | 3.796 | 5.773 | 1.977 |  | 0.187 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 3:22 |  | 1 |  | 4.3603 | 6.8394 | 2.4791 | 0.33083 |  |  | 2.4471 |  |  |  |  | 0.58931 |  |  |  |  | 2.7644 |  |  |  |  | 2.0367 |  |  | 4.4143 | 6.8596 | 2.4453 | 0.32749
166 | 16 | F50 29MAY23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4517 | 6525 |  | 2836 | 3.88 | 5.806 | 1.926 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 3:33 |  | 1 |  | 4.4683 | 6.8798 | 2.4115 | 0.32416
167 | 17 | B50 2MAR23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3159 | 5213 |  | 2953 | 2.838 | 4.695 | 1.857 |  | 0.193 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 3:52 |  | 1 |  | 3.1225 | 5.5199 | 2.3974 | 0.33785
168 | 17 | B50 2MAR23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2836 | 5306 |  | 2927 | 2.59 | 4.774 | 2.184 |  | 0.19 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 4:02 |  | 1 |  | 2.8024 | 5.6163 | 2.8139 | 0.33481 |  |  | 2.0362 |  |  |  |  | 2.4473 |  |  |  |  | 2.8583 |  |  |  |  | 0.63111 |  |  | 2.7742 | 5.5484 | 2.7742 | 0.33376
169 | 17 | B50 2MAR23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2779 | 5175 |  | 2909 | 2.547 | 4.662 | 2.115 |  | 0.188 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 4:12 |  | 1 |  | 2.7459 | 5.4805 | 2.7346 | 0.3327
170 | 18 | B50 24JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3003 | 5926 |  | 2009 | 2.719 | 5.299 | 2.58 |  | 0.094 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 4:31 |  | 1 |  | 2.9679 | 6.2589 | 3.291 | 0.22738
171 | 18 | B50 24JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3012 | 6032 |  | 2114 | 2.725 | 5.389 | 2.663 |  | 0.105 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 4:41 |  | 1 |  | 2.9768 | 6.3688 | 3.392 | 0.23967 |  |  | 1.0597 |  |  |  |  | 0.83349 |  |  |  |  | 2.5249 |  |  |  |  | 5.0047 |  |  | 2.9927 | 6.3424 | 3.3497 | 0.23382
172 | 18 | B50 24JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3044 | 5981 |  | 2014 | 2.75 | 5.346 | 2.595 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 4:51 |  | 1 |  | 3.0085 | 6.3159 | 3.3074 | 0.22797
173 | 19 | B50 29MAY23 3 SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8081 | 18121 |  | 7404 | 6.615 | 15.63 | 9.015 |  | 0.658 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 5:11 |  | 1 |  | 8.0003 | 18.8993 | 10.899 | 0.85871
174 | 19 | B50 29MAY23 3 SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9634 | 18388 |  | 7370 | 7.806 | 15.856 | 8.051 |  | 0.655 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 5:22 |  | 1 |  | 9.5394 | 19.1761 | 9.6367 | 0.85473 |  |  | 0.26006 |  | 174.7505 |  |  | 0.04866 |  | 53.9505 |  |  | 0.16016 |  | -66.8495 |  |  | 0.55976 |  | 171.9001 | 9.527 | 19.1714 | 9.6445 | 0.85713
175 | 19 | B50 29MAY23 3 SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9609 | 18379 |  | 7411 | 7.786 | 15.849 | 8.063 |  | 0.659 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 5:34 |  | 1 |  | 9.5146 | 19.1668 | 9.6522 | 0.85953
176 | 20 | B50 24JUL23 0.1 DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4557 | 6102 |  | 2135 | 3.911 | 5.448 | 1.537 |  | 0.107 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 5:53 |  | 1 |  | 4.508 | 6.4414 | 1.9334 | 0.24213
177 | 20 | B50 24JUL23 0.1 DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3085 | 6211 |  | 2197 | 2.782 | 5.54 | 2.759 |  | 0.114 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 6:03 |  | 1 |  | 3.0492 | 6.5543 | 3.5052 | 0.24938 |  |  | 2.2813 | 3.0169 |  |  |  | 0.41033 | 3.4925 |  |  |  | 1.2463 | 3.9155 |  |  |  | 1.3979 | 7.1425 |  | 3.0844 | 6.5678 | 3.4834 | 0.25114
178 | 20 | B50 24JUL23 0.1 DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3156 | 6237 |  | 2227 | 2.836 | 5.562 | 2.726 |  | 0.117 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 6:14 |  | 1 |  | 3.1195 | 6.5813 | 3.4617 | 0.25289
179 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 867 | 440 |  | 383 | 1.08 | 0.651 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 6:31 |  | 1 |  | 0.8511 | 0.57255 | -0.27855 | 0.037108
180 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 213 | 480 |  | 410 | 0.578 | 0.685 | 0.107 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 6:39 |  | 1 |  | 0.20297 | 0.61401 | 0.41104 | 0.040268 |  |  | 3.8313 |  |  |  |  | 7.0047 |  |  |  |  | 8.5342 |  |  |  |  | 3.1464 |  |  | 0.20693 | 0.63629 | 0.42936 | 0.040911
181 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 221 | 523 |  | 421 | 0.584 | 0.721 | 0.137 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 6:49 |  | 1 |  | 0.2109 | 0.65858 | 0.44768 | 0.041555
182 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4932 | 8281 |  | 4364 | 6.998 | 12.157 | 5.159 |  | 0.567 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 7:06 |  | 1 |  | 8.1327 | 14.4999 | 6.3672 | 0.83828
183 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7117 | 8405 |  | 4328 | 9.791 | 12.332 | 2.541 |  | 0.561 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 7:16 |  | 1 |  | 11.7416 | 14.7141 | 2.9725 | 0.83126 | 113.5341 |  | 2.8154 |  |  |  |  | 0.12923 |  |  | 20.3744 |  | 12.6746 |  |  | 85.5347 |  | 1.259 |  |  | 11.9093 | 14.7046 | 2.7953 | 0.83652
184 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7320 | 8394 |  | 4382 | 10.051 | 12.316 | 2.265 |  | 0.571 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 7:27 |  | 1 |  | 12.0769 | 14.6951 | 2.6182 | 0.84179
185 | 3 |  | Flush | [20 empty] | 2023-10-27 | 7:33
186 | 21 | F50 29MAY23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3774 | 7772 |  | 2787 | 3.31 | 6.863 | 3.553 |  | 0.175 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 7:52 |  | 1 |  | 3.732 | 8.1723 | 4.4404 | 0.31843
187 | 21 | F50 29MAY23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4440 | 7777 |  | 2820 | 3.821 | 6.867 | 3.046 |  | 0.179 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 8:02 |  | 1 |  | 4.392 | 8.1775 | 3.7855 | 0.32229 |  |  | 3.9373 |  |  |  |  | 0.12683 |  |  |  |  | 5.0585 |  |  |  |  | 0.32625 |  |  | 4.4802 | 8.1723 | 3.6921 | 0.32281
188 | 21 | F50 29MAY23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4618 | 7767 |  | 2829 | 3.958 | 6.858 | 2.9 |  | 0.18 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 8:13 |  | 1 |  | 4.5684 | 8.1672 | 3.5988 | 0.32334
189 | 22 | F50 29MAY23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4254 | 7472 |  | 3010 | 3.678 | 6.608 | 2.93 |  | 0.199 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 8:32 |  | 1 |  | 4.2077 | 7.8614 | 3.6537 | 0.34452
190 | 22 | F50 29MAY23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4262 | 7502 |  | 3109 | 3.685 | 6.634 | 2.949 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 8:42 |  | 1 |  | 4.2156 | 7.8925 | 3.6769 | 0.35611 |  |  | 2.8046 |  |  |  |  | 0.18369 |  |  |  |  | 2.9083 |  |  |  |  | 1.3231 |  |  | 4.2756 | 7.8997 | 3.6242 | 0.35377
191 | 22 | F50 29MAY23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4383 | 7516 |  | 3069 | 3.778 | 6.646 | 2.868 |  | 0.205 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 8:53 |  | 1 |  | 4.3355 | 7.907 | 3.5715 | 0.35143
192 | 23 | B50 29MAY23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3184 | 5639 |  | 2056 | 2.857 | 5.056 | 2.198 |  | 0.099 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 9:12 |  | 1 |  | 3.1473 | 5.9614 | 2.8142 | 0.23288
193 | 23 | B50 29MAY23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2852 | 5638 |  | 2003 | 2.603 | 5.055 | 2.452 |  | 0.094 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 9:22 |  | 1 |  | 2.8183 | 5.9604 | 3.1421 | 0.22668 |  |  | 0.35102 |  |  |  |  | 0.85577 |  |  |  |  | 1.9506 |  |  |  |  | 0.5149 |  |  | 2.8232 | 5.935 | 3.1118 | 0.22727
194 | 23 | B50 29MAY23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2862 | 5589 |  | 2013 | 2.611 | 5.013 | 2.403 |  | 0.094 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 9:32 |  | 1 |  | 2.8282 | 5.9096 | 3.0814 | 0.22785
195 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 773 | 505 |  | 413 | 1.008 | 0.706 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 9:49 |  | 1 |  | 0.75794 | 0.63992 | -0.11802 | 0.040619
196 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 223 | 448 |  | 373 | 0.586 | 0.658 | 0.072 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 9:58 |  | 1 |  | 0.21288 | 0.58084 | 0.36796 | 0.035938 |  |  | 18.1966 |  |  |  |  | 6.8921 |  |  |  |  | 0.3138 |  |  |  |  | 11.3634 |  |  | 0.23419 | 0.60157 | 0.36738 | 0.038103
197 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 266 | 488 |  | 410 | 0.619 | 0.692 | 0.073 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 10:07 |  | 1 |  | 0.25549 | 0.6223 | 0.36681 | 0.040268
198 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4773 | 8012 |  | 4366 | 6.795 | 11.777 | 4.982 |  | 0.568 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 10:25 |  | 1 |  | 7.87 | 14.0352 | 6.1652 | 0.83867
199 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6841 | 8045 |  | 4443 | 9.439 | 11.824 | 2.385 |  | 0.581 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 10:35 |  | 1 |  | 11.2857 | 14.0922 | 2.8065 | 0.85369 | 109.7122 |  | 3.8609 |  |  | 66.1237 |  | 0.19633 |  |  | 18.3128 |  | 18.3603 |  |  | 87.489 |  | 0.52408 |  |  | 11.5079 | 14.0784 | 2.5705 | 0.85593
200 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7110 | 8029 |  | 4466 | 9.783 | 11.801 | 2.018 |  | 0.585 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 10:45 |  | 1 |  | 11.7301 | 14.0646 | 2.3345 | 0.85817

## 31oct23
RUN NOTES:
Daily Calibration |  |  | Injection Volume | mgTIC | TIC Area | mgTC | TC Area | mgTNb | TNb Area |  | Volume TIC | TIC inverse prediction mg/L | TC inverse prediction mg/L | TOC inverse prediction mg/L | TNb inverse prediction mg/L
Correct for TOC in Type I water
Diluent signals @ 0.5 mls for corrections to cal curve |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0.035144 | 0.11285 | 0.077706 | 0.0043211
TIC | TC | TNb |  | 0.0006 | 1122.4 | 0.0012 | 2054.5 | 6e-05 | 1035.7 |  | 0.2 | 2.926 | 5.7612 | 2.8352 | 0.32127
46.5 | 242.5 | 277 |  | 0.0018 | 3698.7 | 0.0036 | 6877 | 0.00018 | 2807.1 |  | 0.6 | 3.1468 | 6.2074 | 3.0606 | 0.2841
|  |  |  | 0.002997 | 5636.031 | 0.005994 | 10700.995 | 0.0002997 | 4904.018 |  | 0.333 | 8.6122 | 17.3095 | 8.6973 | 0.88942
|  |  |  | 0.004203 | 8269.569 | 0.008406 | 15761.505 | 0.0004203 | 7162.005 |  | 0.467 | 8.9929 | 18.1225 | 9.1295 | 0.92409
|  |  |  | 0.0054 | 10725.2 | 0.0108 | 20217 | 0.00054 | 8894.1 |  | 0.6 | 9.0693 | 18.0661 | 8.9968 | 0.89233
|  |  | Slope |  | 5.0573e-07 |  | 5.3337e-07 |  | 5.9954e-08
|  |  | Intercept |  | 1.7572e-05 |  | 5.6425e-05 |  | 2.1605e-06
|  |  | RSQ |  | 0.99878 |  | 0.9991 |  | 0.99875
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps
1 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7981 | 12887 |  | 4854 | 10.896 | 18.661 | 7.765 |  | 0.653 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 13:58 |  | 1 |  | 13.5125 | 23.1 | 9.5875 | 0.97726
2 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7500 | 12780 |  | 4928 | 10.281 | 18.51 | 8.229 |  | 0.666 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 14:08 |  | 1 |  | 12.7017 | 22.9098 | 10.2081 | 0.99205 |  |  | 0.56907 |  |  |  |  | 1.0499 |  |  |  |  | 1.6449 |  |  |  |  | 0.44221 |  |  | 12.7379 | 23.0307 | 10.2928 | 0.99425
3 | 1 | RunIn | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7543 | 12916 |  | 4950 | 10.337 | 18.702 | 8.365 |  | 0.67 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 14:18 |  | 1 |  | 12.7742 | 23.1516 | 10.3774 | 0.99644
4 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3177 | 1124 |  | 539 | 2.852 | 1.231 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 14:36 |  | 1 |  | 3.2485 | 1.3119 | -1.9366 | 0.068952
5 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 647 | 1024 |  | 625 | 0.911 | 1.146 | 0.235 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 14:44 |  | 1 |  | 0.68955 | 1.2052 | 0.51565 | 0.079264 |  |  | 17.3775 |  |  |  |  | 0.35468 |  |  |  |  | 18.6377 |  |  |  |  | 6.8854 |  |  | 0.63443 | 1.2031 | 0.56864 | 0.076626
6 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 538 | 1020 |  | 581 | 0.828 | 1.143 | 0.315 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 14:54 |  | 1 |  | 0.5793 | 1.2009 | 0.62163 | 0.073988
7 | 3 |  | Flush | [20 empty] | 2023-10-31 | 14:59 |  | 1
8 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 34 | 191 |  | 240 | 0.441 | 0.44 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 15:14 |  | 1 |  | 0.069533 | 0.3166 | 0.24707 | 0.033099
9 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 50 | 253 |  | 305 | 0.454 | 0.493 | 0.039 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 15:22 |  | 1 |  | 0.085717 | 0.38274 | 0.29702 | 0.040893 |  |  | 8.6158 |  |  |  |  | 6.0295 |  |  |  |  | 5.295 |  |  |  |  | 17.8893 |  |  | 0.082176 | 0.37154 | 0.28936 | 0.037536
10 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 43 | 232 |  | 249 | 0.448 | 0.475 | 0.028 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 15:31 |  | 1 |  | 0.078636 | 0.36034 | 0.2817 | 0.034178
11 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 536 | 2100 |  | 1162 | 2.065 | 5.143 | 3.078 |  | 0.014 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 15:47 |  | 1 |  | 1.4432 | 5.8825 | 4.4393 | 0.35914
12 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1098 | 2154 |  | 1158 | 3.143 | 5.258 | 2.115 |  | 0.013 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 15:56 |  | 1 |  | 2.8643 | 6.0266 | 3.1623 | 0.35794 |  | 0.89926 | 3.7254 |  |  |  | 0.33155 | 0.2215 |  |  |  | 1.3401 | 3.936 |  |  |  | 18.1632 | 1.945 |  |  | 2.9187 | 6.0199 | 3.1012 | 0.35449
13 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1141 | 2149 |  | 1135 | 3.225 | 5.247 | 2.022 |  | 0.007 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 16:05 |  | 1 |  | 2.973 | 6.0132 | 3.0402 | 0.35104
14 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3625 | 7054 |  | 3132 | 2.663 | 5.212 | 2.549 |  | 0.176 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 16:23 |  | 1 |  | 3.0847 | 6.3647 | 3.28 | 0.31656
15 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3675 | 7167 |  | 3132 | 2.695 | 5.292 | 2.597 |  | 0.176 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 16:34 |  | 1 |  | 3.1269 | 6.4652 | 3.3383 | 0.31656 |  | 6.4621 | 4.1961 |  |  |  | 7.768 | 0.027496 |  |  |  | 9.0739 | 4.0413 |  |  |  | 5.7702 | 0.47236 |  |  | 3.1939 | 6.4661 | 3.2722 | 0.31731
16 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3834 | 7169 |  | 3147 | 2.797 | 5.293 | 2.496 |  | 0.178 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 16:45 |  | 1 |  | 3.2609 | 6.467 | 3.2061 | 0.31806
17 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 4546 | 10806 |  | 5091 | 5.86 | 14.164 | 8.304 |  | 0.625 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 17:03 |  | 1 |  | 6.9568 | 17.4777 | 10.5209 | 0.92308
18 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 5636 | 10799 |  | 5033 | 7.115 | 14.156 | 7.041 |  | 0.616 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 17:13 |  | 1 |  | 8.6121 | 17.4665 | 8.8543 | 0.91264 |  | 3.7866 | 1.0874 |  |  |  | 2.3991 | 1.1579 |  |  |  | 1.0116 | 1.2264 |  |  |  | 2.5148 | 2.166 |  |  | 8.6592 | 17.5682 | 8.909 | 0.92263
19 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.333 | 0.333 | 5698 | 10926 |  | 5144 | 7.187 | 14.317 | 7.13 |  | 0.634 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 17:24 |  | 1 |  | 8.7063 | 17.6699 | 8.9636 | 0.93263
20 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8138 | 15963 |  | 7373 | 7.128 | 14.778 | 7.649 |  | 0.701 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 17:44 |  | 1 |  | 8.8505 | 18.3526 | 9.5022 | 0.95118
21 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8276 | 15966 |  | 7357 | 7.242 | 14.78 | 7.538 |  | 0.7 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 17:54 |  | 1 |  | 8.9999 | 18.356 | 9.3561 | 0.94913 |  | 0.4441 | 0.88647 |  |  |  | 2.1176 | 0.2734 |  |  |  | 3.7911 | 0.3199 |  |  |  | 5.908 | 0.84854 |  |  | 9.04 | 18.3812 | 9.3412 | 0.95317
22 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.467 | 0.467 | 8350 | 16010 |  | 7420 | 7.303 | 14.82 | 7.517 |  | 0.707 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 18:06 |  | 1 |  | 9.08 | 18.4063 | 9.3263 | 0.95722
23 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 10484 | 20351 |  | 9180 | 7.048 | 14.6 | 7.551 |  | 0.703 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 18:26 |  | 1 |  | 8.866 | 18.1852 | 9.3192 | 0.9209
24 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 10788 | 20459 |  | 9217 | 7.243 | 14.676 | 7.433 |  | 0.707 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 18:38 |  | 1 |  | 9.1222 | 18.2812 | 9.159 | 0.9246 |  | 1.2925 | 0.12944 |  |  |  | 1.8042 | 0.47541 |  |  |  | 2.3159 | 1.0742 |  |  |  | 2.8383 | 0.20513 |  |  | 9.1163 | 18.3248 | 9.2084 | 0.92554
25 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 10774 | 20557 |  | 9236 | 7.234 | 14.745 | 7.512 |  | 0.708 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 18:50 |  | 1 |  | 9.1104 | 18.3683 | 9.2579 | 0.92649
26 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5862 | 11902 |  | 4608 | 8.187 | 17.27 | 9.083 |  | 0.61 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 19:09 |  | 1 |  | 9.9404 | 21.3488 | 11.4084 | 0.9281
27 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6232 | 11970 |  | 4654 | 8.66 | 17.366 | 8.706 |  | 0.618 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 19:19 |  | 1 |  | 10.5642 | 21.4697 | 10.9055 | 0.93729 | 100 |  | 0.96868 |  |  | 100 |  | 0.38165 |  |  | 100 |  | 1.7073 |  |  | 100 |  | 1.981 |  |  | 10.6156 | 21.4288 | 10.8132 | 0.9281
28 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6293 | 11924 |  | 4562 | 8.738 | 17.301 | 8.563 |  | 0.602 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 19:30 |  | 1 |  | 10.667 | 21.3879 | 10.7209 | 0.9189
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5126 | 6829 |  | 2830 | 4.347 | 6.064 | 1.716 |  | 0.18 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 19:49 |  | 1 |  | 5.2198 | 7.3977 | 2.1778 | 0.34366
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3714 | 6867 |  | 2904 | 3.264 | 6.097 | 2.832 |  | 0.188 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 19:59 |  | 1 |  | 3.7917 | 7.4382 | 3.6465 | 0.35253 |  |  | 0.4259 |  |  |  |  | 0.96552 |  |  |  |  | 2.433 |  |  |  |  | 0.10209 |  |  | 3.7998 | 7.4025 | 3.6027 | 0.35235
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3730 | 6800 |  | 2901 | 3.277 | 6.039 | 2.762 |  | 0.187 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 20:10 |  | 1 |  | 3.8079 | 7.3667 | 3.5589 | 0.35217
32 | 3 |  | Flush | [20 empty] | 2023-10-31 | 20:16 |  | 1
33 | 9 | B50 14AUG23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3614 | 6513 |  | 1695 | 3.188 | 5.796 | 2.608 |  | 0.061 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 20:34 |  | 1 |  | 3.6905 | 7.0606 | 3.3701 | 0.20757
34 | 9 | B50 14AUG23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4413 | 6425 |  | 1697 | 3.8 | 5.722 | 1.922 |  | 0.061 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 20:44 |  | 1 |  | 4.4987 | 6.9667 | 2.468 | 0.20781 |  |  | 4.2256 |  |  |  |  | 0.36682 |  |  |  |  | 7.0728 |  |  |  |  | 5.028 |  |  | 4.5958 | 6.9795 | 2.3837 | 0.20271
35 | 9 | B50 14AUG23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4605 | 6449 |  | 1612 | 3.948 | 5.742 | 1.794 |  | 0.053 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 20:55 |  | 1 |  | 4.6929 | 6.9923 | 2.2994 | 0.19761
36 | 10 | F50 14AUG23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4415 | 9048 |  | 3669 | 3.802 | 7.944 | 4.141 |  | 0.268 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 21:14 |  | 1 |  | 4.5007 | 9.7648 | 5.2641 | 0.44426
37 | 10 | F50 14AUG23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4478 | 9137 |  | 3684 | 3.851 | 8.019 | 4.169 |  | 0.269 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 21:24 |  | 1 |  | 4.5644 | 9.8597 | 5.2953 | 0.44606 |  |  | 0.37742 |  |  |  |  | 1.1316 |  |  |  |  | 1.7862 |  |  |  |  | 1.4895 |  |  | 4.5558 | 9.8043 | 5.2484 | 0.44277
38 | 10 | F50 14AUG23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4461 | 9033 |  | 3629 | 3.837 | 7.931 | 4.094 |  | 0.264 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 21:35 |  | 1 |  | 4.5472 | 9.7488 | 5.2016 | 0.43947
39 | 11 | B50 29may23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4363 | 6791 |  | 6396 | 3.762 | 6.032 | 2.27 |  | 0.553 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 21:54 |  | 1 |  | 4.4481 | 7.3571 | 2.909 | 0.77125
40 | 11 | B50 29may23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4343 | 6824 |  | 6399 | 3.747 | 6.06 | 2.313 |  | 0.553 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 22:04 |  | 1 |  | 4.4279 | 7.3923 | 2.9645 | 0.77161 |  |  | 1.3337 |  |  |  |  | 0.25941 |  |  |  |  | 2.5926 |  |  |  |  | 1.9695 |  |  | 4.3985 | 7.4019 | 3.0034 | 0.77929
41 | 11 | B50 29may23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4285 | 6842 |  | 6527 | 3.702 | 6.075 | 2.373 |  | 0.567 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 22:15 |  | 1 |  | 4.3692 | 7.4115 | 3.0423 | 0.78696
42 | 12 | F50 29may23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3969 | 7908 |  | 3396 | 3.459 | 6.978 | 3.519 |  | 0.239 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 22:34 |  | 1 |  | 4.0496 | 8.5487 | 4.4991 | 0.41153
43 | 12 | F50 29may23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4004 | 7928 |  | 3383 | 3.487 | 6.995 | 3.508 |  | 0.238 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 22:44 |  | 1 |  | 4.085 | 8.57 | 4.485 | 0.40997 |  |  | 2.4563 |  |  |  |  | 0.79982 |  |  |  |  | 0.68549 |  |  |  |  | 1.5326 |  |  | 4.0354 | 8.5359 | 4.5005 | 0.40685
44 | 12 | F50 29may23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3906 | 7864 |  | 3331 | 3.412 | 6.941 | 3.529 |  | 0.232 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 22:55 |  | 1 |  | 3.9859 | 8.5018 | 4.5159 | 0.40374
45 | 13 | F100 24jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4299 | 5474 |  | 1291 | 3.713 | 4.916 | 1.203 |  | 0.019 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 23:13 |  | 1 |  | 4.3834 | 5.9522 | 1.5689 | 0.15912
46 | 13 | F100 24jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4360 | 5468 |  | 1246 | 3.76 | 4.911 | 1.151 |  | 0.014 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 23:23 |  | 1 |  | 4.4451 | 5.9458 | 1.5008 | 0.15373 |  |  | 3.0475 |  |  |  |  | 0.44752 |  |  |  |  | 7.6723 |  |  |  |  | 3.0086 |  |  | 4.5138 | 5.9592 | 1.4453 | 0.15145
47 | 13 | F100 24jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4496 | 5493 |  | 1208 | 3.865 | 4.932 | 1.067 |  | 0.01 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 23:34 |  | 1 |  | 4.5826 | 5.9725 | 1.3899 | 0.14917
48 | 14 | F50 22aug23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4584 | 9930 |  | 4298 | 3.931 | 8.691 | 4.76 |  | 0.334 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-31 | 23:53 |  | 1 |  | 4.6716 | 10.7057 | 6.034 | 0.51969
49 | 14 | F50 22aug23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4633 | 9969 |  | 4376 | 3.969 | 8.724 | 4.755 |  | 0.342 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 0:03 |  | 1 |  | 4.7212 | 10.7473 | 6.0261 | 0.52904 |  |  | 2.7468 |  |  |  |  | 0.48518 |  |  |  |  | 1.3233 |  |  |  |  | 0.27235 |  |  | 4.7869 | 10.7734 | 5.9865 | 0.52832
50 | 14 | F50 22aug23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4763 | 10018 |  | 4364 | 4.069 | 8.766 | 4.697 |  | 0.34 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 0:14 |  | 1 |  | 4.8527 | 10.7995 | 5.9468 | 0.5276
51 | 15 | F50 31jul23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5236 | 7147 |  | 3412 | 4.432 | 6.333 | 1.902 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 0:33 |  | 1 |  | 5.3311 | 7.7369 | 2.4058 | 0.41345
52 | 15 | F50 31jul23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5496 | 7144 |  | 3379 | 4.631 | 6.331 | 1.699 |  | 0.237 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 0:43 |  | 1 |  | 5.5941 | 7.7337 | 2.1396 | 0.40949 |  |  | 1.0908 |  |  |  |  | 0.35928 |  |  |  |  | 1.5283 |  |  |  |  | 2.3155 |  |  | 5.5637 | 7.7198 | 2.1561 | 0.41429
53 | 15 | F50 31jul23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5436 | 7118 |  | 3459 | 4.586 | 6.309 | 1.723 |  | 0.246 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 0:54 |  | 1 |  | 5.5334 | 7.706 | 2.1726 | 0.41908
54 | 16 | F50 24jul23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4503 | 9984 |  | 5426 | 3.87 | 8.737 | 4.867 |  | 0.451 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 1:13 |  | 1 |  | 4.5897 | 10.7633 | 6.1736 | 0.65494
55 | 16 | F50 24jul23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4364 | 9901 |  | 5482 | 3.763 | 8.667 | 4.904 |  | 0.457 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 1:23 |  | 1 |  | 4.4491 | 10.6747 | 6.2256 | 0.66166 |  |  | 1.8023 |  |  |  |  | 0.17003 |  |  |  |  | 1.6038 |  |  |  |  | 1.075 |  |  | 4.4896 | 10.6657 | 6.1761 | 0.65812
56 | 16 | F50 24jul23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4444 | 9884 |  | 5423 | 3.825 | 8.652 | 4.828 |  | 0.451 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 1:34 |  | 1 |  | 4.53 | 10.6566 | 6.1266 | 0.65458
57 | 17 | F50 14aug23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5157 | 7391 |  | 7002 | 4.371 | 6.54 | 2.169 |  | 0.616 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 1:52 |  | 1 |  | 5.2512 | 7.9972 | 2.746 | 0.84392
58 | 17 | F50 14aug23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5457 | 7424 |  | 7057 | 4.601 | 6.568 | 1.967 |  | 0.622 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 2:03 |  | 1 |  | 5.5546 | 8.0324 | 2.4778 | 0.85051 |  |  | 0.82278 |  |  |  |  | 1.0174 |  |  |  |  | 5.0227 |  |  |  |  | 0.46633 |  |  | 5.5319 | 8.0735 | 2.5416 | 0.84853
59 | 17 | F50 14aug23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5412 | 7501 |  | 7024 | 4.567 | 6.633 | 2.066 |  | 0.619 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 2:14 |  | 1 |  | 5.5091 | 8.1145 | 2.6054 | 0.84656
60 | 18 | F50 12sep23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6072 | 7783 |  | 9140 | 5.074 | 6.872 | 1.799 |  | 0.84 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 2:33 |  | 1 |  | 6.1767 | 8.4153 | 2.2387 | 1.1003
61 | 18 | F50 12sep23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6271 | 7767 |  | 9218 | 5.226 | 6.859 | 1.632 |  | 0.848 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 2:43 |  | 1 |  | 6.3779 | 8.3983 | 2.0203 | 1.1096 |  |  | 0.11107 |  |  |  |  | 0.51943 |  |  |  |  | 2.484 |  |  |  |  | 0.47434 |  |  | 6.3744 | 8.4201 | 2.0457 | 1.1123
62 | 18 | F50 12sep23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6264 | 7808 |  | 9262 | 5.221 | 6.893 | 1.672 |  | 0.853 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 2:54 |  | 1 |  | 6.3709 | 8.442 | 2.0711 | 1.1149
63 | 19 | SPK 13 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7730 | 12712 |  | 3993 | 6.345 | 11.048 | 4.702 |  | 0.302 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 3:14 |  | 1 |  | 7.8537 | 13.6733 | 5.8197 | 0.48311
64 | 19 | SPK 13 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8256 | 12628 |  | 4103 | 6.749 | 10.977 | 4.228 |  | 0.313 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 3:24 |  | 1 |  | 8.3857 | 13.5837 | 5.1981 | 0.4963 |  |  | 2.4307 |  | 135.0467 |  |  | 0.39971 |  | 129.5713 |  |  | 2.9662 |  | 124.096 |  |  | 0.67878 |  | 115.8763 | 8.4888 | 13.6109 | 5.1221 | 0.49463
65 | 19 | SPK 13 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8460 | 12679 |  | 4075 | 6.906 | 11.02 | 4.114 |  | 0.31 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 3:36 |  | 1 |  | 8.592 | 13.6381 | 5.0461 | 0.49295
66 | 20 | DUP 18 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6905 | 7933 |  | 8927 | 5.712 | 7 | 1.288 |  | 0.818 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 3:55 |  | 1 |  | 7.0192 | 8.5754 | 1.5561 | 1.0747
67 | 20 | DUP 18 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6353 | 7962 |  | 9046 | 5.289 | 7.024 | 1.735 |  | 0.83 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 4:05 |  | 1 |  | 6.4609 | 8.6063 | 2.1454 | 1.089 |  |  | 0.89633 | 0.90037 |  |  |  | 0.11162 | 2.1308 |  |  |  | 2.215 | 5.869 |  |  |  | 2.1142 | 3.1649 |  | 6.4321 | 8.6015 | 2.1694 | 1.0776
68 | 20 | DUP 18 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6296 | 7953 |  | 8856 | 5.245 | 7.017 | 1.772 |  | 0.81 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 4:16 |  | 1 |  | 6.4032 | 8.5967 | 2.1935 | 1.0662
69 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1659 | 503 |  | 533 | 1.688 | 0.705 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 4:34 |  | 1 |  | 1.7131 | 0.64942 | -1.0637 | 0.068232
70 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 345 | 508 |  | 547 | 0.679 | 0.709 | 0.03 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 4:42 |  | 1 |  | 0.38409 | 0.65476 | 0.27066 | 0.069911 |  |  | 1.0589 |  |  |  |  | 8.4271 |  |  |  |  | 20.4486 |  |  |  |  | 3.6679 |  |  | 0.38207 | 0.68356 | 0.30149 | 0.068652
71 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 341 | 562 |  | 526 | 0.676 | 0.755 | 0.078 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 4:52 |  | 1 |  | 0.38005 | 0.71236 | 0.33231 | 0.067393
72 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4084 | 11050 |  | 4238 | 5.914 | 16.066 | 10.153 |  | 0.545 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 5:10 |  | 1 |  | 6.9432 | 19.834 | 12.8908 | 0.85415
73 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 5981 | 11021 |  | 4244 | 8.339 | 16.026 | 7.686 |  | 0.546 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 5:20 |  | 1 |  | 10.141 | 19.7825 | 9.6414 | 0.85535 | 97.9561 |  | 4.9763 |  |  | 92.3621 |  | 0.24236 |  |  | 86.1993 |  | 4.9914 |  |  | 91.3737 |  | 1.5778 |  |  | 10.3998 | 19.8065 | 9.4066 | 0.84866
74 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6288 | 11048 |  | 4177 | 8.731 | 16.064 | 7.332 |  | 0.535 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 5:31 |  | 1 |  | 10.6586 | 19.8305 | 9.1719 | 0.84196
75 | 3 |  | Flush | [20 empty] | 2023-11-01 | 5:37 |  | 1
76 | 21 | F50 24jul23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4751 | 11047 |  | 5699 | 4.06 | 9.637 | 5.577 |  | 0.48 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 5:57 |  | 1 |  | 4.8405 | 11.8972 | 7.0567 | 0.68768
77 | 21 | F50 24jul23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5609 | 10996 |  | 5640 | 4.718 | 9.594 | 4.876 |  | 0.474 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 6:07 |  | 1 |  | 5.7084 | 11.8428 | 6.1344 | 0.6806 |  |  | 2.6402 |  |  |  |  | 0.41521 |  |  |  |  | 3.3446 |  |  |  |  | 0.78968 |  |  | 5.7847 | 11.8183 | 6.0335 | 0.6833
78 | 21 | F50 24jul23 6.2 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5760 | 10950 |  | 5685 | 4.834 | 9.555 | 4.722 |  | 0.479 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 6:19 |  | 1 |  | 5.8611 | 11.7937 | 5.9326 | 0.686
79 | 22 | B50 22aug23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5639 | 8005 |  | 12699 | 4.741 | 7.06 | 2.319 |  | 1.212 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 6:37 |  | 1 |  | 5.7387 | 8.6522 | 2.9135 | 1.527
80 | 22 | B50 22aug23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5713 | 8005 |  | 12501 | 4.798 | 7.06 | 2.262 |  | 1.191 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 6:47 |  | 1 |  | 5.8136 | 8.6522 | 2.8386 | 1.5033 |  |  | 2.2535 |  |  |  |  | 0.51917 |  |  |  |  | 6.4475 |  |  |  |  | 0.57265 |  |  | 5.8798 | 8.6298 | 2.75 | 1.5076
81 | 22 | B50 22aug23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5844 | 7963 |  | 12573 | 4.898 | 7.024 | 2.126 |  | 1.199 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 6:58 |  | 1 |  | 5.9461 | 8.6074 | 2.6613 | 1.5119
82 | 23 | F50 22aug23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5705 | 8179 |  | 2702 | 4.791 | 7.207 | 2.416 |  | 0.167 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 7:17 |  | 1 |  | 5.8055 | 8.8378 | 3.0323 | 0.32831
83 | 23 | F50 22aug23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5732 | 8193 |  | 2909 | 4.813 | 7.22 | 2.407 |  | 0.188 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 7:28 |  | 1 |  | 5.8328 | 8.8527 | 3.0199 | 0.35313 |  |  | 0.34622 |  |  |  |  | 0.91162 |  |  |  |  | 1.9946 |  |  |  |  | 0.47425 |  |  | 5.8429 | 8.8932 | 3.0504 | 0.35397
84 | 23 | F50 22aug23 3.8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5752 | 8269 |  | 2923 | 4.828 | 7.284 | 2.456 |  | 0.19 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 7:39 |  | 1 |  | 5.853 | 8.9338 | 3.0808 | 0.35481
85 | 24 | F50 31jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4705 | 10132 |  | 4611 | 4.024 | 8.862 | 4.837 |  | 0.366 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 7:58 |  | 1 |  | 4.794 | 10.9211 | 6.1271 | 0.55722
86 | 24 | F50 31jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4460 | 10060 |  | 4616 | 3.837 | 8.801 | 4.964 |  | 0.367 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 8:08 |  | 1 |  | 4.5462 | 10.8443 | 6.2981 | 0.55782 |  |  | 0.73151 |  |  |  |  | 0.019672 |  |  |  |  | 0.49732 |  |  |  |  | 5.2741 |  |  | 4.5629 | 10.8454 | 6.2825 | 0.57293
87 | 24 | F50 31jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4493 | 10062 |  | 4868 | 3.861 | 8.803 | 4.941 |  | 0.393 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 8:19 |  | 1 |  | 4.5796 | 10.8465 | 6.2669 | 0.58803
88 | 25 | F50 24jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4334 | 10305 |  | 5531 | 3.74 | 9.008 | 5.269 |  | 0.462 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 8:38 |  | 1 |  | 4.4188 | 11.1057 | 6.6869 | 0.66753
89 | 25 | F50 24jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4266 | 10384 |  | 5796 | 3.688 | 9.075 | 5.388 |  | 0.49 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 8:48 |  | 1 |  | 4.35 | 11.19 | 6.84 | 0.69931 |  |  | 4.3446 |  |  |  |  | 0.25706 |  |  |  |  | 2.4325 |  |  |  |  | 0.34235 |  |  | 4.4466 | 11.2044 | 6.7578 | 0.70051
90 | 25 | F50 24jul23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4457 | 10411 |  | 5816 | 3.834 | 9.099 | 5.264 |  | 0.492 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 8:59 |  | 1 |  | 4.5432 | 11.2188 | 6.6756 | 0.70171
91 | 26 | F100 14aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4984 | 6640 |  | 1695 | 4.239 | 5.904 | 1.666 |  | 0.061 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 9:18 |  | 1 |  | 5.0762 | 7.1961 | 2.1198 | 0.20757
92 | 26 | F100 14aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5205 | 6640 |  | 1807 | 4.408 | 5.904 | 1.496 |  | 0.073 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 9:28 |  | 1 |  | 5.2997 | 7.1961 | 1.8963 | 0.221 |  |  | 2.5809 |  |  |  |  | 0.56173 |  |  |  |  | 5.3068 |  |  |  |  | 0.37909 |  |  | 5.369 | 7.2163 | 1.8473 | 0.22141
93 | 26 | F100 14aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5342 | 6678 |  | 1814 | 4.513 | 5.936 | 1.423 |  | 0.074 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 9:39 |  | 1 |  | 5.4383 | 7.2366 | 1.7983 | 0.22183
94 | 27 | F100 12sep23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5310 | 6685 |  | 1292 | 4.489 | 5.942 | 1.453 |  | 0.019 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 9:57 |  | 1 |  | 5.4059 | 7.2441 | 1.8381 | 0.15924
95 | 27 | F100 12sep23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5324 | 6692 |  | 1398 | 4.5 | 5.948 | 1.449 |  | 0.03 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 10:08 |  | 1 |  | 5.4201 | 7.2515 | 1.8314 | 0.17195 |  |  | 0.81773 |  |  |  |  | 0.60496 |  |  |  |  | 4.9371 |  |  |  |  | 3.26 |  |  | 5.4424 | 7.2297 | 1.7873 | 0.16919
96 | 27 | F100 12sep23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5368 | 6651 |  | 1352 | 4.533 | 5.913 | 1.38 |  | 0.025 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 10:19 |  | 1 |  | 5.4646 | 7.2078 | 1.7432 | 0.16644
97 | 28 | B50 22aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3731 | 7280 |  | 2740 | 3.278 | 6.446 | 3.169 |  | 0.171 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 10:37 |  | 1 |  | 3.8089 | 7.8788 | 4.0699 | 0.33287
98 | 28 | B50 22aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3322 | 7214 |  | 2696 | 2.963 | 6.39 | 3.427 |  | 0.166 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 10:47 |  | 1 |  | 3.3952 | 7.8084 | 4.4132 | 0.32759 |  |  | 0.53768 |  |  |  |  | 0.78925 |  |  |  |  | 1.7982 |  |  |  |  | 5.5169 |  |  | 3.3861 | 7.8393 | 4.4532 | 0.33689
99 | 28 | B50 22aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3304 | 7272 |  | 2851 | 2.95 | 6.44 | 3.49 |  | 0.182 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 10:58 |  | 1 |  | 3.377 | 7.8702 | 4.4933 | 0.34618
100 | 29 | B50 14aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3186 | 6850 |  | 2431 | 2.859 | 6.082 | 3.222 |  | 0.138 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 11:16 |  | 1 |  | 3.2576 | 7.4201 | 4.1624 | 0.29582
101 | 29 | B50 14aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3198 | 6866 |  | 2469 | 2.868 | 6.095 | 3.227 |  | 0.142 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 11:26 |  | 1 |  | 3.2698 | 7.4371 | 4.1674 | 0.30037 |  |  | 1.5348 |  |  |  |  | 0.12901 |  |  |  |  | 0.98801 |  |  |  |  | 1.2856 |  |  | 3.295 | 7.4419 | 4.1469 | 0.29846
102 | 29 | B50 14aug23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3248 | 6875 |  | 2437 | 2.907 | 6.103 | 3.196 |  | 0.139 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 11:38 |  | 1 |  | 3.3203 | 7.4467 | 4.1264 | 0.29654
103 | 31 | SPK 25 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6936 | 16159 |  | 7580 | 5.736 | 13.969 | 8.232 |  | 0.677 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 11:57 |  | 1 |  | 7.0506 | 17.3504 | 10.2999 | 0.91323
104 | 31 | SPK 25 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8175 | 16365 |  | 7768 | 6.686 | 14.143 | 7.457 |  | 0.696 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 12:08 |  | 1 |  | 8.3038 | 17.5702 | 9.2664 | 0.93577 |  |  | 0.83695 |  |  |  |  | 0.33448 |  |  |  |  | 1.396 |  |  |  |  | 0.79763 |  |  | 8.3386 | 17.5408 | 9.2022 | 0.93205
105 | 31 | SPK 25 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8244 | 16310 |  | 7706 | 6.74 | 14.096 | 7.356 |  | 0.69 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 12:20 |  | 1 |  | 8.3735 | 17.5115 | 9.138 | 0.92833
106 | 32 | DUP 28 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4548 | 7601 |  | 2832 | 3.904 | 6.718 | 2.814 |  | 0.18 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 12:38 |  | 1 |  | 4.6352 | 8.2212 | 3.586 | 0.3439
107 | 32 | DUP 28 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3443 | 7622 |  | 2906 | 3.056 | 6.736 | 3.68 |  | 0.188 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 12:49 |  | 1 |  | 3.5176 | 8.2436 | 4.726 | 0.35277 |  |  | 0.057525 |  | -29.946 |  |  | 0.40035 |  | -47.8314 |  |  | 0.73978 |  | -65.7168 |  |  | 4.2581 |  | -112.2719 | 3.5166 | 8.2601 | 4.7436 | 0.36045
108 | 32 | DUP 28 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3441 | 7653 |  | 3034 | 3.055 | 6.762 | 3.708 |  | 0.201 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 12:59 |  | 1 |  | 3.5155 | 8.2767 | 4.7611 | 0.36812
109 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 958 | 633 |  | 400 | 1.15 | 0.815 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 13:17 |  | 1 |  | 1.0041 | 0.7881 | -0.21601 | 0.052284
110 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 286 | 585 |  | 391 | 0.635 | 0.774 | 0.139 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 13:26 |  | 1 |  | 0.32442 | 0.7369 | 0.41248 | 0.051205 |  |  | 0 | 185.0206 |  |  |  | 0.58073 | 183.9182 |  |  |  | 1.0399 | 182.9246 |  |  |  | 8.9478 | 178.2466 |  | 0.32442 | 0.73476 | 0.41034 | 0.053603
111 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 286 | 581 |  | 431 | 0.634 | 0.771 | 0.136 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 13:35 |  | 1 |  | 0.32442 | 0.73263 | 0.40821 | 0.056002
112 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4351 | 10849 |  | 4392 | 6.255 | 15.783 | 9.528 |  | 0.572 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 13:52 |  | 1 |  | 7.3933 | 19.4767 | 12.0834 | 0.88493
113 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6390 | 10965 |  | 4259 | 8.862 | 15.946 | 7.084 |  | 0.549 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 14:03 |  | 1 |  | 10.8305 | 19.6829 | 8.8524 | 0.85835 | 104.511 |  | 4.7114 |  |  | 91.6548 |  | 0.27135 |  |  | 77.4914 |  | 6.7246 |  |  | 91.569 |  | 1.8564 |  |  | 11.0918 | 19.6562 | 8.5644 | 0.85046
114 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6700 | 10935 |  | 4180 | 9.259 | 15.904 | 6.645 |  | 0.535 | 1 | 0 | 0 |  | 0 |  |  | 2023-11-01 | 14:14 |  | 1 |  | 11.3531 | 19.6296 | 8.2765 | 0.84256
115 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2975 | 880 |  | 543 | 2.697 | 1.024 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 16:26 |  | 2 |  | 3.0442 | 1.0516 | -1.9926 | 0.069431
116 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 546 | 923 |  | 601 | 0.834 | 1.06 | 0.227 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 16:35 |  | 3 |  | 0.5874 | 1.0975 | 0.51006 | 0.076386 |  |  | 5.688 |  |  |  |  | 5.5956 |  |  |  |  | 20.3299 |  |  |  |  | 8.3396 |  |  | 0.60459 | 1.0676 | 0.463 | 0.073328
117 | 3 | Rinse | TIC/TC/TNb |  |  | 0.5 | 0.5 | 580 | 867 |  | 550 | 0.859 | 1.013 | 0.153 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 16:45 |  | 4 |  | 0.62179 | 1.0377 | 0.41593 | 0.070271
118 | 3 |  | Flush | [20 empty] | 2023-10-26 | 16:51
119 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 137 | 184 |  | 345 | 0.52 | 0.434 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:06 |  | 6 |  | 0.17371 | 0.30913 | 0.13542 | 0.045689
120 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 58 | 211 |  | 333 | 0.459 | 0.457 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:14 |  | 7 |  | 0.093808 | 0.33793 | 0.24413 | 0.044251 |  |  | 16.3273 |  |  |  |  | 8.2159 |  |  |  |  | 5.2585 |  |  |  |  | 3.1997 |  |  | 0.086728 | 0.3246 | 0.23787 | 0.04497
121 | 3 | Type I Reagent Grade Water | TIC/TC/TNb |  |  | 0.5 | 0.5 | 44 | 186 |  | 345 | 0.449 | 0.436 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:23 |  | 8 |  | 0.079648 | 0.31126 | 0.23162 | 0.045689
122 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 673 | 2138 |  | 1292 | 2.328 | 5.224 | 2.895 |  | 0.048 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:39 |  | 9 |  | 1.7896 | 5.9839 | 4.1943 | 0.39811
123 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1199 | 2208 |  | 1292 | 3.337 | 5.374 | 2.037 |  | 0.048 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:48 |  | 10 |  | 3.1197 | 6.1706 | 3.0509 | 0.39811 |  |  | 8.6839 |  |  |  |  | 0.82455 |  |  |  |  | 11.5771 |  |  |  |  | 0.45078 |  |  | 3.2613 | 6.1452 | 2.8839 | 0.39901
124 | 4 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.2 | 0.2 | 1311 | 2189 |  | 1298 | 3.552 | 5.332 | 1.78 |  | 0.049 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 17:57 |  | 11 |  | 3.4029 | 6.1199 | 2.717 | 0.3999
125 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3981 | 7398 |  | 3567 | 2.891 | 5.455 | 2.564 |  | 0.214 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 18:17 |  | 12 |  | 3.3848 | 6.6705 | 3.2858 | 0.36003
126 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 3959 | 7606 |  | 3690 | 2.877 | 5.602 | 2.725 |  | 0.225 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 18:27 |  | 13 |  | 3.3662 | 6.8554 | 3.4892 | 0.37232 |  |  | 2.1304 |  |  |  |  | 1.0952 |  |  |  |  | 4.3084 |  |  |  |  | 1.7052 |  |  | 3.4025 | 6.8181 | 3.4156 | 0.36917
127 | 5 | Mixed Check 3/6/0.3 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 4045 | 7522 |  | 3627 | 2.932 | 5.543 | 2.611 |  | 0.219 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 18:39 |  | 14 |  | 3.4387 | 6.7808 | 3.3421 | 0.36602
128 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.33 | 0.33 | 4854 | 11070 |  | 5394 | 6.271 | 14.632 | 8.361 |  | 0.679 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 18:57 |  | 15 |  | 7.492 | 18.0633 | 10.5712 | 0.98652
129 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.33 | 0.33 | 5943 | 11109 |  | 5321 | 7.537 | 14.682 | 7.145 |  | 0.667 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 19:08 |  | 16 |  | 9.1609 | 18.1263 | 8.9654 | 0.97326 |  |  | 1.9219 |  |  |  |  | 0.57792 |  |  |  |  | 0.81433 |  |  |  |  | 0.57701 |  |  | 9.2498 | 18.1788 | 8.929 | 0.97608
130 | 6 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.33 | 0.33 | 6059 | 11174 |  | 5352 | 7.672 | 14.765 | 7.093 |  | 0.672 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 19:19 |  | 17 |  | 9.3387 | 18.2313 | 8.8927 | 0.97889
131 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.47 | 0.47 | 8722 | 16771 |  | 7855 | 7.56 | 15.411 | 7.851 |  | 0.751 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 19:39 |  | 18 |  | 9.4224 | 19.1524 | 9.7301 | 1.0066
132 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.47 | 0.47 | 8909 | 16740 |  | 7922 | 7.713 | 15.384 | 7.671 |  | 0.758 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 19:50 |  | 19 |  | 9.6236 | 19.1172 | 9.4937 | 1.0151 |  |  | 0.57186 |  |  |  |  | 0.65675 |  |  |  |  | 1.8869 |  |  |  |  | 0.60499 |  |  | 9.5961 | 19.1802 | 9.5841 | 1.0121
133 | 7 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.47 | 0.47 | 8858 | 16851 |  | 7874 | 7.671 | 15.483 | 7.812 |  | 0.753 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 20:02 |  | 20 |  | 9.5687 | 19.2432 | 9.6745 | 1.009
134 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11206 | 21419 |  | 10124 | 7.51 | 15.354 | 7.844 |  | 0.786 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 20:23 |  | 21 |  | 9.4745 | 19.1346 | 9.66 | 1.0152
135 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11454 | 21520 |  | 10125 | 7.669 | 15.425 | 7.756 |  | 0.786 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 20:36 |  | 22 |  | 9.6836 | 19.2244 | 9.5408 | 1.0153 |  |  | 0.11309 |  |  |  |  | 0.1571 |  |  |  |  | 0.20174 |  |  |  |  | 0.65721 |  |  | 9.6891 | 19.2395 | 9.5504 | 1.0187
136 | 8 | Mixed Check 9/18/0.9 | TIC/TC/TNb |  |  | 0.6 | 0.6 | 11467 | 21554 |  | 10192 | 7.677 | 15.449 | 7.773 |  | 0.792 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 20:48 |  | 23 |  | 9.6945 | 19.2546 | 9.5601 | 1.022
137 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6569 | 10337 |  | 5002 | 9.091 | 15.06 | 5.969 |  | 0.678 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 21:07 |  | 24 |  | 11.1323 | 18.5664 | 7.4341 | 1.0068
138 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6948 | 10294 |  | 4778 | 9.575 | 14.999 | 5.424 |  | 0.64 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 21:17 |  | 25 |  | 11.7712 | 18.4899 | 6.7187 | 0.96207 | 111.5609 |  | 1.0967 |  |  | 85.9881 |  | 0.40467 |  |  | 57.8151 |  | 3.0904 |  |  | 103.9171 |  | 0.43527 |  |  | 11.8361 | 18.4526 | 6.6165 | 0.96417
139 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7025 | 10252 |  | 4799 | 9.674 | 14.94 | 5.265 |  | 0.643 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 21:28 |  | 26 |  | 11.901 | 18.4152 | 6.5143 | 0.96627
140 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5617 | 7148 |  | 3452 | 4.724 | 6.334 | 1.61 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 21:47 |  | 27 |  | 5.7165 | 7.738 | 2.0215 | 0.41824
141 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3927 | 7107 |  | 3543 | 3.428 | 6.299 | 2.872 |  | 0.255 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 21:57 |  | 28 |  | 4.0071 | 7.6942 | 3.6871 | 0.42916 |  |  | 1.9744 |  |  |  |  | 0.92461 |  |  |  |  | 0.22896 |  |  |  |  | 1.9751 |  |  | 4.0471 | 7.73 | 3.6829 | 0.42496
142 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4006 | 7174 |  | 3473 | 3.488 | 6.356 | 2.868 |  | 0.247 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 22:08 |  | 29 |  | 4.087 | 7.7657 | 3.6787 | 0.42076
143 | 9 | F50 24JUL23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5092 | 7730 |  | 3844 | 4.321 | 6.827 | 2.505 |  | 0.286 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 22:27 |  | 30 |  | 5.1854 | 8.3588 | 3.1734 | 0.46525
144 | 9 | F50 24JUL23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5271 | 7754 |  | 3752 | 4.459 | 6.848 | 2.389 |  | 0.276 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 22:38 |  | 31 |  | 5.3665 | 8.3844 | 3.0179 | 0.45422 |  |  | 1.4965 |  |  |  |  | 1.2418 |  |  |  |  | 6.3024 |  |  |  |  | 1.8114 |  |  | 5.407 | 8.3327 | 2.9257 | 0.45014
145 | 9 | F50 24JUL23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5351 | 7657 |  | 3684 | 4.52 | 6.766 | 2.246 |  | 0.269 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 22:49 |  | 32 |  | 5.4474 | 8.2809 | 2.8335 | 0.44606
146 | 10 | B50 31JUL23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5479 | 8379 |  | 9029 | 4.618 | 7.377 | 2.76 |  | 0.828 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 23:08 |  | 33 |  | 5.5769 | 9.0511 | 3.4742 | 1.087
147 | 10 | B50 31JUL23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5541 | 8456 |  | 9198 | 4.666 | 7.443 | 2.777 |  | 0.846 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 23:18 |  | 34 |  | 5.6396 | 9.1333 | 3.4937 | 1.1072 |  |  | 0.94605 |  |  |  |  | 0.10517 |  |  |  |  | 1.8257 |  |  |  |  | 2.7111 |  |  | 5.6664 | 9.1285 | 3.4621 | 1.0924
148 | 10 | B50 31JUL23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5594 | 8447 |  | 8951 | 4.706 | 7.434 | 2.728 |  | 0.82 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 23:29 |  | 35 |  | 5.6932 | 9.1237 | 3.4305 | 1.0776
149 | 11 | B50 14AUG23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3826 | 7375 |  | 3312 | 3.35 | 6.527 | 3.176 |  | 0.23 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 23:48 |  | 36 |  | 3.905 | 7.9801 | 4.0752 | 0.40146
150 | 11 | B50 14AUG23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3443 | 7435 |  | 3342 | 3.056 | 6.577 | 3.521 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-26 | 23:58 |  | 37 |  | 3.5176 | 8.0441 | 4.5266 | 0.40505 |  |  | 0.6346 |  |  |  |  | 0.14598 |  |  |  |  | 0.23209 |  |  |  |  | 1.0116 |  |  | 3.5064 | 8.0383 | 4.5318 | 0.40302
151 | 11 | B50 14AUG23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3421 | 7424 |  | 3308 | 3.039 | 6.568 | 3.529 |  | 0.23 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 0:09 |  | 38 |  | 3.4953 | 8.0324 | 4.5371 | 0.40098
152 | 12 | B50 29MAY23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2990 | 5822 |  | 2190 | 2.709 | 5.211 | 2.502 |  | 0.113 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 0:27 |  | 39 |  | 3.0594 | 6.3235 | 3.2641 | 0.26692
153 | 12 | B50 29MAY23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2857 | 5904 |  | 2306 | 2.606 | 5.28 | 2.674 |  | 0.125 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 0:38 |  | 40 |  | 2.9249 | 6.4109 | 3.4861 | 0.28083 |  |  | 0.13823 |  |  |  |  | 0.28327 |  |  |  |  | 0.63829 |  |  |  |  | 5.0338 |  |  | 2.9269 | 6.4019 | 3.475 | 0.27393
154 | 12 | B50 29MAY23 3 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2861 | 5887 |  | 2191 | 2.609 | 5.266 | 2.657 |  | 0.113 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 0:48 |  | 41 |  | 2.9289 | 6.3928 | 3.4639 | 0.26704
155 | 13 | F 2MAY22 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4043 | 15402 |  | 3193 | 3.516 | 13.327 | 9.81 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 1:08 |  | 42 |  | 4.1244 | 16.5429 | 12.4185 | 0.38719
156 | 13 | F 2MAY22 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4469 | 15502 |  | 2974 | 3.843 | 13.411 | 9.568 |  | 0.195 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 1:19 |  | 43 |  | 4.5553 | 16.6496 | 12.0943 | 0.36093 |  |  | 2.2907 |  |  |  |  | 0.9657 |  |  |  |  | 0.47112 |  |  |  |  | 4.8938 |  |  | 4.5037 | 16.5696 | 12.0658 | 0.36998
157 | 13 | F 2MAY22 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4367 | 15352 |  | 3125 | 3.765 | 13.284 | 9.519 |  | 0.211 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 1:31 |  | 44 |  | 4.4521 | 16.4896 | 12.0374 | 0.37903
158 | 14 | C50 17NOV22 21 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 11661 | 15921 |  | 2796 | 9.361 | 13.767 | 4.406 |  | 0.176 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 1:50 |  | 45 |  | 11.8297 | 17.0965 | 5.2669 | 0.33958
159 | 14 | C50 17NOV22 21 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 13845 | 15777 |  | 2852 | 11.037 | 13.644 | 2.607 |  | 0.182 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 2:01 |  | 46 |  | 14.0387 | 16.9429 | 2.9043 | 0.3463 |  |  | 0.7608 |  |  |  |  | 0.11327 |  |  |  |  | 3.0771 |  |  |  |  | 0 |  |  | 14.0923 | 16.9525 | 2.8602 | 0.3463
160 | 14 | C50 17NOV22 21 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 13951 | 15795 |  | 2852 | 11.118 | 13.66 | 2.543 |  | 0.182 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 2:13 |  | 47 |  | 14.1459 | 16.9621 | 2.8162 | 0.3463
161 | 15 | F50 14AUG23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 7752 | 7661 |  | 6831 | 6.362 | 6.769 | 0.407 |  | 0.598 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 2:32 |  | 48 |  | 7.8759 | 8.2852 | 0.4093 | 0.82341
162 | 15 | F50 14AUG23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6266 | 7802 |  | 6949 | 5.222 | 6.889 | 1.666 |  | 0.611 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 2:42 |  | 49 |  | 6.3729 | 8.4356 | 2.0627 | 0.83756 |  |  | 0.9408 |  |  |  |  | 0.34085 |  |  |  |  | 4.1993 |  |  |  |  | 0.7701 |  |  | 6.3431 | 8.45 | 2.107 | 0.8408
163 | 15 | F50 14AUG23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6207 | 7829 |  | 7003 | 5.177 | 6.911 | 1.734 |  | 0.616 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 2:53 |  | 50 |  | 6.3132 | 8.4644 | 2.1512 | 0.84404
164 | 16 | F50 29MAY23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4803 | 6465 |  | 2874 | 4.099 | 5.756 | 1.656 |  | 0.185 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 3:12 |  | 51 |  | 4.8931 | 7.0094 | 2.1162 | 0.34894
165 | 16 | F50 29MAY23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4408 | 6486 |  | 2893 | 3.796 | 5.773 | 1.977 |  | 0.187 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 3:22 |  | 52 |  | 4.4936 | 7.0318 | 2.5382 | 0.35122 |  |  | 2.4237 |  |  |  |  | 0.5899 |  |  |  |  | 2.7416 |  |  |  |  | 1.9652 |  |  | 4.5487 | 7.0526 | 2.5038 | 0.3478
166 | 16 | F50 29MAY23 8 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4517 | 6525 |  | 2836 | 3.88 | 5.806 | 1.926 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 3:33 |  | 53 |  | 4.6039 | 7.0734 | 2.4695 | 0.34438
167 | 17 | B50 2MAR23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3159 | 5213 |  | 2953 | 2.838 | 4.695 | 1.857 |  | 0.193 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 3:52 |  | 54 |  | 3.2303 | 5.6738 | 2.4435 | 0.35841
168 | 17 | B50 2MAR23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2836 | 5306 |  | 2927 | 2.59 | 4.774 | 2.184 |  | 0.19 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 4:02 |  | 55 |  | 2.9036 | 5.773 | 2.8694 | 0.35529 |  |  | 2.0055 |  |  |  |  | 2.4503 |  |  |  |  | 2.9024 |  |  |  |  | 0.60934 |  |  | 2.8748 | 5.7031 | 2.8283 | 0.35421
169 | 17 | B50 2MAR23 9 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2779 | 5175 |  | 2909 | 2.547 | 4.662 | 2.115 |  | 0.188 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 4:12 |  | 56 |  | 2.846 | 5.6333 | 2.7873 | 0.35313
170 | 18 | B50 24JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3003 | 5926 |  | 2009 | 2.719 | 5.299 | 2.58 |  | 0.094 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 4:31 |  | 57 |  | 3.0725 | 6.4344 | 3.3619 | 0.24522
171 | 18 | B50 24JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3012 | 6032 |  | 2114 | 2.725 | 5.389 | 2.663 |  | 0.105 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 4:41 |  | 58 |  | 3.0816 | 6.5475 | 3.4658 | 0.25781 |  |  | 1.0448 |  |  |  |  | 0.83438 |  |  |  |  | 2.5353 |  |  |  |  | 4.7618 |  |  | 3.0978 | 6.5203 | 3.4225 | 0.25181
172 | 18 | B50 24JUL23 0.1 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3044 | 5981 |  | 2014 | 2.75 | 5.346 | 2.595 |  | 0.095 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 4:51 |  | 59 |  | 3.114 | 6.4931 | 3.3791 | 0.24582
173 | 19 | B50 29MAY23 3 SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 8081 | 18121 |  | 7404 | 6.615 | 15.63 | 9.015 |  | 0.658 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 5:11 |  | 60 |  | 8.2087 | 19.4434 | 11.2347 | 0.89212
174 | 19 | B50 29MAY23 3 SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9634 | 18388 |  | 7370 | 7.806 | 15.856 | 8.051 |  | 0.655 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 5:22 |  | 61 |  | 9.7795 | 19.7282 | 9.9487 | 0.88804 |  |  | 0.2589 |  |  |  |  | 0.048677 |  |  |  |  | 0.15754 |  |  |  |  | 0.55207 |  |  | 9.7668 | 19.7234 | 9.9566 | 0.8905
175 | 19 | B50 29MAY23 3 SPIKE | TIC/TC/TNb |  |  | 0.5 | 0.5 | 9609 | 18379 |  | 7411 | 7.786 | 15.849 | 8.063 |  | 0.659 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 5:34 |  | 62 |  | 9.7542 | 19.7186 | 9.9644 | 0.89296
176 | 20 | B50 24JUL23 0.1 DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4557 | 6102 |  | 2135 | 3.911 | 5.448 | 1.537 |  | 0.107 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 5:53 |  | 63 |  | 4.6443 | 6.6221 | 1.9778 | 0.26033
177 | 20 | B50 24JUL23 0.1 DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3085 | 6211 |  | 2197 | 2.782 | 5.54 | 2.759 |  | 0.114 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 6:03 |  | 64 |  | 3.1555 | 6.7384 | 3.583 | 0.26776 |  |  | 2.2502 |  | 4.0761 |  |  | 0.41076 |  | 4.8798 |  |  | 1.2378 |  | 5.6836 |  |  | 1.3345 |  | 6.7241 | 3.1914 | 6.7523 | 3.5609 | 0.26956
178 | 20 | B50 24JUL23 0.1 DUP | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3156 | 6237 |  | 2227 | 2.836 | 5.562 | 2.726 |  | 0.117 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 6:14 |  | 65 |  | 3.2273 | 6.7662 | 3.5389 | 0.27136
179 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 867 | 440 |  | 383 | 1.08 | 0.651 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 6:31 |  | 66 |  | 0.91207 | 0.58222 | -0.32985 | 0.050246
180 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 213 | 480 |  | 410 | 0.578 | 0.685 | 0.107 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 6:39 |  | 67 |  | 0.25058 | 0.62489 | 0.37431 | 0.053483 |  |  | 3.1778 |  |  |  |  | 7.0807 |  |  |  |  | 9.6081 |  |  |  |  | 2.4361 |  |  | 0.25463 | 0.64782 | 0.39319 | 0.054143
181 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 221 | 523 |  | 421 | 0.584 | 0.721 | 0.137 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 6:49 |  | 68 |  | 0.25867 | 0.67076 | 0.41208 | 0.054802
182 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4932 | 8281 |  | 4364 | 6.998 | 12.157 | 5.159 |  | 0.567 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 7:06 |  | 69 |  | 8.3727 | 14.911 | 6.5383 | 0.87933
183 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7117 | 8405 |  | 4328 | 9.791 | 12.332 | 2.541 |  | 0.561 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 7:16 |  | 70 |  | 12.0561 | 15.1314 | 3.0754 | 0.87214 | 115.2655 |  | 2.7987 |  |  | 70.3064 |  | 0.12933 |  |  | 20.7758 |  | 12.4984 |  |  | 94.5095 |  | 1.2298 |  |  | 12.2272 | 15.1217 | 2.8945 | 0.87754
184 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7320 | 8394 |  | 4382 | 10.051 | 12.316 | 2.265 |  | 0.571 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 7:27 |  | 71 |  | 12.3983 | 15.1119 | 2.7136 | 0.88293
185 | 3 |  | Flush | [20 empty] | 2023-10-27 | 7:33
186 | 21 | F50 29MAY23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3774 | 7772 |  | 2787 | 3.31 | 6.863 | 3.553 |  | 0.175 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 7:52 |  | 73 |  | 3.8524 | 8.4036 | 4.5513 | 0.33851
187 | 21 | F50 29MAY23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4440 | 7777 |  | 2820 | 3.821 | 6.867 | 3.046 |  | 0.179 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 8:02 |  | 74 |  | 4.526 | 8.4089 | 3.883 | 0.34246 |  |  | 3.9003 |  |  |  |  | 0.12694 |  |  |  |  | 5.035 |  |  |  |  | 0.31463 |  |  | 4.616 | 8.4036 | 3.7876 | 0.343
188 | 21 | F50 29MAY23 1.6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4618 | 7767 |  | 2829 | 3.958 | 6.858 | 2.9 |  | 0.18 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 8:13 |  | 75 |  | 4.706 | 8.3983 | 3.6923 | 0.34354
189 | 22 | F50 29MAY23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4254 | 7472 |  | 3010 | 3.678 | 6.608 | 2.93 |  | 0.199 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 8:32 |  | 76 |  | 4.3379 | 8.0836 | 3.7457 | 0.36524
190 | 22 | F50 29MAY23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4262 | 7502 |  | 3109 | 3.685 | 6.634 | 2.949 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 8:42 |  | 77 |  | 4.3459 | 8.1156 | 3.7696 | 0.37712 |  |  | 2.777 |  |  |  |  | 0.18385 |  |  |  |  | 2.8916 |  |  |  |  | 1.28 |  |  | 4.4071 | 8.1231 | 3.7159 | 0.37472
191 | 22 | F50 29MAY23 5 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4383 | 7516 |  | 3069 | 3.778 | 6.646 | 2.868 |  | 0.205 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 8:53 |  | 78 |  | 4.4683 | 8.1305 | 3.6622 | 0.37232
192 | 23 | B50 29MAY23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3184 | 5639 |  | 2056 | 2.857 | 5.056 | 2.198 |  | 0.099 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 9:12 |  | 79 |  | 3.2556 | 6.1282 | 2.8726 | 0.25085
193 | 23 | B50 29MAY23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2852 | 5638 |  | 2003 | 2.603 | 5.055 | 2.452 |  | 0.094 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 9:22 |  | 80 |  | 2.9198 | 6.1272 | 3.2074 | 0.2445 |  |  | 0.34581 |  |  |  |  | 0.85675 |  |  |  |  | 1.9642 |  |  |  |  | 0.48923 |  |  | 2.9249 | 6.101 | 3.1762 | 0.2451
194 | 23 | B50 29MAY23 6 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2862 | 5589 |  | 2013 | 2.611 | 5.013 | 2.403 |  | 0.094 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 9:32 |  | 81 |  | 2.9299 | 6.0749 | 3.145 | 0.2457
195 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 773 | 505 |  | 413 | 1.008 | 0.706 | 0 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 9:49 |  | 82 |  | 0.817 | 0.65156 | -0.16544 | 0.053843
196 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 223 | 448 |  | 373 | 0.586 | 0.658 | 0.072 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 9:58 |  | 83 |  | 0.2607 | 0.59075 | 0.33006 | 0.049047 |  |  | 15.3986 |  |  |  |  | 6.9712 |  |  |  |  | 0.2495 |  |  |  |  | 8.6542 |  |  | 0.28244 | 0.61209 | 0.32964 | 0.051265
197 | 3 | Water Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 266 | 488 |  | 410 | 0.619 | 0.692 | 0.073 |  | 0 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 10:07 |  | 84 |  | 0.30419 | 0.63342 | 0.32923 | 0.053483
198 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 4773 | 8012 |  | 4366 | 6.795 | 11.777 | 4.982 |  | 0.568 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 10:25 |  | 85 |  | 8.1047 | 14.4327 | 6.3281 | 0.87973
199 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 6841 | 8045 |  | 4443 | 9.439 | 11.824 | 2.385 |  | 0.581 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 10:35 |  | 86 |  | 11.5908 | 14.4914 | 2.9006 | 0.89512 |  |  | 3.8372 |  |  |  |  | 0.19649 |  |  |  |  | 18.1195 |  |  |  |  | 0.51219 |  |  | 11.8175 | 14.4772 | 2.6596 | 0.89742
200 | 1 | Spiked tap as reference 100+1KHP | TIC/TC/TNb |  |  | 0.3 | 0.3 | 7110 | 8029 |  | 4466 | 9.783 | 11.801 | 2.018 |  | 0.585 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-27 | 10:45 |  | 87 |  | 12.0443 | 14.4629 | 2.4187 | 0.89972

## rolling spiked blank
We expect this to run a little above spike contribution because it is not corrected for the contribution of the reagent water
That is what you are seeing after red dot at analysis order #27 where we start accounting for that contribution in the calibration standards BUT NOT in the spiked blanks which are run as samples
|  | EXPECTED CONC | TIC | TC | TOC | TNb
|  | w/out water contribution | 2.9732 | 5.9465 | 2.9732 | 0.29732
|  | with water contribution | 3.184 | 6.363 | 3.179 | 0.32597
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | TIC [mg/l] | TC [mg/l] | TOC (Diff.) [mg/l] | NPOC [mg/l] | TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] | TIC reference signal change over run- initial = 100% | TIC Absolute value Percent error for standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes | TC reference signal change over run- initial = 100% | TC Absolute value Percent error for standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes | T0C reference signal change over run- initial = 100% | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes | TNb reference signal change over run- initial = 100% | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes |  | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order |  |  |  |  | spiked blank conc | spiked blank conc | spiked blank conc | spiked blank conc | TIC Signal per mass injected | TC Signal per mass injected | TOC Signal per mass injected | TNb Signal per mass injected | Calculated TIC mass injected | Calculated TC mass injected | Calculated TOC mass injected | Calculated TNb mass injected | Observed TIC mass injected (ng) | Observed TC mass injected (ng) | Observed TOC mass injected (ng) | Observed TNb mass injected (ng) | Mean of last 2 reps | Mean of last 2 reps | Mean of last 2 reps | Mean of last 2 reps
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | 9 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1933 | 5883 |  | 3010 | 1.898 | 5.263 | 3.365 |  | 0.199 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 16:00 |  | 1 |  | 3.2854 | 5.5471 | 2.2617 | 0.29433 |  |  | 0.50767 |  |  |  |  | 0.052864 |  |  |  |  | 0.60413 |  |  |  |  | 0.24165 |  |  |  | 3.2771 | 5.5456 | 2.2686 | 0.29469 |  | 1 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1300.2647 | 1978.649 | 2657.0333 | 20247.2667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6427 | 2.7735 | 1.1309 | 0.14716 | 1.6385 | 2.7728 | 1.1343 | 0.14734
21 | 9 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1923 | 5880 |  | 3017 | 1.89 | 5.26 | 3.37 |  | 0.2 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 16:08 |  | 1 |  | 3.2687 | 5.5442 | 2.2754 | 0.29504 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1293.538 | 1977.64 | 2661.742 | 20294.3533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6344 | 2.7721 | 1.1377 | 0.14752
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
23 | 10 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1756 | 5822 |  | 2801 | 1.762 | 5.211 | 3.449 |  | 0.177 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 16:31 |  | 1 |  | 2.9909 | 5.4875 | 2.4966 | 0.27307 |  |  | 0.38861 |  |  |  |  | 1.0452 |  |  |  |  | 1.8261 |  |  |  |  | 2.5743 |  |  |  | 2.9967 | 5.5163 | 2.5196 | 0.27663 |  | 2 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1181.2027 | 1958.1327 | 2735.0627 | 18841.3933 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4955 | 2.7437 | 1.2483 | 0.13653 | 1.4984 | 2.7582 | 1.2598 | 0.13831
24 | 10 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1763 | 5881 |  | 2871 | 1.767 | 5.261 | 3.494 |  | 0.184 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 16:39 |  | 1 |  | 3.0026 | 5.5451 | 2.5426 | 0.28019 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1185.9113 | 1977.9763 | 2770.0413 | 19312.26 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5013 | 2.7726 | 1.2713 | 0.14009
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
26 | 11 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1748 | 5839 |  | 2844 | 1.756 | 5.225 | 3.469 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 17:02 |  | 1 |  | 2.9776 | 5.5041 | 2.5265 | 0.27744 |  |  | 2.8095 |  |  |  |  | 0.28447 |  |  |  |  | 4.0578 |  |  |  |  | 1.8527 |  |  |  | 3.02 | 5.4963 | 2.4762 | 0.28004 |  | 3 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1175.8213 | 1963.8503 | 2751.8793 | 19130.64 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4888 | 2.752 | 1.2632 | 0.13872 | 1.51 | 2.7481 | 1.2381 | 0.14002
27 | 11 | spiked blank + 150 uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1799 | 5823 |  | 2895 | 1.795 | 5.212 | 3.417 |  | 0.187 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-09 | 17:10 |  | 1 |  | 3.0625 | 5.4885 | 2.426 | 0.28263 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1210.1273 | 1958.469 | 2706.8107 | 19473.7 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5312 | 2.7442 | 1.213 | 0.14132
113 | 4 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1140 | 6840 |  | 3222 | 1.289 | 6.074 | 4.784 |  | 0.221 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 6:29 |  | 1 |  | 2.3361 | 5.4713 | 3.1352 | 0.2601 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 766.84 | 2300.52 | 3834.2 | 21673.32 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.1681 | 2.7356 | 1.5676 | 0.13005
114 | 4 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1139 | 6832 |  | 3156 | 1.288 | 6.067 | 4.778 |  | 0.214 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 6:38 |  | 1 |  | 2.3339 | 5.4651 | 3.1312 | 0.25479 | [21 empty] | 2.3262 | 5.4554 | 3.1292 | 0.24764 |  | 4 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 766.1673 | 2297.8293 | 3829.4913 | 21229.36 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.167 | 2.7325 | 1.5656 | 0.1274 | 0.58348 | 1.3663 | 0.78279 | 0.063698
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
116 | 5 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1132 | 6787 |  | 3000 | 1.283 | 6.028 | 4.745 |  | 0.198 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 7:01 |  | 1 |  | 2.3185 | 5.4301 | 3.1116 | 0.24225 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 761.4587 | 2282.6943 | 3803.93 | 20180 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.1592 | 2.715 | 1.5558 | 0.12113
117 | 5 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1117 | 6870 |  | 3003 | 1.272 | 6.098 | 4.827 |  | 0.198 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 7:10 |  | 1 |  | 2.2854 | 5.4946 | 3.2092 | 0.24249 | [21 empty] | 2.3019 | 5.4624 | 3.1604 | 0.24237 |  | 5 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 751.3687 | 2310.61 | 3869.8513 | 20200.18 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.1427 | 2.7473 | 1.6046 | 0.12125 | 0.57136 | 1.3737 | 0.8023 | 0.060623
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
119 | 6 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1108 | 6818 |  | 3015 | 1.265 | 6.055 | 4.789 |  | 0.199 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 7:34 |  | 1 |  | 2.2656 | 5.4542 | 3.1886 | 0.24346 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 745.3147 | 2293.1207 | 3840.9267 | 20280.9 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.1328 | 2.7271 | 1.5943 | 0.12173
120 | 6 | spiked blank + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1101 | 6841 |  | 2979 | 1.259 | 6.074 | 4.815 |  | 0.196 | 1 | 0 | 0 |  | 0 |  |  | 2020-10-17 | 7:43 |  | 1 |  | 2.2502 | 5.4721 | 3.2219 | 0.24056 | [21 empty] | 2.2579 | 5.4631 | 3.2053 | 0.24201 |  | 6 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 740.606 | 2300.8563 | 3861.1067 | 20038.74 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.1251 | 2.736 | 1.611 | 0.12028 | 0.56254 | 1.368 | 0.80548 | 0.060141
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
21 | 9 | spiked blank + 150 a | TIC/TC/TNb |  |  | 0.5 | 0.5 | 1000 | 7140 |  | 3455 | 1.182 | 6.327 | 5.145 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 15:57 |  | 1 |  | 2.7334 | 5.5493 | 2.8159 | 0.29195 |  |  | 0.52007 |  |  |  |  | 0.14896 |  |  |  |  | 0.79412 |  |  |  |  | 4.1612 |  |  |  | 2.7263 | 5.5534 | 2.8272 | 0.286 |  | 7 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 672.6667 | 2401.42 | 4130.1733 | 23240.6333 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.3667 | 2.7746 | 1.408 | 0.14597 | 1.3631 | 2.7767 | 1.4136 | 0.143
22 | 9 | spiked blank + 150 a | TIC/TC/TNb |  |  | 0.5 | 0.5 | 995 | 7150 |  | 3327 | 1.178 | 6.336 | 5.158 |  | 0.232 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 16:06 |  | 1 |  | 2.7192 | 5.5576 | 2.8384 | 0.28005 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 669.3033 | 2404.7833 | 4140.2633 | 22379.62 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.3596 | 2.7788 | 1.4192 | 0.14002
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
24 | 10 | spiked blank + 150 b | TIC/TC/TNb |  |  | 0.5 | 0.5 | 937 | 6985 |  | 3370 | 1.134 | 6.197 | 5.063 |  | 0.236 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 16:31 |  | 1 |  | 2.5547 | 5.4211 | 2.8664 | 0.28405 |  |  | 1.4535 |  |  |  |  | 0.51749 |  |  |  |  | 2.2419 |  |  |  |  | 0.098246 |  |  |  | 2.5363 | 5.4351 | 2.8989 | 0.28391 |  | 8 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 630.2887 | 2349.2883 | 4068.288 | 22668.8667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.2774 | 2.7105 | 1.4332 | 0.14202 | 1.2681 | 2.7176 | 1.4494 | 0.14195
25 | 10 | spiked blank + 150 b | TIC/TC/TNb |  |  | 0.5 | 0.5 | 924 | 7019 |  | 3367 | 1.124 | 6.225 | 5.101 |  | 0.236 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 16:40 |  | 1 |  | 2.5178 | 5.4492 | 2.9313 | 0.28377 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 621.544 | 2360.7237 | 4099.9033 | 22648.6867 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.2589 | 2.7246 | 1.4657 | 0.14188
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
27 | 11 | spiked blank + 150 c | TIC/TC/TNb |  |  | 0.5 | 0.5 | 930 | 6877 |  | 3169 | 1.129 | 6.105 | 4.976 |  | 0.215 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 17:04 |  | 1 |  | 2.5349 | 5.3317 | 2.7969 | 0.26536 |  |  | 1.0119 |  |  |  |  | 2.3005 |  |  |  |  | 5.2098 |  |  |  |  | 6.114 |  |  |  | 2.5221 | 5.3938 | 2.8717 | 0.27373 |  | 9 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 625.58 | 2312.9643 | 4000.3487 | 21316.8067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.2674 | 2.6659 | 1.3984 | 0.13268 | 1.261 | 2.6969 | 1.4358 | 0.13686
28 | 11 | spiked blank + 150 c | TIC/TC/TNb |  |  | 0.5 | 0.5 | 921 | 7027 |  | 3349 | 1.121 | 6.232 | 5.11 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2020-11-04 | 17:14 |  | 1 |  | 2.5093 | 5.4558 | 2.9465 | 0.28209 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 619.526 | 2363.4143 | 4107.3027 | 22527.6067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.2547 | 2.7279 | 1.4732 | 0.14105
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
58 | 19 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 902 | 7007 |  | 3333 | 1.107 | 6.215 | 5.108 |  | 0.233 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-16 | 22:19 |  | 1 |  | 3.0295 | 5.9483 | 2.9188 | 0.29141 |  |  | 0.11639 |  |  |  |  | 1.2262 |  |  |  |  | 2.6389 |  |  |  |  | 2.118 |  |  |  | 3.0313 | 5.912 | 2.8808 | 0.28836 |  | 10 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 606.7453 | 2356.6877 | 4106.63 | 22419.98 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5147 | 2.9741 | 1.4594 | 0.14571 | 1.5156 | 2.956 | 1.4404 | 0.14418
59 | 19 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 903 | 6923 |  | 3259 | 1.108 | 6.144 | 5.036 |  | 0.225 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-16 | 22:28 |  | 1 |  | 3.033 | 5.8758 | 2.8427 | 0.28531 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 607.418 | 2328.4357 | 4049.4533 | 21922.2067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5165 | 2.9379 | 1.4214 | 0.14265
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
61 | 20 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 863 | 7065 |  | 3421 | 1.077 | 6.264 | 5.187 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-16 | 22:52 |  | 1 |  | 2.8919 | 5.9983 | 3.1064 | 0.29868 |  |  | 1.4748 |  |  |  |  | 0.77995 |  |  |  |  | 0.13735 |  |  |  |  | 0.44312 |  |  |  | 2.8707 | 5.975 | 3.1043 | 0.29802 |  | 11 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 580.5113 | 2376.195 | 4171.8787 | 23011.9267 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4459 | 2.9992 | 1.5532 | 0.14934 | 1.4354 | 2.9875 | 1.5521 | 0.14901
62 | 20 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 851 | 7011 |  | 3405 | 1.067 | 6.218 | 5.151 |  | 0.24 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-16 | 23:00 |  | 1 |  | 2.8496 | 5.9517 | 3.1022 | 0.29736 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 572.4393 | 2358.033 | 4143.6267 | 22904.3 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4248 | 2.9759 | 1.5511 | 0.14868
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
96 | 28 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 976 | 7106 |  | 3375 | 1.164 | 6.298 | 5.134 |  | 0.237 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-17 | 4:54 |  | 1 |  | 3.2906 | 6.0337 | 2.7431 | 0.29488 |  |  | 2.8272 |  |  |  |  | 0.085855 |  |  |  |  | 3.1064 |  |  |  |  | 0.61388 |  |  |  | 3.2447 | 6.0311 | 2.7864 | 0.29579 |  | 12 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 656.5227 | 2389.9847 | 4123.4467 | 22702.5 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6453 | 3.0168 | 1.3716 | 0.14744 | 1.6224 | 3.0156 | 1.3932 | 0.14789
97 | 28 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 950 | 7100 |  | 3397 | 1.143 | 6.294 | 5.15 |  | 0.239 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-17 | 5:03 |  | 1 |  | 3.1988 | 6.0285 | 2.8297 | 0.2967 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 639.0333 | 2387.9667 | 4136.9 | 22850.4867 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5994 | 3.0143 | 1.4148 | 0.14835
[64 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
99 | 29 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 927 | 6988 |  | 3406 | 1.126 | 6.199 | 5.072 |  | 0.24 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-17 | 5:26 |  | 1 |  | 3.1177 | 5.9319 | 2.8142 | 0.29744 |  |  | 0.22608 |  |  |  |  | 0.97953 |  |  |  |  | 2.3323 |  |  |  |  | 0.52862 |  |  |  | 3.1212 | 5.903 | 2.7817 | 0.29665 |  | 13 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 623.562 | 2350.2973 | 4077.0327 | 22911.0267 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5589 | 2.9659 | 1.4071 | 0.14872 | 1.5606 | 2.9515 | 1.3909 | 0.14833
100 | 29 | spiked blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 929 | 6921 |  | 3387 | 1.128 | 6.142 | 5.015 |  | 0.238 | 1 | 0 | 0 |  | 0 |  |  | 2021-02-17 | 5:35 |  | 1 |  | 3.1248 | 5.874 | 2.7493 | 0.29587 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 624.9073 | 2327.763 | 4030.6187 | 22783.22 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5624 | 2.937 | 1.3746 | 0.14794
27 | 9 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2547 | 6134 |  | 4188 | 2.369 | 5.475 | 3.106 |  | 0.322 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 16:20 |  | 1 |  | 2.4662 | 5.9818 | 3.5157 | 0.42406 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1713.282 | 2063.0687 | 2412.8553 | 28171.28 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.2331 | 2.9909 | 1.7578 | 0.21203
28 | 9 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2923 | 6169 |  | 4244 | 2.658 | 5.504 | 2.847 |  | 0.328 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 16:29 |  | 1 |  | 2.8328 | 6.0171 | 3.1842 | 0.42984 |  |  | 0.85688 |  |  |  |  | 0.40227 |  |  |  |  | 1.5359 |  |  |  |  | 1.5222 |  |  |  | 2.845 | 6.005 | 3.16 | 0.42659 |  | 14 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1966.2047 | 2074.8403 | 2183.476 | 28547.9733 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4164 | 3.0085 | 1.5921 | 0.21492 | 1.4225 | 3.0025 | 1.58 | 0.21329
29 | 9 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2948 | 6145 |  | 4181 | 2.676 | 5.484 | 2.808 |  | 0.321 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 16:38 |  | 1 |  | 2.8572 | 5.9929 | 3.1357 | 0.42334 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1983.0213 | 2066.7683 | 2150.5153 | 28124.1933 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4286 | 2.9965 | 1.5678 | 0.21167
30 | 10 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3134 | 6507 |  | 3152 | 2.819 | 5.791 | 2.972 |  | 0.214 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 16:53 |  | 1 |  | 3.0386 | 6.3573 | 3.3187 | 0.31728 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2108.1373 | 2188.521 | 2268.9047 | 21202.4533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5193 | 3.1786 | 1.6593 | 0.15864
31 | 10 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3230 | 6506 |  | 3085 | 2.893 | 5.79 | 2.897 |  | 0.207 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 17:01 |  | 1 |  | 3.1322 | 6.3563 | 3.2241 | 0.31038 |  |  | 1.5372 |  |  |  |  | 1.2276 |  |  |  |  | 3.8419 |  |  |  |  | 0.099574 |  |  |  | 3.1083 | 6.3955 | 3.2872 | 0.31053 |  | 15 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2172.7133 | 2188.1847 | 2203.656 | 20751.7667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5661 | 3.1781 | 1.612 | 0.15519 | 1.5542 | 3.1978 | 1.6436 | 0.15527
32 | 10 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3181 | 6584 |  | 3088 | 2.855 | 5.856 | 3.001 |  | 0.207 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 17:10 |  | 1 |  | 3.0844 | 6.4348 | 3.3503 | 0.31069 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2139.7527 | 2214.4187 | 2289.0847 | 20771.9467 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5422 | 3.2174 | 1.6752 | 0.15534
33 | 11 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2915 | 6044 |  | 3046 | 2.651 | 5.399 | 2.748 |  | 0.203 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 17:25 |  | 1 |  | 2.825 | 5.8912 | 3.0662 | 0.30636 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1960.8233 | 2032.7987 | 2104.774 | 20489.4267 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4125 | 2.9456 | 1.5331 | 0.15318
34 | 11 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2825 | 5678 |  | 2731 | 2.582 | 5.089 | 2.506 |  | 0.17 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 17:34 |  | 1 |  | 2.7373 | 5.5229 | 2.7856 | 0.27389 |  |  | 0.82274 |  |  |  |  | 6.5913 |  |  |  |  | 13.3622 |  |  |  |  | 9.4646 |  |  |  | 2.7261 | 5.7111 | 2.985 | 0.2875 |  | 16 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1900.2833 | 1909.7007 | 1919.118 | 18370.5267 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.3686 | 2.7614 | 1.3928 | 0.13695 | 1.363 | 2.8555 | 1.4925 | 0.14375
35 | 11 | spiked blank 50 + 150 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2802 | 6052 |  | 2995 | 2.565 | 5.406 | 2.841 |  | 0.197 | 1 | 0 | 0 |  | 0 |  |  | 2021-06-22 | 17:43 |  | 1 |  | 2.7148 | 5.8993 | 3.1845 | 0.3011 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1884.812 | 2035.4893 | 2186.1667 | 20146.3667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.3574 | 2.9496 | 1.5922 | 0.15055
27 | 9 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2995 | 6848 |  | 3199 | 2.713 | 6.08 | 3.367 |  | 0.219 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 16:42 |  | 1 |  | 3.0242 | 6.5175 | 3.4933 | 0.31243 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2014.6367 | 2303.2107 | 2591.7847 | 21518.6067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5121 | 3.2587 | 1.7466 | 0.15621
28 | 9 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3408 | 6813 |  | 3183 | 3.029 | 6.05 | 3.021 |  | 0.217 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 16:50 |  | 1 |  | 3.4332 | 6.4831 | 3.05 | 0.31086 |  |  | 0.81082 |  |  |  |  | 0.30316 |  |  |  |  | 0.26522 |  |  |  |  | 1.7426 |  |  |  | 3.4193 | 6.4733 | 3.054 | 0.30818 |  | 17 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2292.448 | 2291.439 | 2290.43 | 21410.98 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7166 | 3.2416 | 1.525 | 0.15543 | 1.7097 | 3.2367 | 1.527 | 0.15409
29 | 9 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3380 | 6793 |  | 3128 | 3.008 | 6.034 | 3.026 |  | 0.211 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 16:59 |  | 1 |  | 3.4054 | 6.4635 | 3.0581 | 0.30549 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2273.6133 | 2284.7123 | 2295.8113 | 21041.0133 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7027 | 3.2318 | 1.529 | 0.15275
30 | 10 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3335 | 6850 |  | 3451 | 2.974 | 6.082 | 3.108 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 17:15 |  | 1 |  | 3.3609 | 6.5194 | 3.1586 | 0.33703 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2243.3433 | 2303.8833 | 2364.4233 | 23213.7267 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6804 | 3.2597 | 1.5793 | 0.16852
31 | 10 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3264 | 6716 |  | 3425 | 2.919 | 5.968 | 3.049 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 17:23 |  | 1 |  | 3.2906 | 6.388 | 3.0974 | 0.33449 |  |  | 1.4549 |  |  |  |  | 0.84127 |  |  |  |  | 3.224 |  |  |  |  | 0.34967 |  |  |  | 3.2668 | 6.4149 | 3.1481 | 0.33508 |  | 18 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2195.584 | 2258.8147 | 2322.0453 | 23038.8333 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6453 | 3.194 | 1.5487 | 0.16725 | 1.6334 | 3.2075 | 1.5741 | 0.16754
32 | 10 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3216 | 6771 |  | 3437 | 2.882 | 6.015 | 3.133 |  | 0.243 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 17:33 |  | 1 |  | 3.2431 | 6.4419 | 3.1989 | 0.33566 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2163.296 | 2277.313 | 2391.33 | 23119.5533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6215 | 3.221 | 1.5994 | 0.16783
33 | 11 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3252 | 6718 |  | 3107 | 2.91 | 5.97 | 3.061 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 17:48 |  | 1 |  | 3.2787 | 6.3899 | 3.1112 | 0.30344 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2187.512 | 2259.4873 | 2331.4627 | 20899.7533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6393 | 3.195 | 1.5556 | 0.15172
34 | 11 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3225 | 6609 |  | 3103 | 2.889 | 5.877 | 2.988 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 17:57 |  | 1 |  | 3.252 | 6.283 | 3.031 | 0.30305 |  |  | 1.5032 |  |  |  |  | 0.32742 |  |  |  |  | 2.2548 |  |  |  |  | 0.74379 |  |  |  | 3.2277 | 6.2933 | 3.0656 | 0.30193 |  | 19 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2169.35 | 2222.827 | 2276.304 | 20872.8467 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.626 | 3.1415 | 1.5155 | 0.15153 | 1.6139 | 3.1466 | 1.5328 | 0.15096
35 | 11 | spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3176 | 6630 |  | 3080 | 2.851 | 5.895 | 3.044 |  | 0.206 | 1 | 0 | 0 |  | 0 |  |  | 2021-07-26 | 18:06 |  | 1 |  | 3.2034 | 6.3036 | 3.1001 | 0.30081 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2136.3893 | 2229.89 | 2323.3907 | 20718.1333 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6017 | 3.1518 | 1.5501 | 0.1504
109 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3858 | 7028 |  | 3207 | 3.374 | 6.233 | 2.859 |  | 0.219 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 11:09 |  | 1 |  | 3.3538 | 6.0179 | 2.6641 | 0.26132 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2595.148 | 2363.7507 | 2132.3533 | 21572.42 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6769 | 3.009 | 1.3321 | 0.13066
110 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3835 | 7260 |  | 3139 | 3.357 | 6.429 | 3.072 |  | 0.212 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 11:20 |  | 1 |  | 3.331 | 6.2397 | 2.9087 | 0.25524 |  |  | 9.9539 |  |  |  |  | 4.4169 |  |  |  |  | 1.5757 |  |  |  |  | 1.9416 |  |  |  | 3.1731 | 6.1049 | 2.9318 | 0.25775 |  | 20 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2579.6767 | 2441.78 | 2303.8833 | 21115.0067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6655 | 3.1199 | 1.4544 | 0.12762 | 1.5865 | 3.0525 | 1.4659 | 0.12887
111 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3516 | 6978 |  | 3195 | 3.112 | 6.191 | 3.078 |  | 0.218 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-10 | 11:28 |  | 1 |  | 3.0152 | 5.9701 | 2.9549 | 0.26025 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2365.096 | 2346.934 | 2328.772 | 21491.7 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5076 | 2.985 | 1.4775 | 0.13012
112 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3894 | 7129 |  | 3183 | 3.402 | 6.318 | 2.916 |  | 0.217 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 17:08 |  | 1 |  | 3.2906 | 6.0252 | 2.7345 | 0.26282 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2619.364 | 2397.7203 | 2176.0767 | 21410.98 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6453 | 3.0126 | 1.3673 | 0.13141
113 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3574 | 7106 |  | 3186 | 3.157 | 6.299 | 3.142 |  | 0.217 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 17:19 |  | 1 |  | 3 | 6.0041 | 3.0041 | 0.26309 |  |  | 0.24252 |  |  |  |  | 0.045725 |  |  |  |  | 0.15041 |  |  |  |  | 1.656 |  |  |  | 2.9963 | 6.0027 | 3.0064 | 0.26093 |  | 21 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2404.1107 | 2389.9847 | 2375.8587 | 21431.16 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5 | 3.0021 | 1.5021 | 0.13154 | 1.4982 | 3.0014 | 1.5032 | 0.13046
114 | 1 | Spiked Blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3566 | 7103 |  | 3138 | 3.15 | 6.296 | 3.146 |  | 0.212 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-11 | 17:32 |  | 1 |  | 2.9927 | 6.0014 | 3.0087 | 0.25877 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2398.7293 | 2388.9757 | 2379.222 | 21108.28 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4964 | 3.0007 | 1.5043 | 0.12938
109 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4087 | 7360 |  | 3224 | 3.55 | 6.513 | 2.963 |  | 0.221 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 12:39 |  | 1 |  | 3.4645 | 6.3208 | 2.8563 | 0.2789 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2749.1887 | 2475.4133 | 2201.638 | 21686.7733 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7322 | 3.1604 | 1.4282 | 0.13945
110 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3723 | 7343 |  | 3220 | 3.271 | 6.499 | 3.228 |  | 0.221 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 12:50 |  | 1 |  | 3.1353 | 6.3054 | 3.1701 | 0.27854 |  |  | 1.5739 |  |  |  |  | 0.3886 |  |  |  |  | 2.3682 |  |  |  |  | 0.29708 |  |  |  | 3.1602 | 6.2932 | 3.133 | 0.27895 |  | 22 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2504.338 | 2469.6957 | 2435.0533 | 21659.8667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5677 | 3.1527 | 1.585 | 0.13927 | 1.5801 | 3.1466 | 1.5665 | 0.13948
111 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3778 | 7316 |  | 3229 | 3.313 | 6.477 | 3.163 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-14 | 13:01 |  | 1 |  | 3.185 | 6.2809 | 3.0959 | 0.27937 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2541.3347 | 2460.6147 | 2379.8947 | 21720.4067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5925 | 3.1405 | 1.5479 | 0.13968
103 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2703 | 7080 |  | 3178 | 2.489 | 6.277 | 3.788 |  | 0.216 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 13:06 |  | 1 |  | 2.157 | 5.9082 | 3.7512 | 0.28802 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1818.218 | 2381.24 | 2944.262 | 21377.3467 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.0785 | 2.9541 | 1.8756 | 0.14401
104 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3584 | 7146 |  | 3261 | 3.165 | 6.333 | 3.168 |  | 0.225 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 13:16 |  | 1 |  | 2.9422 | 5.9673 | 3.0251 | 0.29539 |  |  | 0.12125 |  |  |  |  | 0.23975 |  |  |  |  | 0.58962 |  |  |  |  | 0.24043 |  |  |  | 2.9404 | 5.9745 | 3.034 | 0.29575 |  | 23 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2410.8373 | 2403.438 | 2396.0387 | 21935.66 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4711 | 2.9836 | 1.5125 | 0.1477 | 1.4702 | 2.9872 | 1.517 | 0.14787
105 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3580 | 7162 |  | 3269 | 3.162 | 6.346 | 3.184 |  | 0.226 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-15 | 13:28 |  | 1 |  | 2.9387 | 5.9816 | 3.043 | 0.2961 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2408.1467 | 2408.8193 | 2409.492 | 21989.4733 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4693 | 2.9908 | 1.5215 | 0.14805
109 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4056 | 7447 |  | 3299 | 3.526 | 6.588 | 3.061 |  | 0.229 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 14:34 |  | 1 |  | 3.3323 | 6.1653 | 2.833 | 0.2814 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2728.336 | 2504.6743 | 2281.0127 | 22191.2733 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6661 | 3.0827 | 1.4165 | 0.1407
110 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3680 | 7469 |  | 3369 | 3.238 | 6.607 | 3.368 |  | 0.236 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 14:45 |  | 1 |  | 3.0003 | 6.1848 | 3.1845 | 0.28781 |  |  | 0.41283 |  |  |  |  | 1.0493 |  |  |  |  | 1.6527 |  |  |  |  | 1.7332 |  |  |  | 2.9941 | 6.1525 | 3.1584 | 0.28533 |  | 24 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2475.4133 | 2512.0737 | 2548.734 | 22662.14 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5002 | 3.0924 | 1.5922 | 0.1439 | 1.4971 | 3.0762 | 1.5792 | 0.14267
111 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3666 | 7396 |  | 3315 | 3.227 | 6.545 | 3.318 |  | 0.231 | 1 | 0 | 0 |  | 0 |  |  | 2021-12-16 | 14:56 |  | 1 |  | 2.9879 | 6.1202 | 3.1323 | 0.28286 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2465.996 | 2487.5213 | 2509.0467 | 22298.9 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.494 | 3.0601 | 1.5661 | 0.14143
103 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4326 | 7658 |  | 3201 | 3.734 | 6.767 | 3.033 |  | 0.219 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 14:56 |  | 1 |  | 3.6357 | 6.4785 | 2.8428 | 0.28576 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2909.956 | 2575.6407 | 2241.3253 | 21532.06 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8178 | 3.2392 | 1.4214 | 0.14288
104 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3951 | 7640 |  | 3231 | 3.446 | 6.751 | 3.306 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 15:07 |  | 1 |  | 3.2899 | 6.4622 | 3.1723 | 0.28859 |  |  | 2.5829 |  |  |  |  | 0.48751 |  |  |  |  | 3.575 |  |  |  |  | 2.1112 |  |  |  | 3.248 | 6.478 | 3.23 | 0.28557 |  | 25 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2657.706 | 2569.5867 | 2481.4673 | 21733.86 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.645 | 3.2311 | 1.5861 | 0.14429 | 1.624 | 3.239 | 1.615 | 0.14279
105 | 1 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3860 | 7675 |  | 3167 | 3.377 | 6.781 | 3.404 |  | 0.215 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-25 | 15:18 |  | 1 |  | 3.206 | 6.4938 | 3.2878 | 0.28256 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2596.4933 | 2581.3583 | 2566.2233 | 21303.3533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.603 | 3.2469 | 1.6439 | 0.14128
163 | 2 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3723 | 7161 |  | 2880 | 3.271 | 6.345 | 3.074 |  | 0.185 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 4:09 |  | 1 |  | 3.1446 | 6.0365 | 2.8919 | 0.26623 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2504.338 | 2408.483 | 2312.628 | 19372.8 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5723 | 3.0182 | 1.446 | 0.13311
164 | 2 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3614 | 7053 |  | 2836 | 3.188 | 6.253 | 3.066 |  | 0.181 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 4:19 |  | 1 |  | 3.0444 | 5.9368 | 2.8924 | 0.26158 |  |  | 1.6442 |  |  |  |  | 0.18643 |  |  |  |  | 2.0777 |  |  |  |  | 1.9613 |  |  |  | 3.0195 | 5.9423 | 2.9228 | 0.26417 |  | 26 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2431.0173 | 2372.159 | 2313.3007 | 19076.8267 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5222 | 2.9684 | 1.4462 | 0.13079 | 1.5098 | 2.9712 | 1.4614 | 0.13208
165 | 2 | Spiked blank 100 + 300 | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3560 | 7065 |  | 2885 | 3.146 | 6.264 | 3.118 |  | 0.186 | 1 | 0 | 0 |  | 0 |  |  | 2022-01-27 | 4:30 |  | 1 |  | 2.9947 | 5.9479 | 2.9532 | 0.26676 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2394.6933 | 2376.195 | 2357.6967 | 19406.4333 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4974 | 2.9739 | 1.4766 | 0.13338
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2555 | 6964 |  | 3157 | 2.375 | 6.179 | 3.804 |  | 0.214 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:37 |  | 1 |  | 2.5319 | 6.7694 | 4.2375 | 0.34669 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1718.6633 | 2342.2253 | 2965.7873 | 21236.0867 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.2659 | 3.3847 | 2.1188 | 0.17335
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3363 | 7048 |  | 3097 | 2.995 | 6.249 | 3.254 |  | 0.208 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:47 |  | 1 |  | 3.2708 | 6.8477 | 3.5769 | 0.34049 |  |  | 0.67328 |  |  |  |  | 0.12255 |  |  |  |  | 0.37839 |  |  |  |  | 3.2102 |  |  |  | 3.2598 | 6.8435 | 3.5837 | 0.33511 |  | 27 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2262.178 | 2370.4773 | 2478.7767 | 20832.4867 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6354 | 3.4238 | 1.7885 | 0.17024 | 1.6299 | 3.4217 | 1.7918 | 0.16755
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3339 | 7039 |  | 2993 | 2.977 | 6.242 | 3.266 |  | 0.197 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-15 | 16:58 |  | 1 |  | 3.2488 | 6.8393 | 3.5905 | 0.32973 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2246.034 | 2367.4503 | 2488.8667 | 20132.9133 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6244 | 3.4196 | 1.7952 | 0.16486
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2565 | 7043 |  | 3289 | 2.383 | 6.246 | 3.863 |  | 0.228 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 19:42 |  | 1 |  | 2.4722 | 6.75 | 4.2778 | 0.31385 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1725.39 | 2368.7957 | 3012.2013 | 22124.0067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.2361 | 3.375 | 2.1389 | 0.15692
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3493 | 7131 |  | 3234 | 3.095 | 6.319 | 3.225 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 19:53 |  | 1 |  | 3.3141 | 6.8305 | 3.5163 | 0.30885 |  |  | 1.6561 |  |  |  |  | 0.026771 |  |  |  |  | 1.5874 |  |  |  |  | 0.70844 |  |  |  | 3.2869 | 6.8314 | 3.5445 | 0.30776 |  | 28 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2349.6247 | 2398.393 | 2447.1613 | 21754.04 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6571 | 3.4152 | 1.7582 | 0.15443 | 1.6435 | 3.4157 | 1.7722 | 0.15388
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3433 | 7133 |  | 3210 | 3.049 | 6.321 | 3.272 |  | 0.22 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-16 | 20:04 |  | 1 |  | 3.2597 | 6.8323 | 3.5726 | 0.30667 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2309.2647 | 2399.0657 | 2488.8667 | 21592.6 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6299 | 3.4162 | 1.7863 | 0.15334
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5721 | 7959 |  | 3684 | 4.804 | 7.022 | 2.218 |  | 0.269 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 17:34 |  | 1 |  | 5.5447 | 7.6343 | 2.0896 | 0.36269 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3848.326 | 2676.877 | 1505.428 | 24781.04 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.7723 | 3.8172 | 1.0448 | 0.18135
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4089 | 7929 |  | 3587 | 3.552 | 6.996 | 3.444 |  | 0.259 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 17:44 |  | 1 |  | 4.006 | 7.6067 | 3.6007 | 0.3543 |  |  | 0.63745 |  |  |  |  | 0.30192 |  |  |  |  | 1.3368 |  |  |  |  | 0.21995 |  |  |  | 3.9933 | 7.6182 | 3.6249 | 0.35391 |  | 29 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2750.534 | 2666.787 | 2583.04 | 24128.5533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.003 | 3.8034 | 1.8003 | 0.17715 | 1.9967 | 3.8091 | 1.8124 | 0.17696
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4062 | 7954 |  | 3578 | 3.531 | 7.017 | 3.486 |  | 0.258 | 1 | 0 | 0 |  | 0 |  |  | 2022-08-18 | 17:55 |  | 1 |  | 3.9806 | 7.6297 | 3.6491 | 0.35352 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2732.372 | 2675.1953 | 2618.0187 | 24068.0133 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.9903 | 3.8149 | 1.8246 | 0.17676
25 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5043 | 6907 |  | 3345 | 4.284 | 6.13 | 1.847 |  | 0.234 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:16 |  | 1 |  | 5.0176 | 7.1112 | 2.0937 | 0.33852 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3392.258 | 2323.0543 | 1253.8507 | 22500.7 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.5088 | 3.5556 | 1.0468 | 0.16926
26 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3545 | 7179 |  | 3324 | 3.135 | 6.36 | 3.225 |  | 0.232 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:26 |  | 1 |  | 3.5293 | 7.3933 | 3.864 | 0.33644 |  |  | 1.7607 |  |  |  |  | 3.9625 |  |  |  |  | 6.016 |  |  |  |  | 1.333 |  |  |  | 3.4985 | 7.2497 | 3.7512 | 0.33421 |  | 30 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2384.6033 | 2414.537 | 2444.4707 | 22359.44 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7646 | 3.6967 | 1.932 | 0.16822 | 1.7492 | 3.6248 | 1.8756 | 0.16711
27 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3483 | 6902 |  | 3279 | 3.087 | 6.126 | 3.039 |  | 0.227 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-06 | 21:37 |  | 1 |  | 3.4677 | 7.106 | 3.6384 | 0.33199 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2342.898 | 2321.3727 | 2299.8473 | 22056.74 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7338 | 3.553 | 1.8192 | 0.16599
25 | 2 | Spiked Blank 100ml + 300uL NOT COVERED | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5148 | 7473 |  | 3415 | 4.365 | 6.61 | 2.245 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:27 |  | 1 |  | 5.1966 | 7.9063 | 2.7097 | 0.37189 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3462.888 | 2513.419 | 1563.95 | 22971.5667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.5983 | 3.9532 | 1.3548 | 0.18595
26 | 2 | Spiked Blank 100ml + 300uL NOT COVERED | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3783 | 7443 |  | 3416 | 3.317 | 6.584 | 3.266 |  | 0.241 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:37 |  | 1 |  | 3.8591 | 7.8771 | 4.018 | 0.372 |  |  | 1.8192 |  |  |  |  | 0.62093 |  |  |  |  | 0.51663 |  |  |  |  | 0.83047 |  |  |  | 3.8243 | 7.8527 | 4.0284 | 0.37355 |  | 31 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2544.698 | 2503.329 | 2461.96 | 22978.2933 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.9296 | 3.9385 | 2.009 | 0.186 | 1.9122 | 3.9264 | 2.0142 | 0.18677
27 | 2 | Spiked Blank 100ml + 300uL NOT COVERED | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3712 | 7393 |  | 3446 | 3.263 | 6.542 | 3.279 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-08 | 21:48 |  | 1 |  | 3.7895 | 7.8283 | 4.0388 | 0.3751 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2496.9387 | 2486.5123 | 2476.086 | 23180.0933 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8948 | 3.9142 | 2.0194 | 0.18755
28 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4938 | 7932 |  | 3522 | 4.203 | 6.998 | 2.795 |  | 0.252 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 17:48 |  | 1 |  | 4.7818 | 7.8664 | 3.0846 | 0.36141 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3321.628 | 2667.796 | 2013.964 | 23691.32 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.3909 | 3.9332 | 1.5423 | 0.18071
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3647 | 7719 |  | 3439 | 3.213 | 6.818 | 3.604 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 17:59 |  | 1 |  | 3.5142 | 7.6572 | 4.143 | 0.35331 |  |  | 0.2798 |  |  |  |  | 1.0447 |  |  |  |  | 1.6981 |  |  |  |  | 0.55453 |  |  |  | 3.5093 | 7.6174 | 4.1081 | 0.35233 |  | 32 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2453.2153 | 2596.157 | 2739.0987 | 23133.0067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7571 | 3.8286 | 2.0715 | 0.17665 | 1.7546 | 3.8087 | 2.054 | 0.17616
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3637 | 7638 |  | 3419 | 3.205 | 6.749 | 3.544 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-20 | 18:10 |  | 1 |  | 3.5044 | 7.5776 | 4.0732 | 0.35135 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2446.4887 | 2568.914 | 2691.3393 | 22998.4733 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7522 | 3.7888 | 2.0366 | 0.17568
28 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5005 | 7715 |  | 3111 | 4.255 | 6.815 | 2.56 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 19:34 |  | 1 |  | 4.9222 | 7.5247 | 2.6025 | 0.33478 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3366.6967 | 2594.8117 | 1822.9267 | 20926.66 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.4611 | 3.7623 | 1.3012 | 0.16739
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3766 | 7635 |  | 3084 | 3.304 | 6.746 | 3.442 |  | 0.207 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 19:45 |  | 1 |  | 3.7104 | 7.4482 | 3.7378 | 0.3319 |  |  | 1.5941 |  |  |  |  | 3.3682 |  |  |  |  | 5.1605 |  |  |  |  | 2.1738 |  |  |  | 3.6811 | 7.3248 | 3.6437 | 0.32833 |  | 33 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2533.2627 | 2567.905 | 2602.5473 | 20745.04 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8552 | 3.7241 | 1.8689 | 0.16595 | 1.8405 | 3.6624 | 1.8219 | 0.16417
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3706 | 7377 |  | 3017 | 3.258 | 6.528 | 3.27 |  | 0.2 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-21 | 19:56 |  | 1 |  | 3.6517 | 7.2014 | 3.5497 | 0.32476 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2492.9027 | 2481.131 | 2469.3593 | 20294.3533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8259 | 3.6007 | 1.7749 | 0.16238
136 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5276 | 7426 |  | 3110 | 4.462 | 6.57 | 2.107 |  | 0.209 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 19:47 |  | 1 |  | 5.5845 | 7.7309 | 2.1465 | 0.35213 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3548.9893 | 2497.6113 | 1446.2333 | 20919.9333 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.7922 | 3.8655 | 1.0732 | 0.17607
137 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4005 | 7293 |  | 3148 | 3.487 | 6.457 | 2.97 |  | 0.213 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 19:57 |  | 1 |  | 4.2973 | 7.5981 | 3.3008 | 0.35617 |  |  | 2.9412 |  |  |  |  | 0.21006 |  |  |  |  | 4.169 |  |  |  |  | 0.47586 |  |  |  | 4.235 | 7.6061 | 3.3711 | 0.35702 |  | 34 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2694.03 | 2452.879 | 2211.728 | 21175.5467 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.1487 | 3.7991 | 1.6504 | 0.17808 | 2.1175 | 3.803 | 1.6855 | 0.17851
138 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3882 | 7309 |  | 3164 | 3.393 | 6.471 | 3.078 |  | 0.215 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-22 | 20:08 |  | 1 |  | 4.1728 | 7.6141 | 3.4413 | 0.35787 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2611.292 | 2458.2603 | 2305.2287 | 21283.1733 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.0864 | 3.807 | 1.7207 | 0.17893
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5559 | 7403 |  | 3849 | 4.68 | 6.55 | 1.87 |  | 0.287 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 20:27 |  | 1 |  | 5.3038 | 7.0926 | 1.7888 | 0.376 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3739.354 | 2489.8757 | 1240.3973 | 25890.94 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.6519 | 3.5463 | 0.89441 | 0.188
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3880 | 7368 |  | 3850 | 3.392 | 6.52 | 3.129 |  | 0.287 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 20:38 |  | 1 |  | 3.7101 | 7.0589 | 3.3487 | 0.3761 |  |  | 0.71891 |  |  |  |  | 2.2226 |  |  |  |  | 3.9152 |  |  |  |  | 2.2929 |  |  |  | 3.6968 | 6.9813 | 3.2844 | 0.37184 |  | 35 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2609.9467 | 2478.104 | 2346.2613 | 25897.6667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8551 | 3.5294 | 1.6744 | 0.18805 | 1.8484 | 3.4906 | 1.6422 | 0.18592
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3852 | 7207 |  | 3763 | 3.37 | 6.384 | 3.014 |  | 0.278 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-23 | 20:49 |  | 1 |  | 3.6835 | 6.9037 | 3.2202 | 0.36758 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2591.112 | 2423.9543 | 2256.7967 | 25312.4467 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8418 | 3.4518 | 1.6101 | 0.18379
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5183 | 7619 |  | 3372 | 4.391 | 6.734 | 2.342 |  | 0.237 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 19:50 |  | 1 |  | 5.1197 | 7.6501 | 2.5303 | 0.3655 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3486.4313 | 2562.5237 | 1638.616 | 22682.32 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.5599 | 3.825 | 1.2652 | 0.18275
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3734 | 7667 |  | 3403 | 3.279 | 6.774 | 3.494 |  | 0.24 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 20:00 |  | 1 |  | 3.7285 | 7.6975 | 3.9689 | 0.36879 |  |  | 1.1775 |  |  |  |  | 0.99245 |  |  |  |  | 3.0746 |  |  |  |  | 0.4881 |  |  |  | 3.7506 | 7.6595 | 3.9088 | 0.3697 |  | 36 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2511.7373 | 2578.6677 | 2645.598 | 22890.8467 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8643 | 3.8487 | 1.9845 | 0.1844 | 1.8753 | 3.8297 | 1.9544 | 0.18485
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3780 | 7590 |  | 3420 | 3.315 | 6.709 | 3.394 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2022-09-26 | 20:11 |  | 1 |  | 3.7727 | 7.6215 | 3.8488 | 0.3706 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2542.68 | 2552.77 | 2562.86 | 23005.2 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8864 | 3.8107 | 1.9244 | 0.1853
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5111 | 7033 |  | 3051 | 4.336 | 6.237 | 1.901 |  | 0.203 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 15:50 |  | 1 |  | 4.9138 | 7.0996 | 2.1857 | 0.33817 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3437.9993 | 2365.4323 | 1292.8653 | 20523.06 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.4569 | 3.5498 | 1.0929 | 0.16909
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3541 | 7051 |  | 3132 | 3.132 | 6.252 | 3.12 |  | 0.212 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 16:01 |  | 1 |  | 3.3962 | 7.1178 | 3.7216 | 0.3469 |  |  | 0.51363 |  |  |  |  | 0.057075 |  |  |  |  | 0.57505 |  |  |  |  | 0.12426 |  |  |  | 3.3875 | 7.1199 | 3.7324 | 0.34668 |  | 37 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2381.9127 | 2371.4863 | 2361.06 | 21067.92 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6981 | 3.5589 | 1.8608 | 0.17345 | 1.6938 | 3.5599 | 1.8662 | 0.17334
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3523 | 7055 |  | 3128 | 3.118 | 6.256 | 3.138 |  | 0.211 | 1 | 0 | 0 |  | 0 |  |  | 2022-10-06 | 16:12 |  | 1 |  | 3.3788 | 7.1219 | 3.7431 | 0.34646 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2369.8047 | 2372.8317 | 2375.8587 | 21041.0133 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6894 | 3.561 | 1.8715 | 0.17323
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5218 | 8133 |  | 3549 | 4.418 | 7.169 | 2.751 |  | 0.255 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 20:28 |  | 1 |  | 5.2211 | 8.0458 | 2.8247 | 0.37314 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3509.9747 | 2735.399 | 1960.8233 | 23872.94 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.6106 | 4.0229 | 1.4124 | 0.18657
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3504 | 7864 |  | 3434 | 3.103 | 6.941 | 3.837 |  | 0.243 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 20:39 |  | 1 |  | 3.5491 | 7.7852 | 4.2361 | 0.36143 |  |  | 0.16478 |  |  |  |  | 0.29915 |  |  |  |  | 0.6895 |  |  |  |  | 0.36555 |  |  |  | 3.552 | 7.7736 | 4.2215 | 0.36209 |  | 38 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2357.024 | 2644.9253 | 2932.8267 | 23099.3733 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7746 | 3.8926 | 2.118 | 0.18072 | 1.776 | 3.8868 | 2.1108 | 0.18105
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3510 | 7840 |  | 3447 | 3.107 | 6.921 | 3.813 |  | 0.244 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-08 | 20:49 |  | 1 |  | 3.555 | 7.7619 | 4.207 | 0.36276 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2361.06 | 2636.8533 | 2912.6467 | 23186.82 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7775 | 3.881 | 2.1035 | 0.18138
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5359 | 7204 |  | 3058 | 4.526 | 6.382 | 1.856 |  | 0.204 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 21:53 |  | 1 |  | 5.4623 | 7.204 | 1.7417 | 0.32611 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3604.8207 | 2422.9453 | 1241.07 | 20570.1467 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.7311 | 3.602 | 0.87087 | 0.16306
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3567 | 7121 |  | 3051 | 3.152 | 6.311 | 3.16 |  | 0.203 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 22:04 |  | 1 |  | 3.6815 | 7.1222 | 3.4408 | 0.32539 |  |  | 1.3588 |  |  |  |  | 0.47161 |  |  |  |  | 0.46908 |  |  |  |  | 1.4375 |  |  |  | 3.6566 | 7.1055 | 3.4488 | 0.32307 |  | 39 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2399.402 | 2395.0297 | 2390.6573 | 20523.06 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8407 | 3.5611 | 1.7204 | 0.16269 | 1.8283 | 3.5527 | 1.7244 | 0.16153
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3517 | 7087 |  | 3006 | 3.113 | 6.282 | 3.169 |  | 0.198 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-09 | 22:15 |  | 1 |  | 3.6318 | 7.0887 | 3.4569 | 0.32074 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2365.7687 | 2383.5943 | 2401.42 | 20220.36 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8159 | 3.5444 | 1.7285 | 0.16037
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5054 | 8053 |  | 3916 | 4.293 | 7.101 | 2.808 |  | 0.294 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 18:59 |  | 1 |  | 4.419 | 7.6745 | 3.2555 | 0.37785 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3399.6573 | 2708.4923 | 2017.3273 | 26341.6267 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.2095 | 3.8373 | 1.6277 | 0.18892
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3666 | 7906 |  | 3981 | 3.227 | 6.977 | 3.749 |  | 0.3 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 19:09 |  | 1 |  | 3.1904 | 7.5367 | 4.3463 | 0.38433 |  |  | 0.027748 |  |  |  |  | 2.8695 |  |  |  |  | 4.9434 |  |  |  |  | 2.7604 |  |  |  | 3.19 | 7.6464 | 4.4564 | 0.37909 |  | 40 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2465.996 | 2659.0513 | 2852.1067 | 26778.86 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5952 | 3.7683 | 2.1731 | 0.19216 | 1.595 | 3.8232 | 2.2282 | 0.18955
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3665 | 8140 |  | 3876 | 3.227 | 7.174 | 3.948 |  | 0.289 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-16 | 19:20 |  | 1 |  | 3.1895 | 7.7561 | 4.5666 | 0.37386 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2465.3233 | 2737.7533 | 3010.1833 | 26072.56 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5948 | 3.878 | 2.2833 | 0.18693
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5049 | 7799 |  | 3421 | 4.288 | 6.885 | 2.597 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 21:11 |  | 1 |  | 4.4619 | 7.4703 | 3.0084 | 0.33331 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3396.294 | 2623.0637 | 1849.8333 | 23011.9267 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.231 | 3.7351 | 1.5042 | 0.16665
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4035 | 7485 |  | 3424 | 3.511 | 6.62 | 3.109 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 21:21 |  | 1 |  | 3.6001 | 7.1764 | 3.5762 | 0.3336 |  |  | 1.3308 |  |  |  |  | 0.44246 |  |  |  |  | 2.1963 |  |  |  |  | 0.84241 |  |  |  | 3.5763 | 7.1923 | 3.616 | 0.33501 |  | 41 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2714.21 | 2517.455 | 2320.7 | 23032.1067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8001 | 3.5882 | 1.7881 | 0.1668 | 1.7882 | 3.5962 | 1.808 | 0.1675
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3979 | 7519 |  | 3453 | 3.467 | 6.648 | 3.181 |  | 0.245 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-17 | 21:32 |  | 1 |  | 3.5526 | 7.2082 | 3.6557 | 0.33642 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2676.5407 | 2528.8903 | 2381.24 | 23227.18 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7763 | 3.6041 | 1.8278 | 0.16821
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5882 | 7415 |  | 3376 | 4.928 | 6.561 | 1.633 |  | 0.237 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 17:48 |  | 1 |  | 5.6398 | 7.1992 | 1.5594 | 0.34809 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3956.6253 | 2493.9117 | 1031.198 | 22709.2267 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.8199 | 3.5996 | 0.7797 | 0.17404
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3911 | 7422 |  | 3312 | 3.416 | 6.566 | 3.151 |  | 0.23 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 17:58 |  | 1 |  | 3.7512 | 7.2059 | 3.4547 | 0.34153 |  |  | 1.0528 |  |  |  |  | 0.3051 |  |  |  |  | 1.7589 |  |  |  |  | 0.81342 |  |  |  | 3.7316 | 7.2169 | 3.4853 | 0.34014 |  | 42 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2630.7993 | 2496.266 | 2361.7327 | 22278.72 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8756 | 3.6029 | 1.7273 | 0.17076 | 1.8658 | 3.6084 | 1.7427 | 0.17007
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3870 | 7445 |  | 3285 | 3.384 | 6.586 | 3.202 |  | 0.228 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-21 | 18:10 |  | 1 |  | 3.7119 | 7.2279 | 3.516 | 0.33876 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2603.22 | 2504.0017 | 2404.7833 | 22097.1 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.856 | 3.614 | 1.758 | 0.16938
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5905 | 9388 |  | 3351 | 4.945 | 8.232 | 3.287 |  | 0.234 | 1 | 0 | 0 |  | 0 |  | TCpe | 2022-11-22 | 22:24 |  | 1 |  | 5.5267 | 8.9526 | 3.426 | 0.34687 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3972.0967 | 3157.4973 | 2342.898 | 22541.06 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.7633 | 4.4763 | 1.713 | 0.17344
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4102 | 7557 |  | 3280 | 3.562 | 6.681 | 3.119 |  | 0.227 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 22:35 |  | 1 |  | 3.8552 | 7.2229 | 3.3677 | 0.33979 |  |  | 0.38549 |  |  |  |  | 1.5952 |  |  |  |  | 2.9982 |  |  |  |  | 0.20553 |  |  |  | 3.8478 | 7.1657 | 3.3179 | 0.33945 |  | 43 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2759.2787 | 2541.671 | 2324.0633 | 22063.4667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.9276 | 3.6114 | 1.6838 | 0.1699 | 1.9239 | 3.5828 | 1.659 | 0.16972
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4086 | 7436 |  | 3273 | 3.55 | 6.578 | 3.028 |  | 0.226 | 1 | 0 | 0 |  | 0 |  |  | 2022-11-22 | 22:46 |  | 1 |  | 3.8404 | 7.1085 | 3.2682 | 0.3391 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2748.516 | 2500.9747 | 2253.4333 | 22016.38 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.9202 | 3.5543 | 1.6341 | 0.16955
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4597 | 7045 |  | 3584 | 3.941 | 6.247 | 2.305 |  | 0.259 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 19:11 |  | 1 |  | 4.6737 | 7.4333 | 2.7595 | 0.3843 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3092.2487 | 2369.4683 | 1646.688 | 24108.3733 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.3369 | 3.7166 | 1.3798 | 0.19215
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3476 | 6983 |  | 3578 | 3.081 | 6.195 | 3.113 |  | 0.258 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 19:21 |  | 1 |  | 3.5465 | 7.369 | 3.8225 | 0.38365 |  |  | 1.7157 |  |  |  |  | 0.77656 |  |  |  |  | 0.087019 |  |  |  |  | 5.0957 |  |  |  | 3.5164 | 7.3405 | 3.8241 | 0.37412 |  | 44 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2338.1893 | 2348.6157 | 2359.042 | 24068.0133 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7733 | 3.6845 | 1.9112 | 0.19183 | 1.7582 | 3.6703 | 1.9121 | 0.18706
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3416 | 6928 |  | 3401 | 3.036 | 6.148 | 3.112 |  | 0.24 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 19:32 |  | 1 |  | 3.4862 | 7.312 | 3.8258 | 0.36459 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2297.8293 | 2330.1173 | 2362.4053 | 22877.3933 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7431 | 3.656 | 1.9129 | 0.18229
33 | 9 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 2751 | 6423 |  | 3236 | 2.525 | 5.72 | 3.194 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 19:56 |  | 1 |  | 2.8175 | 6.7886 | 3.9711 | 0.34682 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 1850.506 | 2160.269 | 2470.032 | 21767.4933 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.4088 | 3.3943 | 1.9855 | 0.17341
34 | 9 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3278 | 6423 |  | 3171 | 2.93 | 5.72 | 2.79 |  | 0.216 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 20:06 |  | 1 |  | 3.3475 | 6.7886 | 3.4412 | 0.33981 |  |  | 1.6048 |  |  |  |  | 0.39616 |  |  |  |  | 2.3049 |  |  |  |  | 2.3084 |  |  |  | 3.3208 | 6.8021 | 3.4813 | 0.33594 |  | 45 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2205.0013 | 2160.269 | 2115.5367 | 21330.26 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6737 | 3.3943 | 1.7206 | 0.16991 | 1.6604 | 3.401 | 1.7406 | 0.16797
35 | 9 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3225 | 6449 |  | 3099 | 2.889 | 5.742 | 2.853 |  | 0.208 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 20:16 |  | 1 |  | 3.2942 | 6.8156 | 3.5214 | 0.33206 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2169.35 | 2169.0137 | 2168.6773 | 20845.94 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6471 | 3.4078 | 1.7607 | 0.16603
36 | 10 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3447 | 6700 |  | 3386 | 3.06 | 5.954 | 2.895 |  | 0.238 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 20:34 |  | 1 |  | 3.5174 | 7.0757 | 3.5583 | 0.36297 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2318.682 | 2253.4333 | 2188.1847 | 22776.4933 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7587 | 3.5379 | 1.7792 | 0.18149
37 | 10 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3509 | 6691 |  | 3352 | 3.107 | 5.947 | 2.841 |  | 0.235 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 20:44 |  | 1 |  | 3.5797 | 7.0664 | 3.4867 | 0.35931 |  |  | 4.1869 |  |  |  |  | 1.152 |  |  |  |  | 6.3506 |  |  |  |  | 0.92495 |  |  |  | 3.5063 | 7.1073 | 3.601 | 0.36098 |  | 46 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2360.3873 | 2250.4063 | 2140.4253 | 22547.7867 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7899 | 3.5332 | 1.7433 | 0.17965 | 1.7532 | 3.5537 | 1.8005 | 0.18049
38 | 10 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3363 | 6770 |  | 3383 | 2.995 | 6.014 | 3.019 |  | 0.238 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 20:55 |  | 1 |  | 3.4329 | 7.1483 | 3.7153 | 0.36265 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2262.178 | 2276.9767 | 2291.7753 | 22756.3133 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7165 | 3.5741 | 1.8577 | 0.18132
39 | 11 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3246 | 6205 |  | 3015 | 2.905 | 5.536 | 2.631 |  | 0.199 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 21:13 |  | 1 |  | 3.3153 | 6.5627 | 3.2474 | 0.32301 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2183.476 | 2086.9483 | 1990.4207 | 20280.9 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6576 | 3.2813 | 1.6237 | 0.16151
40 | 11 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3167 | 6187 |  | 3053 | 2.845 | 5.52 | 2.675 |  | 0.203 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 21:23 |  | 2 |  | 3.2358 | 6.544 | 3.3082 | 0.32711 |  |  | 1.125 |  |  |  |  | 1.1808 |  |  |  |  | 3.3856 |  |  |  |  | 0.13162 |  |  |  | 3.2177 | 6.5829 | 3.3651 | 0.32732 |  | 47 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2130.3353 | 2080.8943 | 2031.4533 | 20536.5133 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6179 | 3.272 | 1.6541 | 0.16355 | 1.6089 | 3.2914 | 1.6826 | 0.16366
41 | 11 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3131 | 6262 |  | 3057 | 2.817 | 5.583 | 2.766 |  | 0.204 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 21:34 |  | 1 |  | 3.1996 | 6.6218 | 3.4221 | 0.32754 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2106.1193 | 2106.1193 | 2106.1193 | 20563.42 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5998 | 3.3109 | 1.7111 | 0.16377
42 | 12 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3138 | 6493 |  | 3063 | 2.823 | 5.779 | 2.956 |  | 0.204 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 21:52 |  | 1 |  | 3.2067 | 6.8612 | 3.6545 | 0.32818 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2110.828 | 2183.8123 | 2256.7967 | 20603.78 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6033 | 3.4306 | 1.8272 | 0.16409
43 | 12 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3184 | 6160 |  | 3009 | 2.858 | 5.498 | 2.64 |  | 0.199 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 22:02 |  | 1 |  | 3.2529 | 6.516 | 3.2631 | 0.32237 |  |  | 4.1975 |  |  |  |  | 0.12717 |  |  |  |  | 4.2598 |  |  |  |  | 0.87248 |  |  |  | 3.1861 | 6.5202 | 3.3341 | 0.32097 |  | 48 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2141.7707 | 2071.8133 | 2001.856 | 20240.54 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6265 | 3.258 | 1.6316 | 0.16118 | 1.593 | 3.2601 | 1.6671 | 0.16048
44 | 12 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3051 | 6168 |  | 2983 | 2.756 | 5.504 | 2.748 |  | 0.196 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 22:13 |  | 1 |  | 3.1192 | 6.5243 | 3.4051 | 0.31957 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2052.306 | 2074.504 | 2096.702 | 20065.6467 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5596 | 3.2622 | 1.7026 | 0.15978
45 | 13 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3281 | 6376 |  | 3205 | 2.932 | 5.68 | 2.748 |  | 0.219 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 22:31 |  | 1 |  | 3.3505 | 6.7399 | 3.3894 | 0.34348 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2207.0193 | 2144.4613 | 2081.9033 | 21558.9667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6752 | 3.37 | 1.6947 | 0.17174
46 | 13 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3273 | 6376 |  | 3215 | 2.926 | 5.68 | 2.754 |  | 0.22 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 22:41 |  | 1 |  | 3.3424 | 6.7399 | 3.3975 | 0.34455 |  |  | 2.1897 |  |  |  |  | 0.47557 |  |  |  |  | 3.03 |  |  |  |  | 2.4683 |  |  |  | 3.3062 | 6.756 | 3.4497 | 0.34035 |  | 49 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2201.638 | 2144.4613 | 2087.2847 | 21626.2333 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6712 | 3.37 | 1.6987 | 0.17228 | 1.6531 | 3.378 | 1.7249 | 0.17018
47 | 13 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3201 | 6407 |  | 3137 | 2.871 | 5.707 | 2.836 |  | 0.212 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 22:52 |  | 1 |  | 3.27 | 6.772 | 3.502 | 0.33615 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2153.206 | 2154.8877 | 2156.5693 | 21101.5533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.635 | 3.386 | 1.751 | 0.16808
48 | 14 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3265 | 6225 |  | 3077 | 2.92 | 5.552 | 2.633 |  | 0.206 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 23:10 |  | 1 |  | 3.3344 | 6.5834 | 3.249 | 0.32969 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2196.2567 | 2093.675 | 1991.0933 | 20697.9533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6672 | 3.2917 | 1.6245 | 0.16485
49 | 14 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3270 | 6173 |  | 3000 | 2.923 | 5.508 | 2.585 |  | 0.198 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 23:20 |  | 1 |  | 3.3394 | 6.5295 | 3.1901 | 0.3214 |  |  | 6.4014 |  |  |  |  | 0.063512 |  |  |  |  | 6.167 |  |  |  |  | 1.4637 |  |  |  | 3.2358 | 6.5274 | 3.2916 | 0.32377 |  | 50 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2199.62 | 2076.1857 | 1952.7513 | 20180 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.6697 | 3.2648 | 1.595 | 0.1607 | 1.6179 | 3.2637 | 1.6458 | 0.16188
50 | 14 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3064 | 6169 |  | 3044 | 2.766 | 5.504 | 2.739 |  | 0.202 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 23:31 |  | 1 |  | 3.1323 | 6.5254 | 3.3931 | 0.32614 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2061.0507 | 2074.8403 | 2088.63 | 20475.9733 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.5661 | 3.2627 | 1.6965 | 0.16307
51 | 15 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3432 | 6760 |  | 3319 | 3.048 | 6.006 | 2.958 |  | 0.231 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 23:49 |  | 1 |  | 3.5023 | 7.1379 | 3.6356 | 0.35576 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2308.592 | 2273.6133 | 2238.6347 | 22325.8067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7512 | 3.5689 | 1.8178 | 0.17788
52 | 15 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3422 | 6757 |  | 3314 | 3.04 | 6.003 | 2.963 |  | 0.231 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-03 | 23:59 |  | 1 |  | 3.4922 | 7.1348 | 3.6425 | 0.35522 |  |  | 1.9773 |  |  |  |  | 0.24665 |  |  |  |  | 2.3333 |  |  |  |  | 1.8326 |  |  |  | 3.4581 | 7.1436 | 3.6855 | 0.3585 |  | 51 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2301.8653 | 2272.6043 | 2243.3433 | 22292.1733 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7461 | 3.5674 | 1.8213 | 0.17761 | 1.729 | 3.5718 | 1.8428 | 0.17925
53 | 15 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3354 | 6774 |  | 3375 | 2.988 | 6.018 | 3.03 |  | 0.237 | 1 | 0 | 0 |  | 0 |  |  | 2023-04-04 | 0:10 |  | 1 |  | 3.4239 | 7.1524 | 3.7285 | 0.36179 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2256.124 | 2278.322 | 2300.52 | 22702.5 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7119 | 3.5762 | 1.8643 | 0.18089
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4964 | 7507 |  | 4058 | 4.223 | 6.638 | 2.415 |  | 0.308 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-08 | 19:01 |  | 1 |  | 4.5795 | 7.0883 | 2.5089 | 0.37351 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3339.1173 | 2524.8543 | 1710.5913 | 27296.8133 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.2897 | 3.5442 | 1.2544 | 0.18676
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3691 | 7502 |  | 4068 | 3.247 | 6.634 | 3.388 |  | 0.309 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-08 | 19:11 |  | 1 |  | 3.4261 | 7.0837 | 3.6576 | 0.37439 |  |  | 0.31682 |  |  |  |  | 1.6216 |  |  |  |  | 2.8284 |  |  |  |  | 0.89121 |  |  |  | 3.4316 | 7.1416 | 3.71 | 0.37607 |  | 52 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2482.8127 | 2523.1727 | 2563.5327 | 27364.08 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7131 | 3.5418 | 1.8288 | 0.1872 | 1.7158 | 3.5708 | 1.855 | 0.18804
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3703 | 7626 |  | 4106 | 3.256 | 6.739 | 3.483 |  | 0.313 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-08 | 19:22 |  | 1 |  | 3.437 | 7.1995 | 3.7625 | 0.37775 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2490.8847 | 2564.878 | 2638.8713 | 27619.6933 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7185 | 3.5997 | 1.8812 | 0.18887
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4782 | 7137 |  | 3936 | 4.084 | 6.324 | 2.241 |  | 0.296 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-09 | 19:28 |  | 1 |  | 4.6752 | 7.1678 | 2.4925 | 0.35473 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3216.692 | 2400.411 | 1584.13 | 26476.16 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.3376 | 3.5839 | 1.2463 | 0.17737
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3574 | 7148 |  | 3840 | 3.157 | 6.334 | 3.177 |  | 0.286 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-09 | 19:38 |  | 1 |  | 3.5202 | 7.1785 | 3.6582 | 0.34641 |  |  | 2.2824 |  |  |  |  | 1.0495 |  |  |  |  | 4.3634 |  |  |  |  | 0.62392 |  |  |  | 3.5609 | 7.141 | 3.5801 | 0.34749 |  | 53 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2404.1107 | 2404.1107 | 2404.1107 | 25830.4 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7601 | 3.5892 | 1.8291 | 0.1732 | 1.7804 | 3.5705 | 1.7901 | 0.17374
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3659 | 7071 |  | 3865 | 3.222 | 6.269 | 3.047 |  | 0.288 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-09 | 19:49 |  | 1 |  | 3.6015 | 7.1035 | 3.502 | 0.34857 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2461.2873 | 2378.213 | 2295.1387 | 25998.5667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8008 | 3.5518 | 1.751 | 0.17429
140 | 2 | Spiked Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4932 | 7542 |  | 3946 | 4.199 | 6.668 | 2.469 |  | 0.297 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-11 | 21:21 |  | 1 |  | 4.5725 | 7.2456 | 2.6731 | 0.35157 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3317.592 | 2536.626 | 1755.66 | 26543.4267 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.2863 | 3.6228 | 1.3366 | 0.17579
141 | 2 | Spiked Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3753 | 7436 |  | 3912 | 3.294 | 6.578 | 3.284 |  | 0.293 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-11 | 21:31 |  | 1 |  | 3.4891 | 7.1468 | 3.6577 | 0.34858 |  |  | 1.2989 |  |  |  |  | 0.58516 |  |  |  |  | 2.3498 |  |  |  |  | 1.602 |  |  |  | 3.4666 | 7.1678 | 3.7012 | 0.34581 |  | 54 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2524.518 | 2500.9747 | 2477.4313 | 26314.72 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7445 | 3.5734 | 1.8289 | 0.17429 | 1.7333 | 3.5839 | 1.8506 | 0.17291
142 | 2 | Spiked Blank | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3704 | 7481 |  | 3849 | 3.257 | 6.617 | 3.36 |  | 0.286 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-11 | 21:42 |  | 1 |  | 3.4441 | 7.1888 | 3.7447 | 0.34304 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2491.5573 | 2516.1097 | 2540.662 | 25890.94 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.722 | 3.5944 | 1.8724 | 0.17152
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5245 | 7587 |  | 3687 | 4.439 | 6.706 | 2.268 |  | 0.27 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-14 | 19:01 |  | 1 |  | 4.7092 | 7.1762 | 2.4671 | 0.32927 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3528.1367 | 2551.761 | 1575.3853 | 24801.22 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.3546 | 3.5881 | 1.2335 | 0.16464
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3839 | 7560 |  | 3760 | 3.36 | 6.683 | 3.323 |  | 0.277 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-14 | 19:11 |  | 1 |  | 3.4541 | 7.1514 | 3.6973 | 0.33564 |  |  | 1.1563 |  |  |  |  | 0.91762 |  |  |  |  | 2.8945 |  |  |  |  | 2.0216 |  |  |  | 3.4742 | 7.1187 | 3.6446 | 0.33228 |  | 55 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2582.3673 | 2542.68 | 2502.9927 | 25292.2667 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.727 | 3.5757 | 1.8487 | 0.16782 | 1.7371 | 3.5594 | 1.8223 | 0.16614
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3884 | 7489 |  | 3683 | 3.395 | 6.623 | 3.228 |  | 0.269 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-14 | 19:22 |  | 1 |  | 3.4942 | 7.0861 | 3.5918 | 0.32892 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2612.6373 | 2518.8003 | 2424.9633 | 24774.3133 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7471 | 3.543 | 1.7959 | 0.16446
103 | 30 | A FRESH SPIKED BLANK | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3851 | 6868 |  | 3279 | 3.37 | 6.097 | 2.728 |  | 0.227 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-15 | 11:25 |  | 1 |  | 3.4648 | 6.5147 | 3.0499 | 0.29368 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2590.4393 | 2309.9373 | 2029.4353 | 22056.74 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7324 | 3.2574 | 1.525 | 0.14684
104 | 30 | A FRESH SPIKED BLANK | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3816 | 6847 |  | 3278 | 3.343 | 6.079 | 2.737 |  | 0.227 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-15 | 11:35 |  | 1 |  | 3.4335 | 6.4954 | 3.0619 | 0.29359 |  |  | 1.086 |  |  |  |  | 0.95857 |  |  |  |  | 0.81546 |  |  |  |  | 1.1815 |  |  |  | 3.4523 | 6.5267 | 3.0744 | 0.29534 |  | 56 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2566.896 | 2302.8743 | 2038.8527 | 22050.0133 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7168 | 3.2477 | 1.5309 | 0.1468 | 1.7261 | 3.2633 | 1.5372 | 0.14767
105 | 30 | A FRESH SPIKED BLANK | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3858 | 6915 |  | 3318 | 3.375 | 6.137 | 2.763 |  | 0.231 | 1 | 0 | 0 |  | 0 |  |  | 2023-08-15 | 11:46 |  | 1 |  | 3.471 | 6.558 | 3.0869 | 0.29708 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2595.148 | 2325.745 | 2056.342 | 22319.08 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7355 | 3.279 | 1.5435 | 0.14854
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5797 | 8150 |  | 3712 | 4.862 | 7.183 | 2.321 |  | 0.272 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-21 | 20:57 |  | 1 |  | 4.9342 | 7.4571 | 2.5228 | 0.35739 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3899.4487 | 2741.1167 | 1582.7847 | 24969.3867 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.4671 | 3.7285 | 1.2614 | 0.17869
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4158 | 8068 |  | 3602 | 3.605 | 7.113 | 3.509 |  | 0.261 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-21 | 21:08 |  | 1 |  | 3.5176 | 7.3835 | 3.8659 | 0.34806 |  |  | 0.29528 |  |  |  |  | 0.54489 |  |  |  |  | 1.3033 |  |  |  |  | 1.7147 |  |  |  | 3.5125 | 7.4037 | 3.8913 | 0.35107 |  | 57 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2796.948 | 2713.5373 | 2630.1267 | 24229.4533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7588 | 3.6918 | 1.933 | 0.17403 | 1.7562 | 3.7019 | 1.9456 | 0.17553
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4146 | 8113 |  | 3673 | 3.596 | 7.152 | 3.556 |  | 0.268 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-21 | 21:19 |  | 1 |  | 3.5073 | 7.4239 | 3.9166 | 0.35408 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2788.876 | 2728.6723 | 2668.4687 | 24707.0467 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7536 | 3.7119 | 1.9583 | 0.17704
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5500 | 8490 |  | 4010 | 4.634 | 7.471 | 2.837 |  | 0.303 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-22 | 22:36 |  | 1 |  | 4.4997 | 7.5955 | 3.0958 | 0.38123 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3699.6667 | 2855.47 | 2011.2733 | 26973.9333 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.2498 | 3.7978 | 1.5479 | 0.19061
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4184 | 7982 |  | 3717 | 3.625 | 7.041 | 3.416 |  | 0.273 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-22 | 22:46 |  | 1 |  | 3.4187 | 7.1533 | 3.7346 | 0.35583 |  |  | 0.31186 |  |  |  |  | 2.1309 |  |  |  |  | 3.7672 |  |  |  |  | 0.34168 |  |  |  | 3.4241 | 7.2304 | 3.8063 | 0.35522 |  | 58 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2814.4373 | 2684.6127 | 2554.788 | 25003.02 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7094 | 3.5767 | 1.8673 | 0.17791 | 1.712 | 3.6152 | 1.9031 | 0.17761
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4197 | 8159 |  | 3703 | 3.634 | 7.191 | 3.556 |  | 0.271 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-22 | 22:58 |  | 1 |  | 3.4294 | 7.3074 | 3.878 | 0.35461 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2823.182 | 2744.1437 | 2665.1053 | 24908.8467 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7147 | 3.6537 | 1.939 | 0.17731
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6314 | 8266 |  | 3682 | 5.259 | 7.282 | 2.023 |  | 0.269 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-25 | 19:58 |  | 1 |  | 5.0504 | 7.0482 | 1.9978 | 0.32034 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 4247.2173 | 2780.1313 | 1313.0453 | 24767.5867 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.5252 | 3.5241 | 0.99889 | 0.16017
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4559 | 8321 |  | 3686 | 3.913 | 7.328 | 3.415 |  | 0.269 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-25 | 20:08 |  | 1 |  | 3.6436 | 7.0937 | 3.4502 | 0.32065 |  |  | 0.74523 |  |  |  |  | 0.7565 |  |  |  |  | 0.76839 |  |  |  |  | 0.45377 |  |  |  | 3.6572 | 7.1207 | 3.4635 | 0.31992 |  | 59 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3066.6873 | 2798.6297 | 2530.572 | 24794.4933 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8218 | 3.5469 | 1.7251 | 0.16032 | 1.8286 | 3.5603 | 1.7317 | 0.15996
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4593 | 8386 |  | 3667 | 3.938 | 7.383 | 3.445 |  | 0.268 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-25 | 20:20 |  | 1 |  | 3.6708 | 7.1476 | 3.4768 | 0.31919 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3089.558 | 2820.4913 | 2551.4247 | 24666.6867 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8354 | 3.5738 | 1.7384 | 0.1596
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 6127 | 8152 |  | 3380 | 5.115 | 7.185 | 2.07 |  | 0.238 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-26 | 21:12 |  | 1 |  | 4.8578 | 6.7599 | 1.9022 | 0.30133 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 4121.4287 | 2741.7893 | 1362.15 | 22736.1333 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.4289 | 3.38 | 0.95109 | 0.15067
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4652 | 8028 |  | 3366 | 3.984 | 7.08 | 3.096 |  | 0.236 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-26 | 21:22 |  | 1 |  | 3.686 | 6.659 | 2.973 | 0.30012 |  |  | 0.54024 |  |  |  |  | 0.2807 |  |  |  |  | 1.2892 |  |  |  |  | 1.7446 |  |  |  | 3.6761 | 6.6684 | 2.9923 | 0.30276 |  | 60 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3129.2453 | 2700.084 | 2270.9227 | 22641.96 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.843 | 3.3295 | 1.4865 | 0.15006 | 1.8381 | 3.3342 | 1.4961 | 0.15138
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4627 | 8051 |  | 3427 | 3.964 | 7.099 | 3.135 |  | 0.242 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-26 | 21:34 |  | 1 |  | 3.6662 | 6.6777 | 3.0116 | 0.3054 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3112.4287 | 2707.8197 | 2303.2107 | 23052.2867 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8331 | 3.3389 | 1.5058 | 0.1527
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 5404 | 6966 |  | 3122 | 4.56 | 6.18 | 1.62 |  | 0.211 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-27 | 23:03 |  | 1 |  | 5.0521 | 6.9593 | 1.9072 | 0.33848 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3635.0907 | 2342.898 | 1050.7053 | 21000.6533 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.526 | 3.4796 | 0.95359 | 0.16924
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3979 | 6888 |  | 3126 | 3.468 | 6.114 | 2.647 |  | 0.211 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-27 | 23:14 |  | 1 |  | 3.7178 | 6.8833 | 3.1655 | 0.3389 |  |  | 1.4501 |  |  |  |  | 0.098939 |  |  |  |  | 1.5117 |  |  |  |  | 1.4561 |  |  |  | 3.745 | 6.8868 | 3.1418 | 0.33645 |  | 61 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2676.5407 | 2316.664 | 1956.7873 | 21027.56 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8589 | 3.4417 | 1.5828 | 0.16945 | 1.8725 | 3.4434 | 1.5709 | 0.16823
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4037 | 6895 |  | 3079 | 3.512 | 6.12 | 2.609 |  | 0.206 | 1 | 0 | 0 |  | 0 |  |  | 2023-09-27 | 23:24 |  | 1 |  | 3.7721 | 6.8902 | 3.118 | 0.334 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2715.5553 | 2319.0183 | 1922.4813 | 20711.4067 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.8861 | 3.4451 | 1.559 | 0.167
29 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 4909 | 7133 |  | 3384 | 4.181 | 6.322 | 2.141 |  | 0.238 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 19:11 |  | 1 |  | 4.9316 | 7.5651 | 2.6335 | 0.38274 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 3302.1207 | 2399.0657 | 1496.0107 | 22763.04 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 2.4658 | 3.7826 | 1.3167 | 0.19137
30 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3455 | 7105 |  | 3285 | 3.065 | 6.298 | 3.233 |  | 0.228 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 19:21 |  | 1 |  | 3.482 | 7.5354 | 4.0534 | 0.37141 |  |  | 2.8786 |  |  |  |  | 0.90644 |  |  |  |  | 4.2759 |  |  |  |  | 1.5206 |  |  |  | 3.5328 | 7.5014 | 3.9686 | 0.36861 |  | 62 |  |  |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2324.0633 | 2389.6483 | 2455.2333 | 22097.1 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.741 | 3.7677 | 2.0267 | 0.18571 | 1.7664 | 3.7507 | 1.9843 | 0.1843
31 | 2 | Spiked Blank 100ml + 300uL | TIC/TC/TNb |  |  | 0.5 | 0.5 | 3557 | 7041 |  | 3236 | 3.144 | 6.243 | 3.1 |  | 0.222 | 1 | 0 | 0 |  | 0 |  |  | 2023-10-23 | 19:32 |  | 1 |  | 3.5837 | 7.4674 | 3.8837 | 0.36581 | [31 empty] | 2.9732 | 5.9465 | 2.9732 | 0.29732 | 2392.6753 | 2368.123 | 2343.5707 | 21767.4933 | 1.4866 | 2.9732 | 1.4866 | 0.14866 | 1.7918 | 3.7337 | 1.9419 | 0.1829
No. | Hole Pos. | Name | Method | Coefficients | NPOC vol. [ml] | TIC vol. [ml] | TC vol. [ml] | TIC Area | TC Area | NPOC Area | TNb Area | Stored Cal TIC [mg/l] | Stored Cal TC [mg/l] | Stored Cal TOC (Diff.) [mg/l] | Stored Cal NPOC [mg/l] | Stored Cal TNb [mg/l] | Dilut. Factor | TC Blank | TIC Blank | NPOC Blank | TNb Blank | Memo | Info | Date | Time | Misc. Notes | BRN Data Quality Code (1=no problems, 2=note, 3=fatal flaws) | BRN Sample Notes | Daily Cal TIC [mg/l] | Daily Cal TC [mg/l] | Daily Cal TOC (Diff.) [mg/l] | Daily Cal TNb [mg/l] |  | TIC Absolute value Percent error for check standards | TIC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TIC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TIC Percent Recovery (PR) of spikes |  | TC Absolute value Percent error for check standards | TC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TC Percent Recovery (PR) of spikes |  | TOC Absolute value Percent error for check standards | TOC Absolute value Relative Percent Difference (RPD) of same vial duplicates | TOC Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TOC Percent Recovery (PR) of spikes |  | TNb Absolute value Percent error for check standards | TNb Absolute value Relative Percent Difference (RPD) of same vial duplicates | TNb Absolute Value Relative Percent Difference (RPD) of independent prep duplicates | TNb Percent Recovery (PR) of spikes |  | TIC Mean of 2 reps | TC Mean of 2 reps | TOC Mean of 2 reps | TNb Mean of 2 reps |  | Analysis Order
[27 empty] | Known | [24 empty] | Known | mg/L | 2.9732 | 5.9465 | 2.9732 | 0.29732 | [20 empty] | known ug injected |  | 1.5816 | 3.1532 | 1.5716 | 0.15866 | How to calculate known?
[5 empty] | Mean |  |  | 3472.7514 | 7119.8786 |  | 3336.9653 | 3.0792 | 6.3104 | 3.2313 |  | 0.23299 | [10 empty] | Mean |  | 3.5201 | 6.7273 | 3.2072 | [20 empty] | Mean | mg/L | 3.2857 | 6.627 | 3.3413 | 0.32075 | Min | 1 | [18 empty] | Mean |  | 1.615 | 3.1293 | 1.5729 | 0.15132 | This is run as an unknown and is NOT corrected for the contribution of water, so that contribution has to be accounted for in the known (along with the stock solutions)
[5 empty] | Std |  |  | 1217.4857 | 613.5216 |  | 303.244 | 0.93405 | 0.51979 | 0.77243 |  | 0.031677 | [10 empty] | Std |  | 0.73559 | 0.71461 | 0.54345 | [20 empty] | Std |  | 0.40295 | 0.71983 | 0.43176 | 0.039153 | Max | 62 | [18 empty] | Std |  | 0.28998 | 0.61829 | 0.3201 | 0.031464 | Just water: |  |  |  |  | baseline conc from median concentration of analyzed water blanks
[5 empty] | CV |  |  | 35.0582 | 8.617 |  | 9.0874 | 30.3346 | 8.2371 | 23.9048 |  | 13.596 | [10 empty] | CV |  | 20.8966 | 10.6225 | 16.9446 | [20 empty] | CV |  | 12.2636 | 10.8621 | 12.922 | 12.2068 | [20 empty] | CV |  | 17.955 | 19.7583 | 20.3513 | 20.7933 | OC |  |  |  |  | 0.17
[84 empty] | IC |  |  |  |  | 0.19
|  |  |  | Warning Limits | Xbar + 2SD |  |  | 5907.7228 | 8346.9217 |  | 3943.4534 | 4.9473 | 7.35 | 4.7762 |  | 0.29634 | [9 empty] | Warning Limits | Xbar + 2SD |  | 4.9913 | 8.1566 | 4.2941 | [19 empty] | Warning Limits | Xbar + 2SD |  | 4.0916 | 8.0667 | 4.2048 | 0.39905 | [19 empty] | Warning Limits | Xbar + 2SD |  | 2.195 | 4.3658 | 2.2131 | 0.21425 | TNb |  |  |  |  | 0.02
[5 empty] | Xbar - 2SD |  |  | 1037.7801 | 5892.8355 |  | 2730.4772 | 1.2111 | 5.2708 | 1.6864 |  | 0.16963 | [10 empty] | Xbar - 2SD |  | 2.049 | 5.2981 | 2.1203 | [20 empty] | Xbar - 2SD |  | 2.4798 | 5.1874 | 2.4778 | 0.24244 | [20 empty] | Xbar - 2SD |  | 1.0351 | 1.8927 | 0.93268 | 0.088391 | TC |  |  |  |  | 0.38
|  |  |  | Control Limits | Xbar + 3SD |  |  | 7125.2084 | 8960.4433 |  | 4246.6975 | 5.8813 | 7.8698 | 5.5486 |  | 0.32802 | [9 empty] | Control Limits | Xbar + 3SD |  | 5.7269 | 8.8712 | 4.8375 | [19 empty] | Control Limits | Xbar + 3SD |  | 4.4946 | 8.7865 | 4.6366 | 0.43821 | [19 empty] | Control Limits | Xbar + 3SD |  | 2.485 | 4.9841 | 2.5332 | 0.24571 | Added stock:
[5 empty] | Xbar - 3SD |  |  | -179.7056 | 5279.3139 |  | 2427.2332 | 0.277 | 4.751 | 0.91399 |  | 0.13796 | [10 empty] | Xbar - 3SD |  | 1.3134 | 4.5835 | 1.5768 | [20 empty] | Xbar - 3SD |  | 2.0769 | 4.4675 | 2.046 | 0.20329 | [20 empty] | Xbar - 3SD |  | 0.74509 | 1.2744 | 0.61257 | 0.056926 |  | ml added | conc of stock | added water | final volume | added conc
[84 empty] | KHP | 0.3 | 1000 | 100 | 100.9 | 2.9732
[26 empty] | Warning Limits as % | Xbar + 2SD |  | 141.7931 | 121.245 | 133.8893 | [45 empty] | COUNT |  |  | 62 | 62 | 62 | 62 | NaHCO3 | 0.3 | 1000 | 100 | 100.9 | 2.9732
[27 empty] | Xbar - 2SD |  | 58.2069 | 78.755 | 66.1107 | [45 empty] | TINV |  |  | 2.389 | 2.389 | 2.389 | 2.389 | TNb | 0.3 | 100 | 100 | 100.9 | 0.29732
[26 empty] | Control Limits as % | Xbar + 3SD |  | 162.6897 | 131.8675 | 150.8339 | [45 empty] | MDL | MDL as ng injected |  | 0.69277 | 1.4771 | 0.76474 | 0.07517 | TC |  |  |  |  | 5.9465
[27 empty] | Xbar - 3SD |  | 37.3103 | 68.1325 | 49.1661 | [46 empty] | MDL as mg/L at 1.0 ml injection volume |  | 0.69277 | 1.4771 | 0.76474 | 0.07517
[78 empty] | MDL as mg/L at 0.5 ml injection volume |  | 1.3855 | 2.9542 | 1.5295 | 0.15034 | Total from water and added stock | [5 empty] | As mg injected in 0.5ml
[26 empty] | Coefficient of Variation % |  |  | 20.8966 | 10.6225 | 16.9446 | [46 empty] | MDL as mg/L at 0.3 ml injection volume |  | 2.3092 | 4.9237 | 2.5491 | 0.25057 | OC |  |  |  |  | 3.1432 | 1.5716
[26 empty] | 3 x Coefficient of Variation % |  |  | 62.6897 | 31.8675 | 50.8339 | [45 empty] | Mean to MDL |  |  | 2.3313 | 2.1185 | 2.0568 | 2.013 | IC |  |  |  |  | 3.1632 | 1.5816
[77 empty] | LOQ as 10sd | LOQ as ng injected |  | 2.8998 | 6.1829 | 3.201 | 0.31464 | TNb |  |  |  |  | 0.31732 | 0.15866
[51 empty] | Count |  |  | 62 | 62 | 62 | 62 | [20 empty] | LOQ as mg/L at 0.5ml injection volume |  | 5.7996 | 12.3658 | 6.402 | 0.62929 | TC |  |  |  |  | 6.3065 | 3.1532
[51 empty] | 99th percentile.inc |  |  | 4.0876 | 7.8044 | 4.3131 | 0.39762 | [20 empty] | LOQ as mg/L at 0.3ml injection volume |  | 9.666 | 20.6097 | 10.6701 | 1.0488
[51 empty] | Maximum |  |  | 4.235 | 7.8527 | 4.4564 | 0.42659
[51 empty] | Known |  |  | 2.9732 | 5.9465 | 2.9732 | 0.29732
[51 empty] | MDL |  |  | 0.96266 | 1.7197 | 1.0315 | 0.093539
[51 empty] | LOQ |  |  | 4.0295 | 7.1983 | 4.3176 | 0.39153
[51 empty] | ratio mean/mdl |  |  | 3.4132 | 3.8536 | 3.2393 | 3.429
[51 empty] | RESTRICTED various ways |  |  |  | Just points after the calibration method change (~July 2022) |  |  |  |  | Add in contribution of reagent grade water | [16 empty] | RESTRICTED various ways |  |  |  | Just points after the calibration method change (~July 2022)
[51 empty] | Known Conc |  |  | 3.184 | 6.363 | 3.179 | 0.32597 |  |  | 0.21077 | 0.41655 | 0.20577 | 0.028651 | [13 empty] | Known Conc | known ug injected |  | 1.5816 | 3.1532 | 1.5716 | 0.15866
[51 empty] | Count |  |  | 36 | 36 | 36 | 36 | [19 empty] | Count |  |  | 36 | 36 | 36 | 36
[51 empty] | Mean |  |  | 3.5377 | 7.1525 | 3.6148 | 0.34395 | [19 empty] | Mean |  |  | 1.7689 | 3.5763 | 1.8074 | 0.17198
[51 empty] | Std Dev |  |  | 0.22969 | 0.3631 | 0.31352 | 0.021294 | [19 empty] | Std Dev |  |  | 0.11484 | 0.18155 | 0.15676 | 0.010647
[51 empty] | CV or %RSD |  |  | 6.4925 | 5.0766 | 8.6731 | 6.191 | [19 empty] | CV or %RSD |  |  | 6.4925 | 5.0766 | 8.6731 | 6.191
[51 empty] | T value |  |  | 2.4377 | 2.4377 | 2.4377 | 2.4377 | [19 empty] | T value |  |  | 2.4377 | 2.4377 | 2.4377 | 2.4377
[51 empty] | MDL | mg/L |  | 0.55991 | 0.88515 | 0.76427 | 0.051909 | [19 empty] | MDL | ug injected |  | 0.27996 | 0.44257 | 0.38214 | 0.025954
[51 empty] | LOQ | mg/L |  | 2.2969 | 3.631 | 3.1352 | 0.21294 | [19 empty] | LOQ | ug injected |  | 1.1484 | 1.8155 | 1.5676 | 0.10647
[51 empty] | Error as % |  |  | 11.1088 | 12.4079 | 13.709 | 5.5148 | [19 empty] | Error as % |  |  | 11.8385 | 13.4158 | 15.0031 | 8.3913
[51 empty] | Known/MDL |  |  | 5.6866 | 7.1887 | 4.1595 | 6.2798 | [19 empty] | Known/MDL |  |  | 5.6495 | 7.1248 | 4.1127 | 6.1131
[51 empty] | Mean PR |  |  | 111.1088 | 112.4079 | 113.709 | 105.5148 | [19 empty] | Mean PR |  |  | 111.8385 | 113.4158 | 115.0031 | 108.3913
[51 empty] | S/N |  |  | 15.4024 | 19.6983 | 11.5298 | 16.1526 | [19 empty] | S/N |  |  | 15.4024 | 19.6983 | 11.5298 | 16.1526
[78 empty] | MDL as ng injected |  | 0.27996 | 0.44257 | 0.38214 | 0.025954
[78 empty] | MDL as mg/L at 1.0 ml injection volume |  | 0.27996 | 0.44257 | 0.38214 | 0.025954
[78 empty] | MDL as mg/L at 0.5 ml injection volume |  | 0.55991 | 0.88515 | 0.76427 | 0.051909
[78 empty] | MDL as mg/L at 0.3 ml injection volume |  | 0.93318 | 1.4752 | 1.2738 | 0.086514
[78 empty] | LOQ as ng injected |  | 1.1484 | 1.8155 | 1.5676 | 0.10647
[78 empty] | LOQ as mg/L at 1.0ml injection volume |  | 1.1484 | 1.8155 | 1.5676 | 0.10647
[78 empty] | LOQ as mg/L at 0.5ml injection volume |  | 2.2969 | 3.631 | 3.1352 | 0.21294
[78 empty] | LOQ as mg/L at 0.3ml injection volume |  | 3.8281 | 6.0517 | 5.2253 | 0.3549